## Summary
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
220822_22-032_HIDS2046
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -10.4074 | -76.1049 |  | 29.3974 |  | 0.33394 | 2.9334 |  | 0.5079
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.20354 | 3.4633
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.29299 | 5.1462
4 |  | EV | -10.3348 | -75.237 |  |  |  | 0.1028 | 1.627
5 |  | OSBS2-D-2 | -2.2358 | -23.5714 |  | 159.8919 |  | 0.13054 | 2.2667 |  | -3.0832
6 |  | OSBS2-C-3 | -0.33742 | -8.9254 |  |  |  | 0.035105 | 0.76247
7 |  | OSBS2-A-4 | -1.1287 | -14.5861 | -1.6045 | 27.3329 |  | 0.037553 | 0.40802 |  | 1.3118
8 |  | OSBS2-A-2 | -1.1331 | -17.6692 | -1.7843 | 103.9774 |  | 0.025113 | 0.042169 |  | -2.0099
9 |  | OSBS2-A-1 | -2.4789 | -25.8455 | -1.5308 | 108.5394 |  | 0.041008 | 0.60417 |  | -0.43465
10 |  | OSBS2-D-3 | -2.8685 | -25.3001 |  | 226.8087 |  | 0.024033 | 0.1786 |  | -5.5272
11 |  | OSBS2-E-1 | -3.4162 | -33.4197 | -2.0728 | 437.4394 |  | 0.056548 | 0.33082 |  | -7.5157
12 |  | OSBS2-C-1 | 1.0539 | -4.8232 |  |  |  | 0.11573 | 1.3562
13 |  | EV | -9.9608 | -69.6295 |  |  |  | 0.23954 | 3.1294
14 |  | OSBS2-C-2 | -0.56712 | -8.9818 | -1.526 | 172.3154 |  | 0.19662 | 2.9854 |  | -3.2099
15 |  | OSBS2-C-4 | -1.0672 | -10.4817 |  |  |  | 0.028202 | 0.12424
16 |  | OSBS2-C-5 | 2.2315 | -11.0661 |  |  |  | 0.073178 | 0.12188
17 |  | OSBS2-B-1 | -1.1056 | -26.2391 | -2.026 | 273.9896 |  | 0.090197 | 0.98388 |  | -5.1122
18 |  | OSBS2-B-2 | -1.9149 | -23.1043 |  | 57.3304 |  | 0.019096 | 0.29327 |  | -1.1982
19 |  | OSBS2-B-3 | -1.8029 | -22.6106 |  | 10.5869 |  | 0.012868 | 0.12206
20 |  | OSBS2-B-4 | -1.7916 | -25.2014 |  | 39.3056 |  | 0.031763 | 0.25622 |  | -0.7824
21 |  | OSBS2-B-5 | -1.042 | -15.6527 | -2.0405 | 539.6067 |  | 0.026564 | 0.53382 |  | -9.3907
22 |  | EV | -10.0056 | -69.9108 |  | 2.4911 |  | 0.17422 | 2.6997
23 |  | OSBS2-A-3 | -2.4503 | -21.5389 |  | 78.9829 |  | 0.16769 | 2.2296 |  | -9.8222
24 |  | OSBS2-A-5 | -3.2522 | -29.6713 |  | 60.8811 |  | 0.03216 | 0.46342 |  | -3.3482
25 |  | OSBS2-E-2 | 0.29657 | -16.3041 | -1.6225 | 6.0529 |  | 0.074677 | 0.6144
26 |  | OSBS2-E-3 | 0.30146 | -12.2862 | -1.641 | 3.6782 |  | 0.017727 | 0.35606
27 |  | OSBS2-E-4 | -0.92969 | -16.8158 | -1.6743 | 7.0336 |  | 0.044606 | 0.14999
28 |  | OSBS2-E-5 | -0.3606 | -16.0594 | -2.2637 | 10.1389 | 13.23 | 0.02609 | 0.34531 |  | 1.9963
29 |  | OSBS2-010-1 | -0.10051 | -3.9234 |  |  |  | 0.025891 | 0.83146
30 |  | OSBS2-010-2 | -0.94952 | -4.2908 |  |  |  | 0.023286 | 0.09909
31 |  | EV | -9.9775 | -69.1912 |  |  |  | 0.17628 | 3.0393
32 |  | OSBS2-D-1 | -2.0437 | -20.284 |  | 149.4109 |  | 0.2038 | 2.4837 |  | -2.4304
33 |  | OSBS2-D-4 | -1.8122 | -16.2919 |  | 52.9364 |  | 0.025902 | 0.335 |  | -13.8797
34 |  | OSBS2-D-5 | -1.636 | -18.5499 | -1.5505 | 110.1179 |  | 0.01967 | 0.1044 |  | -2.8357
35 |  | OSBS2-2030-1 | -1.4952 | -3.8675 |  |  |  | 0.018083 | 0.75829
36 |  | OSBS2-2030-2 | -1.0644 | -5.5896 |  |  |  | 0.030281 | 0.092105
37 |  | OSBS2-1020-1 | -1.419 | -3.7882 |  |  |  | 0.01812 | 0.12816
38 |  | OSBS2-1020-2 | -1.5449 | -4.7855 |  |  |  | 0.024226 | 0.10358
39 |  | OSBS2-deep-1 | -1.9543 | -9.7793 |  |  |  | 0.047268 | 0.2542
40 |  | EV | -10.0097 | -69.3455 |  |  |  | 0.19148 | 3.0174
41 |  | OSBS2-deep-2 | -2.2232 | -10.8042 |  |  |  | 0.18546 | 2.9779
42 |  | OSBS2-3040-1 | -1.224 | -4.0286 |  |  |  | 0.047801 | 0.48115
43 |  | OSBS2-3040-2 | -2.0648 | -6.3169 |  |  |  | 0.029738 | 0.070328
44 |  | EV | -10.0105 | -69.2493 |  |  |  | 0.1782 | 3.1187

## Detail
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
1 |  | EV | -10.178 | -76.906 | -11.3532 | -84.6844 | -1.877 | 298.872 | 2022/08/22 14:53:57
1 |  | EV | -9.394 | -70.711 | -10.5456 | -77.9301 | -1.277 | 98.487 | 2022/08/22 15:02:53
1 |  | EV | -9.175 | -68.643 | -10.32 | -75.6754 | -1.297 | 69.053 | 2022/08/22 15:11:50
1 |  | EV | -9.137 | -68.221 | -10.2809 | -75.2153 | -0.993 | 65.558 | 2022/08/22 15:20:47
1 |  | EV | -9.117 | -68.061 | -10.2603 | -75.0409 | -1.103 | 64.011 | 2022/08/22 15:29:44
1 |  | EV | -9.099 | -67.621 | -10.2417 | -74.5611 | -1.055 | 63.094 | 2022/08/22 15:38:40
1 |  | EV | -9.135 | -67.659 | -10.2788 | -74.6026 | -1.107 | 62.11 | 2022/08/22 15:47:37
1 |  | EV | -9.142 | -67.757 | -10.286 | -74.7094 | -1.071 | 61.795 | 2022/08/22 15:56:33
1 |  | EV | -9.108 | -67.325 | -10.251 | -74.2384 | -1.073 | 62.056 | 2022/08/22 16:05:29
1 |  | EV | -9.113 | -67.465 | -10.2562 | -74.391 | -1.01 | 61.711 | 2022/08/22 16:14:26
Sample 1 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -10.4074 | -76.1049 |  | 29.3974 |  | 0.33394 | 2.9334 |  | 0.5079
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
2 |  | PZ | 2.202 | 8.397 | 1.3987 | 8.3202 | -1.154 | 61.899 | 2022/08/22 16:24:16
2 |  | PZ | 2.573 | 15.173 | 1.7809 | 15.708 | -1.029 | 61.697 | 2022/08/22 16:33:12
2 |  | PZ | 2.675 | 16.9 | 1.8859 | 17.5909 | -1.169 | 62.536 | 2022/08/22 16:42:09
2 |  | PZ | 2.758 | 17.787 | 1.9714 | 18.558 | -1.341 | 62.587 | 2022/08/22 16:51:07
2 |  | PZ | 2.763 | 18.49 | 1.9766 | 19.3245 | -1.149 | 62.348 | 2022/08/22 17:00:03
2 |  | PZ | 2.82 | 18.808 | 2.0353 | 19.6712 | -1.125 | 62.22 | 2022/08/22 17:08:59
2 |  | PZ | 2.831 | 19.204 | 2.0466 | 20.103 | -1.165 | 61.531 | 2022/08/22 17:17:56
2 |  | PZ | 2.807 | 19.363 | 2.0219 | 20.2763 | -0.949 | 61.349 | 2022/08/22 17:26:52
2 |  | PZ | 2.877 | 19.728 | 2.094 | 20.6743 | -0.96 | 61.23 | 2022/08/22 17:35:50
2 |  | PZ | 2.872 | 19.819 | 2.0888 | 20.7735 | -1.123 | 61.391 | 2022/08/22 17:44:47
Sample 2 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.20354 | 3.4633
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
3 |  | UT2 | -13.758 | -94.826 | -15.0407 | -104.2224 | -0.909 | 60.552 | 2022/08/22 17:54:36
3 |  | UT2 | -14.383 | -105.397 | -15.6845 | -115.7478 | -1.034 | 61.414 | 2022/08/22 18:03:33
3 |  | UT2 | -14.504 | -107.942 | -15.8091 | -118.5226 | -0.773 | 60.689 | 2022/08/22 18:12:30
3 |  | UT2 | -14.623 | -109.355 | -15.9317 | -120.0631 | -1.036 | 60.509 | 2022/08/22 18:21:27
3 |  | UT2 | -14.595 | -110.043 | -15.9028 | -120.8133 | -0.888 | 60.821 | 2022/08/22 18:30:24
3 |  | UT2 | -14.656 | -110.713 | -15.9657 | -121.5437 | -0.848 | 60.663 | 2022/08/22 18:39:20
3 |  | UT2 | -14.68 | -110.903 | -15.9904 | -121.7509 | -0.843 | 60.777 | 2022/08/22 18:48:17
3 |  | UT2 | -14.679 | -111.333 | -15.9894 | -122.2197 | -0.762 | 60.235 | 2022/08/22 18:57:14
3 |  | UT2 | -14.721 | -111.644 | -16.0326 | -122.5588 | -0.722 | 60.411 | 2022/08/22 19:06:11
3 |  | UT2 | -14.741 | -111.643 | -16.0532 | -122.5577 | -1 | 60.688 | 2022/08/22 19:15:08
Sample 3 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.29299 | 5.1462
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
4 |  | EV | -9.463 | -72.588 | -10.6167 | -79.9766 | -0.901 | 61.383 | 2022/08/22 19:24:57
4 |  | EV | -9.218 | -69.083 | -10.3643 | -76.1551 | -1.02 | 60.712 | 2022/08/22 19:33:55
4 |  | EV | -9.203 | -68.329 | -10.3489 | -75.3331 | -1.259 | 60.896 | 2022/08/22 19:42:51
4 |  | EV | -9.149 | -67.776 | -10.2932 | -74.7301 | -0.891 | 60.614 | 2022/08/22 19:51:48
4 |  | EV | -9.176 | -67.679 | -10.3211 | -74.6244 | -0.813 | 60.67 | 2022/08/22 20:00:45
4 |  | EV | -9.185 | -67.517 | -10.3303 | -74.4477 | -0.922 | 60.963 | 2022/08/22 20:09:43
4 |  | EV | -9.12 | -67.375 | -10.2634 | -74.2929 | -0.859 | 60.807 | 2022/08/22 20:18:40
4 |  | EV | -9.146 | -67.526 | -10.2902 | -74.4576 | -0.892 | 60.976 | 2022/08/22 20:27:36
4 |  | EV | -9.123 | -67.242 | -10.2665 | -74.1479 | -1.149 | 60.983 | 2022/08/22 20:36:33
4 |  | EV | -9.11 | -67.294 | -10.2531 | -74.2046 | -1.005 | 61.107 | 2022/08/22 20:45:31
Sample 4 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
4 |  | EV | -10.3348 | -75.237 |  |  |  | 0.1028 | 1.627
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
5 |  | OSBS2-D-2 | -1.501 | -24.141 | -2.4155 | -27.1555 | -1.29 | 224.102 | 2022/08/22 20:55:20
5 |  | OSBS2-D-2 | -1.329 | -20.517 | -2.2384 | -23.2043 | -1.467 | 220.402 | 2022/08/22 21:04:18
5 |  | OSBS2-D-2 | -1.288 | -19.645 | -2.1961 | -22.2536 | -1.252 | 219.69 | 2022/08/22 21:13:14
5 |  | OSBS2-D-2 | -1.188 | -19.112 | -2.0931 | -21.6724 | -1.441 | 220.483 | 2022/08/22 21:22:12
Sample 5 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
5 |  | OSBS2-D-2 | -2.2358 | -23.5714 |  | 159.8919 |  | 0.13054 | 2.2667 |  | -3.0832
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
6 |  | OSBS2-C-3 | 0.481 | -8.309 | -0.37399 | -9.8941 | -1.073 | 62.803 | 2022/08/22 21:32:02
6 |  | OSBS2-C-3 | 0.51 | -7.675 | -0.34412 | -9.2028 | -1.171 | 61.778 | 2022/08/22 21:40:59
6 |  | OSBS2-C-3 | 0.51 | -7.193 | -0.34412 | -8.6773 | -0.953 | 61.095 | 2022/08/22 21:49:56
6 |  | OSBS2-C-3 | 0.565 | -6.505 | -0.28746 | -7.9272 | -0.919 | 61.216 | 2022/08/22 21:58:52
Sample 6 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
6 |  | OSBS2-C-3 | -0.33742 | -8.9254 |  |  |  | 0.035105 | 0.76247
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
7 |  | OSBS2-A-4 | -0.197 | -12.022 | -1.0724 | -13.9423 | -1.734 | 89.376 | 2022/08/22 22:08:42
7 |  | OSBS2-A-4 | -0.273 | -12.911 | -1.1506 | -14.9116 | -1.558 | 88.042 | 2022/08/22 22:17:39
7 |  | OSBS2-A-4 | -0.279 | -12.856 | -1.1568 | -14.8516 | -1.539 | 88.38 | 2022/08/22 22:26:35
7 |  | OSBS2-A-4 | -0.258 | -12.661 | -1.1352 | -14.639 | -1.587 | 88.643 | 2022/08/22 22:35:33
Sample 7 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
7 |  | OSBS2-A-4 | -1.1287 | -14.5861 | -1.6045 | 27.3329 |  | 0.037553 | 0.40802 |  | 1.3118
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
8 |  | OSBS2-A-2 | -0.287 | -15.399 | -1.1651 | -17.6242 | -1.816 | 164.858 | 2022/08/22 22:45:20
8 |  | OSBS2-A-2 | -0.231 | -15.417 | -1.1074 | -17.6438 | -1.717 | 165.182 | 2022/08/22 22:54:17
8 |  | OSBS2-A-2 | -0.265 | -15.45 | -1.1424 | -17.6798 | -1.794 | 165.497 | 2022/08/22 23:03:14
8 |  | OSBS2-A-2 | -0.241 | -15.495 | -1.1177 | -17.7289 | -1.81 | 165.482 | 2022/08/22 23:12:11
Sample 8 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
8 |  | OSBS2-A-2 | -1.1331 | -17.6692 | -1.7843 | 103.9774 |  | 0.025113 | 0.042169 |  | -2.0099
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
9 |  | OSBS2-A-1 | -1.507 | -22.072 | -2.4217 | -24.8997 | -1.627 | 171.216 | 2022/08/22 23:21:59
9 |  | OSBS2-A-1 | -1.578 | -22.982 | -2.4948 | -25.8918 | -1.392 | 169.912 | 2022/08/22 23:30:56
9 |  | OSBS2-A-1 | -1.561 | -23.338 | -2.4773 | -26.28 | -1.67 | 169.841 | 2022/08/22 23:39:53
9 |  | OSBS2-A-1 | -1.604 | -23.366 | -2.5216 | -26.3105 | -1.434 | 168.298 | 2022/08/22 23:48:49
Sample 9 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
9 |  | OSBS2-A-1 | -2.4789 | -25.8455 | -1.5308 | 108.5394 |  | 0.041008 | 0.60417 |  | -0.43465
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
10 |  | OSBS2-D-3 | -1.914 | -22.621 | -2.8409 | -25.4983 | -1.277 | 287.096 | 2022/08/22 23:58:37
10 |  | OSBS2-D-3 | -1.963 | -22.53 | -2.8914 | -25.399 | -1.366 | 287.61 | 2022/08/23 00:07:34
10 |  | OSBS2-D-3 | -1.927 | -22.21 | -2.8543 | -25.0501 | -1.127 | 287.894 | 2022/08/23 00:16:30
10 |  | OSBS2-D-3 | -1.959 | -22.396 | -2.8873 | -25.2529 | -1.165 | 289.744 | 2022/08/23 00:25:25
Sample 10 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
10 |  | OSBS2-D-3 | -2.8685 | -25.3001 |  | 226.8087 |  | 0.024033 | 0.1786 |  | -5.5272
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
11 |  | OSBS2-E-1 | -2.39 | -29.411 | -3.3312 | -32.9013 | -2.028 | 504.628 | 2022/08/23 00:35:14
11 |  | OSBS2-E-1 | -2.502 | -29.924 | -3.4466 | -33.4606 | -2.05 | 495.266 | 2022/08/23 00:44:10
11 |  | OSBS2-E-1 | -2.483 | -30.056 | -3.427 | -33.6045 | -2.048 | 497.137 | 2022/08/23 00:53:06
11 |  | OSBS2-E-1 | -2.515 | -30.155 | -3.46 | -33.7125 | -2.165 | 497.836 | 2022/08/23 01:02:02
Sample 11 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
11 |  | OSBS2-E-1 | -3.4162 | -33.4197 | -2.0728 | 437.4394 |  | 0.056548 | 0.33082 |  | -7.5157
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
12 |  | OSBS2-C-1 | 1.695 | -5.64 | 0.87648 | -6.9841 | -1.31 | 66.127 | 2022/08/23 01:11:51
12 |  | OSBS2-C-1 | 1.905 | -3.408 | 1.0928 | -4.5506 | -1.14 | 63.555 | 2022/08/23 01:20:48
12 |  | OSBS2-C-1 | 1.939 | -2.909 | 1.1278 | -4.0065 | -1.008 | 62.39 | 2022/08/23 01:29:45
12 |  | OSBS2-C-1 | 1.93 | -2.675 | 1.1185 | -3.7514 | -0.963 | 61.938 | 2022/08/23 01:38:41
Sample 12 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
12 |  | OSBS2-C-1 | 1.0539 | -4.8232 |  |  |  | 0.11573 | 1.3562
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
13 |  | EV | -8.48 | -58.554 | -9.6042 | -64.6755 | -0.81 | 62.486 | 2022/08/23 01:48:30
13 |  | EV | -8.875 | -63.55 | -10.011 | -70.1226 | -0.99 | 61.786 | 2022/08/23 01:57:27
13 |  | EV | -8.922 | -64.861 | -10.0594 | -71.5519 | -0.931 | 61.753 | 2022/08/23 02:06:24
13 |  | EV | -9.028 | -65.426 | -10.1686 | -72.168 | -0.781 | 61.669 | 2022/08/23 02:15:20
Sample 13 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
13 |  | EV | -9.9608 | -69.6295 |  |  |  | 0.23954 | 3.1294
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
14 |  | OSBS2-C-2 | 0.022 | -11.737 | -0.84677 | -13.6316 | -1.62 | 233.624 | 2022/08/23 02:25:10
14 |  | OSBS2-C-2 | 0.288 | -7.239 | -0.57278 | -8.7275 | -1.608 | 235.147 | 2022/08/23 02:34:06
14 |  | OSBS2-C-2 | 0.388 | -5.886 | -0.46978 | -7.2523 | -1.503 | 231.882 | 2022/08/23 02:43:03
14 |  | OSBS2-C-2 | 0.476 | -5.027 | -0.37914 | -6.3158 | -1.373 | 233.718 | 2022/08/23 02:51:59
Sample 14 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
14 |  | OSBS2-C-2 | -0.56712 | -8.9818 | -1.526 | 172.3154 |  | 0.19662 | 2.9854 |  | -3.2099
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
15 |  | OSBS2-C-4 | -0.197 | -8.665 | -1.0724 | -10.2822 | -1.197 | 64.312 | 2022/08/23 03:01:47
15 |  | OSBS2-C-4 | -0.206 | -8.876 | -1.0816 | -10.5123 | -1.285 | 63.113 | 2022/08/23 03:10:43
15 |  | OSBS2-C-4 | -0.214 | -8.922 | -1.0899 | -10.5624 | -1.105 | 62.411 | 2022/08/23 03:19:41
15 |  | OSBS2-C-4 | -0.151 | -8.929 | -1.025 | -10.5701 | -1.113 | 62.182 | 2022/08/23 03:28:37
Sample 15 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
15 |  | OSBS2-C-4 | -1.0672 | -10.4817 |  |  |  | 0.028202 | 0.12424
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
16 |  | OSBS2-C-5 | 2.912 | -9.563 | 2.13 | -11.2613 | -0.913 | 61.691 | 2022/08/23 03:38:26
16 |  | OSBS2-C-5 | 2.999 | -9.316 | 2.2196 | -10.992 | -1.016 | 61.559 | 2022/08/23 03:47:22
16 |  | OSBS2-C-5 | 3.075 | -9.358 | 2.2979 | -11.0378 | -1.078 | 61.608 | 2022/08/23 03:56:19
16 |  | OSBS2-C-5 | 3.056 | -9.299 | 2.2784 | -10.9735 | -1.181 | 61.901 | 2022/08/23 04:05:15
Sample 16 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
16 |  | OSBS2-C-5 | 2.2315 | -11.0661 |  |  |  | 0.073178 | 0.12188
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
17 |  | OSBS2-B-1 | -0.1 | -21.889 | -0.97244 | -24.7002 | -2.056 | 327.086 | 2022/08/23 04:15:03
17 |  | OSBS2-B-1 | -0.252 | -23.393 | -1.129 | -26.34 | -2.109 | 339.206 | 2022/08/23 04:24:00
17 |  | OSBS2-B-1 | -0.255 | -23.831 | -1.1321 | -26.8175 | -2.063 | 335.163 | 2022/08/23 04:32:56
17 |  | OSBS2-B-1 | -0.31 | -24.089 | -1.1887 | -27.0988 | -1.876 | 339.613 | 2022/08/23 04:41:53
Sample 17 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
17 |  | OSBS2-B-1 | -1.1056 | -26.2391 | -2.026 | 273.9896 |  | 0.090197 | 0.98388 |  | -5.1122
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
18 |  | OSBS2-B-2 | -1.002 | -20.816 | -1.9015 | -23.5303 | -1.423 | 120.237 | 2022/08/23 04:51:40
18 |  | OSBS2-B-2 | -1.001 | -20.42 | -1.9005 | -23.0985 | -1.48 | 118.627 | 2022/08/23 05:00:37
18 |  | OSBS2-B-2 | -1.042 | -20.108 | -1.9427 | -22.7584 | -1.335 | 117.805 | 2022/08/23 05:09:34
18 |  | OSBS2-B-2 | -1.015 | -20.357 | -1.9149 | -23.0298 | -1.453 | 117.762 | 2022/08/23 05:18:31
Sample 18 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
18 |  | OSBS2-B-2 | -1.9149 | -23.1043 |  | 57.3304 |  | 0.019096 | 0.29327 |  | -1.1982
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
19 |  | OSBS2-B-3 | -0.89 | -19.84 | -1.7862 | -22.4662 | -1.622 | 72.299 | 2022/08/23 05:28:20
19 |  | OSBS2-B-3 | -0.92 | -20.085 | -1.8171 | -22.7333 | -1.35 | 71.848 | 2022/08/23 05:37:17
19 |  | OSBS2-B-3 | -0.903 | -19.898 | -1.7996 | -22.5294 | -1.381 | 71.832 | 2022/08/23 05:46:13
19 |  | OSBS2-B-3 | -0.912 | -20.067 | -1.8088 | -22.7137 | -1.36 | 71.478 | 2022/08/23 05:55:09
Sample 19 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
19 |  | OSBS2-B-3 | -1.8029 | -22.6106 |  | 10.5869 |  | 0.012868 | 0.12206
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
20 |  | OSBS2-B-4 | -0.872 | -22 | -1.7676 | -24.8212 | -1.617 | 100.29 | 2022/08/23 06:04:58
20 |  | OSBS2-B-4 | -0.919 | -22.397 | -1.816 | -25.254 | -1.435 | 99.843 | 2022/08/23 06:13:54
20 |  | OSBS2-B-4 | -0.864 | -22.381 | -1.7594 | -25.2366 | -1.453 | 101.205 | 2022/08/23 06:22:51
20 |  | OSBS2-B-4 | -0.926 | -22.617 | -1.8233 | -25.4939 | -1.322 | 100.994 | 2022/08/23 06:31:48
Sample 20 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
20 |  | OSBS2-B-4 | -1.7916 | -25.2014 |  | 39.3056 |  | 0.031763 | 0.25622 |  | -0.7824
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
21 |  | OSBS2-B-5 | -0.204 | -14.371 | -1.0796 | -16.5034 | -1.998 | 609.611 | 2022/08/23 06:41:37
21 |  | OSBS2-B-5 | -0.145 | -13.5 | -1.0188 | -15.5538 | -1.867 | 602.014 | 2022/08/23 06:50:33
21 |  | OSBS2-B-5 | -0.151 | -13.255 | -1.025 | -15.2866 | -2.241 | 595.525 | 2022/08/23 06:59:30
21 |  | OSBS2-B-5 | -0.17 | -13.237 | -1.0445 | -15.267 | -2.056 | 596.386 | 2022/08/23 07:08:26
Sample 21 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
21 |  | OSBS2-B-5 | -1.042 | -15.6527 | -2.0405 | 539.6067 |  | 0.026564 | 0.53382 |  | -9.3907
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
22 |  | EV | -8.622 | -59.422 | -9.7504 | -65.6219 | -1.103 | 67.185 | 2022/08/23 07:18:12
22 |  | EV | -8.875 | -63.799 | -10.011 | -70.3941 | -0.881 | 63.765 | 2022/08/23 07:27:07
22 |  | EV | -8.98 | -64.874 | -10.1192 | -71.5661 | -0.816 | 62.314 | 2022/08/23 07:36:03
22 |  | EV | -9.002 | -65.328 | -10.1418 | -72.0611 | -1.017 | 61.81 | 2022/08/23 07:44:58
Sample 22 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
22 |  | EV | -10.0056 | -69.9108 |  | 2.4911 |  | 0.17422 | 2.6997
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
23 |  | OSBS2-A-3 | -1.779 | -22.247 | -2.7019 | -25.0905 | -1.354 | 174.488 | 2022/08/23 07:54:44
23 |  | OSBS2-A-3 | -1.51 | -18.58 | -2.4248 | -21.0924 | -1.22 | 124.366 | 2022/08/23 08:03:41
23 |  | OSBS2-A-3 | -1.426 | -17.764 | -2.3383 | -20.2027 | -1.259 | 135.56 | 2022/08/23 08:12:38
23 |  | OSBS2-A-3 | -1.424 | -17.367 | -2.3362 | -19.7699 | -1.138 | 126.627 | 2022/08/23 08:21:34
Sample 23 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
23 |  | OSBS2-A-3 | -2.4503 | -21.5389 |  | 78.9829 |  | 0.16769 | 2.2296 |  | -9.8222
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
24 |  | OSBS2-A-5 | -2.275 | -25.754 | -3.2128 | -28.9141 | -1.359 | 122.691 | 2022/08/23 08:31:21
24 |  | OSBS2-A-5 | -2.326 | -26.673 | -3.2653 | -29.9161 | -1.39 | 121.417 | 2022/08/23 08:40:17
24 |  | OSBS2-A-5 | -2.35 | -26.66 | -3.29 | -29.9019 | -1.419 | 122.148 | 2022/08/23 08:49:14
24 |  | OSBS2-A-5 | -2.302 | -26.707 | -3.2406 | -29.9532 | -1.781 | 122.378 | 2022/08/23 08:58:10
Sample 24 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
24 |  | OSBS2-A-5 | -3.2522 | -29.6713 |  | 60.8811 |  | 0.03216 | 0.46342 |  | -3.3482
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
25 |  | OSBS2-E-2 | 1.02 | -15.089 | 0.1812 | -17.2862 | -2.094 | 78.161 | 2022/08/23 09:07:58
25 |  | OSBS2-E-2 | 1.169 | -14.069 | 0.33468 | -16.1741 | -1.332 | 63.085 | 2022/08/23 09:16:53
25 |  | OSBS2-E-2 | 1.168 | -13.77 | 0.33365 | -15.8481 | -1.546 | 63.652 | 2022/08/23 09:25:49
25 |  | OSBS2-E-2 | 1.171 | -13.825 | 0.33674 | -15.9081 | -1.518 | 64.423 | 2022/08/23 09:34:45
Sample 25 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
25 |  | OSBS2-E-2 | 0.29657 | -16.3041 | -1.6225 | 6.0529 |  | 0.074677 | 0.6144
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
26 |  | OSBS2-E-3 | 1.112 | -10.958 | 0.27597 | -12.7822 | -1.608 | 64.274 | 2022/08/23 09:44:33
26 |  | OSBS2-E-3 | 1.139 | -10.385 | 0.30378 | -12.1575 | -1.699 | 65.192 | 2022/08/23 09:53:29
26 |  | OSBS2-E-3 | 1.142 | -10.561 | 0.30687 | -12.3494 | -1.766 | 65.151 | 2022/08/23 10:02:25
26 |  | OSBS2-E-3 | 1.154 | -10.108 | 0.31923 | -11.8555 | -1.491 | 65.205 | 2022/08/23 10:11:22
Sample 26 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
26 |  | OSBS2-E-3 | 0.30146 | -12.2862 | -1.641 | 3.6782 |  | 0.017727 | 0.35606
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
27 |  | OSBS2-E-4 | 0 | -14.443 | -0.86944 | -16.5819 | -1.568 | 67.842 | 2022/08/23 10:21:10
27 |  | OSBS2-E-4 | -0.052 | -14.696 | -0.923 | -16.8577 | -1.496 | 68.481 | 2022/08/23 10:30:06
27 |  | OSBS2-E-4 | -0.077 | -14.793 | -0.94875 | -16.9635 | -1.857 | 68.608 | 2022/08/23 10:39:02
27 |  | OSBS2-E-4 | -0.105 | -14.698 | -0.97759 | -16.8599 | -1.776 | 68.313 | 2022/08/23 10:47:58
Sample 27 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
27 |  | OSBS2-E-4 | -0.92969 | -16.8158 | -1.6743 | 7.0336 |  | 0.044606 | 0.14999
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
28 |  | OSBS2-E-5 | 0.529 | -13.595 | -0.32455 | -15.6573 | -2.197 | 71.776 | 2022/08/23 10:57:46
28 |  | OSBS2-E-5 | 0.479 | -14.383 | -0.37605 | -16.5165 | -2.247 | 70.493 | 2022/08/23 11:06:42
28 |  | OSBS2-E-5 | 0.498 | -14.088 | -0.35648 | -16.1948 | -2.209 | 71.902 | 2022/08/23 11:15:38
28 |  | OSBS2-E-5 | 0.47 | -13.789 | -0.38532 | -15.8688 | -2.402 | 71.494 | 2022/08/23 11:24:35
Sample 28 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
28 |  | OSBS2-E-5 | -0.3606 | -16.0594 | -2.2637 | 10.1389 | 13.23 | 0.02609 | 0.34531 |  | 1.9963
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
29 |  | OSBS2-010-1 | 0.708 | -4.015 | -0.14017 | -5.2124 | -1.265 | 61.603 | 2022/08/23 11:34:18
29 |  | OSBS2-010-1 | 0.756 | -2.771 | -0.090727 | -3.8561 | -1.235 | 61.133 | 2022/08/23 11:43:15
29 |  | OSBS2-010-1 | 0.764 | -2.422 | -0.082487 | -3.4756 | -0.915 | 61.013 | 2022/08/23 11:52:11
29 |  | OSBS2-010-1 | 0.758 | -2.123 | -0.088667 | -3.1496 | -1.12 | 60.993 | 2022/08/23 12:01:06
Sample 29 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
29 |  | OSBS2-010-1 | -0.10051 | -3.9234 |  |  |  | 0.025891 | 0.83146
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
30 |  | OSBS2-010-2 | -0.087 | -3.269 | -0.95905 | -4.399 | -0.958 | 60.407 | 2022/08/23 12:10:52
30 |  | OSBS2-010-2 | -0.102 | -3.172 | -0.9745 | -4.2933 | -0.966 | 59.861 | 2022/08/23 12:19:48
30 |  | OSBS2-010-2 | -0.075 | -3.204 | -0.94669 | -4.3282 | -0.928 | 60.039 | 2022/08/23 12:28:43
30 |  | OSBS2-010-2 | -0.047 | -3.034 | -0.91785 | -4.1428 | -1.027 | 60.115 | 2022/08/23 12:37:38
Sample 30 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
30 |  | OSBS2-010-2 | -0.94952 | -4.2908 |  |  |  | 0.023286 | 0.09909
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
31 |  | EV | -8.59 | -58.34 | -9.7175 | -64.4422 | -1.039 | 61.611 | 2022/08/23 12:47:24
31 |  | EV | -8.856 | -62.942 | -9.9914 | -69.4597 | -0.896 | 60.694 | 2022/08/23 12:56:20
31 |  | EV | -8.943 | -64.408 | -10.0811 | -71.058 | -0.846 | 60.483 | 2022/08/23 13:05:16
31 |  | EV | -8.981 | -65.093 | -10.1202 | -71.8049 | -0.915 | 60.461 | 2022/08/23 13:14:12
Sample 31 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
31 |  | EV | -9.9775 | -69.1912 |  |  |  | 0.17628 | 3.0393
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
32 |  | OSBS2-D-1 | -1.445 | -21.482 | -2.3578 | -24.2564 | -1.4 | 209.896 | 2022/08/23 13:23:58
32 |  | OSBS2-D-1 | -1.05 | -17.276 | -1.951 | -19.6707 | -1.352 | 210.92 | 2022/08/23 13:32:54
32 |  | OSBS2-D-1 | -1.046 | -16.588 | -1.9469 | -18.9206 | -1.448 | 210.965 | 2022/08/23 13:41:50
32 |  | OSBS2-D-1 | -1.019 | -16.008 | -1.919 | -18.2882 | -1.538 | 210.972 | 2022/08/23 13:50:45
Sample 32 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
32 |  | OSBS2-D-1 | -2.0437 | -20.284 |  | 149.4109 |  | 0.2038 | 2.4837 |  | -2.4304
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
33 |  | OSBS2-D-4 | -0.921 | -14.673 | -1.8181 | -16.8327 | -0.95 | 110.681 | 2022/08/23 14:00:32
33 |  | OSBS2-D-4 | -0.901 | -14.061 | -1.7975 | -16.1654 | -0.914 | 108.674 | 2022/08/23 14:09:28
33 |  | OSBS2-D-4 | -0.89 | -14.037 | -1.7862 | -16.1392 | -1.022 | 119.127 | 2022/08/23 14:18:24
33 |  | OSBS2-D-4 | -0.949 | -13.937 | -1.8469 | -16.0302 | -1.124 | 118.373 | 2022/08/23 14:27:20
Sample 33 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
33 |  | OSBS2-D-4 | -1.8122 | -16.2919 |  | 52.9364 |  | 0.025902 | 0.335 |  | -13.8797
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
34 |  | OSBS2-D-5 | -0.733 | -16.118 | -1.6245 | -18.4081 | -1.682 | 171.917 | 2022/08/23 14:37:08
34 |  | OSBS2-D-5 | -0.724 | -16.273 | -1.6152 | -18.5771 | -1.501 | 170.964 | 2022/08/23 14:46:03
34 |  | OSBS2-D-5 | -0.752 | -16.37 | -1.644 | -18.6829 | -1.421 | 171.838 | 2022/08/23 14:55:00
34 |  | OSBS2-D-5 | -0.768 | -16.231 | -1.6605 | -18.5313 | -1.598 | 170.862 | 2022/08/23 15:03:56
Sample 34 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
34 |  | OSBS2-D-5 | -1.636 | -18.5499 | -1.5505 | 110.1179 |  | 0.01967 | 0.1044 |  | -2.8357
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
35 |  | OSBS2-2030-1 | -0.626 | -3.913 | -1.5142 | -5.1012 | -1.071 | 62.323 | 2022/08/23 15:13:44
35 |  | OSBS2-2030-1 | -0.591 | -2.511 | -1.4782 | -3.5726 | -1.139 | 61.371 | 2022/08/23 15:22:40
35 |  | OSBS2-2030-1 | -0.62 | -2.33 | -1.5081 | -3.3753 | -0.919 | 61.404 | 2022/08/23 15:31:36
35 |  | OSBS2-2030-1 | -0.593 | -2.372 | -1.4802 | -3.4211 | -0.935 | 60.823 | 2022/08/23 15:40:33
Sample 35 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
35 |  | OSBS2-2030-1 | -1.4952 | -3.8675 |  |  |  | 0.018083 | 0.75829
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
36 |  | OSBS2-2030-2 | -0.226 | -4.314 | -1.1022 | -5.5384 | -0.908 | 60.359 | 2022/08/23 15:50:16
36 |  | OSBS2-2030-2 | -0.202 | -4.485 | -1.0775 | -5.7248 | -1.153 | 61.005 | 2022/08/23 15:59:12
36 |  | OSBS2-2030-2 | -0.167 | -4.372 | -1.0415 | -5.6016 | -0.957 | 60.514 | 2022/08/23 16:08:09
36 |  | OSBS2-2030-2 | -0.162 | -4.273 | -1.0363 | -5.4937 | -0.933 | 60.935 | 2022/08/23 16:17:05
Sample 36 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
36 |  | OSBS2-2030-2 | -1.0644 | -5.5896 |  |  |  | 0.030281 | 0.092105
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
37 |  | OSBS2-1020-1 | -0.543 | -2.639 | -1.4287 | -3.7122 | -0.915 | 60.268 | 2022/08/23 16:26:50
37 |  | OSBS2-1020-1 | -0.515 | -2.566 | -1.3999 | -3.6326 | -1.028 | 60.775 | 2022/08/23 16:35:45
37 |  | OSBS2-1020-1 | -0.522 | -2.788 | -1.4071 | -3.8746 | -1.084 | 60.639 | 2022/08/23 16:44:41
37 |  | OSBS2-1020-1 | -0.554 | -2.842 | -1.4401 | -3.9335 | -1.152 | 61.274 | 2022/08/23 16:53:37
Sample 37 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
37 |  | OSBS2-1020-1 | -1.419 | -3.7882 |  |  |  | 0.01812 | 0.12816
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
38 |  | OSBS2-1020-2 | -0.636 | -3.473 | -1.5245 | -4.6215 | -1.102 | 60.019 | 2022/08/23 17:03:21
38 |  | OSBS2-1020-2 | -0.691 | -3.646 | -1.5812 | -4.8101 | -1.152 | 60.049 | 2022/08/23 17:12:17
38 |  | OSBS2-1020-2 | -0.646 | -3.667 | -1.5348 | -4.833 | -1.093 | 60.885 | 2022/08/23 17:21:13
38 |  | OSBS2-1020-2 | -0.65 | -3.708 | -1.539 | -4.8777 | -1.092 | 60.337 | 2022/08/23 17:30:09
Sample 38 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
38 |  | OSBS2-1020-2 | -1.5449 | -4.7855 |  |  |  | 0.024226 | 0.10358
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
39 |  | OSBS2-deep-1 | -0.994 | -7.859 | -1.8933 | -9.4035 | -1.057 | 59.685 | 2022/08/23 17:39:54
39 |  | OSBS2-deep-1 | -1.039 | -8.289 | -1.9396 | -9.8723 | -1.065 | 59.826 | 2022/08/23 17:48:50
39 |  | OSBS2-deep-1 | -1.078 | -8.203 | -1.9798 | -9.7785 | -1.052 | 59.736 | 2022/08/23 17:57:46
39 |  | OSBS2-deep-1 | -1.102 | -8.464 | -2.0045 | -10.0631 | -0.994 | 59.85 | 2022/08/23 18:06:42
Sample 39 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
39 |  | OSBS2-deep-1 | -1.9543 | -9.7793 |  |  |  | 0.047268 | 0.2542
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
40 |  | EV | -8.604 | -58.517 | -9.7319 | -64.6352 | -1.009 | 60.637 | 2022/08/23 18:16:26
40 |  | EV | -8.871 | -63.039 | -10.0069 | -69.5654 | -0.962 | 60.223 | 2022/08/23 18:25:21
40 |  | EV | -9.003 | -64.643 | -10.1429 | -71.3143 | -0.888 | 60.123 | 2022/08/23 18:34:18
40 |  | EV | -9.017 | -65.15 | -10.1573 | -71.867 | -0.958 | 60.392 | 2022/08/23 18:43:14
Sample 40 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
40 |  | EV | -10.0097 | -69.3455 |  |  |  | 0.19148 | 3.0174
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
41 |  | OSBS2-deep-2 | -1.579 | -13.443 | -2.4959 | -15.4916 | -1.103 | 60.281 | 2022/08/23 18:52:58
41 |  | OSBS2-deep-2 | -1.285 | -8.82 | -2.193 | -10.4512 | -1.104 | 60.441 | 2022/08/23 19:01:54
41 |  | OSBS2-deep-2 | -1.244 | -7.393 | -2.1508 | -8.8954 | -1.077 | 60.24 | 2022/08/23 19:10:51
41 |  | OSBS2-deep-2 | -1.149 | -6.919 | -2.0529 | -8.3786 | -0.979 | 59.771 | 2022/08/23 19:19:47
Sample 41 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
41 |  | OSBS2-deep-2 | -2.2232 | -10.8042 |  |  |  | 0.18546 | 2.9779
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
42 |  | OSBS2-3040-1 | -0.405 | -3.541 | -1.2866 | -4.6956 | -1.061 | 60.989 | 2022/08/23 19:29:33
42 |  | OSBS2-3040-1 | -0.317 | -2.886 | -1.196 | -3.9815 | -1.117 | 60.683 | 2022/08/23 19:38:28
42 |  | OSBS2-3040-1 | -0.358 | -2.925 | -1.2382 | -4.024 | -1.121 | 60.476 | 2022/08/23 19:47:24
42 |  | OSBS2-3040-1 | -0.297 | -2.365 | -1.1754 | -3.4134 | -1.176 | 60.945 | 2022/08/23 19:56:19
Sample 42 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
42 |  | OSBS2-3040-1 | -1.224 | -4.0286 |  |  |  | 0.047801 | 0.48115
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
43 |  | OSBS2-3040-2 | -1.148 | -5.045 | -2.0519 | -6.3354 | -1.176 | 59.873 | 2022/08/23 20:06:03
43 |  | OSBS2-3040-2 | -1.143 | -4.928 | -2.0468 | -6.2078 | -0.988 | 59.883 | 2022/08/23 20:14:58
43 |  | OSBS2-3040-2 | -1.205 | -5.093 | -2.1106 | -6.3877 | -1.014 | 60.276 | 2022/08/23 20:23:53
43 |  | OSBS2-3040-2 | -1.146 | -5.046 | -2.0499 | -6.3365 | -1.043 | 60.09 | 2022/08/23 20:32:49
Sample 43 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
43 |  | OSBS2-3040-2 | -2.0648 | -6.3169 |  |  |  | 0.029738 | 0.070328
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
44 |  | EV | -8.61 | -58.175 | -9.7381 | -64.2623 | -1.028 | 60.753 | 2022/08/23 20:42:34
44 |  | EV | -8.931 | -63.406 | -10.0687 | -69.9656 | -0.96 | 60.997 | 2022/08/23 20:51:29
44 |  | EV | -8.963 | -64.429 | -10.1017 | -71.0809 | -1.065 | 60.925 | 2022/08/23 21:00:25
44 |  | EV | -8.994 | -64.986 | -10.1336 | -71.6882 | -0.86 | 60.797 | 2022/08/23 21:09:21
Sample 44 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
44 |  | EV | -10.0105 | -69.2493 |  |  |  | 0.1782 | 3.1187

## Source
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Line | Analysis | Time Code | Port | Inj Nr | d(18_16)Mean | d(D_H)Mean | H2O_Mean | Ignore | Good | Identifier 1 | Identifier 2 | Gas Configuration | Timestamp Mean | d(18_16)_SD | d(D_H)_SD | H2O_SD | d(18_16)_Sl | d(D_H)_Sl | H2O_Sl | baseline_shift | slope_shift | residuals | baseline_curvature | interval | ch4_ppm | h16od_adjust | h16od_shift | n2_flag | DAS Temp | Tray | Sample | Job | Method | Error Code
1 | P-20886 | 2022/08/22 14:53:57 | 1-01 | 1 | -10.178 | -76.906 | 15286 | -1 | 1 | EV | standard | H2O | 1661201942.29 | 0.232 | 0.943 | 83.158 | 0.002 | 0.007 | 2.025 | 298.872 | 29.243 | 2.555 | 69.456 | 0.858 | -1.877 | 0 | 0 | 1 | 42.562 | 1 | 1 | 1 | GJB_Rinse | 0
2 | P-20886 | 2022/08/22 15:02:53 | 1-01 | 2 | -9.394 | -70.711 | 12716 | -1 | 1 | EV | standard | H2O | 1661202478.93 | 0.207 | 0.748 | 70.702 | 0.002 | 0 | 1.748 | 98.487 | 28.018 | 2.17 | 63.582 | 0.856 | -1.277 | 0 | 0 | 1 | 41.812 | 1 | 1 | 1 | GJB_Rinse | 0
3 | P-20886 | 2022/08/22 15:11:50 | 1-01 | 3 | -9.175 | -68.643 | 12801 | -1 | 1 | EV | standard | H2O | 1661203015.75 | 0.23 | 0.699 | 67.378 | 0.001 | -0.001 | 1.667 | 69.053 | 29.367 | 2.171 | 63.539 | 0.856 | -1.297 | 0 | 0 | 1 | 41.688 | 1 | 1 | 1 | GJB_Rinse | 0
4 | P-20886 | 2022/08/22 15:20:47 | 1-01 | 4 | -9.137 | -68.221 | 12910 | 0 | 1 | EV | standard | H2O | 1661203552.25 | 0.207 | 0.783 | 65.853 | 0.001 | -0.001 | 1.618 | 65.558 | 29.185 | 2.199 | 65.302 | 0.856 | -0.993 | 0 | 0 | 1 | 41.5 | 1 | 1 | 1 | GJB_Rinse | 0
5 | P-20886 | 2022/08/22 15:29:44 | 1-01 | 5 | -9.117 | -68.061 | 12853 | 0 | 1 | EV | standard | H2O | 1661204089.11 | 0.214 | 0.659 | 64.07 | 0 | -0.003 | 1.576 | 64.011 | 29.639 | 2.188 | 64.468 | 0.856 | -1.103 | 0 | 0 | 1 | 41.438 | 1 | 1 | 1 | GJB_Rinse | 0
6 | P-20886 | 2022/08/22 15:38:40 | 1-01 | 6 | -9.099 | -67.621 | 12724 | 0 | 1 | EV | standard | H2O | 1661204625.82 | 0.201 | 0.757 | 62.251 | 0.001 | -0.001 | 1.525 | 63.094 | 29.318 | 2.173 | 64.266 | 0.856 | -1.055 | 0 | 0 | 1 | 41.5 | 1 | 1 | 1 | GJB_Rinse | 0
7 | P-20886 | 2022/08/22 15:47:37 | 1-01 | 7 | -9.135 | -67.659 | 12342 | 0 | 1 | EV | standard | H2O | 1661205162.87 | 0.195 | 0.612 | 66.289 | 0.001 | 0.001 | 1.641 | 62.11 | 28.025 | 2.108 | 62.976 | 0.856 | -1.107 | 0 | 0 | 1 | 41.625 | 1 | 1 | 1 | GJB_Rinse | 0
8 | P-20886 | 2022/08/22 15:56:33 | 1-01 | 8 | -9.142 | -67.757 | 12302 | 0 | 1 | EV | standard | H2O | 1661205699.02 | 0.177 | 0.832 | 61.557 | 0.001 | 0 | 1.52 | 61.795 | 27.739 | 2.079 | 63.111 | 0.856 | -1.071 | 0 | 0 | 1 | 41.688 | 1 | 1 | 1 | GJB_Rinse | 0
9 | P-20886 | 2022/08/22 16:05:29 | 1-01 | 9 | -9.108 | -67.325 | 12405 | 0 | 1 | EV | standard | H2O | 1661206234.99 | 0.202 | 0.814 | 60.718 | 0.001 | 0.003 | 1.499 | 62.056 | 28.064 | 2.114 | 63.163 | 0.856 | -1.073 | 0 | 0 | 1 | 41.625 | 1 | 1 | 1 | GJB_Rinse | 0
10 | P-20886 | 2022/08/22 16:14:26 | 1-01 | 10 | -9.113 | -67.465 | 12293 | 0 | 1 | EV | standard | H2O | 1661206771.7 | 0.188 | 0.737 | 60.605 | 0.001 | 0.003 | 1.495 | 61.711 | 27.56 | 2.089 | 63.33 | 0.856 | -1.01 | 0 | 0 | 1 | 41.75 | 1 | 1 | 1 | GJB_Rinse | 0
11 | P-20887 | 2022/08/22 16:24:16 | 1-02 | 1 | 2.202 | 8.397 | 12625 | -1 | 1 | PZ | standard | H2O | 1661207361.19 | 0.211 | 0.736 | 61.989 | 0 | 0.002 | 1.53 | 61.899 | 28.798 | 2.267 | 66.525 | 0.856 | -1.154 | 0 | 0 | 1 | 41.688 | 1 | 2 | 1 | GJB_Rinse | 0
12 | P-20887 | 2022/08/22 16:33:12 | 1-02 | 2 | 2.573 | 15.173 | 12530 | -1 | 1 | PZ | standard | H2O | 1661207897.93 | 0.211 | 0.768 | 65.072 | 0 | -0.003 | 1.609 | 61.697 | 28.738 | 2.268 | 66.833 | 0.856 | -1.029 | 0 | 0 | 1 | 41.938 | 1 | 2 | 1 | GJB_Rinse | 0
13 | P-20887 | 2022/08/22 16:42:09 | 1-02 | 3 | 2.675 | 16.9 | 12862 | -1 | 1 | PZ | standard | H2O | 1661208434.85 | 0.194 | 0.844 | 62.087 | 0.001 | -0.004 | 1.53 | 62.536 | 29.725 | 2.312 | 66.938 | 0.856 | -1.169 | 0 | 0 | 1 | 42.062 | 1 | 2 | 1 | GJB_Rinse | 0
14 | P-20887 | 2022/08/22 16:51:07 | 1-02 | 4 | 2.758 | 17.787 | 12868 | 0 | 1 | PZ | standard | H2O | 1661208972.59 | 0.204 | 0.812 | 59.591 | 0.001 | -0.005 | 1.451 | 62.587 | 29.862 | 2.304 | 66.266 | 0.857 | -1.341 | 0 | 0 | 1 | 41.75 | 1 | 2 | 1 | GJB_Rinse | 0
15 | P-20887 | 2022/08/22 17:00:03 | 1-02 | 5 | 2.763 | 18.49 | 12843 | 0 | 1 | PZ | standard | H2O | 1661209508.54 | 0.204 | 0.844 | 66.428 | 0.001 | -0.003 | 1.634 | 62.348 | 29.537 | 2.324 | 67.807 | 0.856 | -1.149 | 0 | 0 | 1 | 41.75 | 1 | 2 | 1 | GJB_Rinse | 0
16 | P-20887 | 2022/08/22 17:08:59 | 1-02 | 6 | 2.82 | 18.808 | 12836 | 0 | 1 | PZ | standard | H2O | 1661210044.66 | 0.194 | 0.841 | 61.393 | 0.001 | -0.003 | 1.506 | 62.22 | 29.754 | 2.321 | 66.697 | 0.856 | -1.125 | 0 | 0 | 1 | 41.812 | 1 | 2 | 1 | GJB_Rinse | 0
17 | P-20887 | 2022/08/22 17:17:56 | 1-02 | 7 | 2.831 | 19.204 | 12470 | 0 | 1 | PZ | standard | H2O | 1661210581.67 | 0.188 | 0.913 | 61.827 | 0 | -0.006 | 1.523 | 61.531 | 28.663 | 2.263 | 65.348 | 0.856 | -1.165 | 0 | 0 | 1 | 41.75 | 1 | 2 | 1 | GJB_Rinse | 0
18 | P-20887 | 2022/08/22 17:26:52 | 1-02 | 8 | 2.807 | 19.363 | 12489 | 0 | 1 | PZ | standard | H2O | 1661211118.16 | 0.19 | 0.898 | 62.575 | 0.001 | -0.006 | 1.557 | 61.349 | 28.795 | 2.263 | 66.175 | 0.856 | -0.949 | 0 | 0 | 1 | 41.938 | 1 | 2 | 1 | GJB_Rinse | 0
19 | P-20887 | 2022/08/22 17:35:50 | 1-02 | 9 | 2.877 | 19.728 | 12420 | 0 | 1 | PZ | standard | H2O | 1661211655.64 | 0.194 | 0.903 | 61.542 | 0 | 0 | 1.513 | 61.23 | 28.524 | 2.252 | 65.601 | 0.856 | -0.96 | 0 | 0 | 1 | 42.125 | 1 | 2 | 1 | GJB_Rinse | 0
20 | P-20887 | 2022/08/22 17:44:47 | 1-02 | 10 | 2.872 | 19.819 | 12391 | 0 | 1 | PZ | standard | H2O | 1661212192.52 | 0.212 | 0.884 | 61.51 | 0.001 | -0.006 | 1.516 | 61.391 | 28.074 | 2.242 | 65.68 | 0.856 | -1.123 | 0 | 0 | 1 | 42 | 1 | 2 | 1 | GJB_Rinse | 0
21 | P-20888 | 2022/08/22 17:54:36 | 1-03 | 1 | -13.758 | -94.826 | 12072 | -1 | 1 | UT2 | standard | H2O | 1661212781.35 | 0.213 | 0.817 | 60.735 | 0.002 | -0.004 | 1.522 | 60.552 | 27.355 | 2.035 | 61.332 | 0.856 | -0.909 | 0 | 0 | 1 | 41.875 | 1 | 3 | 1 | GJB_Rinse | 0
22 | P-20888 | 2022/08/22 18:03:33 | 1-03 | 2 | -14.383 | -105.397 | 12459 | -1 | 1 | UT2 | standard | H2O | 1661213319.35 | 0.206 | 0.823 | 62.502 | 0.001 | 0.004 | 1.545 | 61.414 | 28.585 | 2.059 | 61.436 | 0.856 | -1.034 | 0 | 0 | 1 | 42 | 1 | 3 | 1 | GJB_Rinse | 0
23 | P-20888 | 2022/08/22 18:12:30 | 1-03 | 3 | -14.504 | -107.942 | 12126 | -1 | 1 | UT2 | standard | H2O | 1661213855.94 | 0.208 | 0.791 | 62.703 | 0.001 | 0.007 | 1.547 | 60.689 | 27.156 | 2.008 | 61.874 | 0.856 | -0.773 | 0 | 0 | 1 | 41.938 | 1 | 3 | 1 | GJB_Rinse | 0
24 | P-20888 | 2022/08/22 18:21:27 | 1-03 | 4 | -14.623 | -109.355 | 12002 | 0 | 1 | UT2 | standard | H2O | 1661214393.11 | 0.198 | 0.712 | 59.982 | 0.002 | 0.004 | 1.495 | 60.509 | 27.214 | 1.98 | 60.129 | 0.855 | -1.036 | 0 | 0 | 1 | 42.125 | 1 | 3 | 1 | GJB_Rinse | 0
25 | P-20888 | 2022/08/22 18:30:24 | 1-03 | 5 | -14.595 | -110.043 | 12164 | 0 | 1 | UT2 | standard | H2O | 1661214929.57 | 0.201 | 0.881 | 58.071 | 0.001 | 0.006 | 1.445 | 60.821 | 27.472 | 2.005 | 61.451 | 0.856 | -0.888 | 0 | 0 | 1 | 42.062 | 1 | 3 | 1 | GJB_Rinse | 0
26 | P-20888 | 2022/08/22 18:39:20 | 1-03 | 6 | -14.656 | -110.713 | 12070 | 0 | 1 | UT2 | standard | H2O | 1661215466.42 | 0.213 | 0.77 | 58.938 | 0.001 | 0.006 | 1.467 | 60.663 | 27.042 | 1.996 | 61.062 | 0.856 | -0.848 | 0 | 0 | 1 | 42.062 | 1 | 3 | 1 | GJB_Rinse | 0
27 | P-20888 | 2022/08/22 18:48:17 | 1-03 | 7 | -14.68 | -110.903 | 12145 | 0 | 1 | UT2 | standard | H2O | 1661216002.38 | 0.195 | 0.784 | 62.369 | 0.001 | 0.005 | 1.552 | 60.777 | 27.198 | 2.014 | 62.152 | 0.856 | -0.843 | 0 | 0 | 1 | 42.062 | 1 | 3 | 1 | GJB_Rinse | 0
28 | P-20888 | 2022/08/22 18:57:14 | 1-03 | 8 | -14.679 | -111.333 | 11958 | 0 | 1 | UT2 | standard | H2O | 1661216539.84 | 0.191 | 0.744 | 58.146 | 0.001 | 0.002 | 1.456 | 60.235 | 26.928 | 1.994 | 61.034 | 0.855 | -0.762 | 0 | 0 | 1 | 42 | 1 | 3 | 1 | GJB_Rinse | 0
29 | P-20888 | 2022/08/22 19:06:11 | 1-03 | 9 | -14.721 | -111.644 | 12081 | 0 | 1 | UT2 | standard | H2O | 1661217076.59 | 0.206 | 0.763 | 61.465 | 0.001 | 0.004 | 1.519 | 60.411 | 27.574 | 1.999 | 61.088 | 0.856 | -0.722 | 0 | 0 | 1 | 42.125 | 1 | 3 | 1 | GJB_Rinse | 0
30 | P-20888 | 2022/08/22 19:15:08 | 1-03 | 10 | -14.741 | -111.643 | 12082 | 0 | 1 | UT2 | standard | H2O | 1661217613.73 | 0.2 | 0.807 | 62.603 | 0.002 | 0.003 | 1.552 | 60.688 | 27.223 | 1.989 | 60.8 | 0.856 | -1 | 0 | 0 | 1 | 42.125 | 1 | 3 | 1 | GJB_Rinse | 0
31 | P-20889 | 2022/08/22 19:24:57 | 1-04 | 1 | -9.463 | -72.588 | 12527 | -1 | 1 | EV | standard | H2O | 1661218202.98 | 0.196 | 0.844 | 64.299 | 0.002 | 0.002 | 1.589 | 61.383 | 28.592 | 2.134 | 63.971 | 0.856 | -0.901 | 0 | 0 | 1 | 42.062 | 1 | 4 | 1 | GJB_Rinse | 0
32 | P-20889 | 2022/08/22 19:33:55 | 1-04 | 2 | -9.218 | -69.083 | 12131 | -1 | 1 | EV | standard | H2O | 1661218740.54 | 0.2 | 0.748 | 61.831 | 0.001 | 0.002 | 1.522 | 60.712 | 27.569 | 2.065 | 61.347 | 0.856 | -1.02 | 0 | 0 | 1 | 42.125 | 1 | 4 | 1 | GJB_Rinse | 0
33 | P-20889 | 2022/08/22 19:42:51 | 1-04 | 3 | -9.203 | -68.329 | 12165 | -1 | 1 | EV | standard | H2O | 1661219277.23 | 0.193 | 0.781 | 59.525 | 0 | 0.002 | 1.472 | 60.896 | 27.861 | 2.055 | 60.288 | 0.856 | -1.259 | 0 | 0 | 1 | 42.125 | 1 | 4 | 1 | GJB_Rinse | 0
34 | P-20889 | 2022/08/22 19:51:48 | 1-04 | 4 | -9.149 | -67.776 | 12136 | 0 | 1 | EV | standard | H2O | 1661219814.6 | 0.211 | 0.822 | 62.569 | 0.002 | -0.001 | 1.556 | 60.614 | 27.629 | 2.08 | 62.24 | 0.856 | -0.891 | 0 | 0 | 1 | 41.938 | 1 | 4 | 1 | GJB_Rinse | 0
35 | P-20889 | 2022/08/22 20:00:45 | 1-04 | 5 | -9.176 | -67.679 | 12178 | 0 | 1 | EV | standard | H2O | 1661220351.03 | 0.195 | 0.861 | 61.913 | 0.001 | 0 | 1.533 | 60.67 | 27.394 | 2.09 | 63.161 | 0.855 | -0.813 | 0 | 0 | 1 | 42.125 | 1 | 4 | 1 | GJB_Rinse | 0
36 | P-20889 | 2022/08/22 20:09:43 | 1-04 | 6 | -9.185 | -67.517 | 12308 | 0 | 1 | EV | standard | H2O | 1661220888.27 | 0.205 | 0.695 | 59.548 | 0.001 | 0.001 | 1.468 | 60.963 | 28.278 | 2.098 | 62.436 | 0.856 | -0.922 | 0 | 0 | 1 | 42.062 | 1 | 4 | 1 | GJB_Rinse | 0
37 | P-20889 | 2022/08/22 20:18:40 | 1-04 | 7 | -9.12 | -67.375 | 12208 | 0 | 1 | EV | standard | H2O | 1661221424.73 | 0.202 | 0.791 | 61.122 | 0.001 | -0.001 | 1.515 | 60.807 | 27.482 | 2.095 | 63.133 | 0.856 | -0.859 | 0 | 0 | 1 | 42.125 | 1 | 4 | 1 | GJB_Rinse | 0
38 | P-20889 | 2022/08/22 20:27:36 | 1-04 | 8 | -9.146 | -67.526 | 12257 | 0 | 1 | EV | standard | H2O | 1661221962.29 | 0.204 | 0.752 | 62.204 | 0.001 | 0.002 | 1.548 | 60.976 | 27.673 | 2.083 | 63.013 | 0.856 | -0.892 | 0 | 0 | 1 | 42.312 | 1 | 4 | 1 | GJB_Rinse | 0
39 | P-20889 | 2022/08/22 20:36:33 | 1-04 | 9 | -9.123 | -67.242 | 12274 | 0 | 1 | EV | standard | H2O | 1661222499.1 | 0.232 | 0.775 | 62.354 | 0.001 | 0.002 | 1.55 | 60.983 | 28.073 | 2.093 | 62.051 | 0.856 | -1.149 | 0 | 0 | 1 | 42.312 | 1 | 4 | 1 | GJB_Rinse | 0
40 | P-20889 | 2022/08/22 20:45:31 | 1-04 | 10 | -9.11 | -67.294 | 12331 | 0 | 1 | EV | standard | H2O | 1661223036.3 | 0.209 | 0.74 | 64.377 | 0.001 | 0.001 | 1.594 | 61.107 | 28 | 2.107 | 63.046 | 0.856 | -1.005 | 0 | 0 | 1 | 42.125 | 1 | 4 | 1 | GJB_Rinse | 0
41 | P-20890 | 2022/08/22 20:55:20 | 1-05 | 1 | -1.501 | -24.141 | 12616 | -1 | 1 | OSBS2-D-2 | 22-032 | H2O | 1661223626.12 | 0.218 | 0.748 | 63.683 | 0.001 | 0.002 | 1.568 | 224.102 | 23.94 | 2.242 | 62.144 | 0.856 | -1.29 | 0 | 0 | 1 | 42.25 | 1 | 5 | 1 | GJB_Rinse | 0
42 | P-20890 | 2022/08/22 21:04:18 | 1-05 | 2 | -1.329 | -20.517 | 12304 | -1 | 1 | OSBS2-D-2 | 22-032 | H2O | 1661224163.14 | 0.224 | 0.814 | 62.531 | 0.001 | -0.004 | 1.553 | 220.402 | 23.012 | 2.185 | 60.047 | 0.856 | -1.467 | 0 | 0 | 1 | 42.125 | 1 | 5 | 1 | GJB_Rinse | 0
43 | P-20890 | 2022/08/22 21:13:14 | 1-05 | 3 | -1.288 | -19.645 | 12232 | -1 | 1 | OSBS2-D-2 | 22-032 | H2O | 1661224699.49 | 0.213 | 0.852 | 61.03 | 0 | -0.004 | 1.509 | 219.69 | 22.787 | 2.185 | 60.846 | 0.856 | -1.252 | 0 | 0 | 1 | 42.125 | 1 | 5 | 1 | GJB_Rinse | 0
44 | P-20890 | 2022/08/22 21:22:12 | 1-05 | 4 | -1.188 | -19.112 | 12262 | 0 | 1 | OSBS2-D-2 | 22-032 | H2O | 1661225237.35 | 0.223 | 0.745 | 61.744 | 0 | 0 | 1.527 | 220.483 | 23.088 | 2.184 | 59.876 | 0.857 | -1.441 | 0 | 0 | 1 | 42.062 | 1 | 5 | 1 | GJB_Rinse | 0
45 | P-20891 | 2022/08/22 21:32:02 | 1-06 | 1 | 0.481 | -8.309 | 11956 | -1 | 1 | OSBS2-C-3 | 22-032 | H2O | 1661225827.96 | 0.192 | 0.792 | 60.357 | 0.001 | -0.005 | 1.499 | 62.803 | 26.457 | 2.112 | 63.817 | 0.856 | -1.073 | 0 | 0 | 1 | 42.188 | 1 | 6 | 1 | GJB_Rinse | 0
46 | P-20891 | 2022/08/22 21:40:59 | 1-06 | 2 | 0.51 | -7.675 | 11959 | -1 | 1 | OSBS2-C-3 | 22-032 | H2O | 1661226364.67 | 0.213 | 0.803 | 57.074 | 0 | 0.001 | 1.415 | 61.778 | 26.628 | 2.13 | 62.864 | 0.856 | -1.171 | 0 | 0 | 1 | 42.25 | 1 | 6 | 1 | GJB_Rinse | 0
47 | P-20891 | 2022/08/22 21:49:56 | 1-06 | 3 | 0.51 | -7.193 | 11905 | -1 | 1 | OSBS2-C-3 | 22-032 | H2O | 1661226901.59 | 0.214 | 0.702 | 59.577 | 0.001 | -0.002 | 1.485 | 61.095 | 26.492 | 2.132 | 64.277 | 0.855 | -0.953 | 0 | 0 | 1 | 42.375 | 1 | 6 | 1 | GJB_Rinse | 0
48 | P-20891 | 2022/08/22 21:58:52 | 1-06 | 4 | 0.565 | -6.505 | 12044 | 0 | 1 | OSBS2-C-3 | 22-032 | H2O | 1661227437.38 | 0.195 | 0.793 | 60.51 | 0 | -0.002 | 1.503 | 61.216 | 26.869 | 2.162 | 64.838 | 0.856 | -0.919 | 0 | 0 | 1 | 42.312 | 1 | 6 | 1 | GJB_Rinse | 0
49 | P-20892 | 2022/08/22 22:08:42 | 1-07 | 1 | -0.197 | -12.022 | 12620 | -1 | 1 | OSBS2-A-4 | 22-032 | H2O | 1661228027.84 | 0.205 | 0.841 | 60.513 | 0.001 | -0.005 | 1.493 | 89.376 | 26.693 | 2.334 | 65.482 | 0.856 | -1.734 | 0 | 0 | 1 | 42.25 | 1 | 7 | 1 | GJB_Rinse | 0
50 | P-20892 | 2022/08/22 22:17:39 | 1-07 | 2 | -0.273 | -12.911 | 12259 | -1 | 1 | OSBS2-A-4 | 22-032 | H2O | 1661228564.48 | 0.202 | 0.838 | 62.716 | 0.001 | -0.003 | 1.555 | 88.042 | 25.906 | 2.282 | 64.927 | 0.856 | -1.558 | 0 | 0 | 1 | 42.188 | 1 | 7 | 1 | GJB_Rinse | 0
51 | P-20892 | 2022/08/22 22:26:35 | 1-07 | 3 | -0.279 | -12.856 | 12330 | -1 | 1 | OSBS2-A-4 | 22-032 | H2O | 1661229100.81 | 0.206 | 0.829 | 59.2 | 0 | -0.001 | 1.458 | 88.38 | 26.045 | 2.296 | 64.779 | 0.856 | -1.539 | 0 | 0 | 1 | 42.188 | 1 | 7 | 1 | GJB_Rinse | 0
52 | P-20892 | 2022/08/22 22:35:33 | 1-07 | 4 | -0.258 | -12.661 | 12400 | 0 | 1 | OSBS2-A-4 | 22-032 | H2O | 1661229638.01 | 0.198 | 0.785 | 62.669 | 0.001 | -0.001 | 1.56 | 88.643 | 26.163 | 2.294 | 65.305 | 0.856 | -1.587 | 0 | 0 | 1 | 42.312 | 1 | 7 | 1 | GJB_Rinse | 0
53 | P-20893 | 2022/08/22 22:45:20 | 1-08 | 1 | -0.287 | -15.399 | 11996 | -1 | 1 | OSBS2-A-2 | 22-032 | H2O | 1661230226.05 | 0.231 | 0.847 | 60.171 | 0.001 | -0.003 | 1.496 | 164.858 | 22.501 | 2.255 | 61.534 | 0.856 | -1.816 | 0 | 0 | 1 | 42.25 | 1 | 8 | 1 | GJB_Rinse | 0
54 | P-20893 | 2022/08/22 22:54:17 | 1-08 | 2 | -0.231 | -15.417 | 12004 | -1 | 1 | OSBS2-A-2 | 22-032 | H2O | 1661230762.89 | 0.184 | 0.805 | 57.438 | 0.001 | -0.002 | 1.428 | 165.182 | 22.636 | 2.258 | 62.243 | 0.856 | -1.717 | 0 | 0 | 1 | 42.25 | 1 | 8 | 1 | GJB_Rinse | 0
55 | P-20893 | 2022/08/22 23:03:14 | 1-08 | 3 | -0.265 | -15.45 | 12005 | -1 | 1 | OSBS2-A-2 | 22-032 | H2O | 1661231300.59 | 0.219 | 0.904 | 66.417 | 0.002 | 0 | 1.652 | 165.497 | 22.575 | 2.258 | 61.374 | 0.856 | -1.794 | 0 | 0 | 1 | 42 | 1 | 8 | 1 | GJB_Rinse | 0
56 | P-20893 | 2022/08/22 23:12:11 | 1-08 | 4 | -0.241 | -15.495 | 11989 | 0 | 1 | OSBS2-A-2 | 22-032 | H2O | 1661231836.91 | 0.219 | 0.842 | 58.047 | 0 | 0.001 | 1.451 | 165.482 | 22.621 | 2.254 | 62.055 | 0.856 | -1.81 | 0 | 0 | 1 | 41.938 | 1 | 8 | 1 | GJB_Rinse | 0
57 | P-20894 | 2022/08/22 23:21:59 | 1-09 | 1 | -1.507 | -22.072 | 12528 | -1 | 1 | OSBS2-A-1 | 22-032 | H2O | 1661232424.92 | 0.206 | 0.869 | 62.005 | 0.001 | -0.002 | 1.536 | 171.216 | 24.511 | 2.285 | 63.662 | 0.857 | -1.627 | 0 | 0 | 1 | 42 | 1 | 9 | 1 | GJB_Rinse | 0
58 | P-20894 | 2022/08/22 23:30:56 | 1-09 | 2 | -1.578 | -22.982 | 12398 | -1 | 1 | OSBS2-A-1 | 22-032 | H2O | 1661232961.81 | 0.21 | 0.857 | 59.735 | 0.001 | 0 | 1.482 | 169.912 | 23.578 | 2.28 | 63.869 | 0.856 | -1.392 | 0 | 0 | 1 | 41.938 | 1 | 9 | 1 | GJB_Rinse | 0
59 | P-20894 | 2022/08/22 23:39:53 | 1-09 | 3 | -1.561 | -23.338 | 12390 | -1 | 1 | OSBS2-A-1 | 22-032 | H2O | 1661233498.11 | 0.205 | 0.891 | 60.375 | 0 | 0.002 | 1.5 | 169.841 | 23.82 | 2.259 | 63.228 | 0.856 | -1.67 | 0 | 0 | 1 | 42 | 1 | 9 | 1 | GJB_Rinse | 0
60 | P-20894 | 2022/08/22 23:48:49 | 1-09 | 4 | -1.604 | -23.366 | 12260 | 0 | 1 | OSBS2-A-1 | 22-032 | H2O | 1661234034.41 | 0.203 | 0.783 | 59.094 | 0 | 0 | 1.47 | 168.298 | 23.715 | 2.244 | 62.748 | 0.856 | -1.434 | 0 | 0 | 1 | 42 | 1 | 9 | 1 | GJB_Rinse | 0
61 | P-20895 | 2022/08/22 23:58:37 | 1-10 | 1 | -1.914 | -22.621 | 11762 | -1 | 1 | OSBS2-D-3 | 22-032 | H2O | 1661234622.93 | 0.234 | 0.867 | 65.682 | 0 | 0.001 | 1.645 | 287.096 | 18.945 | 2.129 | 57.827 | 0.857 | -1.277 | 0 | 0 | 1 | 41.875 | 1 | 10 | 1 | GJB_Rinse | 0
62 | P-20895 | 2022/08/23 00:07:34 | 1-10 | 2 | -1.963 | -22.53 | 11760 | -1 | 1 | OSBS2-D-3 | 22-032 | H2O | 1661235159.4 | 0.218 | 0.881 | 60.745 | 0 | -0.001 | 1.518 | 287.61 | 19.095 | 2.12 | 57.199 | 0.856 | -1.366 | 0 | 0 | 1 | 42 | 1 | 10 | 1 | GJB_Rinse | 0
63 | P-20895 | 2022/08/23 00:16:30 | 1-10 | 3 | -1.927 | -22.21 | 11759 | -1 | 1 | OSBS2-D-3 | 22-032 | H2O | 1661235695.71 | 0.224 | 0.904 | 58.301 | 0.001 | 0.002 | 1.46 | 287.894 | 18.916 | 2.115 | 58.316 | 0.856 | -1.127 | 0 | 0 | 1 | 42 | 1 | 10 | 1 | GJB_Rinse | 0
64 | P-20895 | 2022/08/23 00:25:25 | 1-10 | 4 | -1.959 | -22.396 | 11833 | 0 | 1 | OSBS2-D-3 | 22-032 | H2O | 1661236231.29 | 0.226 | 0.866 | 62.582 | 0.001 | -0.003 | 1.557 | 289.744 | 18.646 | 2.132 | 59.795 | 0.856 | -1.165 | 0 | 0 | 1 | 42.125 | 1 | 10 | 1 | GJB_Rinse | 0
65 | P-20896 | 2022/08/23 00:35:14 | 1-11 | 1 | -2.39 | -29.411 | 12310 | -1 | 1 | OSBS2-E-1 | 22-032 | H2O | 1661236819.66 | 0.245 | 1.052 | 62.252 | 0 | -0.001 | 1.547 | 504.628 | 12.658 | 2.319 | 56.83 | 0.858 | -2.028 | 0 | 0 | 1 | 42.312 | 1 | 11 | 1 | GJB_Rinse | 0
66 | P-20896 | 2022/08/23 00:44:10 | 1-11 | 2 | -2.502 | -29.924 | 12036 | -1 | 1 | OSBS2-E-1 | 22-032 | H2O | 1661237355.59 | 0.252 | 1.015 | 59.95 | 0.002 | 0.001 | 1.495 | 495.266 | 12.29 | 2.286 | 55.769 | 0.859 | -2.05 | 0 | 0 | 1 | 42.312 | 1 | 11 | 1 | GJB_Rinse | 0
67 | P-20896 | 2022/08/23 00:53:06 | 1-11 | 3 | -2.483 | -30.056 | 12066 | -1 | 1 | OSBS2-E-1 | 22-032 | H2O | 1661237891.44 | 0.257 | 0.982 | 63.968 | 0.002 | 0 | 1.577 | 497.137 | 11.992 | 2.279 | 56.422 | 0.857 | -2.048 | 0 | 0 | 1 | 42 | 1 | 11 | 1 | GJB_Rinse | 0
68 | P-20896 | 2022/08/23 01:02:02 | 1-11 | 4 | -2.515 | -30.155 | 12078 | 0 | 1 | OSBS2-E-1 | 22-032 | H2O | 1661238428.23 | 0.213 | 0.989 | 60.567 | 0 | 0 | 1.511 | 497.836 | 11.871 | 2.276 | 56.162 | 0.857 | -2.165 | 0 | 0 | 1 | 41.75 | 1 | 11 | 1 | GJB_Rinse | 0
69 | P-20897 | 2022/08/23 01:11:51 | 1-12 | 1 | 1.695 | -5.64 | 12040 | -1 | 1 | OSBS2-C-1 | 22-032 | H2O | 1661239016.92 | 0.201 | 0.743 | 62.353 | 0.001 | 0.003 | 1.551 | 66.127 | 27.246 | 2.142 | 62.085 | 0.856 | -1.31 | 0 | 0 | 1 | 41.688 | 1 | 12 | 1 | GJB_Rinse | 0
70 | P-20897 | 2022/08/23 01:20:48 | 1-12 | 2 | 1.905 | -3.408 | 11973 | -1 | 1 | OSBS2-C-1 | 22-032 | H2O | 1661239553.79 | 0.171 | 0.737 | 60.389 | 0 | 0 | 1.497 | 63.555 | 27.09 | 2.146 | 63.208 | 0.856 | -1.14 | 0 | 0 | 1 | 41.625 | 1 | 12 | 1 | GJB_Rinse | 0
71 | P-20897 | 2022/08/23 01:29:45 | 1-12 | 3 | 1.939 | -2.909 | 11919 | -1 | 1 | OSBS2-C-1 | 22-032 | H2O | 1661240090.43 | 0.206 | 0.854 | 61.483 | 0.001 | -0.003 | 1.528 | 62.39 | 26.864 | 2.138 | 63.288 | 0.856 | -1.008 | 0 | 0 | 1 | 41.688 | 1 | 12 | 1 | GJB_Rinse | 0
72 | P-20897 | 2022/08/23 01:38:41 | 1-12 | 4 | 1.93 | -2.675 | 11956 | 0 | 1 | OSBS2-C-1 | 22-032 | H2O | 1661240626.69 | 0.204 | 0.825 | 61.899 | 0.001 | 0.002 | 1.54 | 61.938 | 26.988 | 2.152 | 63.654 | 0.856 | -0.963 | 0 | 0 | 1 | 42.062 | 1 | 12 | 1 | GJB_Rinse | 0
73 | P-20898 | 2022/08/23 01:48:30 | 1-13 | 1 | -8.48 | -58.554 | 12757 | -1 | 1 | EV | standard | H2O | 1661241215.81 | 0.21 | 0.729 | 61.492 | 0.002 | -0.003 | 1.519 | 62.486 | 28.997 | 2.191 | 66.317 | 0.856 | -0.81 | 0 | 0 | 1 | 42.062 | 1 | 13 | 1 | GJB_Rinse | 0
74 | P-20898 | 2022/08/23 01:57:27 | 1-13 | 2 | -8.875 | -63.55 | 12311 | -1 | 1 | EV | standard | H2O | 1661241752.66 | 0.202 | 0.767 | 64.315 | 0.002 | 0.005 | 1.591 | 61.786 | 28.059 | 2.114 | 62.995 | 0.856 | -0.99 | 0 | 0 | 1 | 42 | 1 | 13 | 1 | GJB_Rinse | 0
75 | P-20898 | 2022/08/23 02:06:24 | 1-13 | 3 | -8.922 | -64.861 | 12321 | -1 | 1 | EV | standard | H2O | 1661242289.08 | 0.191 | 0.803 | 61.656 | 0.001 | 0.004 | 1.522 | 61.753 | 27.838 | 2.102 | 63.407 | 0.856 | -0.931 | 0 | 0 | 1 | 42.062 | 1 | 13 | 1 | GJB_Rinse | 0
76 | P-20898 | 2022/08/23 02:15:20 | 1-13 | 4 | -9.028 | -65.426 | 12325 | 0 | 1 | EV | standard | H2O | 1661242825.88 | 0.217 | 0.831 | 60.46 | 0.001 | 0.003 | 1.504 | 61.669 | 27.463 | 2.104 | 64.23 | 0.856 | -0.781 | 0 | 0 | 1 | 42.062 | 1 | 13 | 1 | GJB_Rinse | 0
77 | P-20899 | 2022/08/23 02:25:10 | 1-14 | 1 | 0.022 | -11.737 | 12012 | -1 | 1 | OSBS2-C-2 | 22-032 | H2O | 1661243416.21 | 0.207 | 0.871 | 60.494 | 0.001 | 0.007 | 1.505 | 233.624 | 20.627 | 2.22 | 60.379 | 0.856 | -1.62 | 0 | 0 | 1 | 42.062 | 1 | 14 | 1 | GJB_Rinse | 0
78 | P-20899 | 2022/08/23 02:34:06 | 1-14 | 2 | 0.288 | -7.239 | 12023 | -1 | 1 | OSBS2-C-2 | 22-032 | H2O | 1661243952.56 | 0.221 | 0.857 | 62.493 | 0 | -0.001 | 1.545 | 235.147 | 20.599 | 2.228 | 60.773 | 0.856 | -1.608 | 0 | 0 | 1 | 42.188 | 1 | 14 | 1 | GJB_Rinse | 0
79 | P-20899 | 2022/08/23 02:43:03 | 1-14 | 3 | 0.388 | -5.886 | 11806 | -1 | 1 | OSBS2-C-2 | 22-032 | H2O | 1661244489.09 | 0.22 | 0.845 | 59.828 | 0.001 | 0.002 | 1.497 | 231.882 | 20.265 | 2.197 | 60.289 | 0.856 | -1.503 | 0 | 0 | 1 | 42.25 | 1 | 14 | 1 | GJB_Rinse | 0
80 | P-20899 | 2022/08/23 02:51:59 | 1-14 | 4 | 0.476 | -5.027 | 11912 | 0 | 1 | OSBS2-C-2 | 22-032 | H2O | 1661245024.82 | 0.212 | 0.878 | 64.888 | 0.001 | 0 | 1.611 | 233.718 | 20.41 | 2.214 | 60.965 | 0.856 | -1.373 | 0 | 0 | 1 | 42.25 | 1 | 14 | 1 | GJB_Rinse | 0
81 | P-20900 | 2022/08/23 03:01:47 | 1-15 | 1 | -0.197 | -8.665 | 12443 | -1 | 1 | OSBS2-C-4 | 22-032 | H2O | 1661245612.12 | 0.199 | 0.799 | 61.761 | 0 | 0.004 | 1.534 | 64.312 | 27.987 | 2.216 | 65.385 | 0.856 | -1.197 | 0 | 0 | 1 | 42.312 | 1 | 15 | 1 | GJB_Rinse | 0
82 | P-20900 | 2022/08/23 03:10:43 | 1-15 | 2 | -0.206 | -8.876 | 12379 | -1 | 1 | OSBS2-C-4 | 22-032 | H2O | 1661246149.02 | 0.199 | 0.811 | 63.64 | 0.001 | -0.001 | 1.582 | 63.113 | 27.929 | 2.215 | 64.559 | 0.856 | -1.285 | 0 | 0 | 1 | 42.25 | 1 | 15 | 1 | GJB_Rinse | 0
83 | P-20900 | 2022/08/23 03:19:41 | 1-15 | 3 | -0.214 | -8.922 | 12353 | -1 | 1 | OSBS2-C-4 | 22-032 | H2O | 1661246686.57 | 0.21 | 0.806 | 58.711 | 0.001 | -0.002 | 1.451 | 62.411 | 28.013 | 2.208 | 65.372 | 0.856 | -1.105 | 0 | 0 | 1 | 42.25 | 1 | 15 | 1 | GJB_Rinse | 0
84 | P-20900 | 2022/08/23 03:28:37 | 1-15 | 4 | -0.151 | -8.929 | 12324 | 0 | 1 | OSBS2-C-4 | 22-032 | H2O | 1661247223.01 | 0.225 | 0.929 | 61.064 | 0.001 | 0 | 1.518 | 62.182 | 27.845 | 2.201 | 64.767 | 0.856 | -1.113 | 0 | 0 | 1 | 42.188 | 1 | 15 | 1 | GJB_Rinse | 0
85 | P-20901 | 2022/08/23 03:38:26 | 1-16 | 1 | 2.912 | -9.563 | 12297 | -1 | 1 | OSBS2-C-5 | 22-032 | H2O | 1661247811.55 | 0.218 | 0.844 | 66.495 | 0.001 | 0.006 | 1.637 | 61.691 | 27.878 | 2.201 | 65.3 | 0.856 | -0.913 | 0 | 0 | 1 | 42.188 | 1 | 16 | 1 | GJB_Rinse | 0
86 | P-20901 | 2022/08/23 03:47:22 | 1-16 | 2 | 2.999 | -9.316 | 12110 | -1 | 1 | OSBS2-C-5 | 22-032 | H2O | 1661248347.2 | 0.206 | 0.849 | 64.308 | 0.001 | 0 | 1.599 | 61.559 | 27.179 | 2.158 | 63.854 | 0.856 | -1.016 | 0 | 0 | 1 | 42.375 | 1 | 16 | 1 | GJB_Rinse | 0
87 | P-20901 | 2022/08/23 03:56:19 | 1-16 | 3 | 3.075 | -9.358 | 12122 | -1 | 1 | OSBS2-C-5 | 22-032 | H2O | 1661248884.47 | 0.215 | 0.734 | 66.072 | 0.001 | 0.001 | 1.637 | 61.608 | 27.192 | 2.161 | 64.331 | 0.855 | -1.078 | 0 | 0 | 1 | 42.375 | 1 | 16 | 1 | GJB_Rinse | 0
88 | P-20901 | 2022/08/23 04:05:15 | 1-16 | 4 | 3.056 | -9.299 | 12257 | 0 | 1 | OSBS2-C-5 | 22-032 | H2O | 1661249420.83 | 0.196 | 0.856 | 62.695 | 0 | 0.003 | 1.56 | 61.901 | 27.7 | 2.168 | 63.744 | 0.856 | -1.181 | 0 | 0 | 1 | 42.375 | 1 | 16 | 1 | GJB_Rinse | 0
89 | P-20902 | 2022/08/23 04:15:03 | 1-17 | 1 | -0.1 | -21.889 | 11754 | -1 | 1 | OSBS2-B-1 | 22-032 | H2O | 1661250009.26 | 0.217 | 1.02 | 58.483 | 0 | -0.004 | 1.453 | 327.086 | 16.539 | 2.234 | 57.546 | 0.857 | -2.056 | 0 | 0 | 1 | 42.5 | 1 | 17 | 1 | GJB_Rinse | 0
90 | P-20902 | 2022/08/23 04:24:00 | 1-17 | 2 | -0.252 | -23.393 | 12166 | -1 | 1 | OSBS2-B-1 | 22-032 | H2O | 1661250546 | 0.234 | 0.907 | 59.922 | 0.001 | -0.002 | 1.485 | 339.206 | 17.348 | 2.311 | 59.008 | 0.857 | -2.109 | 0 | 0 | 1 | 42.5 | 1 | 17 | 1 | GJB_Rinse | 0
91 | P-20902 | 2022/08/23 04:32:56 | 1-17 | 3 | -0.255 | -23.831 | 11984 | -1 | 1 | OSBS2-B-1 | 22-032 | H2O | 1661251081.38 | 0.22 | 0.835 | 63.601 | 0.001 | 0 | 1.582 | 335.163 | 16.85 | 2.274 | 58.478 | 0.856 | -2.063 | 0 | 0 | 1 | 42.438 | 1 | 17 | 1 | GJB_Rinse | 0
92 | P-20902 | 2022/08/23 04:41:53 | 1-17 | 4 | -0.31 | -24.089 | 12148 | 0 | 1 | OSBS2-B-1 | 22-032 | H2O | 1661251618.36 | 0.229 | 0.97 | 61.549 | 0.001 | 0.004 | 1.525 | 339.613 | 17.113 | 2.316 | 59.765 | 0.856 | -1.876 | 0 | 0 | 1 | 42.5 | 1 | 17 | 1 | GJB_Rinse | 0
93 | P-20903 | 2022/08/23 04:51:40 | 1-18 | 1 | -1.002 | -20.816 | 12297 | -1 | 1 | OSBS2-B-2 | 22-032 | H2O | 1661252205.89 | 0.223 | 0.767 | 59.976 | 0.002 | -0.002 | 1.491 | 120.237 | 25.794 | 2.217 | 63.198 | 0.856 | -1.423 | 0 | 0 | 1 | 42.375 | 1 | 18 | 1 | GJB_Rinse | 0
94 | P-20903 | 2022/08/23 05:00:37 | 1-18 | 2 | -1.001 | -20.42 | 12221 | -1 | 1 | OSBS2-B-2 | 22-032 | H2O | 1661252742.71 | 0.205 | 0.886 | 58.471 | 0.001 | 0.002 | 1.446 | 118.627 | 25.717 | 2.203 | 62.23 | 0.856 | -1.48 | 0 | 0 | 1 | 42.25 | 1 | 18 | 1 | GJB_Rinse | 0
95 | P-20903 | 2022/08/23 05:09:34 | 1-18 | 3 | -1.042 | -20.108 | 12171 | -1 | 1 | OSBS2-B-2 | 22-032 | H2O | 1661253279.68 | 0.19 | 0.826 | 61.86 | 0.001 | 0.002 | 1.54 | 117.805 | 25.401 | 2.193 | 62.841 | 0.856 | -1.335 | 0 | 0 | 1 | 42.375 | 1 | 18 | 1 | GJB_Rinse | 0
96 | P-20903 | 2022/08/23 05:18:31 | 1-18 | 4 | -1.015 | -20.357 | 12186 | 0 | 1 | OSBS2-B-2 | 22-032 | H2O | 1661253816.34 | 0.201 | 0.85 | 61.517 | 0.001 | 0 | 1.512 | 117.762 | 25.593 | 2.184 | 62.184 | 0.856 | -1.453 | 0 | 0 | 1 | 42.5 | 1 | 18 | 1 | GJB_Rinse | 0
97 | P-20904 | 2022/08/23 05:28:20 | 1-19 | 1 | -0.89 | -19.84 | 11977 | -1 | 1 | OSBS2-B-3 | 22-032 | H2O | 1661254406.04 | 0.214 | 0.873 | 61.191 | 0.001 | 0.002 | 1.514 | 72.299 | 26.884 | 2.165 | 59.866 | 0.856 | -1.622 | 0 | 0 | 1 | 42.188 | 1 | 19 | 1 | GJB_Rinse | 0
98 | P-20904 | 2022/08/23 05:37:17 | 1-19 | 2 | -0.92 | -20.085 | 11910 | -1 | 1 | OSBS2-B-3 | 22-032 | H2O | 1661254942.52 | 0.214 | 0.901 | 59.686 | 0 | 0.001 | 1.493 | 71.848 | 25.995 | 2.159 | 62.432 | 0.856 | -1.35 | 0 | 0 | 1 | 42.062 | 1 | 19 | 1 | GJB_Rinse | 0
99 | P-20904 | 2022/08/23 05:46:13 | 1-19 | 3 | -0.903 | -19.898 | 11889 | -1 | 1 | OSBS2-B-3 | 22-032 | H2O | 1661255478.82 | 0.187 | 0.864 | 62.164 | 0.001 | 0.001 | 1.547 | 71.832 | 25.476 | 2.154 | 62.85 | 0.855 | -1.381 | 0 | 0 | 1 | 41.938 | 1 | 19 | 1 | GJB_Rinse | 0
100 | P-20904 | 2022/08/23 05:55:09 | 1-19 | 4 | -0.912 | -20.067 | 11857 | 0 | 1 | OSBS2-B-3 | 22-032 | H2O | 1661256015.16 | 0.197 | 0.923 | 59.426 | 0.001 | 0 | 1.478 | 71.478 | 25.943 | 2.154 | 61.573 | 0.856 | -1.36 | 0 | 0 | 1 | 41.938 | 1 | 19 | 1 | GJB_Rinse | 0
101 | P-20905 | 2022/08/23 06:04:58 | 1-20 | 1 | -0.872 | -22 | 12037 | -1 | 1 | OSBS2-B-4 | 22-032 | H2O | 1661256604.05 | 0.19 | 0.735 | 58.876 | 0.001 | -0.001 | 1.461 | 100.29 | 25.61 | 2.198 | 62.071 | 0.856 | -1.617 | 0 | 0 | 1 | 42.062 | 1 | 20 | 1 | GJB_Rinse | 0
102 | P-20905 | 2022/08/23 06:13:54 | 1-20 | 2 | -0.919 | -22.397 | 11900 | -1 | 1 | OSBS2-B-4 | 22-032 | H2O | 1661257139.71 | 0.212 | 0.815 | 58.846 | 0.001 | -0.001 | 1.463 | 99.843 | 24.979 | 2.175 | 61.876 | 0.856 | -1.435 | 0 | 0 | 1 | 42.062 | 1 | 20 | 1 | GJB_Rinse | 0
103 | P-20905 | 2022/08/23 06:22:51 | 1-20 | 3 | -0.864 | -22.381 | 12111 | -1 | 1 | OSBS2-B-4 | 22-032 | H2O | 1661257676.98 | 0.215 | 0.773 | 59.054 | 0 | 0.001 | 1.467 | 101.205 | 24.778 | 2.214 | 63.79 | 0.856 | -1.453 | 0 | 0 | 1 | 42.125 | 1 | 20 | 1 | GJB_Rinse | 0
104 | P-20905 | 2022/08/23 06:31:48 | 1-20 | 4 | -0.926 | -22.617 | 12125 | 0 | 1 | OSBS2-B-4 | 22-032 | H2O | 1661258213.73 | 0.209 | 0.86 | 61.482 | 0.001 | 0.002 | 1.528 | 100.994 | 24.986 | 2.229 | 64.379 | 0.856 | -1.322 | 0 | 0 | 1 | 42.25 | 1 | 20 | 1 | GJB_Rinse | 0
105 | P-20906 | 2022/08/23 06:41:37 | 1-21 | 1 | -0.204 | -14.371 | 12133 | -1 | 1 | OSBS2-B-5 | 22-032 | H2O | 1661258802.5 | 0.252 | 1.081 | 61.516 | 0 | 0.002 | 1.532 | 609.611 | 8.087 | 2.331 | 56.252 | 0.858 | -1.998 | 0 | 0 | 1 | 42.438 | 1 | 21 | 1 | GJB_Rinse | 0
106 | P-20906 | 2022/08/23 06:50:33 | 1-21 | 2 | -0.145 | -13.5 | 11912 | -1 | 1 | OSBS2-B-5 | 22-032 | H2O | 1661259338.97 | 0.262 | 1.03 | 61.286 | 0.002 | -0.001 | 1.523 | 602.014 | 7.711 | 2.306 | 55.696 | 0.857 | -1.867 | 0 | 0 | 1 | 42.438 | 1 | 21 | 1 | GJB_Rinse | 0
107 | P-20906 | 2022/08/23 06:59:30 | 1-21 | 3 | -0.151 | -13.255 | 11749 | -1 | 1 | OSBS2-B-5 | 22-032 | H2O | 1661259875.81 | 0.228 | 1.021 | 57.014 | 0.001 | 0.003 | 1.424 | 595.525 | 8.106 | 2.241 | 52.743 | 0.858 | -2.241 | 0 | 0 | 1 | 42.5 | 1 | 21 | 1 | GJB_Rinse | 0
108 | P-20906 | 2022/08/23 07:08:26 | 1-21 | 4 | -0.17 | -13.237 | 11762 | 0 | 1 | OSBS2-B-5 | 22-032 | H2O | 1661260412.13 | 0.254 | 1.125 | 55.164 | 0 | -0.005 | 1.372 | 596.386 | 8.193 | 2.252 | 52.992 | 0.857 | -2.056 | 0 | 0 | 1 | 42.25 | 1 | 21 | 1 | GJB_Rinse | 0
109 | P-20907 | 2022/08/23 07:18:12 | 1-22 | 1 | -8.622 | -59.422 | 12394 | -1 | 1 | EV | standard | H2O | 1661260997.12 | 0.205 | 0.74 | 62.092 | 0.001 | 0.001 | 1.536 | 67.185 | 28.142 | 2.11 | 63.252 | 0.856 | -1.103 | 0 | 0 | 1 | 42 | 1 | 22 | 1 | GJB_Rinse | 0
110 | P-20907 | 2022/08/23 07:27:07 | 1-22 | 2 | -8.875 | -63.799 | 12088 | -1 | 1 | EV | standard | H2O | 1661261533.13 | 0.188 | 0.781 | 63.852 | 0.001 | 0.004 | 1.589 | 63.765 | 27.012 | 2.068 | 62.674 | 0.856 | -0.881 | 0 | 0 | 1 | 41.875 | 1 | 22 | 1 | GJB_Rinse | 0
111 | P-20907 | 2022/08/23 07:36:03 | 1-22 | 3 | -8.98 | -64.874 | 12057 | -1 | 1 | EV | standard | H2O | 1661262068.36 | 0.212 | 0.815 | 60.315 | 0.001 | 0.001 | 1.49 | 62.314 | 26.998 | 2.065 | 62.711 | 0.855 | -0.816 | 0 | 0 | 1 | 41.938 | 1 | 22 | 1 | GJB_Rinse | 0
112 | P-20907 | 2022/08/23 07:44:58 | 1-22 | 4 | -9.002 | -65.328 | 12080 | 0 | 1 | EV | standard | H2O | 1661262603.88 | 0.219 | 0.846 | 63.663 | 0.002 | 0.004 | 1.583 | 61.81 | 27.228 | 2.062 | 62.345 | 0.856 | -1.017 | 0 | 0 | 1 | 41.938 | 1 | 22 | 1 | GJB_Rinse | 0
113 | P-20908 | 2022/08/23 07:54:44 | 1-23 | 1 | -1.779 | -22.247 | 13391 | -1 | 1 | OSBS2-A-3 | 22-032 | H2O | 1661263190.08 | 0.221 | 0.832 | 61.549 | 0 | 0.001 | 1.508 | 174.488 | 26.498 | 2.436 | 67.917 | 0.857 | -1.354 | 0 | 0 | 1 | 41.938 | 1 | 23 | 1 | GJB_Rinse | 0
114 | P-20908 | 2022/08/23 08:03:41 | 1-23 | 2 | -1.51 | -18.58 | 8484 | -1 | 1 | OSBS2-A-3 | 22-032 | H2O | 1661263731.68 | 0.206 | 0.996 | 51.942 | 0.001 | 0.001 | 1.407 | 124.366 | 14.464 | 1.58 | 47.048 | 0.855 | -1.22 | 0 | 0 | 1 | 42.062 | 1 | 23 | 1 | GJB_Rinse | 0
115 | P-20908 | 2022/08/23 08:12:38 | 1-23 | 3 | -1.426 | -17.764 | 9600 | -1 | 1 | OSBS2-A-3 | 22-032 | H2O | 1661264265.69 | 0.203 | 1.106 | 54.335 | 0.001 | 0 | 1.407 | 135.56 | 16.972 | 1.781 | 52.21 | 0.855 | -1.259 | 0 | 0 | 1 | 42.188 | 1 | 23 | 1 | GJB_Rinse | 0
116 | P-20908 | 2022/08/23 08:21:34 | 1-23 | 4 | -1.424 | -17.367 | 8656 | 0 | 1 | OSBS2-A-3 | 22-032 | H2O | 1661264803.88 | 0.225 | 1.009 | 48.341 | 0.001 | 0.003 | 1.282 | 126.627 | 14.425 | 1.613 | 48.782 | 0.855 | -1.138 | 0 | 0 | 1 | 42.25 | 1 | 23 | 1 | GJB_Rinse | 0
117 | P-20909 | 2022/08/23 08:31:21 | 1-24 | 1 | -2.275 | -25.754 | 11763 | -1 | 1 | OSBS2-A-5 | 22-032 | H2O | 1661265387.28 | 0.218 | 0.826 | 59.966 | 0.001 | 0.001 | 1.492 | 122.691 | 23.469 | 2.151 | 61.38 | 0.856 | -1.359 | 0 | 0 | 1 | 42.312 | 1 | 24 | 1 | GJB_Rinse | 0
118 | P-20909 | 2022/08/23 08:40:17 | 1-24 | 2 | -2.326 | -26.673 | 11623 | -1 | 1 | OSBS2-A-5 | 22-032 | H2O | 1661265923 | 0.194 | 0.84 | 56.614 | 0.001 | 0 | 1.406 | 121.417 | 23.173 | 2.132 | 61.03 | 0.856 | -1.39 | 0 | 0 | 1 | 42.25 | 1 | 24 | 1 | GJB_Rinse | 0
119 | P-20909 | 2022/08/23 08:49:14 | 1-24 | 3 | -2.35 | -26.66 | 11732 | -1 | 1 | OSBS2-A-5 | 22-032 | H2O | 1661266459.35 | 0.207 | 0.876 | 58.598 | 0 | 0.001 | 1.462 | 122.148 | 23.877 | 2.141 | 60.373 | 0.856 | -1.419 | 0 | 0 | 1 | 42.062 | 1 | 24 | 1 | GJB_Rinse | 0
120 | P-20909 | 2022/08/23 08:58:10 | 1-24 | 4 | -2.302 | -26.707 | 11774 | 0 | 1 | OSBS2-A-5 | 22-032 | H2O | 1661266996.4 | 0.211 | 0.748 | 56.316 | 0.001 | -0.001 | 1.408 | 122.378 | 24.52 | 2.163 | 59.07 | 0.856 | -1.781 | 0 | 0 | 1 | 41.875 | 1 | 24 | 1 | GJB_Rinse | 0
121 | P-20910 | 2022/08/23 09:07:58 | 1-25 | 1 | 1.02 | -15.089 | 13988 | -1 | 1 | OSBS2-E-2 | 22-032 | H2O | 1661267583.3 | 0.217 | 0.756 | 57.345 | 0 | -0.001 | 1.394 | 78.161 | 30.391 | 2.667 | 71.777 | 0.857 | -2.094 | 0 | 0 | 1 | 41.75 | 1 | 25 | 1 | GJB_Rinse | 0
122 | P-20910 | 2022/08/23 09:16:53 | 1-25 | 2 | 1.169 | -14.069 | 8878 | -1 | 1 | OSBS2-E-2 | 22-032 | H2O | 1661268122.17 | 0.214 | 0.978 | 54.257 | 0 | -0.006 | 1.439 | 63.085 | 16.097 | 1.727 | 52.167 | 0.855 | -1.332 | 0 | 0 | 1 | 41.688 | 1 | 25 | 1 | GJB_Rinse | 0
123 | P-20910 | 2022/08/23 09:25:49 | 1-25 | 3 | 1.168 | -13.77 | 9148 | -1 | 1 | OSBS2-E-2 | 22-032 | H2O | 1661268657.2 | 0.233 | 0.833 | 49.514 | 0.001 | -0.001 | 1.298 | 63.652 | 17.456 | 1.769 | 51.768 | 0.855 | -1.546 | 0 | 0 | 1 | 41.812 | 1 | 25 | 1 | GJB_Rinse | 0
124 | P-20910 | 2022/08/23 09:34:45 | 1-25 | 4 | 1.171 | -13.825 | 9423 | 0 | 1 | OSBS2-E-2 | 22-032 | H2O | 1661269193.36 | 0.211 | 0.926 | 50.558 | 0 | -0.002 | 1.315 | 64.423 | 18.06 | 1.825 | 53.258 | 0.854 | -1.518 | 0 | 0 | 1 | 41.812 | 1 | 25 | 1 | GJB_Rinse | 0
125 | P-20911 | 2022/08/23 09:44:33 | 1-26 | 1 | 1.112 | -10.958 | 11854 | -1 | 1 | OSBS2-E-3 | 22-032 | H2O | 1661269778.7 | 0.205 | 0.864 | 55.212 | 0.001 | -0.002 | 1.38 | 64.274 | 25.496 | 2.231 | 62.797 | 0.856 | -1.608 | 0 | 0 | 1 | 41.812 | 1 | 26 | 1 | GJB_Rinse | 0
126 | P-20911 | 2022/08/23 09:53:29 | 1-26 | 2 | 1.139 | -10.385 | 12194 | -1 | 1 | OSBS2-E-3 | 22-032 | H2O | 1661270314.53 | 0.187 | 0.806 | 56.919 | 0 | -0.002 | 1.411 | 65.192 | 25.91 | 2.296 | 64.922 | 0.856 | -1.699 | 0 | 0 | 1 | 41.812 | 1 | 26 | 1 | GJB_Rinse | 0
127 | P-20911 | 2022/08/23 10:02:25 | 1-26 | 3 | 1.142 | -10.561 | 12205 | -1 | 1 | OSBS2-E-3 | 22-032 | H2O | 1661270851.21 | 0.217 | 0.756 | 60.513 | 0.001 | -0.001 | 1.487 | 65.151 | 26.215 | 2.299 | 64.502 | 0.856 | -1.766 | 0 | 0 | 1 | 41.75 | 1 | 26 | 1 | GJB_Rinse | 0
128 | P-20911 | 2022/08/23 10:11:22 | 1-26 | 4 | 1.154 | -10.108 | 12355 | 0 | 1 | OSBS2-E-3 | 22-032 | H2O | 1661271387.49 | 0.182 | 0.778 | 60.319 | 0.001 | -0.001 | 1.5 | 65.205 | 26.733 | 2.339 | 66.823 | 0.856 | -1.491 | 0 | 0 | 1 | 41.75 | 1 | 26 | 1 | GJB_Rinse | 0
129 | P-20912 | 2022/08/23 10:21:10 | 1-27 | 1 | 0 | -14.443 | 11952 | -1 | 1 | OSBS2-E-4 | 22-032 | H2O | 1661271975.98 | 0.205 | 0.8 | 61.2 | 0 | 0.002 | 1.523 | 67.842 | 25.715 | 2.213 | 63.198 | 0.856 | -1.568 | 0 | 0 | 1 | 41.75 | 1 | 27 | 1 | GJB_Rinse | 0
130 | P-20912 | 2022/08/23 10:30:06 | 1-27 | 2 | -0.052 | -14.696 | 12247 | -1 | 1 | OSBS2-E-4 | 22-032 | H2O | 1661272512.5 | 0.193 | 0.871 | 66.494 | 0.001 | 0.001 | 1.653 | 68.481 | 26.704 | 2.268 | 64.599 | 0.856 | -1.496 | 0 | 0 | 1 | 41.812 | 1 | 27 | 1 | GJB_Rinse | 0
131 | P-20912 | 2022/08/23 10:39:02 | 1-27 | 3 | -0.077 | -14.793 | 12159 | -1 | 1 | OSBS2-E-4 | 22-032 | H2O | 1661273047.52 | 0.197 | 0.814 | 60.391 | 0.001 | -0.003 | 1.5 | 68.608 | 26.349 | 2.241 | 62.81 | 0.856 | -1.857 | 0 | 0 | 1 | 41.75 | 1 | 27 | 1 | GJB_Rinse | 0
132 | P-20912 | 2022/08/23 10:47:58 | 1-27 | 4 | -0.105 | -14.698 | 12117 | 0 | 1 | OSBS2-E-4 | 22-032 | H2O | 1661273583.75 | 0.22 | 0.862 | 65.344 | 0.002 | 0.003 | 1.625 | 68.313 | 26.454 | 2.233 | 63 | 0.856 | -1.776 | 0 | 0 | 1 | 41.75 | 1 | 27 | 1 | GJB_Rinse | 0
133 | P-20913 | 2022/08/23 10:57:46 | 1-28 | 1 | 0.529 | -13.595 | 12481 | -1 | 1 | OSBS2-E-5 | 22-032 | H2O | 1661274172.23 | 0.221 | 0.937 | 63.577 | 0 | -0.001 | 1.57 | 71.776 | 25.819 | 2.46 | 66.572 | 0.856 | -2.197 | 0 | 0 | 1 | 41.75 | 1 | 28 | 1 | GJB_Rinse | 0
134 | P-20913 | 2022/08/23 11:06:42 | 1-28 | 2 | 0.479 | -14.383 | 11980 | -1 | 1 | OSBS2-E-5 | 22-032 | H2O | 1661274707.79 | 0.201 | 0.798 | 60.039 | 0.001 | -0.002 | 1.492 | 70.493 | 24.387 | 2.361 | 64.002 | 0.856 | -2.247 | 0 | 0 | 1 | 41.938 | 1 | 28 | 1 | GJB_Rinse | 0
135 | P-20913 | 2022/08/23 11:15:38 | 1-28 | 3 | 0.498 | -14.088 | 12463 | -1 | 1 | OSBS2-E-5 | 22-032 | H2O | 1661275244.21 | 0.209 | 0.844 | 58.165 | 0 | -0.002 | 1.436 | 71.902 | 25.46 | 2.465 | 66.792 | 0.856 | -2.209 | 0 | 0 | 1 | 41.875 | 1 | 28 | 1 | GJB_Rinse | 0
136 | P-20913 | 2022/08/23 11:24:35 | 1-28 | 4 | 0.47 | -13.789 | 12301 | 0 | 1 | OSBS2-E-5 | 22-032 | H2O | 1661275780.22 | 0.192 | 0.752 | 63.52 | 0.001 | 0.005 | 1.572 | 71.494 | 25.052 | 2.428 | 65.865 | 0.856 | -2.402 | 0 | 0 | 1 | 41.812 | 1 | 28 | 1 | GJB_Rinse | 0
137 | P-20914 | 2022/08/23 11:34:18 | 1-29 | 1 | 0.708 | -4.015 | 12414 | -1 | 1 | OSBS2-010-1 | 22-032 | H2O | 1661276363.83 | 0.185 | 0.821 | 65.805 | 0.001 | 0 | 1.63 | 61.603 | 28.44 | 2.215 | 63.78 | 0.856 | -1.265 | 0 | 0 | 1 | 41.688 | 1 | 29 | 1 | GJB_Rinse | 0
138 | P-20914 | 2022/08/23 11:43:15 | 1-29 | 2 | 0.756 | -2.771 | 12291 | -1 | 1 | OSBS2-010-1 | 22-032 | H2O | 1661276900.19 | 0.202 | 0.906 | 61.857 | 0.001 | 0.004 | 1.537 | 61.133 | 28.141 | 2.185 | 63.955 | 0.856 | -1.235 | 0 | 0 | 1 | 41.75 | 1 | 29 | 1 | GJB_Rinse | 0
139 | P-20914 | 2022/08/23 11:52:11 | 1-29 | 3 | 0.764 | -2.422 | 12364 | -1 | 1 | OSBS2-010-1 | 22-032 | H2O | 1661277436.43 | 0.21 | 0.809 | 63.075 | 0 | 0.002 | 1.563 | 61.013 | 28.334 | 2.22 | 65.237 | 0.856 | -0.915 | 0 | 0 | 1 | 41.75 | 1 | 29 | 1 | GJB_Rinse | 0
140 | P-20914 | 2022/08/23 12:01:06 | 1-29 | 4 | 0.758 | -2.123 | 12285 | 0 | 1 | OSBS2-010-1 | 22-032 | H2O | 1661277971.88 | 0.213 | 0.852 | 62.365 | 0.001 | 0 | 1.548 | 60.993 | 28.117 | 2.196 | 63.931 | 0.856 | -1.12 | 0 | 0 | 1 | 41.875 | 1 | 29 | 1 | GJB_Rinse | 0
141 | P-20915 | 2022/08/23 12:10:52 | 1-30 | 1 | -0.087 | -3.269 | 12061 | -1 | 1 | OSBS2-010-2 | 22-032 | H2O | 1661278557.85 | 0.217 | 0.886 | 58.364 | 0.001 | -0.001 | 1.45 | 60.407 | 27.503 | 2.171 | 63.403 | 0.856 | -0.958 | 0 | 0 | 1 | 41.875 | 1 | 30 | 1 | GJB_Rinse | 0
142 | P-20915 | 2022/08/23 12:19:48 | 1-30 | 2 | -0.102 | -3.172 | 11778 | -1 | 1 | OSBS2-010-2 | 22-032 | H2O | 1661279093.83 | 0.207 | 0.825 | 58.939 | 0.001 | -0.001 | 1.462 | 59.861 | 26.607 | 2.106 | 63.057 | 0.856 | -0.966 | 0 | 0 | 1 | 41.812 | 1 | 30 | 1 | GJB_Rinse | 0
143 | P-20915 | 2022/08/23 12:28:43 | 1-30 | 3 | -0.075 | -3.204 | 11894 | -1 | 1 | OSBS2-010-2 | 22-032 | H2O | 1661279628.89 | 0.191 | 0.769 | 63.454 | 0 | 0.002 | 1.578 | 60.039 | 26.923 | 2.14 | 63.601 | 0.856 | -0.928 | 0 | 0 | 1 | 41.938 | 1 | 30 | 1 | GJB_Rinse | 0
144 | P-20915 | 2022/08/23 12:37:38 | 1-30 | 4 | -0.047 | -3.034 | 11925 | 0 | 1 | OSBS2-010-2 | 22-032 | H2O | 1661280163.3 | 0.208 | 0.77 | 57.408 | 0.001 | 0.001 | 1.423 | 60.115 | 27.139 | 2.141 | 63.247 | 0.856 | -1.027 | 0 | 0 | 1 | 42.062 | 1 | 30 | 1 | GJB_Rinse | 0
145 | P-20916 | 2022/08/23 12:47:24 | 1-31 | 1 | -8.59 | -58.34 | 12640 | -1 | 1 | EV | standard | H2O | 1661280749.1 | 0.197 | 0.83 | 62.826 | 0.001 | -0.001 | 1.55 | 61.611 | 29.339 | 2.15 | 63.173 | 0.857 | -1.039 | 0 | 0 | 1 | 42.062 | 1 | 31 | 1 | GJB_Rinse | 0
146 | P-20916 | 2022/08/23 12:56:20 | 1-31 | 2 | -8.856 | -62.942 | 12116 | -1 | 1 | EV | standard | H2O | 1661281285.55 | 0.185 | 0.633 | 60.314 | 0.001 | 0.003 | 1.49 | 60.694 | 26.879 | 2.07 | 63.651 | 0.856 | -0.896 | 0 | 0 | 1 | 42 | 1 | 31 | 1 | GJB_Rinse | 0
147 | P-20916 | 2022/08/23 13:05:16 | 1-31 | 3 | -8.943 | -64.408 | 12066 | -1 | 1 | EV | standard | H2O | 1661281821.83 | 0.222 | 0.903 | 63.575 | 0.002 | 0.005 | 1.581 | 60.483 | 27.178 | 2.066 | 62.279 | 0.856 | -0.846 | 0 | 0 | 1 | 41.875 | 1 | 31 | 1 | GJB_Rinse | 0
148 | P-20916 | 2022/08/23 13:14:12 | 1-31 | 4 | -8.981 | -65.093 | 12060 | 0 | 1 | EV | standard | H2O | 1661282357.45 | 0.183 | 0.805 | 60.643 | 0.001 | 0.008 | 1.498 | 60.461 | 27.273 | 2.061 | 62.017 | 0.856 | -0.915 | 0 | 0 | 1 | 41.938 | 1 | 31 | 1 | GJB_Rinse | 0
149 | P-20917 | 2022/08/23 13:23:58 | 1-32 | 1 | -1.445 | -21.482 | 12219 | -1 | 1 | OSBS2-D-1 | 22-032 | H2O | 1661282943.14 | 0.194 | 0.876 | 60.77 | 0 | 0.004 | 1.511 | 209.896 | 22.739 | 2.224 | 60.816 | 0.857 | -1.4 | 0 | 0 | 1 | 41.938 | 1 | 32 | 1 | GJB_Rinse | 0
150 | P-20917 | 2022/08/23 13:32:54 | 1-32 | 2 | -1.05 | -17.276 | 12214 | -1 | 1 | OSBS2-D-1 | 22-032 | H2O | 1661283479.86 | 0.209 | 0.765 | 60.839 | 0 | 0.002 | 1.515 | 210.92 | 22.153 | 2.231 | 62.782 | 0.856 | -1.352 | 0 | 0 | 1 | 41.938 | 1 | 32 | 1 | GJB_Rinse | 0
151 | P-20917 | 2022/08/23 13:41:50 | 1-32 | 3 | -1.046 | -16.588 | 12188 | -1 | 1 | OSBS2-D-1 | 22-032 | H2O | 1661284015.46 | 0.202 | 0.78 | 64.468 | 0.001 | 0.001 | 1.595 | 210.965 | 22.324 | 2.219 | 61.848 | 0.856 | -1.448 | 0 | 0 | 1 | 41.875 | 1 | 32 | 1 | GJB_Rinse | 0
152 | P-20917 | 2022/08/23 13:50:45 | 1-32 | 4 | -1.019 | -16.008 | 12188 | 0 | 1 | OSBS2-D-1 | 22-032 | H2O | 1661284551.25 | 0.22 | 0.823 | 58.948 | 0 | 0.001 | 1.46 | 210.972 | 22.968 | 2.204 | 60.078 | 0.856 | -1.538 | 0 | 0 | 1 | 41.938 | 1 | 32 | 1 | GJB_Rinse | 0
153 | P-20918 | 2022/08/23 14:00:32 | 1-33 | 1 | -0.921 | -14.673 | 8562 | -1 | 1 | OSBS2-D-4 | 22-032 | H2O | 1661285142.23 | 0.206 | 0.931 | 51.922 | 0 | -0.001 | 1.398 | 110.681 | 15.303 | 1.556 | 47.815 | 0.855 | -0.95 | 0 | 0 | 1 | 41.812 | 1 | 33 | 1 | GJB_Rinse | 0
154 | P-20918 | 2022/08/23 14:09:28 | 1-33 | 2 | -0.901 | -14.061 | 8364 | -1 | 1 | OSBS2-D-4 | 22-032 | H2O | 1661285678.26 | 0.213 | 0.835 | 50.088 | 0.001 | -0.001 | 1.357 | 108.674 | 14.686 | 1.521 | 47.806 | 0.855 | -0.914 | 0 | 0 | 1 | 41.875 | 1 | 33 | 1 | GJB_Rinse | 0
155 | P-20918 | 2022/08/23 14:18:24 | 1-33 | 3 | -0.89 | -14.037 | 9614 | -1 | 1 | OSBS2-D-4 | 22-032 | H2O | 1661286211.97 | 0.206 | 0.864 | 52.406 | 0.001 | -0.002 | 1.353 | 119.127 | 18.139 | 1.744 | 52.485 | 0.855 | -1.022 | 0 | 0 | 1 | 41.875 | 1 | 33 | 1 | GJB_Rinse | 0
156 | P-20918 | 2022/08/23 14:27:20 | 1-33 | 4 | -0.949 | -13.937 | 9534 | 0 | 1 | OSBS2-D-4 | 22-032 | H2O | 1661286748.75 | 0.228 | 0.874 | 52.359 | 0.001 | 0.001 | 1.364 | 118.373 | 18.126 | 1.725 | 51.621 | 0.855 | -1.124 | 0 | 0 | 1 | 41.938 | 1 | 33 | 1 | GJB_Rinse | 0
157 | P-20919 | 2022/08/23 14:37:08 | 1-34 | 1 | -0.733 | -16.118 | 11941 | -1 | 1 | OSBS2-D-5 | 22-032 | H2O | 1661287333.25 | 0.226 | 0.856 | 61.803 | 0.002 | 0 | 1.542 | 171.917 | 22.768 | 2.212 | 60.852 | 0.856 | -1.682 | 0 | 0 | 1 | 42 | 1 | 34 | 1 | GJB_Rinse | 0
158 | P-20919 | 2022/08/23 14:46:03 | 1-34 | 2 | -0.724 | -16.273 | 11840 | -1 | 1 | OSBS2-D-5 | 22-032 | H2O | 1661287869.2 | 0.194 | 0.825 | 63.2 | 0 | 0.003 | 1.573 | 170.964 | 22.41 | 2.202 | 60.751 | 0.856 | -1.501 | 0 | 0 | 1 | 42.062 | 1 | 34 | 1 | GJB_Rinse | 0
159 | P-20919 | 2022/08/23 14:55:00 | 1-34 | 3 | -0.752 | -16.37 | 11922 | -1 | 1 | OSBS2-D-5 | 22-032 | H2O | 1661288405.85 | 0.23 | 0.866 | 61.055 | 0.001 | 0 | 1.508 | 171.838 | 22.234 | 2.224 | 62.098 | 0.856 | -1.421 | 0 | 0 | 1 | 41.938 | 1 | 34 | 1 | GJB_Rinse | 0
160 | P-20919 | 2022/08/23 15:03:56 | 1-34 | 4 | -0.768 | -16.231 | 11812 | 0 | 1 | OSBS2-D-5 | 22-032 | H2O | 1661288942.01 | 0.219 | 0.88 | 62.075 | 0.001 | 0.003 | 1.545 | 170.862 | 22.16 | 2.181 | 60.202 | 0.855 | -1.598 | 0 | 0 | 1 | 41.75 | 1 | 34 | 1 | GJB_Rinse | 0
161 | P-20920 | 2022/08/23 15:13:44 | 1-35 | 1 | -0.626 | -3.913 | 12224 | -1 | 1 | OSBS2-2030-1 | 22-032 | H2O | 1661289529.4 | 0.2 | 0.833 | 60.244 | 0.001 | 0 | 1.5 | 62.323 | 27.725 | 2.176 | 64.259 | 0.856 | -1.071 | 0 | 0 | 1 | 41.75 | 1 | 35 | 1 | GJB_Rinse | 0
162 | P-20920 | 2022/08/23 15:22:40 | 1-35 | 2 | -0.591 | -2.511 | 12090 | -1 | 1 | OSBS2-2030-1 | 22-032 | H2O | 1661290065.92 | 0.208 | 0.815 | 60.174 | 0.001 | -0.002 | 1.495 | 61.371 | 27.593 | 2.149 | 62.93 | 0.856 | -1.139 | 0 | 0 | 1 | 41.875 | 1 | 35 | 1 | GJB_Rinse | 0
163 | P-20920 | 2022/08/23 15:31:36 | 1-35 | 3 | -0.62 | -2.33 | 12246 | -1 | 1 | OSBS2-2030-1 | 22-032 | H2O | 1661290601.77 | 0.217 | 0.789 | 58.915 | 0.002 | -0.003 | 1.469 | 61.404 | 27.592 | 2.193 | 65.627 | 0.855 | -0.919 | 0 | 0 | 1 | 41.812 | 1 | 35 | 1 | GJB_Rinse | 0
164 | P-20920 | 2022/08/23 15:40:33 | 1-35 | 4 | -0.593 | -2.372 | 12129 | 0 | 1 | OSBS2-2030-1 | 22-032 | H2O | 1661291138.23 | 0.208 | 0.831 | 60.202 | 0 | 0.002 | 1.498 | 60.823 | 27.761 | 2.168 | 63.571 | 0.856 | -0.935 | 0 | 0 | 1 | 41.875 | 1 | 35 | 1 | GJB_Rinse | 0
165 | P-20921 | 2022/08/23 15:50:16 | 1-36 | 1 | -0.226 | -4.314 | 11941 | -1 | 1 | OSBS2-2030-2 | 22-032 | H2O | 1661291721.38 | 0.185 | 0.891 | 61.402 | 0 | 0.003 | 1.528 | 60.359 | 27.142 | 2.133 | 63.162 | 0.855 | -0.908 | 0 | 0 | 1 | 41.938 | 1 | 36 | 1 | GJB_Rinse | 0
166 | P-20921 | 2022/08/23 15:59:12 | 1-36 | 2 | -0.202 | -4.485 | 12119 | -1 | 1 | OSBS2-2030-2 | 22-032 | H2O | 1661292258.19 | 0.192 | 0.773 | 59.951 | 0 | 0 | 1.491 | 61.005 | 27.593 | 2.151 | 63.729 | 0.856 | -1.153 | 0 | 0 | 1 | 41.812 | 1 | 36 | 1 | GJB_Rinse | 0
167 | P-20921 | 2022/08/23 16:08:09 | 1-36 | 3 | -0.167 | -4.372 | 11851 | -1 | 1 | OSBS2-2030-2 | 22-032 | H2O | 1661292794.96 | 0.173 | 0.831 | 62.082 | 0.001 | 0.001 | 1.542 | 60.514 | 26.599 | 2.119 | 63.22 | 0.855 | -0.957 | 0 | 0 | 1 | 41.875 | 1 | 36 | 1 | GJB_Rinse | 0
168 | P-20921 | 2022/08/23 16:17:05 | 1-36 | 4 | -0.162 | -4.273 | 12079 | 0 | 1 | OSBS2-2030-2 | 22-032 | H2O | 1661293330.8 | 0.221 | 0.929 | 59.641 | 0.001 | 0 | 1.479 | 60.935 | 27.145 | 2.149 | 63.575 | 0.856 | -0.933 | 0 | 0 | 1 | 41.812 | 1 | 36 | 1 | GJB_Rinse | 0
169 | P-20922 | 2022/08/23 16:26:50 | 1-37 | 1 | -0.543 | -2.639 | 11952 | -1 | 1 | OSBS2-1020-1 | 22-032 | H2O | 1661293915.77 | 0.211 | 0.774 | 61.194 | 0.001 | 0.002 | 1.513 | 60.268 | 27.061 | 2.128 | 63.318 | 0.856 | -0.915 | 0 | 0 | 1 | 41.75 | 1 | 37 | 1 | GJB_Rinse | 0
170 | P-20922 | 2022/08/23 16:35:45 | 1-37 | 2 | -0.515 | -2.566 | 12112 | -1 | 1 | OSBS2-1020-1 | 22-032 | H2O | 1661294451.41 | 0.198 | 0.788 | 62.441 | 0.001 | 0.001 | 1.552 | 60.775 | 27.623 | 2.151 | 63.466 | 0.856 | -1.028 | 0 | 0 | 1 | 41.75 | 1 | 37 | 1 | GJB_Rinse | 0
171 | P-20922 | 2022/08/23 16:44:41 | 1-37 | 3 | -0.522 | -2.788 | 12028 | -1 | 1 | OSBS2-1020-1 | 22-032 | H2O | 1661294986.51 | 0.202 | 0.765 | 60.471 | 0 | 0 | 1.505 | 60.639 | 27.233 | 2.137 | 63.476 | 0.856 | -1.084 | 0 | 0 | 1 | 41.812 | 1 | 37 | 1 | GJB_Rinse | 0
172 | P-20922 | 2022/08/23 16:53:37 | 1-37 | 4 | -0.554 | -2.842 | 12377 | 0 | 1 | OSBS2-1020-1 | 22-032 | H2O | 1661295522.45 | 0.191 | 0.95 | 62.589 | 0 | 0.002 | 1.551 | 61.274 | 28.305 | 2.201 | 64.673 | 0.856 | -1.152 | 0 | 0 | 1 | 41.688 | 1 | 37 | 1 | GJB_Rinse | 0
173 | P-20923 | 2022/08/23 17:03:21 | 1-38 | 1 | -0.636 | -3.473 | 11831 | -1 | 1 | OSBS2-1020-2 | 22-032 | H2O | 1661296106.31 | 0.219 | 0.796 | 58.387 | 0.002 | -0.002 | 1.459 | 60.019 | 26.959 | 2.112 | 62.145 | 0.856 | -1.102 | 0 | 0 | 1 | 41.75 | 1 | 38 | 1 | GJB_Rinse | 0
174 | P-20923 | 2022/08/23 17:12:17 | 1-38 | 2 | -0.691 | -3.646 | 11782 | -1 | 1 | OSBS2-1020-2 | 22-032 | H2O | 1661296642.91 | 0.191 | 0.74 | 58.66 | 0.001 | -0.001 | 1.461 | 60.049 | 26.685 | 2.112 | 62.307 | 0.856 | -1.152 | 0 | 0 | 1 | 41.938 | 1 | 38 | 1 | GJB_Rinse | 0
175 | P-20923 | 2022/08/23 17:21:13 | 1-38 | 3 | -0.646 | -3.667 | 12177 | -1 | 1 | OSBS2-1020-2 | 22-032 | H2O | 1661297178.22 | 0.205 | 0.876 | 62.162 | 0 | 0.006 | 1.554 | 60.885 | 27.51 | 2.166 | 64.131 | 0.856 | -1.093 | 0 | 0 | 1 | 42 | 1 | 38 | 1 | GJB_Rinse | 0
176 | P-20923 | 2022/08/23 17:30:09 | 1-38 | 4 | -0.65 | -3.708 | 11947 | 0 | 1 | OSBS2-1020-2 | 22-032 | H2O | 1661297714.61 | 0.211 | 0.797 | 59.969 | 0.001 | -0.003 | 1.485 | 60.337 | 27.152 | 2.132 | 63.021 | 0.855 | -1.092 | 0 | 0 | 1 | 42 | 1 | 38 | 1 | GJB_Rinse | 0
177 | P-20924 | 2022/08/23 17:39:54 | 1-39 | 1 | -0.994 | -7.859 | 11580 | -1 | 1 | OSBS2-deep-1 | 22-032 | H2O | 1661298299.5 | 0.221 | 0.873 | 60.633 | 0.001 | 0.002 | 1.521 | 59.685 | 25.965 | 2.062 | 61.484 | 0.856 | -1.057 | 0 | 0 | 1 | 41.938 | 1 | 39 | 1 | GJB_Rinse | 0
178 | P-20924 | 2022/08/23 17:48:50 | 1-39 | 2 | -1.039 | -8.289 | 11656 | -1 | 1 | OSBS2-deep-1 | 22-032 | H2O | 1661298835.95 | 0.2 | 0.872 | 65.24 | 0.001 | 0.001 | 1.624 | 59.826 | 26.234 | 2.066 | 61.842 | 0.855 | -1.065 | 0 | 0 | 1 | 41.875 | 1 | 39 | 1 | GJB_Rinse | 0
179 | P-20924 | 2022/08/23 17:57:46 | 1-39 | 3 | -1.078 | -8.203 | 11663 | -1 | 1 | OSBS2-deep-1 | 22-032 | H2O | 1661299371.94 | 0.215 | 0.759 | 60.391 | 0.002 | -0.001 | 1.513 | 59.736 | 26.313 | 2.08 | 61.61 | 0.856 | -1.052 | 0 | 0 | 1 | 41.875 | 1 | 39 | 1 | GJB_Rinse | 0
180 | P-20924 | 2022/08/23 18:06:42 | 1-39 | 4 | -1.102 | -8.464 | 11678 | 0 | 1 | OSBS2-deep-1 | 22-032 | H2O | 1661299907.84 | 0.194 | 0.943 | 58.595 | 0.001 | 0.002 | 1.467 | 59.85 | 25.999 | 2.082 | 62.036 | 0.856 | -0.994 | 0 | 0 | 1 | 41.812 | 1 | 39 | 1 | GJB_Rinse | 0
181 | P-20925 | 2022/08/23 18:16:26 | 1-40 | 1 | -8.604 | -58.517 | 12140 | -1 | 1 | EV | standard | H2O | 1661300491.25 | 0.191 | 0.875 | 62.693 | 0.002 | 0 | 1.561 | 60.637 | 27.572 | 2.088 | 61.764 | 0.856 | -1.009 | 0 | 0 | 1 | 41.875 | 1 | 40 | 1 | GJB_Rinse | 0
182 | P-20925 | 2022/08/23 18:25:21 | 1-40 | 2 | -8.871 | -63.039 | 11980 | -1 | 1 | EV | standard | H2O | 1661301027.48 | 0.212 | 0.711 | 54.821 | 0.002 | 0.002 | 1.364 | 60.223 | 27.105 | 2.047 | 61.412 | 0.855 | -0.962 | 0 | 0 | 1 | 41.938 | 1 | 40 | 1 | GJB_Rinse | 0
183 | P-20925 | 2022/08/23 18:34:18 | 1-40 | 3 | -9.003 | -64.643 | 11954 | -1 | 1 | EV | standard | H2O | 1661301563.97 | 0.212 | 0.826 | 60.496 | 0.001 | 0.005 | 1.503 | 60.123 | 26.924 | 2.06 | 62.42 | 0.856 | -0.888 | 0 | 0 | 1 | 42 | 1 | 40 | 1 | GJB_Rinse | 0
184 | P-20925 | 2022/08/23 18:43:14 | 1-40 | 4 | -9.017 | -65.15 | 12034 | 0 | 1 | EV | standard | H2O | 1661302099.87 | 0.204 | 0.833 | 64.9 | 0.002 | 0.007 | 1.617 | 60.392 | 26.989 | 2.064 | 62.599 | 0.856 | -0.958 | 0 | 0 | 1 | 42 | 1 | 40 | 1 | GJB_Rinse | 0
185 | P-20926 | 2022/08/23 18:52:58 | 1-41 | 1 | -1.579 | -13.443 | 11936 | -1 | 1 | OSBS2-deep-2 | 22-032 | H2O | 1661302684.41 | 0.211 | 0.818 | 58.198 | 0.001 | 0.002 | 1.459 | 60.281 | 27.082 | 2.105 | 62.38 | 0.856 | -1.103 | 0 | 0 | 1 | 42 | 1 | 41 | 1 | GJB_Rinse | 0
186 | P-20926 | 2022/08/23 19:01:54 | 1-41 | 2 | -1.285 | -8.82 | 12012 | -1 | 1 | OSBS2-deep-2 | 22-032 | H2O | 1661303220.12 | 0.2 | 0.697 | 62.374 | 0 | -0.001 | 1.549 | 60.441 | 27.295 | 2.123 | 62.868 | 0.856 | -1.104 | 0 | 0 | 1 | 42.062 | 1 | 41 | 1 | GJB_Rinse | 0
187 | P-20926 | 2022/08/23 19:10:51 | 1-41 | 3 | -1.244 | -7.393 | 11922 | -1 | 1 | OSBS2-deep-2 | 22-032 | H2O | 1661303756.81 | 0.218 | 0.844 | 59.576 | 0.001 | -0.002 | 1.483 | 60.24 | 26.889 | 2.11 | 62.2 | 0.856 | -1.077 | 0 | 0 | 1 | 42.062 | 1 | 41 | 1 | GJB_Rinse | 0
188 | P-20926 | 2022/08/23 19:19:47 | 1-41 | 4 | -1.149 | -6.919 | 11760 | 0 | 1 | OSBS2-deep-2 | 22-032 | H2O | 1661304292.67 | 0.218 | 0.804 | 61.757 | 0.001 | -0.001 | 1.538 | 59.771 | 26.831 | 2.094 | 61.819 | 0.856 | -0.979 | 0 | 0 | 1 | 42.062 | 1 | 41 | 1 | GJB_Rinse | 0
189 | P-20927 | 2022/08/23 19:29:33 | 1-42 | 1 | -0.405 | -3.541 | 12325 | -1 | 1 | OSBS2-3040-1 | 22-032 | H2O | 1661304878.39 | 0.185 | 0.83 | 60.789 | 0.001 | 0.002 | 1.509 | 60.989 | 27.896 | 2.192 | 65.004 | 0.856 | -1.061 | 0 | 0 | 1 | 42 | 1 | 42 | 1 | GJB_Rinse | 0
190 | P-20927 | 2022/08/23 19:38:28 | 1-42 | 2 | -0.317 | -2.886 | 12165 | -1 | 1 | OSBS2-3040-1 | 22-032 | H2O | 1661305414.22 | 0.209 | 0.786 | 62.298 | 0.001 | 0 | 1.536 | 60.683 | 27.627 | 2.161 | 64.03 | 0.856 | -1.117 | 0 | 0 | 1 | 42.188 | 1 | 42 | 1 | GJB_Rinse | 0
191 | P-20927 | 2022/08/23 19:47:24 | 1-42 | 3 | -0.358 | -2.925 | 12099 | -1 | 1 | OSBS2-3040-1 | 22-032 | H2O | 1661305949.24 | 0.191 | 0.837 | 60.165 | 0.001 | -0.001 | 1.493 | 60.476 | 27.801 | 2.157 | 62.496 | 0.856 | -1.121 | 0 | 0 | 1 | 42.188 | 1 | 42 | 1 | GJB_Rinse | 0
192 | P-20927 | 2022/08/23 19:56:19 | 1-42 | 4 | -0.297 | -2.365 | 12264 | 0 | 1 | OSBS2-3040-1 | 22-032 | H2O | 1661306485.06 | 0.208 | 0.843 | 64.957 | 0.001 | 0.001 | 1.604 | 60.945 | 27.84 | 2.195 | 64.622 | 0.856 | -1.176 | 0 | 0 | 1 | 42.062 | 1 | 42 | 1 | GJB_Rinse | 0
193 | P-20928 | 2022/08/23 20:06:03 | 1-43 | 1 | -1.148 | -5.045 | 11676 | -1 | 1 | OSBS2-3040-2 | 22-032 | H2O | 1661307068.72 | 0.223 | 0.932 | 58.514 | 0.001 | -0.002 | 1.465 | 59.873 | 26.026 | 2.077 | 61.762 | 0.855 | -1.176 | 0 | 0 | 1 | 42.188 | 1 | 43 | 1 | GJB_Rinse | 0
194 | P-20928 | 2022/08/23 20:14:58 | 1-43 | 2 | -1.143 | -4.928 | 11828 | -1 | 1 | OSBS2-3040-2 | 22-032 | H2O | 1661307604.09 | 0.189 | 0.801 | 62.995 | 0.001 | 0.001 | 1.57 | 59.883 | 26.816 | 2.109 | 63.09 | 0.855 | -0.988 | 0 | 0 | 1 | 42.188 | 1 | 43 | 1 | GJB_Rinse | 0
195 | P-20928 | 2022/08/23 20:23:53 | 1-43 | 3 | -1.205 | -5.093 | 12006 | -1 | 1 | OSBS2-3040-2 | 22-032 | H2O | 1661308138.89 | 0.195 | 0.827 | 61.141 | 0.001 | 0.001 | 1.516 | 60.276 | 27.38 | 2.137 | 63.311 | 0.856 | -1.014 | 0 | 0 | 1 | 42.125 | 1 | 43 | 1 | GJB_Rinse | 0
196 | P-20928 | 2022/08/23 20:32:49 | 1-43 | 4 | -1.146 | -5.046 | 11842 | 0 | 1 | OSBS2-3040-2 | 22-032 | H2O | 1661308674.31 | 0.212 | 0.828 | 61.826 | 0.001 | 0.002 | 1.535 | 60.09 | 26.736 | 2.123 | 62.202 | 0.856 | -1.043 | 0 | 0 | 1 | 42.125 | 1 | 43 | 1 | GJB_Rinse | 0
197 | P-20929 | 2022/08/23 20:42:34 | 1-44 | 1 | -8.61 | -58.175 | 12189 | -1 | 1 | EV | standard | H2O | 1661309259.79 | 0.234 | 0.744 | 62.387 | 0.002 | -0.003 | 1.547 | 60.753 | 27.546 | 2.09 | 62.683 | 0.856 | -1.028 | 0 | 0 | 1 | 42 | 1 | 44 | 1 | GJB_Rinse | 0
198 | P-20929 | 2022/08/23 20:51:29 | 1-44 | 2 | -8.931 | -63.406 | 12304 | -1 | 1 | EV | standard | H2O | 1661309795.26 | 0.206 | 0.857 | 64.658 | 0.001 | 0.008 | 1.601 | 60.997 | 27.884 | 2.097 | 62.902 | 0.856 | -0.96 | 0 | 0 | 1 | 42.188 | 1 | 44 | 1 | GJB_Rinse | 0
199 | P-20929 | 2022/08/23 21:00:25 | 1-44 | 3 | -8.963 | -64.429 | 12268 | -1 | 1 | EV | standard | H2O | 1661310331.12 | 0.215 | 0.808 | 61.995 | 0.001 | 0.006 | 1.534 | 60.925 | 27.93 | 2.089 | 62.175 | 0.856 | -1.065 | 0 | 0 | 1 | 42.312 | 1 | 44 | 1 | GJB_Rinse | 0
200 | P-20929 | 2022/08/23 21:09:21 | 1-44 | 4 | -8.994 | -64.986 | 12269 | 0 | 1 | EV | standard | H2O | 1661310866.28 | 0.21 | 0.75 | 61.954 | 0.001 | 0.004 | 1.539 | 60.797 | 27.765 | 2.096 | 63.61 | 0.856 | -0.86 | 0 | 0 | 1 | 42.312 | 1 | 44 | 1 | GJB_Rinse | 0

## Instructions
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Variable | Instruction | Comment
StandardsFile | standards file.csv
OutputBase | ChemCorrect
CorrectionType | M
timecode | Time Code
identifier | Identifier 1
sample | Sample
injection | Inj Nr | Inj Nr
line | Line
job | Job
Variable | Instruction | Comment
baseline_shift_MEAN | mean_standards(baseline_shift) | The mean of the organic baseline of the standards.
baseline_shift_STD | std_standards(baseline_shift) | The standard deviation of the organic baseline of the standards.
ch4_ppm_MEAN | mean_standards(ch4_ppm) | The mean methane concentration of the standards.
ch4_ppm_STD | std_standards(ch4_ppm) | The standard deviation of the methane concentrations of the standards.
residuals_MEAN | mean_standards(residuals) | The mean difference between the measured and modeled spectra of the standards.
residuals_STD | std_standards(residuals) | The standard deviation of the difference between measured and modeled spectra of the standards.
slope_shift_MEAN | mean_standards(slope_shift) | The mean organic slope of the standards.
slope_shift_STD | std_standards(slope_shift) | The standard deviation of the organic slope of the standards.
baseline_curvature_MEAN | mean_standards(baseline_curvature) | The mean of the baseline curvature of the standards.
baseline_curvature_STD | std_standards(baseline_curvature) | The standard deviation of the baseline curvature of the standards.
residuals_MEAN_ABS | abs(residuals_MEAN) | The absolute value of the mean difference between the measured and modeled spectra of the standards.
baseline_shift_MEAN_ABS | abs(baseline_shift_MEAN) | The absolute value of the mean of the organic baseline of the standards.
ch4_ppm_MEAN_ABS | abs(ch4_ppm_MEAN) | The absolute value of the mean methane concentration of the standards.
slope_shift_MEAN_ABS | abs(slope_shift_MEAN) | The absolute value of the mean organic slope of the standards.
residuals_15SIGMA | evaluation(residuals_STD*1.5) | One and half times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_15SIGMA | evaluation(residuals_MEAN+residuals_15SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_2SIGMA | evaluation(residuals_STD*2) | Twice the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_2SIGMA | evaluation(residuals_MEAN+residuals_2SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_3SIGMA | evaluation(residuals_STD*3) | Three times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_3SIGMA | evaluation(residuals_MEAN+residuals_3SIGMA) | The mean difference between measured and modeled spectra of the standards plus three standard deviation.
residuals_10PCT | evaluation(residuals_MEAN*0.1) | 30% of the difference between measured and modeled spectra of the standards.
ch4_ppm_3SIGMA | evaluation(ch4_ppm_STD*3) | Three times the standard deviation of the methanol amplitudes of the standards.
ch4_ppm_PLUS_3SIGMA | evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA) | The mean methanol amplitude of the standards plus three standard deviation.
HALF | float(0.5) | A fixed value of 0.5 or half.
baseline_shift_HALFSIGMA | evaluation(baseline_shift_STD*HALF) | Half the standard deviation of the organic baseline of the standards.
baseline_shift_PLUS_HALFSIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA) | The absolute value of the mean organic baseline of the standards plus one-half standard deviation.
baseline_shift_PLUS_1SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD) | The absolute value of the mean organic baseline of the standards plus one standard deviation.
baseline_shift_3SIGMA | evaluation(baseline_shift_STD*3) | Three times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_3SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA) | The absolute value of the mean organic baseline of the standards plus three standard deviation.
baseline_shift_18SIGMA | evaluation(baseline_shift_STD*18) | Eighteen times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_18SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA) | The absolute value of the mean organic baseline of the standards plus eighteen standard deviation.
slope_shift_5SIGMA | evaluation(slope_shift_STD*5) | Five times the standard deviation of the organic slope of the standards.
slope_shift_PLUS_5SIGMA | evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA) | The absolute value of the mean organic slope of the standards plus five standard deviation.
baseline_curvature_3SIGMA | evaluation(baseline_curvature_STD*3) | Three times the standard deviation of the baseline curvature of the standards.
baseline_curvature_PLUS_3SIGMA | evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
TRUE | evaluation(1>0)
FALSE | evaluation(1<0)
PROC_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the uncalibrated d18O value in order to calibrate it.
PROC_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the uncalibrated dD value in order to calibrate it.
Variable | Instruction | Comment
IS_STANDARD | is_standard_sample() | Standard recognition.
IS_NOT_STANDARD | bool_xor(IS_STANDARD,TRUE)
SAMP_residuals_MEAN | mean_all(residuals) | The mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_residuals_STD | std_all(residuals) | One standard deviation of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN | mean_all(ch4_ppm) | The mean methane concentration of a sample.
SAMP_ch4_ppm_STD | std_all(ch4_ppm) | The standard deviation of the methane concentrations of a sample.
SAMP_baseline_shift_MEAN | mean_all(baseline_shift) | The mean of the organic baseline of a sample.
SAMP_baseline_shift_STD | std_all(baseline_shift) | The standard deviation of the organic baseline of a sample.
SAMP_slope_shift_MEAN | mean_all(slope_shift) | The mean organic slope of a sample.
SAMP_slope_shift_STD | std_all(slope_shift) | The standard deviation of the organic slope of a sample.
SAMP_d18O_STD | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
SAMP_dD_STD | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_baseline_curvature_MEAN | mean_all(baseline_curvature) | The mean of the baseline curvature of the ssamples
SAMP_baseline_curvature_STD | std_all(baseline_curvature) | The standard deviation of the baseline curvature of the samples
SAMP_residuals_MEAN_ABS | abs(SAMP_residuals_MEAN) | Absolute value of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN_ABS | abs(SAMP_ch4_ppm_MEAN) | Absolute value of the mean methane concentration of a sample.
SAMP_baseline_shift_MEAN_ABS | abs(SAMP_baseline_shift_MEAN) | Absolute value of the mean organic baseline of a sample.
SAMP_slope_shift_MEAN_ABS | abs(SAMP_slope_shift_MEAN) | Absolute value of the mean organic slope of a sample.
SAMP_residuals_DIFFERENTIAL | evaluation(SAMP_residuals_MEAN-residuals_MEAN)
SAMP_residuals_DIFFERENTIAL_ABS | abs(SAMP_residuals_DIFFERENTIAL) | Absolute value of the difference between the mean of the sample residual and the mean of all standards residuals
SAMP_residuals_DIFFERENTIAL_PCT | evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
SAMP_baseline_curvature_DIFFERENTIAL | evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
SAMP_STD_DELTA_18_16O_CURR_MEAN | current_standards_mean(d(18_16)Mean) | The d18O mean of the last run standard.
SAMP_STD_DELTA_DH_CURR_MEAN | current_standards_mean(d(D_H)Mean) | The dD mean of the last run standard.
known_dD | known_standard(d(D_H)Mean) | The dD value of the known Standard.
known_d18O | known_standard(d(18_16)Mean) | The d18O value of the known Standard.
d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
measured_d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
measured_dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
measured_d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
measured_dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the mean uncalibrated d18O value in order to calibrate it.
SAMP_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the mean uncalibrated dD value in order to calibrate it.
SAMP_TEST_VALUE | evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
residuals_GT_3SIGMA | evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA) | Threshold of red trigger. If the mean difference between the measured spectra of the sample and the modeled spectra of the standard is greater than that of the mean difference of the standard plus three standard deviations, this will trigger.
residuals_GE_2SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GE_15SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GT_10PCT | evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
ch4_ppm_GE_3SIGMA | evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA) | True if the mean methane concentration of the sample is greater than or equal to that of the standard plus three standard deviations, otherwise false.
baseline_shift_GT_1SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus one standard deviation, otherwise false.
baseline_shift_GT_3SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus three standard deviation, otherwise false. Sets yellow trigger for "other", trigger threshold is approximately sample to sample dD precision specification
baseline_shift_GT_18SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus eighteen standard deviation, otherwise false. Sets red trigger for "other", trigger threshold is approximately 24 hour dD drift specification
slope_shift_GT_5SIGMA | evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA) | True if the mean organic slope of the sample is greater than that of the standard plus five standard deviations, otherwise false.
baseline_curvature_GT_3SIGMA | evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
RED | truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA)) | Red trigger. If the sample is not a standard and the threshold is triggered (measured-modeled > standard+3stdev's), then RED.
NOT_RED | bool_xor(RED,TRUE) | If the sample is determined to be "red" than it can not be "not red."
YELLOW | truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA))) | Truth evaluation used to determine whether or not a given sample is yellow.
NOT_YELLOW | bool_xor(YELLOW,TRUE) | If sample is determined to be "yellow" than it can not be "not yellow."
GREEN | truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW)) | This is the pre-requisite for a sample to be colored green. It must not be a standard, must not be yellow, and must not be red.
NOT_GREEN | bool_xor(GREEN,TRUE) | Pre-requisite for being "not-green."
CYAN | truth_evaluation((IS_STANDARD)) | This is the pre-requisite for a sample to be a standard and therefore cyan colored.
NOT_CYAN | bool_xor(CYAN,TRUE) | Pre-requisite for being not a standard.
METHANE | truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA)) | This is the threshold for amounts of methane being included in the main ChemCorrect spreadsheet. Methane will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured methane concentration must be greater than or equal to that of the standard plus three standard deviations.
OTHER1 | truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA)) | This is the threshold for amounts of "OTHER1" being included in the main ChemCorrect spreadsheet. This column (OTHER1) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic baseline must be greater than that of the standard plus one standard deviation.
OTHER2 | truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA)) | This is the threshold for amounts of "OTHER2" being included in the main ChemCorrect spreadsheet. This column (OTHER2) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic slope must be greater than that of the standard plus three standard deviations.
OTHER3 | truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | truth_evaluation(RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN)) | Organic baseline differential showing the difference between the measured baseline and the standard baseline.
Variable | Instruction | Comment
DTL_TEST | math_evaluation((1 + 2)) | Test evaluation of 1+2 being equal to 3.
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample | FALSE | FALSE | FALSE | FALSE | Testing Sample
B | identifier | TRUE | Name | RED | YELLOW | GREEN | CYAN | Testing Name
C | SAMP_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d<sup>18</sup>O Mean | RED | YELLOW | GREEN | CYAN | Testing18 16 Mean
D | SAMP_DELTA_DH_CALIBRATE | TRUE | Calibrated d<sup>2</sup>H Mean | RED | YELLOW | GREEN | CYAN | This is a test to see that I can use tool tips
E | SAMP_ch4_ppm_MEAN | METHANE | CH<sub>4</sub> | RED | YELLOW | FALSE | FALSE
F | baseline_shift_DIFF | OTHER1 | C<sub>2+</sub> alcohols | RED | YELLOW | FALSE | FALSE
G | SAMP_residuals_DIFFERENTIAL_PCT | OTHER3 | relative %deviation | RED | YELLOW | FALSE | FALSE
H | SAMP_d18O_STD | TRUE | Uncalibrated d<sup>18</sup>O precision | RED | YELLOW | GREEN | CYAN
I | SAMP_dD_STD | TRUE | Uncalibrated d<sup>2</sup>H precision | RED | YELLOW | GREEN | CYAN
J | SAMP_slope_shift_MEAN | OTHER2 | slope | RED | YELLOW | FALSE | FALSE
K | SAMP_baseline_curvature_DIFFERENTIAL | OTHER4 | curvature | RED | YELLOW | FALSE | FALSE
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample
B | identifier | TRUE | Name
C | col(d(18_16)Mean) | TRUE | Measured d(18_16) Mean
D | col(d(D_H)Mean) | TRUE | Measured d(D_H) Mean
E | DTL_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d(18_16) Mean
F | DTL_DELTA_DH_CALIBRATE | TRUE | Calibrated d(D_H) Mean
G | col(ch4_ppm) | METHANE | CH<sub>4</sub>
H | col(baseline_shift) | OTHER1 | Other
I | col(Time Code) | TRUE | Time Code
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | d18O_mean | TRUE | d(18_16) Mean | RED | YELLOW | GREEN | CYAN
B | dD_mean | TRUE | d(D_H) Mean | RED | YELLOW | GREEN | CYAN
C | XY_X(d18O_mean) XY_Y(dD_mean) | TRUE | d18O x dD | gmwl | best_fit

## Process Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Process Variables
Variable | Value | Operations
baseline_shift_MEAN | 61.27735
varray: [61.899, 61.697, 62.536, 62.587, 62.348, 62.22, 61.531, 61.349, 61.23, 61.391, 60.552, 61.414, 60.689, 60.509, 60.821, 60.663, 60.777, 60.235, 60.411, 60.688]
inst: mean_standards
operation: mean_standards(baseline_shift)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0, 10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: baseline_shift
baseline_shift_STD | 0.742796726107
varray: [61.899, 61.697, 62.536, 62.587, 62.348, 62.22, 61.531, 61.349, 61.23, 61.391, 60.552, 61.414, 60.689, 60.509, 60.821, 60.663, 60.777, 60.235, 60.411, 60.688]
inst: std_standards
operation: std_standards(baseline_shift)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0, 10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: baseline_shift
ch4_ppm_MEAN | -0.99895
varray: [-1.154, -1.029, -1.169, -1.341, -1.149, -1.125, -1.165, -0.949, -0.96, -1.123, -0.909, -1.034, -0.773, -1.036, -0.888, -0.848, -0.843, -0.762, -0.722, -1.0]
inst: mean_standards
operation: mean_standards(ch4_ppm)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0, 10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: ch4_ppm
ch4_ppm_STD | 0.163874944297
varray: [-1.154, -1.029, -1.169, -1.341, -1.149, -1.125, -1.165, -0.949, -0.96, -1.123, -0.909, -1.034, -0.773, -1.036, -0.888, -0.848, -0.843, -0.762, -0.722, -1.0]
inst: std_standards
operation: std_standards(ch4_ppm)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0, 10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: ch4_ppm
residuals_MEAN | 2.14475
varray: [2.267, 2.268, 2.312, 2.304, 2.324, 2.321, 2.263, 2.263, 2.252, 2.242, 2.035, 2.059, 2.008, 1.98, 2.005, 1.996, 2.014, 1.994, 1.999, 1.989]
inst: mean_standards
operation: mean_standards(residuals)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0, 10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: residuals
residuals_STD | 0.142867209232
varray: [2.267, 2.268, 2.312, 2.304, 2.324, 2.321, 2.263, 2.263, 2.252, 2.242, 2.035, 2.059, 2.008, 1.98, 2.005, 1.996, 2.014, 1.994, 1.999, 1.989]
inst: std_standards
operation: std_standards(residuals)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0, 10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: residuals
slope_shift_MEAN | 28.21085
varray: [28.798, 28.738, 29.725, 29.862, 29.537, 29.754, 28.663, 28.795, 28.524, 28.074, 27.355, 28.585, 27.156, 27.214, 27.472, 27.042, 27.198, 26.928, 27.574, 27.223]
inst: mean_standards
operation: mean_standards(slope_shift)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0, 10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: slope_shift
slope_shift_STD | 1.01015935559
varray: [28.798, 28.738, 29.725, 29.862, 29.537, 29.754, 28.663, 28.795, 28.524, 28.074, 27.355, 28.585, 27.156, 27.214, 27.472, 27.042, 27.198, 26.928, 27.574, 27.223]
inst: std_standards
operation: std_standards(slope_shift)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0, 10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: slope_shift
baseline_curvature_MEAN | 63.8114
varray: [66.525, 66.833, 66.938, 66.266, 67.807, 66.697, 65.348, 66.175, 65.601, 65.68, 61.332, 61.436, 61.874, 60.129, 61.451, 61.062, 62.152, 61.034, 61.088, 60.8]
inst: mean_standards
operation: mean_standards(baseline_curvature)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0, 10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: baseline_curvature
baseline_curvature_STD | 2.71850567311
varray: [66.525, 66.833, 66.938, 66.266, 67.807, 66.697, 65.348, 66.175, 65.601, 65.68, 61.332, 61.436, 61.874, 60.129, 61.451, 61.062, 62.152, 61.034, 61.088, 60.8]
inst: std_standards
operation: std_standards(baseline_curvature)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0, 10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: baseline_curvature
residuals_MEAN_ABS | 2.14475
operation: abs(residuals_MEAN)
inst: abs
source_col: residuals_MEAN
baseline_shift_MEAN_ABS | 61.27735
operation: abs(baseline_shift_MEAN)
inst: abs
source_col: baseline_shift_MEAN
ch4_ppm_MEAN_ABS | 0.99895
operation: abs(ch4_ppm_MEAN)
inst: abs
source_col: ch4_ppm_MEAN
slope_shift_MEAN_ABS | 28.21085
operation: abs(slope_shift_MEAN)
inst: abs
source_col: slope_shift_MEAN
residuals_15SIGMA | 0.214300813848
operation: evaluation(residuals_STD*1.5)
inst: evaluation
factors: residuals_STD*1.5
residuals_PLUS_15SIGMA | 2.35905081385
operation: evaluation(residuals_MEAN+residuals_15SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_15SIGMA
residuals_2SIGMA | 0.285734418464
operation: evaluation(residuals_STD*2)
inst: evaluation
factors: residuals_STD*2
residuals_PLUS_2SIGMA | 2.43048441846
operation: evaluation(residuals_MEAN+residuals_2SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_2SIGMA
residuals_3SIGMA | 0.428601627695
operation: evaluation(residuals_STD*3)
inst: evaluation
factors: residuals_STD*3
residuals_PLUS_3SIGMA | 2.5733516277
operation: evaluation(residuals_MEAN+residuals_3SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_3SIGMA
residuals_10PCT | 0.214475
operation: evaluation(residuals_MEAN*0.1)
inst: evaluation
factors: residuals_MEAN*0.1
ch4_ppm_3SIGMA | 0.491624832892
operation: evaluation(ch4_ppm_STD*3)
inst: evaluation
factors: ch4_ppm_STD*3
ch4_ppm_PLUS_3SIGMA | 1.49057483289
operation: evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA)
inst: evaluation
factors: ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA
HALF | 0.5
operation: float(0.5)
inst: float
baseline_shift_HALFSIGMA | 0.371398363054
operation: evaluation(baseline_shift_STD*HALF)
inst: evaluation
factors: baseline_shift_STD*HALF
baseline_shift_PLUS_HALFSIGMA | 61.6487483631
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA
baseline_shift_PLUS_1SIGMA | 62.0201467261
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_STD
baseline_shift_3SIGMA | 2.22839017832
operation: evaluation(baseline_shift_STD*3)
inst: evaluation
factors: baseline_shift_STD*3
baseline_shift_PLUS_3SIGMA | 63.5057401783
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_3SIGMA
baseline_shift_18SIGMA | 13.3703410699
operation: evaluation(baseline_shift_STD*18)
inst: evaluation
factors: baseline_shift_STD*18
baseline_shift_PLUS_18SIGMA | 74.6476910699
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_18SIGMA
slope_shift_5SIGMA | 5.05079677795
operation: evaluation(slope_shift_STD*5)
inst: evaluation
factors: slope_shift_STD*5
slope_shift_PLUS_5SIGMA | 33.2616467779
operation: evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA)
inst: evaluation
factors: slope_shift_MEAN_ABS+slope_shift_5SIGMA
baseline_curvature_3SIGMA | 8.15551701933
operation: evaluation(baseline_curvature_STD*3)
inst: evaluation
factors: baseline_curvature_STD*3
baseline_curvature_PLUS_3SIGMA | 71.9669170193
operation: evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
inst: evaluation
factors: baseline_curvature_3SIGMA+baseline_curvature_MEAN
TRUE | True
operation: evaluation(1>0)
inst: evaluation
factors: 1>0
FALSE | False
operation: evaluation(1<0)
inst: evaluation
factors: 1<0
PROC_DELTA_18_16O_CALIBRATE | -3.64250772963
actual_array: [ 1.93 -15.84]
varray: [-10.178, -9.394, -9.175, -9.137, -9.117, -9.099, -9.135, -9.142, -9.108, -9.113, 2.202, 2.573, 2.675, 2.758, 2.763, 2.82, 2.831, 2.807, 2.877, 2.872, -13.758, -14.383, -14.504, -14.623, -14.595, -14.656, -14.68, -14.679, -14.721, -14.741, -9.463, -9.218, -9.203, -9.149, -9.176, -9.185, -9.12, -9.146, -9.123, -9.11, -1.501, -1.329, -1.288, -1.188, 0.481, 0.51, 0.51, 0.565, -0.197, -0.273, -0.279, -0.258, -0.287, -0.231, -0.265, -0.241, -1.507, -1.578, -1.561, -1.604, -1.914, -1.963, -1.927, -1.959, -2.39, -2.502, -2.483, -2.515, 1.695, 1.905, 1.939, 1.93, -8.48, -8.875, -8.922, -9.028, 0.022, 0.288, 0.388, 0.476, -0.197, -0.206, -0.214, -0.151, 2.912, 2.999, 3.075, 3.056, -0.1, -0.252, -0.255, -0.31, -1.002, -1.001, -1.042, -1.015, -0.89, -0.92, -0.903, -0.912, -0.872, -0.919, -0.864, -0.926, -0.204, -0.145, -0.151, -0.17, -8.622, -8.875, -8.98, -9.002, -1.779, -1.51, -1.426, -1.424, -2.275, -2.326, -2.35, -2.302, 1.02, 1.169, 1.168, 1.171, 1.112, 1.139, 1.142, 1.154, 0.0, -0.052, -0.077, -0.105, 0.529, 0.479, 0.498, 0.47, 0.708, 0.756, 0.764, 0.758, -0.087, -0.102, -0.075, -0.047, -8.59, -8.856, -8.943, -8.981, -1.445, -1.05, -1.046, -1.019, -0.921, -0.901, -0.89, -0.949, -0.733, -0.724, -0.752, -0.768, -0.626, -0.591, -0.62, -0.593, -0.226, -0.202, -0.167, -0.162, -0.543, -0.515, -0.522, -0.554, -0.636, -0.691, -0.646, -0.65, -0.994, -1.039, -1.078, -1.102, -8.604, -8.871, -9.003, -9.017, -1.579, -1.285, -1.244, -1.149, -0.405, -0.317, -0.358, -0.297, -1.148, -1.143, -1.205, -1.146, -8.61, -8.931, -8.963, -8.994]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -2.692205
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [1.0, 537.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3757.0, 4293.0, 4830.0, 5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0, 10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0, 16261.0, 16799.0, 17335.0, 17872.0, 18409.0, 18947.0, 19484.0, 20020.0, 20557.0, 21095.0, 21684.0, 22222.0, 22758.0, 23296.0, 23886.0, 24423.0, 24960.0, 25496.0, 26086.0, 26623.0, 27159.0, 27697.0, 28284.0, 28821.0, 29358.0, 29895.0, 30483.0, 31020.0, 31557.0, 32093.0, 32681.0, 33218.0, 33754.0, 34289.0, 34878.0, 35414.0, 35950.0, 36486.0, 37075.0, 37612.0, 38149.0, 38685.0, 39274.0, 39811.0, 40348.0, 40884.0, 41474.0, 42010.0, 42547.0, 43083.0, 43671.0, 44207.0, 44745.0, 45281.0, 45870.0, 46406.0, 46943.0, 47479.0, 48067.0, 48604.0, 49140.0, 49677.0, 50264.0, 50801.0, 51338.0, 51875.0, 52464.0, 53001.0, 53537.0, 54073.0, 54662.0, 55198.0, 55735.0, 56272.0, 56861.0, 57397.0, 57934.0, 58470.0, 59056.0, 59591.0, 60127.0, 60662.0, 61248.0, 61785.0, 62322.0, 62858.0, 63445.0, 63981.0, 64518.0, 65054.0, 65642.0, 66177.0, 66713.0, 67249.0, 67837.0, 68373.0, 68909.0, 69446.0, 70034.0, 70570.0, 71106.0, 71642.0, 72230.0, 72766.0, 73302.0, 73839.0, 74422.0, 74959.0, 75495.0, 76030.0, 76616.0, 77152.0, 77687.0, 78222.0, 78808.0, 79344.0, 79880.0, 80416.0, 81002.0, 81538.0, 82074.0, 82609.0, 83196.0, 83732.0, 84268.0, 84804.0, 85392.0, 85927.0, 86464.0, 87000.0, 87588.0, 88124.0, 88660.0, 89197.0, 89780.0, 90316.0, 90853.0, 91389.0, 91974.0, 92509.0, 93045.0, 93581.0, 94165.0, 94701.0, 95237.0, 95773.0, 96358.0, 96894.0, 97430.0, 97966.0, 98550.0, 99085.0, 99622.0, 100158.0, 100742.0, 101278.0, 101815.0, 102351.0, 102937.0, 103472.0, 104008.0, 104543.0, 105127.0, 105662.0, 106197.0, 106733.0, 107318.0, 107853.0, 108389.0, 108925.0]
source_col: d(18_16)Mean
PROC_DELTA_DH_CALIBRATE | -30.0076953569
actual_array: [ 18.1 -119. ]
varray: [-76.906, -70.711, -68.643, -68.221, -68.061, -67.621, -67.659, -67.757, -67.325, -67.465, 8.397, 15.173, 16.9, 17.787, 18.49, 18.808, 19.204, 19.363, 19.728, 19.819, -94.826, -105.397, -107.942, -109.355, -110.043, -110.713, -110.903, -111.333, -111.644, -111.643, -72.588, -69.083, -68.329, -67.776, -67.679, -67.517, -67.375, -67.526, -67.242, -67.294, -24.141, -20.517, -19.645, -19.112, -8.309, -7.675, -7.193, -6.505, -12.022, -12.911, -12.856, -12.661, -15.399, -15.417, -15.45, -15.495, -22.072, -22.982, -23.338, -23.366, -22.621, -22.53, -22.21, -22.396, -29.411, -29.924, -30.056, -30.155, -5.64, -3.408, -2.909, -2.675, -58.554, -63.55, -64.861, -65.426, -11.737, -7.239, -5.886, -5.027, -8.665, -8.876, -8.922, -8.929, -9.563, -9.316, -9.358, -9.299, -21.889, -23.393, -23.831, -24.089, -20.816, -20.42, -20.108, -20.357, -19.84, -20.085, -19.898, -20.067, -22.0, -22.397, -22.381, -22.617, -14.371, -13.5, -13.255, -13.237, -59.422, -63.799, -64.874, -65.328, -22.247, -18.58, -17.764, -17.367, -25.754, -26.673, -26.66, -26.707, -15.089, -14.069, -13.77, -13.825, -10.958, -10.385, -10.561, -10.108, -14.443, -14.696, -14.793, -14.698, -13.595, -14.383, -14.088, -13.789, -4.015, -2.771, -2.422, -2.123, -3.269, -3.172, -3.204, -3.034, -58.34, -62.942, -64.408, -65.093, -21.482, -17.276, -16.588, -16.008, -14.673, -14.061, -14.037, -13.937, -16.118, -16.273, -16.37, -16.231, -3.913, -2.511, -2.33, -2.372, -4.314, -4.485, -4.372, -4.273, -2.639, -2.566, -2.788, -2.842, -3.473, -3.646, -3.667, -3.708, -7.859, -8.289, -8.203, -8.464, -58.517, -63.039, -64.643, -65.15, -13.443, -8.82, -7.393, -6.919, -3.541, -2.886, -2.925, -2.365, -5.045, -4.928, -5.093, -5.046, -58.175, -63.406, -64.429, -64.986]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -26.757015
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [1.0, 537.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3757.0, 4293.0, 4830.0, 5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0, 10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0, 16261.0, 16799.0, 17335.0, 17872.0, 18409.0, 18947.0, 19484.0, 20020.0, 20557.0, 21095.0, 21684.0, 22222.0, 22758.0, 23296.0, 23886.0, 24423.0, 24960.0, 25496.0, 26086.0, 26623.0, 27159.0, 27697.0, 28284.0, 28821.0, 29358.0, 29895.0, 30483.0, 31020.0, 31557.0, 32093.0, 32681.0, 33218.0, 33754.0, 34289.0, 34878.0, 35414.0, 35950.0, 36486.0, 37075.0, 37612.0, 38149.0, 38685.0, 39274.0, 39811.0, 40348.0, 40884.0, 41474.0, 42010.0, 42547.0, 43083.0, 43671.0, 44207.0, 44745.0, 45281.0, 45870.0, 46406.0, 46943.0, 47479.0, 48067.0, 48604.0, 49140.0, 49677.0, 50264.0, 50801.0, 51338.0, 51875.0, 52464.0, 53001.0, 53537.0, 54073.0, 54662.0, 55198.0, 55735.0, 56272.0, 56861.0, 57397.0, 57934.0, 58470.0, 59056.0, 59591.0, 60127.0, 60662.0, 61248.0, 61785.0, 62322.0, 62858.0, 63445.0, 63981.0, 64518.0, 65054.0, 65642.0, 66177.0, 66713.0, 67249.0, 67837.0, 68373.0, 68909.0, 69446.0, 70034.0, 70570.0, 71106.0, 71642.0, 72230.0, 72766.0, 73302.0, 73839.0, 74422.0, 74959.0, 75495.0, 76030.0, 76616.0, 77152.0, 77687.0, 78222.0, 78808.0, 79344.0, 79880.0, 80416.0, 81002.0, 81538.0, 82074.0, 82609.0, 83196.0, 83732.0, 84268.0, 84804.0, 85392.0, 85927.0, 86464.0, 87000.0, 87588.0, 88124.0, 88660.0, 89197.0, 89780.0, 90316.0, 90853.0, 91389.0, 91974.0, 92509.0, 93045.0, 93581.0, 94165.0, 94701.0, 95237.0, 95773.0, 96358.0, 96894.0, 97430.0, 97966.0, 98550.0, 99085.0, 99622.0, 100158.0, 100742.0, 101278.0, 101815.0, 102351.0, 102937.0, 103472.0, 104008.0, 104543.0, 105127.0, 105662.0, 106197.0, 106733.0, 107318.0, 107853.0, 108389.0, 108925.0]
source_col: d(D_H)Mean

## 1 - Sample 1 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 1
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1846
varray: [2.555, 2.17, 2.171, 2.199, 2.188, 2.173, 2.108, 2.079, 2.114, 2.089]
inst: mean_all
operation: mean_all(residuals)
tarray: [1.0, 537.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3757.0, 4293.0, 4830.0]
source_col: residuals
SAMP_residuals_STD | 0.137068677028
varray: [2.555, 2.17, 2.171, 2.199, 2.188, 2.173, 2.108, 2.079, 2.114, 2.089]
inst: std_all
operation: std_all(residuals)
tarray: [1.0, 537.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3757.0, 4293.0, 4830.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.1863
varray: [-1.877, -1.277, -1.297, -0.993, -1.103, -1.055, -1.107, -1.071, -1.073, -1.01]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [1.0, 537.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3757.0, 4293.0, 4830.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.263015018583
varray: [-1.877, -1.277, -1.297, -0.993, -1.103, -1.055, -1.107, -1.071, -1.073, -1.01]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [1.0, 537.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3757.0, 4293.0, 4830.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 90.6747
varray: [298.872, 98.487, 69.053, 65.558, 64.011, 63.094, 62.11, 61.795, 62.056, 61.711]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [1.0, 537.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3757.0, 4293.0, 4830.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 74.0007862301
varray: [298.872, 98.487, 69.053, 65.558, 64.011, 63.094, 62.11, 61.795, 62.056, 61.711]
inst: std_all
operation: std_all(baseline_shift)
tarray: [1.0, 537.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3757.0, 4293.0, 4830.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.6158
varray: [29.243, 28.018, 29.367, 29.185, 29.639, 29.318, 28.025, 27.739, 28.064, 27.56]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [1.0, 537.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3757.0, 4293.0, 4830.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.796943996359
varray: [29.243, 28.018, 29.367, 29.185, 29.639, 29.318, 28.025, 27.739, 28.064, 27.56]
inst: std_all
operation: std_all(slope_shift)
tarray: [1.0, 537.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3757.0, 4293.0, 4830.0]
source_col: slope_shift
SAMP_d18O_STD | 0.333943707831
varray: [-10.178, -9.394, -9.175, -9.137, -9.117, -9.099, -9.135, -9.142, -9.108, -9.113]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3757.0, 4293.0, 4830.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.93336312295
varray: [-76.906, -70.711, -68.643, -68.221, -68.061, -67.621, -67.659, -67.757, -67.325, -67.465]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3757.0, 4293.0, 4830.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.3193
varray: [69.456, 63.582, 63.539, 65.302, 64.468, 64.266, 62.976, 63.111, 63.163, 63.33]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [1.0, 537.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3757.0, 4293.0, 4830.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.94676387029
varray: [69.456, 63.582, 63.539, 65.302, 64.468, 64.266, 62.976, 63.111, 63.163, 63.33]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [1.0, 537.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3757.0, 4293.0, 4830.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1846
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.1863
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 90.6747
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.6158
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.03985
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.03985
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.85802541089
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.5079
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 0
varray: [-10.178, -9.394, -9.175, -9.137, -9.117, -9.099, -9.135, -9.142, -9.108, -9.113]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [1.0, 537.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3757.0, 4293.0, 4830.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 0
varray: [-76.906, -70.711, -68.643, -68.221, -68.061, -67.621, -67.659, -67.757, -67.325, -67.465]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [1.0, 537.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3757.0, 4293.0, 4830.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.2598
varray: [-10.178, -9.394, -9.175, -9.137, -9.117, -9.099, -9.135, -9.142, -9.108, -9.113]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3757.0, 4293.0, 4830.0]
source_col: d(18_16)Mean
dD_mean | -69.0369
varray: [-76.906, -70.711, -68.643, -68.221, -68.061, -67.621, -67.659, -67.757, -67.325, -67.465]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3757.0, 4293.0, 4830.0]
source_col: d(D_H)Mean
d18O_stdev | 0.333943707831
varray: [-10.178, -9.394, -9.175, -9.137, -9.117, -9.099, -9.135, -9.142, -9.108, -9.113]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3757.0, 4293.0, 4830.0]
source_col: d(18_16)Mean
dD_stdev | 2.93336312295
varray: [-76.906, -70.711, -68.643, -68.221, -68.061, -67.621, -67.659, -67.757, -67.325, -67.465]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3757.0, 4293.0, 4830.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.2598
varray: [-10.178, -9.394, -9.175, -9.137, -9.117, -9.099, -9.135, -9.142, -9.108, -9.113]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3757.0, 4293.0, 4830.0]
source_col: d(18_16)Mean
measured_dD_mean | -69.0369
varray: [-76.906, -70.711, -68.643, -68.221, -68.061, -67.621, -67.659, -67.757, -67.325, -67.465]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3757.0, 4293.0, 4830.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.333943707831
varray: [-10.178, -9.394, -9.175, -9.137, -9.117, -9.099, -9.135, -9.142, -9.108, -9.113]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3757.0, 4293.0, 4830.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.93336312295
varray: [-76.906, -70.711, -68.643, -68.221, -68.061, -67.621, -67.659, -67.757, -67.325, -67.465]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3757.0, 4293.0, 4830.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.4073765056
actual_array: [ 1.93 -15.84]
varray: [-10.178, -9.394, -9.175, -9.137, -9.117, -9.099, -9.135, -9.142, -9.108, -9.113]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -9.2598
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [1.0, 537.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3757.0, 4293.0, 4830.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -76.1048702631
actual_array: [ 18.1 -119. ]
varray: [-76.906, -70.711, -68.643, -68.221, -68.061, -67.621, -67.659, -67.757, -67.325, -67.465]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -69.0369
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [1.0, 537.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3757.0, 4293.0, 4830.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.09275
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 29.39735
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 2 - Sample 2 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 2
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.2816
varray: [2.267, 2.268, 2.312, 2.304, 2.324, 2.321, 2.263, 2.263, 2.252, 2.242]
inst: mean_all
operation: mean_all(residuals)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0]
source_col: residuals
SAMP_residuals_STD | 0.0303871317794
varray: [2.267, 2.268, 2.312, 2.304, 2.324, 2.321, 2.263, 2.263, 2.252, 2.242]
inst: std_all
operation: std_all(residuals)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.1164
varray: [-1.154, -1.029, -1.169, -1.341, -1.149, -1.125, -1.165, -0.949, -0.96, -1.123]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.114644571515
varray: [-1.154, -1.029, -1.169, -1.341, -1.149, -1.125, -1.165, -0.949, -0.96, -1.123]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 61.8788
varray: [61.899, 61.697, 62.536, 62.587, 62.348, 62.22, 61.531, 61.349, 61.23, 61.391]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.512513154303
varray: [61.899, 61.697, 62.536, 62.587, 62.348, 62.22, 61.531, 61.349, 61.23, 61.391]
inst: std_all
operation: std_all(baseline_shift)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.047
varray: [28.798, 28.738, 29.725, 29.862, 29.537, 29.754, 28.663, 28.795, 28.524, 28.074]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.619368674987
varray: [28.798, 28.738, 29.725, 29.862, 29.537, 29.754, 28.663, 28.795, 28.524, 28.074]
inst: std_all
operation: std_all(slope_shift)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0]
source_col: slope_shift
SAMP_d18O_STD | 0.203536947233
varray: [2.202, 2.573, 2.675, 2.758, 2.763, 2.82, 2.831, 2.807, 2.877, 2.872]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.46332924126
varray: [8.397, 15.173, 16.9, 17.787, 18.49, 18.808, 19.204, 19.363, 19.728, 19.819]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 66.387
varray: [66.525, 66.833, 66.938, 66.266, 67.807, 66.697, 65.348, 66.175, 65.601, 65.68]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.737172676892
varray: [66.525, 66.833, 66.938, 66.266, 67.807, 66.697, 65.348, 66.175, 65.601, 65.68]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.2816
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.1164
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 61.8788
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.047
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.13685
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.13685
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 6.38069705094
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 2.5756
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 2.7178
varray: [2.202, 2.573, 2.675, 2.758, 2.763, 2.82, 2.831, 2.807, 2.877, 2.872]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 17.3669
varray: [8.397, 15.173, 16.9, 17.787, 18.49, 18.808, 19.204, 19.363, 19.728, 19.819]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0]
source_col: d(D_H)Mean
known_dD | 18.1
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 1.93
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.7178
varray: [2.202, 2.573, 2.675, 2.758, 2.763, 2.82, 2.831, 2.807, 2.877, 2.872]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0]
source_col: d(18_16)Mean
dD_mean | 17.3669
varray: [8.397, 15.173, 16.9, 17.787, 18.49, 18.808, 19.204, 19.363, 19.728, 19.819]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0]
source_col: d(D_H)Mean
d18O_stdev | 0.203536947233
varray: [2.202, 2.573, 2.675, 2.758, 2.763, 2.82, 2.831, 2.807, 2.877, 2.872]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0]
source_col: d(18_16)Mean
dD_stdev | 3.46332924126
varray: [8.397, 15.173, 16.9, 17.787, 18.49, 18.808, 19.204, 19.363, 19.728, 19.819]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.7178
varray: [2.202, 2.573, 2.675, 2.758, 2.763, 2.82, 2.831, 2.807, 2.877, 2.872]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0]
source_col: d(18_16)Mean
measured_dD_mean | 17.3669
varray: [8.397, 15.173, 16.9, 17.787, 18.49, 18.808, 19.204, 19.363, 19.728, 19.819]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.203536947233
varray: [2.202, 2.573, 2.675, 2.758, 2.763, 2.82, 2.831, 2.807, 2.877, 2.872]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.46332924126
varray: [8.397, 15.173, 16.9, 17.787, 18.49, 18.808, 19.204, 19.363, 19.728, 19.819]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.93
actual_array: [ 1.93 -15.84]
varray: [2.202, 2.573, 2.675, 2.758, 2.763, 2.82, 2.831, 2.807, 2.877, 2.872]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 2.7178
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 18.1
actual_array: [ 18.1 -119. ]
varray: [8.397, 15.173, 16.9, 17.787, 18.49, 18.808, 19.204, 19.363, 19.728, 19.819]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: 17.3669
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [5420.0, 5956.0, 6493.0, 7031.0, 7567.0, 8103.0, 8640.0, 9176.0, 9714.0, 10251.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 58.99575
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.60145
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 3 - Sample 3 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 3
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0079
varray: [2.035, 2.059, 2.008, 1.98, 2.005, 1.996, 2.014, 1.994, 1.999, 1.989]
inst: mean_all
operation: mean_all(residuals)
tarray: [10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: residuals
SAMP_residuals_STD | 0.0234305119212
varray: [2.035, 2.059, 2.008, 1.98, 2.005, 1.996, 2.014, 1.994, 1.999, 1.989]
inst: std_all
operation: std_all(residuals)
tarray: [10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.8815
varray: [-0.909, -1.034, -0.773, -1.036, -0.888, -0.848, -0.843, -0.762, -0.722, -1.0]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.11356080114
varray: [-0.909, -1.034, -0.773, -1.036, -0.888, -0.848, -0.843, -0.762, -0.722, -1.0]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 60.6759
varray: [60.552, 61.414, 60.689, 60.509, 60.821, 60.663, 60.777, 60.235, 60.411, 60.688]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.313461480887
varray: [60.552, 61.414, 60.689, 60.509, 60.821, 60.663, 60.777, 60.235, 60.411, 60.688]
inst: std_all
operation: std_all(baseline_shift)
tarray: [10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.3747
varray: [27.355, 28.585, 27.156, 27.214, 27.472, 27.042, 27.198, 26.928, 27.574, 27.223]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.465775589743
varray: [27.355, 28.585, 27.156, 27.214, 27.472, 27.042, 27.198, 26.928, 27.574, 27.223]
inst: std_all
operation: std_all(slope_shift)
tarray: [10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: slope_shift
SAMP_d18O_STD | 0.292985020474
varray: [-13.758, -14.383, -14.504, -14.623, -14.595, -14.656, -14.68, -14.679, -14.721, -14.741]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 5.14616075773
varray: [-94.826, -105.397, -107.942, -109.355, -110.043, -110.713, -110.903, -111.333, -111.644, -111.643]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.2358
varray: [61.332, 61.436, 61.874, 60.129, 61.451, 61.062, 62.152, 61.034, 61.088, 60.8]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.562738305076
varray: [61.332, 61.436, 61.874, 60.129, 61.451, 61.062, 62.152, 61.034, 61.088, 60.8]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0079
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.8815
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 60.6759
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.3747
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.13685
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.13685
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.38069705094
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.5756
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-13.758, -14.383, -14.504, -14.623, -14.595, -14.656, -14.68, -14.679, -14.721, -14.741]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-94.826, -105.397, -107.942, -109.355, -110.043, -110.713, -110.903, -111.333, -111.644, -111.643]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: d(D_H)Mean
known_dD | -119.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | -15.84
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -14.534
varray: [-13.758, -14.383, -14.504, -14.623, -14.595, -14.656, -14.68, -14.679, -14.721, -14.741]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: d(18_16)Mean
dD_mean | -108.3799
varray: [-94.826, -105.397, -107.942, -109.355, -110.043, -110.713, -110.903, -111.333, -111.644, -111.643]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: d(D_H)Mean
d18O_stdev | 0.292985020474
varray: [-13.758, -14.383, -14.504, -14.623, -14.595, -14.656, -14.68, -14.679, -14.721, -14.741]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: d(18_16)Mean
dD_stdev | 5.14616075773
varray: [-94.826, -105.397, -107.942, -109.355, -110.043, -110.713, -110.903, -111.333, -111.644, -111.643]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: d(D_H)Mean
measured_d18O_mean | -14.534
varray: [-13.758, -14.383, -14.504, -14.623, -14.595, -14.656, -14.68, -14.679, -14.721, -14.741]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: d(18_16)Mean
measured_dD_mean | -108.3799
varray: [-94.826, -105.397, -107.942, -109.355, -110.043, -110.713, -110.903, -111.333, -111.644, -111.643]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.292985020474
varray: [-13.758, -14.383, -14.504, -14.623, -14.595, -14.656, -14.68, -14.679, -14.721, -14.741]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: d(18_16)Mean
measured_dD_stdev | 5.14616075773
varray: [-94.826, -105.397, -107.942, -109.355, -110.043, -110.713, -110.903, -111.333, -111.644, -111.643]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -15.84
actual_array: [ 1.93 -15.84]
varray: [-13.758, -14.383, -14.504, -14.623, -14.595, -14.656, -14.68, -14.679, -14.721, -14.741]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -14.534
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -119.0
actual_array: [ 18.1 -119. ]
varray: [-94.826, -105.397, -107.942, -109.355, -110.043, -110.713, -110.903, -111.333, -111.644, -111.643]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -108.3799
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [10840.0, 11377.0, 11914.0, 12451.0, 12988.0, 13524.0, 14061.0, 14598.0, 15135.0, 15672.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.26945
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.60145
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 4 - Sample 4 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 4
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.09
varray: [2.134, 2.065, 2.055, 2.08, 2.09, 2.098, 2.095, 2.083, 2.093, 2.107]
inst: mean_all
operation: mean_all(residuals)
tarray: [16261.0, 16799.0, 17335.0, 17872.0, 18409.0, 18947.0, 19484.0, 20020.0, 20557.0, 21095.0]
source_col: residuals
SAMP_residuals_STD | 0.0219139732185
varray: [2.134, 2.065, 2.055, 2.08, 2.09, 2.098, 2.095, 2.083, 2.093, 2.107]
inst: std_all
operation: std_all(residuals)
tarray: [16261.0, 16799.0, 17335.0, 17872.0, 18409.0, 18947.0, 19484.0, 20020.0, 20557.0, 21095.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9711
varray: [-0.901, -1.02, -1.259, -0.891, -0.813, -0.922, -0.859, -0.892, -1.149, -1.005]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [16261.0, 16799.0, 17335.0, 17872.0, 18409.0, 18947.0, 19484.0, 20020.0, 20557.0, 21095.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.13969605896
varray: [-0.901, -1.02, -1.259, -0.891, -0.813, -0.922, -0.859, -0.892, -1.149, -1.005]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [16261.0, 16799.0, 17335.0, 17872.0, 18409.0, 18947.0, 19484.0, 20020.0, 20557.0, 21095.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 60.9111
varray: [61.383, 60.712, 60.896, 60.614, 60.67, 60.963, 60.807, 60.976, 60.983, 61.107]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [16261.0, 16799.0, 17335.0, 17872.0, 18409.0, 18947.0, 19484.0, 20020.0, 20557.0, 21095.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.228581349487
varray: [61.383, 60.712, 60.896, 60.614, 60.67, 60.963, 60.807, 60.976, 60.983, 61.107]
inst: std_all
operation: std_all(baseline_shift)
tarray: [16261.0, 16799.0, 17335.0, 17872.0, 18409.0, 18947.0, 19484.0, 20020.0, 20557.0, 21095.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.8551
varray: [28.592, 27.569, 27.861, 27.629, 27.394, 28.278, 27.482, 27.673, 28.073, 28.0]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [16261.0, 16799.0, 17335.0, 17872.0, 18409.0, 18947.0, 19484.0, 20020.0, 20557.0, 21095.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.38127401928
varray: [28.592, 27.569, 27.861, 27.629, 27.394, 28.278, 27.482, 27.673, 28.073, 28.0]
inst: std_all
operation: std_all(slope_shift)
tarray: [16261.0, 16799.0, 17335.0, 17872.0, 18409.0, 18947.0, 19484.0, 20020.0, 20557.0, 21095.0]
source_col: slope_shift
SAMP_d18O_STD | 0.102796508804
varray: [-9.463, -9.218, -9.203, -9.149, -9.176, -9.185, -9.12, -9.146, -9.123, -9.11]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16261.0, 16799.0, 17335.0, 17872.0, 18409.0, 18947.0, 19484.0, 20020.0, 20557.0, 21095.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.62704069812
varray: [-72.588, -69.083, -68.329, -67.776, -67.679, -67.517, -67.375, -67.526, -67.242, -67.294]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16261.0, 16799.0, 17335.0, 17872.0, 18409.0, 18947.0, 19484.0, 20020.0, 20557.0, 21095.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.4686
varray: [63.971, 61.347, 60.288, 62.24, 63.161, 62.436, 63.133, 63.013, 62.051, 63.046]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [16261.0, 16799.0, 17335.0, 17872.0, 18409.0, 18947.0, 19484.0, 20020.0, 20557.0, 21095.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.056647024
varray: [63.971, 61.347, 60.288, 62.24, 63.161, 62.436, 63.133, 63.013, 62.051, 63.046]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [16261.0, 16799.0, 17335.0, 17872.0, 18409.0, 18947.0, 19484.0, 20020.0, 20557.0, 21095.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.09
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9711
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 60.9111
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.8551
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.05475
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.05475
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.55274507518
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.3428
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-9.463, -9.218, -9.203, -9.149, -9.176, -9.185, -9.12, -9.146, -9.123, -9.11]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [16261.0, 16799.0, 17335.0, 17872.0, 18409.0, 18947.0, 19484.0, 20020.0, 20557.0, 21095.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-72.588, -69.083, -68.329, -67.776, -67.679, -67.517, -67.375, -67.526, -67.242, -67.294]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [16261.0, 16799.0, 17335.0, 17872.0, 18409.0, 18947.0, 19484.0, 20020.0, 20557.0, 21095.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.1893
varray: [-9.463, -9.218, -9.203, -9.149, -9.176, -9.185, -9.12, -9.146, -9.123, -9.11]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16261.0, 16799.0, 17335.0, 17872.0, 18409.0, 18947.0, 19484.0, 20020.0, 20557.0, 21095.0]
source_col: d(18_16)Mean
dD_mean | -68.2409
varray: [-72.588, -69.083, -68.329, -67.776, -67.679, -67.517, -67.375, -67.526, -67.242, -67.294]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16261.0, 16799.0, 17335.0, 17872.0, 18409.0, 18947.0, 19484.0, 20020.0, 20557.0, 21095.0]
source_col: d(D_H)Mean
d18O_stdev | 0.102796508804
varray: [-9.463, -9.218, -9.203, -9.149, -9.176, -9.185, -9.12, -9.146, -9.123, -9.11]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16261.0, 16799.0, 17335.0, 17872.0, 18409.0, 18947.0, 19484.0, 20020.0, 20557.0, 21095.0]
source_col: d(18_16)Mean
dD_stdev | 1.62704069812
varray: [-72.588, -69.083, -68.329, -67.776, -67.679, -67.517, -67.375, -67.526, -67.242, -67.294]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16261.0, 16799.0, 17335.0, 17872.0, 18409.0, 18947.0, 19484.0, 20020.0, 20557.0, 21095.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.1893
varray: [-9.463, -9.218, -9.203, -9.149, -9.176, -9.185, -9.12, -9.146, -9.123, -9.11]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16261.0, 16799.0, 17335.0, 17872.0, 18409.0, 18947.0, 19484.0, 20020.0, 20557.0, 21095.0]
source_col: d(18_16)Mean
measured_dD_mean | -68.2409
varray: [-72.588, -69.083, -68.329, -67.776, -67.679, -67.517, -67.375, -67.526, -67.242, -67.294]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16261.0, 16799.0, 17335.0, 17872.0, 18409.0, 18947.0, 19484.0, 20020.0, 20557.0, 21095.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.102796508804
varray: [-9.463, -9.218, -9.203, -9.149, -9.176, -9.185, -9.12, -9.146, -9.123, -9.11]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16261.0, 16799.0, 17335.0, 17872.0, 18409.0, 18947.0, 19484.0, 20020.0, 20557.0, 21095.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.62704069812
varray: [-72.588, -69.083, -68.329, -67.776, -67.679, -67.517, -67.375, -67.526, -67.242, -67.294]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16261.0, 16799.0, 17335.0, 17872.0, 18409.0, 18947.0, 19484.0, 20020.0, 20557.0, 21095.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.3347588657
actual_array: [ 1.93 -15.84]
varray: [-9.463, -9.218, -9.203, -9.149, -9.176, -9.185, -9.12, -9.146, -9.123, -9.11]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -9.1893
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [16261.0, 16799.0, 17335.0, 17872.0, 18409.0, 18947.0, 19484.0, 20020.0, 20557.0, 21095.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -75.2370024525
actual_array: [ 18.1 -119. ]
varray: [-72.588, -69.083, -68.329, -67.776, -67.679, -67.517, -67.375, -67.526, -67.242, -67.294]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -68.2409
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [16261.0, 16799.0, 17335.0, 17872.0, 18409.0, 18947.0, 19484.0, 20020.0, 20557.0, 21095.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.18735
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.36625
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 5 - Sample 5 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 5
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.199
varray: [2.242, 2.185, 2.185, 2.184]
inst: mean_all
operation: mean_all(residuals)
tarray: [21684.0, 22222.0, 22758.0, 23296.0]
source_col: residuals
SAMP_residuals_STD | 0.0286705423737
varray: [2.242, 2.185, 2.185, 2.184]
inst: std_all
operation: std_all(residuals)
tarray: [21684.0, 22222.0, 22758.0, 23296.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.3625
varray: [-1.29, -1.467, -1.252, -1.441]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [21684.0, 22222.0, 22758.0, 23296.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.107314180486
varray: [-1.29, -1.467, -1.252, -1.441]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [21684.0, 22222.0, 22758.0, 23296.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 221.16925
varray: [224.102, 220.402, 219.69, 220.483]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [21684.0, 22222.0, 22758.0, 23296.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.98736129495
varray: [224.102, 220.402, 219.69, 220.483]
inst: std_all
operation: std_all(baseline_shift)
tarray: [21684.0, 22222.0, 22758.0, 23296.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 23.20675
varray: [23.94, 23.012, 22.787, 23.088]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [21684.0, 22222.0, 22758.0, 23296.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.505263874954
varray: [23.94, 23.012, 22.787, 23.088]
inst: std_all
operation: std_all(slope_shift)
tarray: [21684.0, 22222.0, 22758.0, 23296.0]
source_col: slope_shift
SAMP_d18O_STD | 0.130538627744
varray: [-1.501, -1.329, -1.288, -1.188]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21684.0, 22222.0, 22758.0, 23296.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.26672912291
varray: [-24.141, -20.517, -19.645, -19.112]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21684.0, 22222.0, 22758.0, 23296.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 60.72825
varray: [62.144, 60.047, 60.846, 59.876]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [21684.0, 22222.0, 22758.0, 23296.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.03418965862
varray: [62.144, 60.047, 60.846, 59.876]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [21684.0, 22222.0, 22758.0, 23296.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.199
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.3625
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 221.16925
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 23.20675
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.05425
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.05425
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.52943233477
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.08315
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-1.501, -1.329, -1.288, -1.188]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [21684.0, 22222.0, 22758.0, 23296.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-24.141, -20.517, -19.645, -19.112]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [21684.0, 22222.0, 22758.0, 23296.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.3265
varray: [-1.501, -1.329, -1.288, -1.188]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21684.0, 22222.0, 22758.0, 23296.0]
source_col: d(18_16)Mean
dD_mean | -20.85375
varray: [-24.141, -20.517, -19.645, -19.112]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21684.0, 22222.0, 22758.0, 23296.0]
source_col: d(D_H)Mean
d18O_stdev | 0.130538627744
varray: [-1.501, -1.329, -1.288, -1.188]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21684.0, 22222.0, 22758.0, 23296.0]
source_col: d(18_16)Mean
dD_stdev | 2.26672912291
varray: [-24.141, -20.517, -19.645, -19.112]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21684.0, 22222.0, 22758.0, 23296.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.3265
varray: [-1.501, -1.329, -1.288, -1.188]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21684.0, 22222.0, 22758.0, 23296.0]
source_col: d(18_16)Mean
measured_dD_mean | -20.85375
varray: [-24.141, -20.517, -19.645, -19.112]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21684.0, 22222.0, 22758.0, 23296.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.130538627744
varray: [-1.501, -1.329, -1.288, -1.188]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21684.0, 22222.0, 22758.0, 23296.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.26672912291
varray: [-24.141, -20.517, -19.645, -19.112]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21684.0, 22222.0, 22758.0, 23296.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.2357804403
actual_array: [ 1.93 -15.84]
varray: [-1.501, -1.329, -1.288, -1.188]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.3265
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [21684.0, 22222.0, 22758.0, 23296.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -23.5714470269
actual_array: [ 18.1 -119. ]
varray: [-24.141, -20.517, -19.645, -19.112]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -20.85375
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [21684.0, 22222.0, 22758.0, 23296.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.07835
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 159.8919
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 6 - Sample 6 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 6
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.134
varray: [2.112, 2.13, 2.132, 2.162]
inst: mean_all
operation: mean_all(residuals)
tarray: [23886.0, 24423.0, 24960.0, 25496.0]
source_col: residuals
SAMP_residuals_STD | 0.0207203603572
varray: [2.112, 2.13, 2.132, 2.162]
inst: std_all
operation: std_all(residuals)
tarray: [23886.0, 24423.0, 24960.0, 25496.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.029
varray: [-1.073, -1.171, -0.953, -0.919]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [23886.0, 24423.0, 24960.0, 25496.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.115435407624
varray: [-1.073, -1.171, -0.953, -0.919]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [23886.0, 24423.0, 24960.0, 25496.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 61.723
varray: [62.803, 61.778, 61.095, 61.216]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [23886.0, 24423.0, 24960.0, 25496.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.779071669788
varray: [62.803, 61.778, 61.095, 61.216]
inst: std_all
operation: std_all(baseline_shift)
tarray: [23886.0, 24423.0, 24960.0, 25496.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.6115
varray: [26.457, 26.628, 26.492, 26.869]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [23886.0, 24423.0, 24960.0, 25496.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.186841287372
varray: [26.457, 26.628, 26.492, 26.869]
inst: std_all
operation: std_all(slope_shift)
tarray: [23886.0, 24423.0, 24960.0, 25496.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0351046055858
varray: [0.481, 0.51, 0.51, 0.565]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23886.0, 24423.0, 24960.0, 25496.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.762473824687
varray: [-8.309, -7.675, -7.193, -6.505]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23886.0, 24423.0, 24960.0, 25496.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.949
varray: [63.817, 62.864, 64.277, 64.838]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [23886.0, 24423.0, 24960.0, 25496.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.835175430673
varray: [63.817, 62.864, 64.277, 64.838]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [23886.0, 24423.0, 24960.0, 25496.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.134
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.029
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 61.723
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.6115
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.01075
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.01075
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.501223918872
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.1376
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [0.481, 0.51, 0.51, 0.565]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [23886.0, 24423.0, 24960.0, 25496.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-8.309, -7.675, -7.193, -6.505]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [23886.0, 24423.0, 24960.0, 25496.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.5165
varray: [0.481, 0.51, 0.51, 0.565]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23886.0, 24423.0, 24960.0, 25496.0]
source_col: d(18_16)Mean
dD_mean | -7.4205
varray: [-8.309, -7.675, -7.193, -6.505]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23886.0, 24423.0, 24960.0, 25496.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0351046055858
varray: [0.481, 0.51, 0.51, 0.565]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23886.0, 24423.0, 24960.0, 25496.0]
source_col: d(18_16)Mean
dD_stdev | 0.762473824687
varray: [-8.309, -7.675, -7.193, -6.505]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23886.0, 24423.0, 24960.0, 25496.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.5165
varray: [0.481, 0.51, 0.51, 0.565]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23886.0, 24423.0, 24960.0, 25496.0]
source_col: d(18_16)Mean
measured_dD_mean | -7.4205
varray: [-8.309, -7.675, -7.193, -6.505]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23886.0, 24423.0, 24960.0, 25496.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0351046055858
varray: [0.481, 0.51, 0.51, 0.565]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23886.0, 24423.0, 24960.0, 25496.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.762473824687
varray: [-8.309, -7.675, -7.193, -6.505]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23886.0, 24423.0, 24960.0, 25496.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.337421428489
actual_array: [ 1.93 -15.84]
varray: [0.481, 0.51, 0.51, 0.565]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 0.5165
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [23886.0, 24423.0, 24960.0, 25496.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -8.92536000916
actual_array: [ 18.1 -119. ]
varray: [-8.309, -7.675, -7.193, -6.505]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -7.4205
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [23886.0, 24423.0, 24960.0, 25496.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.14335
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.44565
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 7 - Sample 7 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 7
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.3015
varray: [2.334, 2.282, 2.296, 2.294]
inst: mean_all
operation: mean_all(residuals)
tarray: [26086.0, 26623.0, 27159.0, 27697.0]
source_col: residuals
SAMP_residuals_STD | 0.0225314594882
varray: [2.334, 2.282, 2.296, 2.294]
inst: std_all
operation: std_all(residuals)
tarray: [26086.0, 26623.0, 27159.0, 27697.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.6045
varray: [-1.734, -1.558, -1.539, -1.587]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [26086.0, 26623.0, 27159.0, 27697.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0885607136376
varray: [-1.734, -1.558, -1.539, -1.587]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [26086.0, 26623.0, 27159.0, 27697.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 88.61025
varray: [89.376, 88.042, 88.38, 88.643]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [26086.0, 26623.0, 27159.0, 27697.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.566677083238
varray: [89.376, 88.042, 88.38, 88.643]
inst: std_all
operation: std_all(baseline_shift)
tarray: [26086.0, 26623.0, 27159.0, 27697.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.20175
varray: [26.693, 25.906, 26.045, 26.163]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [26086.0, 26623.0, 27159.0, 27697.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.343931558114
varray: [26.693, 25.906, 26.045, 26.163]
inst: std_all
operation: std_all(slope_shift)
tarray: [26086.0, 26623.0, 27159.0, 27697.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0375532954613
varray: [-0.197, -0.273, -0.279, -0.258]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26086.0, 26623.0, 27159.0, 27697.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.408018381939
varray: [-12.022, -12.911, -12.856, -12.661]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26086.0, 26623.0, 27159.0, 27697.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 65.12325
varray: [65.482, 64.927, 64.779, 65.305]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [26086.0, 26623.0, 27159.0, 27697.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.325963571584
varray: [65.482, 64.927, 64.779, 65.305]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [26086.0, 26623.0, 27159.0, 27697.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.3015
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.6045
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 88.61025
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.20175
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.15675
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.15675
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 7.30854411936
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.31185
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-0.197, -0.273, -0.279, -0.258]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [26086.0, 26623.0, 27159.0, 27697.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-12.022, -12.911, -12.856, -12.661]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [26086.0, 26623.0, 27159.0, 27697.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.25175
varray: [-0.197, -0.273, -0.279, -0.258]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26086.0, 26623.0, 27159.0, 27697.0]
source_col: d(18_16)Mean
dD_mean | -12.6125
varray: [-12.022, -12.911, -12.856, -12.661]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26086.0, 26623.0, 27159.0, 27697.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0375532954613
varray: [-0.197, -0.273, -0.279, -0.258]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26086.0, 26623.0, 27159.0, 27697.0]
source_col: d(18_16)Mean
dD_stdev | 0.408018381939
varray: [-12.022, -12.911, -12.856, -12.661]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26086.0, 26623.0, 27159.0, 27697.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.25175
varray: [-0.197, -0.273, -0.279, -0.258]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26086.0, 26623.0, 27159.0, 27697.0]
source_col: d(18_16)Mean
measured_dD_mean | -12.6125
varray: [-12.022, -12.911, -12.856, -12.661]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26086.0, 26623.0, 27159.0, 27697.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0375532954613
varray: [-0.197, -0.273, -0.279, -0.258]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26086.0, 26623.0, 27159.0, 27697.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.408018381939
varray: [-12.022, -12.911, -12.856, -12.661]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26086.0, 26623.0, 27159.0, 27697.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.12874769589
actual_array: [ 1.93 -15.84]
varray: [-0.197, -0.273, -0.279, -0.258]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.25175
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [26086.0, 26623.0, 27159.0, 27697.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -14.5861259292
actual_array: [ 18.1 -119. ]
varray: [-12.022, -12.911, -12.856, -12.661]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -12.6125
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [26086.0, 26623.0, 27159.0, 27697.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 58.97585
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 27.3329
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 8 - Sample 8 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 8
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.25625
varray: [2.255, 2.258, 2.258, 2.254]
inst: mean_all
operation: mean_all(residuals)
tarray: [28284.0, 28821.0, 29358.0, 29895.0]
source_col: residuals
SAMP_residuals_STD | 0.00206155281281
varray: [2.255, 2.258, 2.258, 2.254]
inst: std_all
operation: std_all(residuals)
tarray: [28284.0, 28821.0, 29358.0, 29895.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.78425
varray: [-1.816, -1.717, -1.794, -1.81]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [28284.0, 28821.0, 29358.0, 29895.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0457848228128
varray: [-1.816, -1.717, -1.794, -1.81]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [28284.0, 28821.0, 29358.0, 29895.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 165.25475
varray: [164.858, 165.182, 165.497, 165.482]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [28284.0, 28821.0, 29358.0, 29895.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.301679051311
varray: [164.858, 165.182, 165.497, 165.482]
inst: std_all
operation: std_all(baseline_shift)
tarray: [28284.0, 28821.0, 29358.0, 29895.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 22.58325
varray: [22.501, 22.636, 22.575, 22.621]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [28284.0, 28821.0, 29358.0, 29895.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0606650640814
varray: [22.501, 22.636, 22.575, 22.621]
inst: std_all
operation: std_all(slope_shift)
tarray: [28284.0, 28821.0, 29358.0, 29895.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0251130776024
varray: [-0.287, -0.231, -0.265, -0.241]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28284.0, 28821.0, 29358.0, 29895.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0421693016305
varray: [-15.399, -15.417, -15.45, -15.495]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28284.0, 28821.0, 29358.0, 29895.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.8015
varray: [61.534, 62.243, 61.374, 62.055]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [28284.0, 28821.0, 29358.0, 29895.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.413721726124
varray: [61.534, 62.243, 61.374, 62.055]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [28284.0, 28821.0, 29358.0, 29895.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.25625
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.78425
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 165.25475
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 22.58325
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.1115
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1115
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 5.19874111202
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.0099
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-0.287, -0.231, -0.265, -0.241]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [28284.0, 28821.0, 29358.0, 29895.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-15.399, -15.417, -15.45, -15.495]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [28284.0, 28821.0, 29358.0, 29895.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.256
varray: [-0.287, -0.231, -0.265, -0.241]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28284.0, 28821.0, 29358.0, 29895.0]
source_col: d(18_16)Mean
dD_mean | -15.44025
varray: [-15.399, -15.417, -15.45, -15.495]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28284.0, 28821.0, 29358.0, 29895.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0251130776024
varray: [-0.287, -0.231, -0.265, -0.241]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28284.0, 28821.0, 29358.0, 29895.0]
source_col: d(18_16)Mean
dD_stdev | 0.0421693016305
varray: [-15.399, -15.417, -15.45, -15.495]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28284.0, 28821.0, 29358.0, 29895.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.256
varray: [-0.287, -0.231, -0.265, -0.241]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28284.0, 28821.0, 29358.0, 29895.0]
source_col: d(18_16)Mean
measured_dD_mean | -15.44025
varray: [-15.399, -15.417, -15.45, -15.495]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28284.0, 28821.0, 29358.0, 29895.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0251130776024
varray: [-0.287, -0.231, -0.265, -0.241]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28284.0, 28821.0, 29358.0, 29895.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0421693016305
varray: [-15.399, -15.417, -15.45, -15.495]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28284.0, 28821.0, 29358.0, 29895.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.13312535504
actual_array: [ 1.93 -15.84]
varray: [-0.287, -0.231, -0.265, -0.241]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.256
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [28284.0, 28821.0, 29358.0, 29895.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -17.6691827148
actual_array: [ 18.1 -119. ]
varray: [-15.399, -15.417, -15.45, -15.495]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -15.44025
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [28284.0, 28821.0, 29358.0, 29895.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.0211
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 103.9774
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 9 - Sample 9 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 9
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.267
varray: [2.285, 2.28, 2.259, 2.244]
inst: mean_all
operation: mean_all(residuals)
tarray: [30483.0, 31020.0, 31557.0, 32093.0]
source_col: residuals
SAMP_residuals_STD | 0.0190262975904
varray: [2.285, 2.28, 2.259, 2.244]
inst: std_all
operation: std_all(residuals)
tarray: [30483.0, 31020.0, 31557.0, 32093.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.53075
varray: [-1.627, -1.392, -1.67, -1.434]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [30483.0, 31020.0, 31557.0, 32093.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.138162645699
varray: [-1.627, -1.392, -1.67, -1.434]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [30483.0, 31020.0, 31557.0, 32093.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 169.81675
varray: [171.216, 169.912, 169.841, 168.298]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [30483.0, 31020.0, 31557.0, 32093.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.19361673776
varray: [171.216, 169.912, 169.841, 168.298]
inst: std_all
operation: std_all(baseline_shift)
tarray: [30483.0, 31020.0, 31557.0, 32093.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 23.906
varray: [24.511, 23.578, 23.82, 23.715]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [30483.0, 31020.0, 31557.0, 32093.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.415325575101
varray: [24.511, 23.578, 23.82, 23.715]
inst: std_all
operation: std_all(slope_shift)
tarray: [30483.0, 31020.0, 31557.0, 32093.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0410081292754
varray: [-1.507, -1.578, -1.561, -1.604]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30483.0, 31020.0, 31557.0, 32093.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.604170781595
varray: [-22.072, -22.982, -23.338, -23.366]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30483.0, 31020.0, 31557.0, 32093.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.37675
varray: [63.662, 63.869, 63.228, 62.748]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [30483.0, 31020.0, 31557.0, 32093.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.497034790868
varray: [63.662, 63.869, 63.228, 62.748]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [30483.0, 31020.0, 31557.0, 32093.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.267
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.53075
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 169.81675
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 23.906
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.12225
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.12225
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 5.69996503089
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.43465
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-1.507, -1.578, -1.561, -1.604]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [30483.0, 31020.0, 31557.0, 32093.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-22.072, -22.982, -23.338, -23.366]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [30483.0, 31020.0, 31557.0, 32093.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.5625
varray: [-1.507, -1.578, -1.561, -1.604]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30483.0, 31020.0, 31557.0, 32093.0]
source_col: d(18_16)Mean
dD_mean | -22.9395
varray: [-22.072, -22.982, -23.338, -23.366]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30483.0, 31020.0, 31557.0, 32093.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0410081292754
varray: [-1.507, -1.578, -1.561, -1.604]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30483.0, 31020.0, 31557.0, 32093.0]
source_col: d(18_16)Mean
dD_stdev | 0.604170781595
varray: [-22.072, -22.982, -23.338, -23.366]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30483.0, 31020.0, 31557.0, 32093.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.5625
varray: [-1.507, -1.578, -1.561, -1.604]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30483.0, 31020.0, 31557.0, 32093.0]
source_col: d(18_16)Mean
measured_dD_mean | -22.9395
varray: [-22.072, -22.982, -23.338, -23.366]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30483.0, 31020.0, 31557.0, 32093.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0410081292754
varray: [-1.507, -1.578, -1.561, -1.604]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30483.0, 31020.0, 31557.0, 32093.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.604170781595
varray: [-22.072, -22.982, -23.338, -23.366]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30483.0, 31020.0, 31557.0, 32093.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.47886927741
actual_array: [ 1.93 -15.84]
varray: [-1.507, -1.578, -1.561, -1.604]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.5625
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [30483.0, 31020.0, 31557.0, 32093.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -25.8455114564
actual_array: [ 18.1 -119. ]
varray: [-22.072, -22.982, -23.338, -23.366]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -22.9395
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [30483.0, 31020.0, 31557.0, 32093.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.01035
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 108.5394
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 10 - Sample 10 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 10
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.124
varray: [2.129, 2.12, 2.115, 2.132]
inst: mean_all
operation: mean_all(residuals)
tarray: [32681.0, 33218.0, 33754.0, 34289.0]
source_col: residuals
SAMP_residuals_STD | 0.00787400787401
varray: [2.129, 2.12, 2.115, 2.132]
inst: std_all
operation: std_all(residuals)
tarray: [32681.0, 33218.0, 33754.0, 34289.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.23375
varray: [-1.277, -1.366, -1.127, -1.165]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [32681.0, 33218.0, 33754.0, 34289.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.108754693385
varray: [-1.277, -1.366, -1.127, -1.165]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [32681.0, 33218.0, 33754.0, 34289.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 288.086
varray: [287.096, 287.61, 287.894, 289.744]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [32681.0, 33218.0, 33754.0, 34289.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.1536180766
varray: [287.096, 287.61, 287.894, 289.744]
inst: std_all
operation: std_all(baseline_shift)
tarray: [32681.0, 33218.0, 33754.0, 34289.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 18.9005
varray: [18.945, 19.095, 18.916, 18.646]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [32681.0, 33218.0, 33754.0, 34289.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.186923335444
varray: [18.945, 19.095, 18.916, 18.646]
inst: std_all
operation: std_all(slope_shift)
tarray: [32681.0, 33218.0, 33754.0, 34289.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0240329634738
varray: [-1.914, -1.963, -1.927, -1.959]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32681.0, 33218.0, 33754.0, 34289.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.178600811868
varray: [-22.621, -22.53, -22.21, -22.396]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32681.0, 33218.0, 33754.0, 34289.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 58.28425
varray: [57.827, 57.199, 58.316, 59.795]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [32681.0, 33218.0, 33754.0, 34289.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.10607696387
varray: [57.827, 57.199, 58.316, 59.795]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [32681.0, 33218.0, 33754.0, 34289.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.124
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.23375
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 288.086
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 18.9005
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.02075
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.02075
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.967478727124
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.52715
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-1.914, -1.963, -1.927, -1.959]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [32681.0, 33218.0, 33754.0, 34289.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-22.621, -22.53, -22.21, -22.396]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [32681.0, 33218.0, 33754.0, 34289.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.94075
varray: [-1.914, -1.963, -1.927, -1.959]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32681.0, 33218.0, 33754.0, 34289.0]
source_col: d(18_16)Mean
dD_mean | -22.43925
varray: [-22.621, -22.53, -22.21, -22.396]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32681.0, 33218.0, 33754.0, 34289.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0240329634738
varray: [-1.914, -1.963, -1.927, -1.959]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32681.0, 33218.0, 33754.0, 34289.0]
source_col: d(18_16)Mean
dD_stdev | 0.178600811868
varray: [-22.621, -22.53, -22.21, -22.396]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32681.0, 33218.0, 33754.0, 34289.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.94075
varray: [-1.914, -1.963, -1.927, -1.959]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32681.0, 33218.0, 33754.0, 34289.0]
source_col: d(18_16)Mean
measured_dD_mean | -22.43925
varray: [-22.621, -22.53, -22.21, -22.396]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32681.0, 33218.0, 33754.0, 34289.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0240329634738
varray: [-1.914, -1.963, -1.927, -1.959]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32681.0, 33218.0, 33754.0, 34289.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.178600811868
varray: [-22.621, -22.53, -22.21, -22.396]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32681.0, 33218.0, 33754.0, 34289.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.86848094112
actual_array: [ 1.93 -15.84]
varray: [-1.914, -1.963, -1.927, -1.959]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.94075
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [32681.0, 33218.0, 33754.0, 34289.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -25.3000957877
actual_array: [ 18.1 -119. ]
varray: [-22.621, -22.53, -22.21, -22.396]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -22.43925
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [32681.0, 33218.0, 33754.0, 34289.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.15335
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 226.80865
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 11 - Sample 11 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 11
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.29
varray: [2.319, 2.286, 2.279, 2.276]
inst: mean_all
operation: mean_all(residuals)
tarray: [34878.0, 35414.0, 35950.0, 36486.0]
source_col: residuals
SAMP_residuals_STD | 0.0197821468333
varray: [2.319, 2.286, 2.279, 2.276]
inst: std_all
operation: std_all(residuals)
tarray: [34878.0, 35414.0, 35950.0, 36486.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.07275
varray: [-2.028, -2.05, -2.048, -2.165]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [34878.0, 35414.0, 35950.0, 36486.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0622970036733
varray: [-2.028, -2.05, -2.048, -2.165]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [34878.0, 35414.0, 35950.0, 36486.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 498.71675
varray: [504.628, 495.266, 497.137, 497.836]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [34878.0, 35414.0, 35950.0, 36486.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 4.08745571841
varray: [504.628, 495.266, 497.137, 497.836]
inst: std_all
operation: std_all(baseline_shift)
tarray: [34878.0, 35414.0, 35950.0, 36486.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 12.20275
varray: [12.658, 12.29, 11.992, 11.871]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [34878.0, 35414.0, 35950.0, 36486.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.350874502731
varray: [12.658, 12.29, 11.992, 11.871]
inst: std_all
operation: std_all(slope_shift)
tarray: [34878.0, 35414.0, 35950.0, 36486.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0565479147862
varray: [-2.39, -2.502, -2.483, -2.515]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34878.0, 35414.0, 35950.0, 36486.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.330821704246
varray: [-29.411, -29.924, -30.056, -30.155]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34878.0, 35414.0, 35950.0, 36486.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 56.29575
varray: [56.83, 55.769, 56.422, 56.162]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [34878.0, 35414.0, 35950.0, 36486.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.44598832197
varray: [56.83, 55.769, 56.422, 56.162]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [34878.0, 35414.0, 35950.0, 36486.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.29
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.07275
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 498.71675
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 12.20275
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.14525
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.14525
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 6.77235108987
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -7.51565
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-2.39, -2.502, -2.483, -2.515]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [34878.0, 35414.0, 35950.0, 36486.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-29.411, -29.924, -30.056, -30.155]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [34878.0, 35414.0, 35950.0, 36486.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.4725
varray: [-2.39, -2.502, -2.483, -2.515]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [34878.0, 35414.0, 35950.0, 36486.0]
source_col: d(18_16)Mean
dD_mean | -29.8865
varray: [-29.411, -29.924, -30.056, -30.155]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [34878.0, 35414.0, 35950.0, 36486.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0565479147862
varray: [-2.39, -2.502, -2.483, -2.515]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34878.0, 35414.0, 35950.0, 36486.0]
source_col: d(18_16)Mean
dD_stdev | 0.330821704246
varray: [-29.411, -29.924, -30.056, -30.155]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34878.0, 35414.0, 35950.0, 36486.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.4725
varray: [-2.39, -2.502, -2.483, -2.515]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [34878.0, 35414.0, 35950.0, 36486.0]
source_col: d(18_16)Mean
measured_dD_mean | -29.8865
varray: [-29.411, -29.924, -30.056, -30.155]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [34878.0, 35414.0, 35950.0, 36486.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0565479147862
varray: [-2.39, -2.502, -2.483, -2.515]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34878.0, 35414.0, 35950.0, 36486.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.330821704246
varray: [-29.411, -29.924, -30.056, -30.155]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34878.0, 35414.0, 35950.0, 36486.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.41620335269
actual_array: [ 1.93 -15.84]
varray: [-2.39, -2.502, -2.483, -2.515]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -2.4725
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [34878.0, 35414.0, 35950.0, 36486.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -33.4197296472
actual_array: [ 18.1 -119. ]
varray: [-29.411, -29.924, -30.056, -30.155]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -29.8865
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [34878.0, 35414.0, 35950.0, 36486.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 58.98735
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 437.4394
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 12 - Sample 12 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 12
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1445
varray: [2.142, 2.146, 2.138, 2.152]
inst: mean_all
operation: mean_all(residuals)
tarray: [37075.0, 37612.0, 38149.0, 38685.0]
source_col: residuals
SAMP_residuals_STD | 0.00597215762239
varray: [2.142, 2.146, 2.138, 2.152]
inst: std_all
operation: std_all(residuals)
tarray: [37075.0, 37612.0, 38149.0, 38685.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.10525
varray: [-1.31, -1.14, -1.008, -0.963]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [37075.0, 37612.0, 38149.0, 38685.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.155801957626
varray: [-1.31, -1.14, -1.008, -0.963]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [37075.0, 37612.0, 38149.0, 38685.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 63.5025
varray: [66.127, 63.555, 62.39, 61.938]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [37075.0, 37612.0, 38149.0, 38685.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.877593584
varray: [66.127, 63.555, 62.39, 61.938]
inst: std_all
operation: std_all(baseline_shift)
tarray: [37075.0, 37612.0, 38149.0, 38685.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.047
varray: [27.246, 27.09, 26.864, 26.988]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [37075.0, 37612.0, 38149.0, 38685.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.161678693711
varray: [27.246, 27.09, 26.864, 26.988]
inst: std_all
operation: std_all(slope_shift)
tarray: [37075.0, 37612.0, 38149.0, 38685.0]
source_col: slope_shift
SAMP_d18O_STD | 0.11573064993
varray: [1.695, 1.905, 1.939, 1.93]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37075.0, 37612.0, 38149.0, 38685.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.35623424722
varray: [-5.64, -3.408, -2.909, -2.675]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37075.0, 37612.0, 38149.0, 38685.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.05875
varray: [62.085, 63.208, 63.288, 63.654]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [37075.0, 37612.0, 38149.0, 38685.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.677579700109
varray: [62.085, 63.208, 63.288, 63.654]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [37075.0, 37612.0, 38149.0, 38685.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1445
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.10525
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 63.5025
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.047
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.00025
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.00025
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.0116563702063
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.75265
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [1.695, 1.905, 1.939, 1.93]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [37075.0, 37612.0, 38149.0, 38685.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-5.64, -3.408, -2.909, -2.675]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [37075.0, 37612.0, 38149.0, 38685.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.86725
varray: [1.695, 1.905, 1.939, 1.93]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37075.0, 37612.0, 38149.0, 38685.0]
source_col: d(18_16)Mean
dD_mean | -3.658
varray: [-5.64, -3.408, -2.909, -2.675]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37075.0, 37612.0, 38149.0, 38685.0]
source_col: d(D_H)Mean
d18O_stdev | 0.11573064993
varray: [1.695, 1.905, 1.939, 1.93]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37075.0, 37612.0, 38149.0, 38685.0]
source_col: d(18_16)Mean
dD_stdev | 1.35623424722
varray: [-5.64, -3.408, -2.909, -2.675]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37075.0, 37612.0, 38149.0, 38685.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.86725
varray: [1.695, 1.905, 1.939, 1.93]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37075.0, 37612.0, 38149.0, 38685.0]
source_col: d(18_16)Mean
measured_dD_mean | -3.658
varray: [-5.64, -3.408, -2.909, -2.675]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37075.0, 37612.0, 38149.0, 38685.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.11573064993
varray: [1.695, 1.905, 1.939, 1.93]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37075.0, 37612.0, 38149.0, 38685.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.35623424722
varray: [-5.64, -3.408, -2.909, -2.675]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37075.0, 37612.0, 38149.0, 38685.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.05390165084
actual_array: [ 1.93 -15.84]
varray: [1.695, 1.905, 1.939, 1.93]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 1.86725
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [37075.0, 37612.0, 38149.0, 38685.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -4.82315820363
actual_array: [ 18.1 -119. ]
varray: [-5.64, -3.408, -2.909, -2.675]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -3.658
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [37075.0, 37612.0, 38149.0, 38685.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.13285
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 2.22515
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 13 - Sample 13 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 13
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.12775
varray: [2.191, 2.114, 2.102, 2.104]
inst: mean_all
operation: mean_all(residuals)
tarray: [39274.0, 39811.0, 40348.0, 40884.0]
source_col: residuals
SAMP_residuals_STD | 0.0424921561389
varray: [2.191, 2.114, 2.102, 2.104]
inst: std_all
operation: std_all(residuals)
tarray: [39274.0, 39811.0, 40348.0, 40884.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.878
varray: [-0.81, -0.99, -0.931, -0.781]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [39274.0, 39811.0, 40348.0, 40884.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0989713763334
varray: [-0.81, -0.99, -0.931, -0.781]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [39274.0, 39811.0, 40348.0, 40884.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 61.9235
varray: [62.486, 61.786, 61.753, 61.669]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [39274.0, 39811.0, 40348.0, 40884.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.378220834963
varray: [62.486, 61.786, 61.753, 61.669]
inst: std_all
operation: std_all(baseline_shift)
tarray: [39274.0, 39811.0, 40348.0, 40884.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.08925
varray: [28.997, 28.059, 27.838, 27.463]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [39274.0, 39811.0, 40348.0, 40884.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.653258690464
varray: [28.997, 28.059, 27.838, 27.463]
inst: std_all
operation: std_all(slope_shift)
tarray: [39274.0, 39811.0, 40348.0, 40884.0]
source_col: slope_shift
SAMP_d18O_STD | 0.239538966907
varray: [-8.48, -8.875, -8.922, -9.028]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39274.0, 39811.0, 40348.0, 40884.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.12942980696
varray: [-58.554, -63.55, -64.861, -65.426]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39274.0, 39811.0, 40348.0, 40884.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.23725
varray: [66.317, 62.995, 63.407, 64.23]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [39274.0, 39811.0, 40348.0, 40884.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.4785029309
varray: [66.317, 62.995, 63.407, 64.23]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [39274.0, 39811.0, 40348.0, 40884.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.12775
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.878
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 61.9235
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.08925
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.017
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.017
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.79263317403
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.42585
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-8.48, -8.875, -8.922, -9.028]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [39274.0, 39811.0, 40348.0, 40884.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-58.554, -63.55, -64.861, -65.426]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [39274.0, 39811.0, 40348.0, 40884.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.82625
varray: [-8.48, -8.875, -8.922, -9.028]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39274.0, 39811.0, 40348.0, 40884.0]
source_col: d(18_16)Mean
dD_mean | -63.09775
varray: [-58.554, -63.55, -64.861, -65.426]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39274.0, 39811.0, 40348.0, 40884.0]
source_col: d(D_H)Mean
d18O_stdev | 0.239538966907
varray: [-8.48, -8.875, -8.922, -9.028]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39274.0, 39811.0, 40348.0, 40884.0]
source_col: d(18_16)Mean
dD_stdev | 3.12942980696
varray: [-58.554, -63.55, -64.861, -65.426]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39274.0, 39811.0, 40348.0, 40884.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.82625
varray: [-8.48, -8.875, -8.922, -9.028]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39274.0, 39811.0, 40348.0, 40884.0]
source_col: d(18_16)Mean
measured_dD_mean | -63.09775
varray: [-58.554, -63.55, -64.861, -65.426]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39274.0, 39811.0, 40348.0, 40884.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.239538966907
varray: [-8.48, -8.875, -8.922, -9.028]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39274.0, 39811.0, 40348.0, 40884.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.12942980696
varray: [-58.554, -63.55, -64.861, -65.426]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39274.0, 39811.0, 40348.0, 40884.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.9608037712
actual_array: [ 1.93 -15.84]
varray: [-8.48, -8.875, -8.922, -9.028]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -8.82625
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [39274.0, 39811.0, 40348.0, 40884.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -69.629497013
actual_array: [ 18.1 -119. ]
varray: [-58.554, -63.55, -64.861, -65.426]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -63.09775
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [39274.0, 39811.0, 40348.0, 40884.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.1496
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.64615
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 14 - Sample 14 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 14
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.21475
varray: [2.22, 2.228, 2.197, 2.214]
inst: mean_all
operation: mean_all(residuals)
tarray: [41474.0, 42010.0, 42547.0, 43083.0]
source_col: residuals
SAMP_residuals_STD | 0.0131497781984
varray: [2.22, 2.228, 2.197, 2.214]
inst: std_all
operation: std_all(residuals)
tarray: [41474.0, 42010.0, 42547.0, 43083.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.526
varray: [-1.62, -1.608, -1.503, -1.373]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [41474.0, 42010.0, 42547.0, 43083.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.114743191519
varray: [-1.62, -1.608, -1.503, -1.373]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [41474.0, 42010.0, 42547.0, 43083.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 233.59275
varray: [233.624, 235.147, 231.882, 233.718]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [41474.0, 42010.0, 42547.0, 43083.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.33654065283
varray: [233.624, 235.147, 231.882, 233.718]
inst: std_all
operation: std_all(baseline_shift)
tarray: [41474.0, 42010.0, 42547.0, 43083.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 20.47525
varray: [20.627, 20.599, 20.265, 20.41]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [41474.0, 42010.0, 42547.0, 43083.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.170102665078
varray: [20.627, 20.599, 20.265, 20.41]
inst: std_all
operation: std_all(slope_shift)
tarray: [41474.0, 42010.0, 42547.0, 43083.0]
source_col: slope_shift
SAMP_d18O_STD | 0.196620616077
varray: [0.022, 0.288, 0.388, 0.476]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41474.0, 42010.0, 42547.0, 43083.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.98540531866
varray: [-11.737, -7.239, -5.886, -5.027]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41474.0, 42010.0, 42547.0, 43083.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 60.6015
varray: [60.379, 60.773, 60.289, 60.965]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [41474.0, 42010.0, 42547.0, 43083.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.320783935591
varray: [60.379, 60.773, 60.289, 60.965]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [41474.0, 42010.0, 42547.0, 43083.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.21475
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.526
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 233.59275
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 20.47525
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.07
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.07
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 3.26378365777
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.2099
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [0.022, 0.288, 0.388, 0.476]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [41474.0, 42010.0, 42547.0, 43083.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-11.737, -7.239, -5.886, -5.027]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [41474.0, 42010.0, 42547.0, 43083.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.2935
varray: [0.022, 0.288, 0.388, 0.476]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [41474.0, 42010.0, 42547.0, 43083.0]
source_col: d(18_16)Mean
dD_mean | -7.47225
varray: [-11.737, -7.239, -5.886, -5.027]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [41474.0, 42010.0, 42547.0, 43083.0]
source_col: d(D_H)Mean
d18O_stdev | 0.196620616077
varray: [0.022, 0.288, 0.388, 0.476]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41474.0, 42010.0, 42547.0, 43083.0]
source_col: d(18_16)Mean
dD_stdev | 2.98540531866
varray: [-11.737, -7.239, -5.886, -5.027]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41474.0, 42010.0, 42547.0, 43083.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.2935
varray: [0.022, 0.288, 0.388, 0.476]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [41474.0, 42010.0, 42547.0, 43083.0]
source_col: d(18_16)Mean
measured_dD_mean | -7.47225
varray: [-11.737, -7.239, -5.886, -5.027]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [41474.0, 42010.0, 42547.0, 43083.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.196620616077
varray: [0.022, 0.288, 0.388, 0.476]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41474.0, 42010.0, 42547.0, 43083.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.98540531866
varray: [-11.737, -7.239, -5.886, -5.027]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41474.0, 42010.0, 42547.0, 43083.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.567119778806
actual_array: [ 1.93 -15.84]
varray: [0.022, 0.288, 0.388, 0.476]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 0.2935
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [41474.0, 42010.0, 42547.0, 43083.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -8.98178231971
actual_array: [ 18.1 -119. ]
varray: [-11.737, -7.239, -5.886, -5.027]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -7.47225
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [41474.0, 42010.0, 42547.0, 43083.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.0626
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 172.3154
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 15 - Sample 15 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 15
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.21
varray: [2.216, 2.215, 2.208, 2.201]
inst: mean_all
operation: mean_all(residuals)
tarray: [43671.0, 44207.0, 44745.0, 45281.0]
source_col: residuals
SAMP_residuals_STD | 0.00697614984549
varray: [2.216, 2.215, 2.208, 2.201]
inst: std_all
operation: std_all(residuals)
tarray: [43671.0, 44207.0, 44745.0, 45281.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.175
varray: [-1.197, -1.285, -1.105, -1.113]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [43671.0, 44207.0, 44745.0, 45281.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0843168626867
varray: [-1.197, -1.285, -1.105, -1.113]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [43671.0, 44207.0, 44745.0, 45281.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 63.0045
varray: [64.312, 63.113, 62.411, 62.182]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [43671.0, 44207.0, 44745.0, 45281.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.957440511642
varray: [64.312, 63.113, 62.411, 62.182]
inst: std_all
operation: std_all(baseline_shift)
tarray: [43671.0, 44207.0, 44745.0, 45281.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.9435
varray: [27.987, 27.929, 28.013, 27.845]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [43671.0, 44207.0, 44745.0, 45281.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0744647567645
varray: [27.987, 27.929, 28.013, 27.845]
inst: std_all
operation: std_all(slope_shift)
tarray: [43671.0, 44207.0, 44745.0, 45281.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0282016547978
varray: [-0.197, -0.206, -0.214, -0.151]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43671.0, 44207.0, 44745.0, 45281.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.124244382837
varray: [-8.665, -8.876, -8.922, -8.929]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43671.0, 44207.0, 44745.0, 45281.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 65.02075
varray: [65.385, 64.559, 65.372, 64.767]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [43671.0, 44207.0, 44745.0, 45281.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.421764843643
varray: [65.385, 64.559, 65.372, 64.767]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [43671.0, 44207.0, 44745.0, 45281.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.21
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.175
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 63.0045
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.9435
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.06525
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.06525
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 3.04231262385
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.20935
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-0.197, -0.206, -0.214, -0.151]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [43671.0, 44207.0, 44745.0, 45281.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-8.665, -8.876, -8.922, -8.929]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [43671.0, 44207.0, 44745.0, 45281.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.192
varray: [-0.197, -0.206, -0.214, -0.151]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [43671.0, 44207.0, 44745.0, 45281.0]
source_col: d(18_16)Mean
dD_mean | -8.848
varray: [-8.665, -8.876, -8.922, -8.929]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [43671.0, 44207.0, 44745.0, 45281.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0282016547978
varray: [-0.197, -0.206, -0.214, -0.151]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43671.0, 44207.0, 44745.0, 45281.0]
source_col: d(18_16)Mean
dD_stdev | 0.124244382837
varray: [-8.665, -8.876, -8.922, -8.929]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43671.0, 44207.0, 44745.0, 45281.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.192
varray: [-0.197, -0.206, -0.214, -0.151]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [43671.0, 44207.0, 44745.0, 45281.0]
source_col: d(18_16)Mean
measured_dD_mean | -8.848
varray: [-8.665, -8.876, -8.922, -8.929]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [43671.0, 44207.0, 44745.0, 45281.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0282016547978
varray: [-0.197, -0.206, -0.214, -0.151]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43671.0, 44207.0, 44745.0, 45281.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.124244382837
varray: [-8.665, -8.876, -8.922, -8.929]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43671.0, 44207.0, 44745.0, 45281.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.06720295853
actual_array: [ 1.93 -15.84]
varray: [-0.197, -0.206, -0.214, -0.151]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.192
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [43671.0, 44207.0, 44745.0, 45281.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -10.4817435513
actual_array: [ 18.1 -119. ]
varray: [-8.665, -8.876, -8.922, -8.929]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -8.848
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [43671.0, 44207.0, 44745.0, 45281.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.06735
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1.72715
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 16 - Sample 16 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 16
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.172
varray: [2.201, 2.158, 2.161, 2.168]
inst: mean_all
operation: mean_all(residuals)
tarray: [45870.0, 46406.0, 46943.0, 47479.0]
source_col: residuals
SAMP_residuals_STD | 0.0197821468333
varray: [2.201, 2.158, 2.161, 2.168]
inst: std_all
operation: std_all(residuals)
tarray: [45870.0, 46406.0, 46943.0, 47479.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.047
varray: [-0.913, -1.016, -1.078, -1.181]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [45870.0, 46406.0, 46943.0, 47479.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.112300192935
varray: [-0.913, -1.016, -1.078, -1.181]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [45870.0, 46406.0, 46943.0, 47479.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 61.68975
varray: [61.691, 61.559, 61.608, 61.901]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [45870.0, 46406.0, 46943.0, 47479.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.151004139016
varray: [61.691, 61.559, 61.608, 61.901]
inst: std_all
operation: std_all(baseline_shift)
tarray: [45870.0, 46406.0, 46943.0, 47479.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.48725
varray: [27.878, 27.179, 27.192, 27.7]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [45870.0, 46406.0, 46943.0, 47479.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.355967578112
varray: [27.878, 27.179, 27.192, 27.7]
inst: std_all
operation: std_all(slope_shift)
tarray: [45870.0, 46406.0, 46943.0, 47479.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0731778655059
varray: [2.912, 2.999, 3.075, 3.056]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45870.0, 46406.0, 46943.0, 47479.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.121882457037
varray: [-9.563, -9.316, -9.358, -9.299]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45870.0, 46406.0, 46943.0, 47479.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.30725
varray: [65.3, 63.854, 64.331, 63.744]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [45870.0, 46406.0, 46943.0, 47479.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.709178574126
varray: [65.3, 63.854, 64.331, 63.744]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [45870.0, 46406.0, 46943.0, 47479.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.172
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.047
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 61.68975
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.48725
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.02725
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.02725
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.27054435249
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.49585
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [2.912, 2.999, 3.075, 3.056]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [45870.0, 46406.0, 46943.0, 47479.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-9.563, -9.316, -9.358, -9.299]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [45870.0, 46406.0, 46943.0, 47479.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 3.0105
varray: [2.912, 2.999, 3.075, 3.056]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45870.0, 46406.0, 46943.0, 47479.0]
source_col: d(18_16)Mean
dD_mean | -9.384
varray: [-9.563, -9.316, -9.358, -9.299]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45870.0, 46406.0, 46943.0, 47479.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0731778655059
varray: [2.912, 2.999, 3.075, 3.056]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45870.0, 46406.0, 46943.0, 47479.0]
source_col: d(18_16)Mean
dD_stdev | 0.121882457037
varray: [-9.563, -9.316, -9.358, -9.299]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45870.0, 46406.0, 46943.0, 47479.0]
source_col: d(D_H)Mean
measured_d18O_mean | 3.0105
varray: [2.912, 2.999, 3.075, 3.056]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45870.0, 46406.0, 46943.0, 47479.0]
source_col: d(18_16)Mean
measured_dD_mean | -9.384
varray: [-9.563, -9.316, -9.358, -9.299]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45870.0, 46406.0, 46943.0, 47479.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0731778655059
varray: [2.912, 2.999, 3.075, 3.056]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45870.0, 46406.0, 46943.0, 47479.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.121882457037
varray: [-9.563, -9.316, -9.358, -9.299]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45870.0, 46406.0, 46943.0, 47479.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 2.23149196026
actual_array: [ 1.93 -15.84]
varray: [2.912, 2.999, 3.075, 3.056]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 3.0105
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [45870.0, 46406.0, 46943.0, 47479.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -11.0661369514
actual_array: [ 18.1 -119. ]
varray: [-9.563, -9.316, -9.358, -9.299]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -9.384
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [45870.0, 46406.0, 46943.0, 47479.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.10535
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.4124
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 17 - Sample 17 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 17
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.28375
varray: [2.234, 2.311, 2.274, 2.316]
inst: mean_all
operation: mean_all(residuals)
tarray: [48067.0, 48604.0, 49140.0, 49677.0]
source_col: residuals
SAMP_residuals_STD | 0.0380909000506
varray: [2.234, 2.311, 2.274, 2.316]
inst: std_all
operation: std_all(residuals)
tarray: [48067.0, 48604.0, 49140.0, 49677.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.026
varray: [-2.056, -2.109, -2.063, -1.876]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [48067.0, 48604.0, 49140.0, 49677.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.102726173231
varray: [-2.056, -2.109, -2.063, -1.876]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [48067.0, 48604.0, 49140.0, 49677.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 335.267
varray: [327.086, 339.206, 335.163, 339.613]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [48067.0, 48604.0, 49140.0, 49677.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 5.81214286358
varray: [327.086, 339.206, 335.163, 339.613]
inst: std_all
operation: std_all(baseline_shift)
tarray: [48067.0, 48604.0, 49140.0, 49677.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 16.9625
varray: [16.539, 17.348, 16.85, 17.113]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [48067.0, 48604.0, 49140.0, 49677.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.347979405521
varray: [16.539, 17.348, 16.85, 17.113]
inst: std_all
operation: std_all(slope_shift)
tarray: [48067.0, 48604.0, 49140.0, 49677.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0901974685528
varray: [-0.1, -0.252, -0.255, -0.31]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48067.0, 48604.0, 49140.0, 49677.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.983878549416
varray: [-21.889, -23.393, -23.831, -24.089]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48067.0, 48604.0, 49140.0, 49677.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 58.69925
varray: [57.546, 59.008, 58.478, 59.765]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [48067.0, 48604.0, 49140.0, 49677.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.93275340614
varray: [57.546, 59.008, 58.478, 59.765]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [48067.0, 48604.0, 49140.0, 49677.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.28375
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.026
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 335.267
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 16.9625
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.139
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.139
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 6.48094183471
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.11215
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-0.1, -0.252, -0.255, -0.31]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [48067.0, 48604.0, 49140.0, 49677.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-21.889, -23.393, -23.831, -24.089]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [48067.0, 48604.0, 49140.0, 49677.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.22925
varray: [-0.1, -0.252, -0.255, -0.31]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48067.0, 48604.0, 49140.0, 49677.0]
source_col: d(18_16)Mean
dD_mean | -23.3005
varray: [-21.889, -23.393, -23.831, -24.089]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48067.0, 48604.0, 49140.0, 49677.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0901974685528
varray: [-0.1, -0.252, -0.255, -0.31]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48067.0, 48604.0, 49140.0, 49677.0]
source_col: d(18_16)Mean
dD_stdev | 0.983878549416
varray: [-21.889, -23.393, -23.831, -24.089]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48067.0, 48604.0, 49140.0, 49677.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.22925
varray: [-0.1, -0.252, -0.255, -0.31]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48067.0, 48604.0, 49140.0, 49677.0]
source_col: d(18_16)Mean
measured_dD_mean | -23.3005
varray: [-21.889, -23.393, -23.831, -24.089]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48067.0, 48604.0, 49140.0, 49677.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0901974685528
varray: [-0.1, -0.252, -0.255, -0.31]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48067.0, 48604.0, 49140.0, 49677.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.983878549416
varray: [-21.889, -23.393, -23.831, -24.089]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48067.0, 48604.0, 49140.0, 49677.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.10557185337
actual_array: [ 1.93 -15.84]
varray: [-0.1, -0.252, -0.255, -0.31]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.22925
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [48067.0, 48604.0, 49140.0, 49677.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -26.2391047724
actual_array: [ 18.1 -119. ]
varray: [-21.889, -23.393, -23.831, -24.089]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -23.3005
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [48067.0, 48604.0, 49140.0, 49677.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 58.9936
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 273.98965
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 18 - Sample 18 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 18
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.19925
varray: [2.217, 2.203, 2.193, 2.184]
inst: mean_all
operation: mean_all(residuals)
tarray: [50264.0, 50801.0, 51338.0, 51875.0]
source_col: residuals
SAMP_residuals_STD | 0.0141509716981
varray: [2.217, 2.203, 2.193, 2.184]
inst: std_all
operation: std_all(residuals)
tarray: [50264.0, 50801.0, 51338.0, 51875.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.42275
varray: [-1.423, -1.48, -1.335, -1.453]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [50264.0, 50801.0, 51338.0, 51875.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0629622903014
varray: [-1.423, -1.48, -1.335, -1.453]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [50264.0, 50801.0, 51338.0, 51875.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 118.60775
varray: [120.237, 118.627, 117.805, 117.762]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [50264.0, 50801.0, 51338.0, 51875.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.15679539389
varray: [120.237, 118.627, 117.805, 117.762]
inst: std_all
operation: std_all(baseline_shift)
tarray: [50264.0, 50801.0, 51338.0, 51875.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 25.62625
varray: [25.794, 25.717, 25.401, 25.593]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [50264.0, 50801.0, 51338.0, 51875.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.171482506396
varray: [25.794, 25.717, 25.401, 25.593]
inst: std_all
operation: std_all(slope_shift)
tarray: [50264.0, 50801.0, 51338.0, 51875.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0190962474499
varray: [-1.002, -1.001, -1.042, -1.015]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50264.0, 50801.0, 51338.0, 51875.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.29326822194
varray: [-20.816, -20.42, -20.108, -20.357]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50264.0, 50801.0, 51338.0, 51875.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.61325
varray: [63.198, 62.23, 62.841, 62.184]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [50264.0, 50801.0, 51338.0, 51875.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.491575273992
varray: [63.198, 62.23, 62.841, 62.184]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [50264.0, 50801.0, 51338.0, 51875.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.19925
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.42275
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 118.60775
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 25.62625
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0545
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0545
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.54108870498
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.19815
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-1.002, -1.001, -1.042, -1.015]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [50264.0, 50801.0, 51338.0, 51875.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-20.816, -20.42, -20.108, -20.357]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [50264.0, 50801.0, 51338.0, 51875.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.015
varray: [-1.002, -1.001, -1.042, -1.015]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50264.0, 50801.0, 51338.0, 51875.0]
source_col: d(18_16)Mean
dD_mean | -20.42525
varray: [-20.816, -20.42, -20.108, -20.357]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50264.0, 50801.0, 51338.0, 51875.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0190962474499
varray: [-1.002, -1.001, -1.042, -1.015]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50264.0, 50801.0, 51338.0, 51875.0]
source_col: d(18_16)Mean
dD_stdev | 0.29326822194
varray: [-20.816, -20.42, -20.108, -20.357]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50264.0, 50801.0, 51338.0, 51875.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.015
varray: [-1.002, -1.001, -1.042, -1.015]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50264.0, 50801.0, 51338.0, 51875.0]
source_col: d(18_16)Mean
measured_dD_mean | -20.42525
varray: [-20.816, -20.42, -20.108, -20.357]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50264.0, 50801.0, 51338.0, 51875.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0190962474499
varray: [-1.002, -1.001, -1.042, -1.015]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50264.0, 50801.0, 51338.0, 51875.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.29326822194
varray: [-20.816, -20.42, -20.108, -20.357]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50264.0, 50801.0, 51338.0, 51875.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.91492377607
actual_array: [ 1.93 -15.84]
varray: [-1.002, -1.001, -1.042, -1.015]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.015
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [50264.0, 50801.0, 51338.0, 51875.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -23.1042593927
actual_array: [ 18.1 -119. ]
varray: [-20.816, -20.42, -20.108, -20.357]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -20.42525
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [50264.0, 50801.0, 51338.0, 51875.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.0781
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 57.3304
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 19 - Sample 19 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 19
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.158
varray: [2.165, 2.159, 2.154, 2.154]
inst: mean_all
operation: mean_all(residuals)
tarray: [52464.0, 53001.0, 53537.0, 54073.0]
source_col: residuals
SAMP_residuals_STD | 0.00522812904712
varray: [2.165, 2.159, 2.154, 2.154]
inst: std_all
operation: std_all(residuals)
tarray: [52464.0, 53001.0, 53537.0, 54073.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.42825
varray: [-1.622, -1.35, -1.381, -1.36]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [52464.0, 53001.0, 53537.0, 54073.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.129811080678
varray: [-1.622, -1.35, -1.381, -1.36]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [52464.0, 53001.0, 53537.0, 54073.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 71.86425
varray: [72.299, 71.848, 71.832, 71.478]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [52464.0, 53001.0, 53537.0, 54073.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.336402908232
varray: [72.299, 71.848, 71.832, 71.478]
inst: std_all
operation: std_all(baseline_shift)
tarray: [52464.0, 53001.0, 53537.0, 54073.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.0745
varray: [26.884, 25.995, 25.476, 25.943]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [52464.0, 53001.0, 53537.0, 54073.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.587964001166
varray: [26.884, 25.995, 25.476, 25.943]
inst: std_all
operation: std_all(slope_shift)
tarray: [52464.0, 53001.0, 53537.0, 54073.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0128679187646
varray: [-0.89, -0.92, -0.903, -0.912]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52464.0, 53001.0, 53537.0, 54073.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.122055998077
varray: [-19.84, -20.085, -19.898, -20.067]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52464.0, 53001.0, 53537.0, 54073.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.68025
varray: [59.866, 62.432, 62.85, 61.573]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [52464.0, 53001.0, 53537.0, 54073.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.32116725537
varray: [59.866, 62.432, 62.85, 61.573]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [52464.0, 53001.0, 53537.0, 54073.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.158
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.42825
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 71.86425
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.0745
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.01325
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.01325
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.617787620935
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.13115
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-0.89, -0.92, -0.903, -0.912]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [52464.0, 53001.0, 53537.0, 54073.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-19.84, -20.085, -19.898, -20.067]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [52464.0, 53001.0, 53537.0, 54073.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.90625
varray: [-0.89, -0.92, -0.903, -0.912]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [52464.0, 53001.0, 53537.0, 54073.0]
source_col: d(18_16)Mean
dD_mean | -19.9725
varray: [-19.84, -20.085, -19.898, -20.067]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [52464.0, 53001.0, 53537.0, 54073.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0128679187646
varray: [-0.89, -0.92, -0.903, -0.912]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52464.0, 53001.0, 53537.0, 54073.0]
source_col: d(18_16)Mean
dD_stdev | 0.122055998077
varray: [-19.84, -20.085, -19.898, -20.067]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52464.0, 53001.0, 53537.0, 54073.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.90625
varray: [-0.89, -0.92, -0.903, -0.912]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [52464.0, 53001.0, 53537.0, 54073.0]
source_col: d(18_16)Mean
measured_dD_mean | -19.9725
varray: [-19.84, -20.085, -19.898, -20.067]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [52464.0, 53001.0, 53537.0, 54073.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0128679187646
varray: [-0.89, -0.92, -0.903, -0.912]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52464.0, 53001.0, 53537.0, 54073.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.122055998077
varray: [-19.84, -20.085, -19.898, -20.067]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52464.0, 53001.0, 53537.0, 54073.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.80290720389
actual_array: [ 1.93 -15.84]
varray: [-0.89, -0.92, -0.903, -0.912]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.90625
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [52464.0, 53001.0, 53537.0, 54073.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -22.6106323183
actual_array: [ 18.1 -119. ]
varray: [-19.84, -20.085, -19.898, -20.067]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -19.9725
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [52464.0, 53001.0, 53537.0, 54073.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.11935
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 10.5869
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 20 - Sample 20 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 20
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.204
varray: [2.198, 2.175, 2.214, 2.229]
inst: mean_all
operation: mean_all(residuals)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: residuals
SAMP_residuals_STD | 0.0231084400166
varray: [2.198, 2.175, 2.214, 2.229]
inst: std_all
operation: std_all(residuals)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.45675
varray: [-1.617, -1.435, -1.453, -1.322]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.121552115023
varray: [-1.617, -1.435, -1.453, -1.322]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 100.583
varray: [100.29, 99.843, 101.205, 100.994]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.629617344107
varray: [100.29, 99.843, 101.205, 100.994]
inst: std_all
operation: std_all(baseline_shift)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 25.08825
varray: [25.61, 24.979, 24.778, 24.986]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.360956483988
varray: [25.61, 24.979, 24.778, 24.986]
inst: std_all
operation: std_all(slope_shift)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0317634485953
varray: [-0.872, -0.919, -0.864, -0.926]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.25622434831
varray: [-22.0, -22.397, -22.381, -22.617]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.029
varray: [62.071, 61.876, 63.79, 64.379]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.24482850224
varray: [62.071, 61.876, 63.79, 64.379]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.204
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.45675
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 100.583
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 25.08825
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.05925
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.05925
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.7625597389
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.7824
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-0.872, -0.919, -0.864, -0.926]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-22.0, -22.397, -22.381, -22.617]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.89525
varray: [-0.872, -0.919, -0.864, -0.926]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: d(18_16)Mean
dD_mean | -22.34875
varray: [-22.0, -22.397, -22.381, -22.617]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0317634485953
varray: [-0.872, -0.919, -0.864, -0.926]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: d(18_16)Mean
dD_stdev | 0.25622434831
varray: [-22.0, -22.397, -22.381, -22.617]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.89525
varray: [-0.872, -0.919, -0.864, -0.926]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: d(18_16)Mean
measured_dD_mean | -22.34875
varray: [-22.0, -22.397, -22.381, -22.617]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0317634485953
varray: [-0.872, -0.919, -0.864, -0.926]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.25622434831
varray: [-22.0, -22.397, -22.381, -22.617]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.79157679199
actual_array: [ 1.93 -15.84]
varray: [-0.872, -0.919, -0.864, -0.926]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.89525
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -25.2014248872
actual_array: [ 18.1 -119. ]
varray: [-22.0, -22.397, -22.381, -22.617]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -22.34875
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.07335
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 39.30565
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 21 - Sample 21 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 21
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.2825
varray: [2.331, 2.306, 2.241, 2.252]
inst: mean_all
operation: mean_all(residuals)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: residuals
SAMP_residuals_STD | 0.0430387422369
varray: [2.331, 2.306, 2.241, 2.252]
inst: std_all
operation: std_all(residuals)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.0405
varray: [-1.998, -1.867, -2.241, -2.056]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.155294344177
varray: [-1.998, -1.867, -2.241, -2.056]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 600.884
varray: [609.611, 602.014, 595.525, 596.386]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 6.49071937461
varray: [609.611, 602.014, 595.525, 596.386]
inst: std_all
operation: std_all(baseline_shift)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 8.02425
varray: [8.087, 7.711, 8.106, 8.193]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.2138712619
varray: [8.087, 7.711, 8.106, 8.193]
inst: std_all
operation: std_all(slope_shift)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0265643871879
varray: [-0.204, -0.145, -0.151, -0.17]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.533820428609
varray: [-14.371, -13.5, -13.255, -13.237]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 54.42075
varray: [56.252, 55.696, 52.743, 52.992]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.81070067009
varray: [56.252, 55.696, 52.743, 52.992]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.2825
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.0405
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 600.884
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 8.02425
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.13775
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.13775
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 6.42265998368
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -9.39065
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-0.204, -0.145, -0.151, -0.17]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-14.371, -13.5, -13.255, -13.237]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.1675
varray: [-0.204, -0.145, -0.151, -0.17]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: d(18_16)Mean
dD_mean | -13.59075
varray: [-14.371, -13.5, -13.255, -13.237]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0265643871879
varray: [-0.204, -0.145, -0.151, -0.17]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: d(18_16)Mean
dD_stdev | 0.533820428609
varray: [-14.371, -13.5, -13.255, -13.237]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.1675
varray: [-0.204, -0.145, -0.151, -0.17]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: d(18_16)Mean
measured_dD_mean | -13.59075
varray: [-14.371, -13.5, -13.255, -13.237]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0265643871879
varray: [-0.204, -0.145, -0.151, -0.17]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.533820428609
varray: [-14.371, -13.5, -13.255, -13.237]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.04196704112
actual_array: [ 1.93 -15.84]
varray: [-0.204, -0.145, -0.151, -0.17]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.1675
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -15.6526983987
actual_array: [ 18.1 -119. ]
varray: [-14.371, -13.5, -13.255, -13.237]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -13.59075
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 58.99485
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 539.60665
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 22 - Sample 22 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 22
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.07625
varray: [2.11, 2.068, 2.065, 2.062]
inst: mean_all
operation: mean_all(residuals)
tarray: [59056.0, 59591.0, 60127.0, 60662.0]
source_col: residuals
SAMP_residuals_STD | 0.0226329405955
varray: [2.11, 2.068, 2.065, 2.062]
inst: std_all
operation: std_all(residuals)
tarray: [59056.0, 59591.0, 60127.0, 60662.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.95425
varray: [-1.103, -0.881, -0.816, -1.017]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [59056.0, 59591.0, 60127.0, 60662.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.129798240871
varray: [-1.103, -0.881, -0.816, -1.017]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [59056.0, 59591.0, 60127.0, 60662.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 63.7685
varray: [67.185, 63.765, 62.314, 61.81]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [59056.0, 59591.0, 60127.0, 60662.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.42375583204
varray: [67.185, 63.765, 62.314, 61.81]
inst: std_all
operation: std_all(baseline_shift)
tarray: [59056.0, 59591.0, 60127.0, 60662.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.345
varray: [28.142, 27.012, 26.998, 27.228]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [59056.0, 59591.0, 60127.0, 60662.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.541662871782
varray: [28.142, 27.012, 26.998, 27.228]
inst: std_all
operation: std_all(slope_shift)
tarray: [59056.0, 59591.0, 60127.0, 60662.0]
source_col: slope_shift
SAMP_d18O_STD | 0.174215144768
varray: [-8.622, -8.875, -8.98, -9.002]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59056.0, 59591.0, 60127.0, 60662.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.69973534197
varray: [-59.422, -63.799, -64.874, -65.328]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59056.0, 59591.0, 60127.0, 60662.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.7455
varray: [63.252, 62.674, 62.711, 62.345]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [59056.0, 59591.0, 60127.0, 60662.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.375608395362
varray: [63.252, 62.674, 62.711, 62.345]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [59056.0, 59591.0, 60127.0, 60662.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.07625
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.95425
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 63.7685
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.345
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0685
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0685
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.19384543653
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.0659
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-8.622, -8.875, -8.98, -9.002]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [59056.0, 59591.0, 60127.0, 60662.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-59.422, -63.799, -64.874, -65.328]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [59056.0, 59591.0, 60127.0, 60662.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.86975
varray: [-8.622, -8.875, -8.98, -9.002]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [59056.0, 59591.0, 60127.0, 60662.0]
source_col: d(18_16)Mean
dD_mean | -63.35575
varray: [-59.422, -63.799, -64.874, -65.328]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [59056.0, 59591.0, 60127.0, 60662.0]
source_col: d(D_H)Mean
d18O_stdev | 0.174215144768
varray: [-8.622, -8.875, -8.98, -9.002]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59056.0, 59591.0, 60127.0, 60662.0]
source_col: d(18_16)Mean
dD_stdev | 2.69973534197
varray: [-59.422, -63.799, -64.874, -65.328]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59056.0, 59591.0, 60127.0, 60662.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.86975
varray: [-8.622, -8.875, -8.98, -9.002]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [59056.0, 59591.0, 60127.0, 60662.0]
source_col: d(18_16)Mean
measured_dD_mean | -63.35575
varray: [-59.422, -63.799, -64.874, -65.328]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [59056.0, 59591.0, 60127.0, 60662.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.174215144768
varray: [-8.622, -8.875, -8.98, -9.002]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59056.0, 59591.0, 60127.0, 60662.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.69973534197
varray: [-59.422, -63.799, -64.874, -65.328]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59056.0, 59591.0, 60127.0, 60662.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.0056104001
actual_array: [ 1.93 -15.84]
varray: [-8.622, -8.875, -8.98, -9.002]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -8.86975
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [59056.0, 59591.0, 60127.0, 60662.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -69.9107908511
actual_array: [ 18.1 -119. ]
varray: [-59.422, -63.799, -64.874, -65.328]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -63.35575
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [59056.0, 59591.0, 60127.0, 60662.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.2011
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 2.49115
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 23 - Sample 23 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 23
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.8525
varray: [2.436, 1.58, 1.781, 1.613]
inst: mean_all
operation: mean_all(residuals)
tarray: [61248.0, 61785.0, 62322.0, 62858.0]
source_col: residuals
SAMP_residuals_STD | 0.398832044851
varray: [2.436, 1.58, 1.781, 1.613]
inst: std_all
operation: std_all(residuals)
tarray: [61248.0, 61785.0, 62322.0, 62858.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.24275
varray: [-1.354, -1.22, -1.259, -1.138]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [61248.0, 61785.0, 62322.0, 62858.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0896860264107
varray: [-1.354, -1.22, -1.259, -1.138]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [61248.0, 61785.0, 62322.0, 62858.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 140.26025
varray: [174.488, 124.366, 135.56, 126.627]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [61248.0, 61785.0, 62322.0, 62858.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 23.3246929865
varray: [174.488, 124.366, 135.56, 126.627]
inst: std_all
operation: std_all(baseline_shift)
tarray: [61248.0, 61785.0, 62322.0, 62858.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 18.08975
varray: [26.498, 14.464, 16.972, 14.425]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [61248.0, 61785.0, 62322.0, 62858.0]
source_col: slope_shift
SAMP_slope_shift_STD | 5.73075005998
varray: [26.498, 14.464, 16.972, 14.425]
inst: std_all
operation: std_all(slope_shift)
tarray: [61248.0, 61785.0, 62322.0, 62858.0]
source_col: slope_shift
SAMP_d18O_STD | 0.167692923723
varray: [-1.779, -1.51, -1.426, -1.424]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61248.0, 61785.0, 62322.0, 62858.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.22960033788
varray: [-22.247, -18.58, -17.764, -17.367]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61248.0, 61785.0, 62322.0, 62858.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 53.98925
varray: [67.917, 47.048, 52.21, 48.782]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [61248.0, 61785.0, 62322.0, 62858.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 9.52967916127
varray: [67.917, 47.048, 52.21, 48.782]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [61248.0, 61785.0, 62322.0, 62858.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.8525
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.24275
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 140.26025
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 18.08975
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.29225
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.29225
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -13.6262967712
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -9.82215
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-1.779, -1.51, -1.426, -1.424]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [61248.0, 61785.0, 62322.0, 62858.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-22.247, -18.58, -17.764, -17.367]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [61248.0, 61785.0, 62322.0, 62858.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.53475
varray: [-1.779, -1.51, -1.426, -1.424]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [61248.0, 61785.0, 62322.0, 62858.0]
source_col: d(18_16)Mean
dD_mean | -18.9895
varray: [-22.247, -18.58, -17.764, -17.367]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [61248.0, 61785.0, 62322.0, 62858.0]
source_col: d(D_H)Mean
d18O_stdev | 0.167692923723
varray: [-1.779, -1.51, -1.426, -1.424]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61248.0, 61785.0, 62322.0, 62858.0]
source_col: d(18_16)Mean
dD_stdev | 2.22960033788
varray: [-22.247, -18.58, -17.764, -17.367]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61248.0, 61785.0, 62322.0, 62858.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.53475
varray: [-1.779, -1.51, -1.426, -1.424]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [61248.0, 61785.0, 62322.0, 62858.0]
source_col: d(18_16)Mean
measured_dD_mean | -18.9895
varray: [-22.247, -18.58, -17.764, -17.367]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [61248.0, 61785.0, 62322.0, 62858.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.167692923723
varray: [-1.779, -1.51, -1.426, -1.424]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61248.0, 61785.0, 62322.0, 62858.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.22960033788
varray: [-22.247, -18.58, -17.764, -17.367]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61248.0, 61785.0, 62322.0, 62858.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.4502857383
actual_array: [ 1.93 -15.84]
varray: [-1.779, -1.51, -1.426, -1.424]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.53475
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [61248.0, 61785.0, 62322.0, 62858.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -21.5388809894
actual_array: [ 18.1 -119. ]
varray: [-22.247, -18.58, -17.764, -17.367]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -18.9895
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [61248.0, 61785.0, 62322.0, 62858.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.42485
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 78.9829
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 24 - Sample 24 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 24
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.14675
varray: [2.151, 2.132, 2.141, 2.163]
inst: mean_all
operation: mean_all(residuals)
tarray: [63445.0, 63981.0, 64518.0, 65054.0]
source_col: residuals
SAMP_residuals_STD | 0.0133260396718
varray: [2.151, 2.132, 2.141, 2.163]
inst: std_all
operation: std_all(residuals)
tarray: [63445.0, 63981.0, 64518.0, 65054.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.48725
varray: [-1.359, -1.39, -1.419, -1.781]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [63445.0, 63981.0, 64518.0, 65054.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.197359865896
varray: [-1.359, -1.39, -1.419, -1.781]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [63445.0, 63981.0, 64518.0, 65054.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 122.1585
varray: [122.691, 121.417, 122.148, 122.378]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [63445.0, 63981.0, 64518.0, 65054.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.542115916264
varray: [122.691, 121.417, 122.148, 122.378]
inst: std_all
operation: std_all(baseline_shift)
tarray: [63445.0, 63981.0, 64518.0, 65054.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 23.75975
varray: [23.469, 23.173, 23.877, 24.52]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [63445.0, 63981.0, 64518.0, 65054.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.583249160594
varray: [23.469, 23.173, 23.877, 24.52]
inst: std_all
operation: std_all(slope_shift)
tarray: [63445.0, 63981.0, 64518.0, 65054.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0321597574618
varray: [-2.275, -2.326, -2.35, -2.302]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63445.0, 63981.0, 64518.0, 65054.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.463423852069
varray: [-25.754, -26.673, -26.66, -26.707]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63445.0, 63981.0, 64518.0, 65054.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 60.46325
varray: [61.38, 61.03, 60.373, 59.07]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [63445.0, 63981.0, 64518.0, 65054.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.01831998082
varray: [61.38, 61.03, 60.373, 59.07]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [63445.0, 63981.0, 64518.0, 65054.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.14675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.48725
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 122.1585
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 23.75975
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.002
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.002
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.0932509616505
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.34815
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-2.275, -2.326, -2.35, -2.302]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [63445.0, 63981.0, 64518.0, 65054.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-25.754, -26.673, -26.66, -26.707]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [63445.0, 63981.0, 64518.0, 65054.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.31325
varray: [-2.275, -2.326, -2.35, -2.302]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63445.0, 63981.0, 64518.0, 65054.0]
source_col: d(18_16)Mean
dD_mean | -26.4485
varray: [-25.754, -26.673, -26.66, -26.707]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63445.0, 63981.0, 64518.0, 65054.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0321597574618
varray: [-2.275, -2.326, -2.35, -2.302]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63445.0, 63981.0, 64518.0, 65054.0]
source_col: d(18_16)Mean
dD_stdev | 0.463423852069
varray: [-25.754, -26.673, -26.66, -26.707]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63445.0, 63981.0, 64518.0, 65054.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.31325
varray: [-2.275, -2.326, -2.35, -2.302]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63445.0, 63981.0, 64518.0, 65054.0]
source_col: d(18_16)Mean
measured_dD_mean | -26.4485
varray: [-25.754, -26.673, -26.66, -26.707]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63445.0, 63981.0, 64518.0, 65054.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0321597574618
varray: [-2.275, -2.326, -2.35, -2.302]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63445.0, 63981.0, 64518.0, 65054.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.463423852069
varray: [-25.754, -26.673, -26.66, -26.707]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63445.0, 63981.0, 64518.0, 65054.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.25216988952
actual_array: [ 1.93 -15.84]
varray: [-2.275, -2.326, -2.35, -2.302]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -2.31325
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [63445.0, 63981.0, 64518.0, 65054.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -29.6713257117
actual_array: [ 18.1 -119. ]
varray: [-25.754, -26.673, -26.66, -26.707]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -26.4485
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [63445.0, 63981.0, 64518.0, 65054.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.1306
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 60.88115
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 25 - Sample 25 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 25
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.997
varray: [2.667, 1.727, 1.769, 1.825]
inst: mean_all
operation: mean_all(residuals)
tarray: [65642.0, 66177.0, 66713.0, 67249.0]
source_col: residuals
SAMP_residuals_STD | 0.448467018483
varray: [2.667, 1.727, 1.769, 1.825]
inst: std_all
operation: std_all(residuals)
tarray: [65642.0, 66177.0, 66713.0, 67249.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.6225
varray: [-2.094, -1.332, -1.546, -1.518]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [65642.0, 66177.0, 66713.0, 67249.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.328367172537
varray: [-2.094, -1.332, -1.546, -1.518]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [65642.0, 66177.0, 66713.0, 67249.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 67.33025
varray: [78.161, 63.085, 63.652, 64.423]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [65642.0, 66177.0, 66713.0, 67249.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 7.24129175286
varray: [78.161, 63.085, 63.652, 64.423]
inst: std_all
operation: std_all(baseline_shift)
tarray: [65642.0, 66177.0, 66713.0, 67249.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 20.501
varray: [30.391, 16.097, 17.456, 18.06]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [65642.0, 66177.0, 66713.0, 67249.0]
source_col: slope_shift
SAMP_slope_shift_STD | 6.6442411656
varray: [30.391, 16.097, 17.456, 18.06]
inst: std_all
operation: std_all(slope_shift)
tarray: [65642.0, 66177.0, 66713.0, 67249.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0746770826068
varray: [1.02, 1.169, 1.168, 1.171]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65642.0, 66177.0, 66713.0, 67249.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.614398011607
varray: [-15.089, -14.069, -13.77, -13.825]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65642.0, 66177.0, 66713.0, 67249.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 57.2425
varray: [71.777, 52.167, 51.768, 53.258]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [65642.0, 66177.0, 66713.0, 67249.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 9.71011124207
varray: [71.777, 52.167, 51.768, 53.258]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [65642.0, 66177.0, 66713.0, 67249.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.997
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.6225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 67.33025
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 20.501
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.14775
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.14775
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.88891479193
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -6.5689
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [1.02, 1.169, 1.168, 1.171]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [65642.0, 66177.0, 66713.0, 67249.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-15.089, -14.069, -13.77, -13.825]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [65642.0, 66177.0, 66713.0, 67249.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.132
varray: [1.02, 1.169, 1.168, 1.171]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65642.0, 66177.0, 66713.0, 67249.0]
source_col: d(18_16)Mean
dD_mean | -14.18825
varray: [-15.089, -14.069, -13.77, -13.825]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65642.0, 66177.0, 66713.0, 67249.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0746770826068
varray: [1.02, 1.169, 1.168, 1.171]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65642.0, 66177.0, 66713.0, 67249.0]
source_col: d(18_16)Mean
dD_stdev | 0.614398011607
varray: [-15.089, -14.069, -13.77, -13.825]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65642.0, 66177.0, 66713.0, 67249.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.132
varray: [1.02, 1.169, 1.168, 1.171]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65642.0, 66177.0, 66713.0, 67249.0]
source_col: d(18_16)Mean
measured_dD_mean | -14.18825
varray: [-15.089, -14.069, -13.77, -13.825]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65642.0, 66177.0, 66713.0, 67249.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0746770826068
varray: [1.02, 1.169, 1.168, 1.171]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65642.0, 66177.0, 66713.0, 67249.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.614398011607
varray: [-15.089, -14.069, -13.77, -13.825]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65642.0, 66177.0, 66713.0, 67249.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 0.296566619135
actual_array: [ 1.93 -15.84]
varray: [1.02, 1.169, 1.168, 1.171]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 1.132
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [65642.0, 66177.0, 66713.0, 67249.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -16.3041443997
actual_array: [ 18.1 -119. ]
varray: [-15.089, -14.069, -13.77, -13.825]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -14.18825
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [65642.0, 66177.0, 66713.0, 67249.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.28035
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 6.0529
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 26 - Sample 26 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 26
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.29125
varray: [2.231, 2.296, 2.299, 2.339]
inst: mean_all
operation: mean_all(residuals)
tarray: [67837.0, 68373.0, 68909.0, 69446.0]
source_col: residuals
SAMP_residuals_STD | 0.0446943322283
varray: [2.231, 2.296, 2.299, 2.339]
inst: std_all
operation: std_all(residuals)
tarray: [67837.0, 68373.0, 68909.0, 69446.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.641
varray: [-1.608, -1.699, -1.766, -1.491]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [67837.0, 68373.0, 68909.0, 69446.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.119132978921
varray: [-1.608, -1.699, -1.766, -1.491]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [67837.0, 68373.0, 68909.0, 69446.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 64.9555
varray: [64.274, 65.192, 65.151, 65.205]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [67837.0, 68373.0, 68909.0, 69446.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.454915743114
varray: [64.274, 65.192, 65.151, 65.205]
inst: std_all
operation: std_all(baseline_shift)
tarray: [67837.0, 68373.0, 68909.0, 69446.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.0885
varray: [25.496, 25.91, 26.215, 26.733]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [67837.0, 68373.0, 68909.0, 69446.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.520992962205
varray: [25.496, 25.91, 26.215, 26.733]
inst: std_all
operation: std_all(slope_shift)
tarray: [67837.0, 68373.0, 68909.0, 69446.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0177270979012
varray: [1.112, 1.139, 1.142, 1.154]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67837.0, 68373.0, 68909.0, 69446.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.356060856222
varray: [-10.958, -10.385, -10.561, -10.108]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67837.0, 68373.0, 68909.0, 69446.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.761
varray: [62.797, 64.922, 64.502, 66.823]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [67837.0, 68373.0, 68909.0, 69446.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.65349548936
varray: [62.797, 64.922, 64.502, 66.823]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [67837.0, 68373.0, 68909.0, 69446.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.29125
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.641
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 64.9555
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.0885
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.1465
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1465
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 6.8306329409
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.9496
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [1.112, 1.139, 1.142, 1.154]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [67837.0, 68373.0, 68909.0, 69446.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-10.958, -10.385, -10.561, -10.108]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [67837.0, 68373.0, 68909.0, 69446.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.13675
varray: [1.112, 1.139, 1.142, 1.154]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [67837.0, 68373.0, 68909.0, 69446.0]
source_col: d(18_16)Mean
dD_mean | -10.503
varray: [-10.958, -10.385, -10.561, -10.108]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [67837.0, 68373.0, 68909.0, 69446.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0177270979012
varray: [1.112, 1.139, 1.142, 1.154]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67837.0, 68373.0, 68909.0, 69446.0]
source_col: d(18_16)Mean
dD_stdev | 0.356060856222
varray: [-10.958, -10.385, -10.561, -10.108]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67837.0, 68373.0, 68909.0, 69446.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.13675
varray: [1.112, 1.139, 1.142, 1.154]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [67837.0, 68373.0, 68909.0, 69446.0]
source_col: d(18_16)Mean
measured_dD_mean | -10.503
varray: [-10.958, -10.385, -10.561, -10.108]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [67837.0, 68373.0, 68909.0, 69446.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0177270979012
varray: [1.112, 1.139, 1.142, 1.154]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67837.0, 68373.0, 68909.0, 69446.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.356060856222
varray: [-10.958, -10.385, -10.561, -10.108]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67837.0, 68373.0, 68909.0, 69446.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 0.301459297001
actual_array: [ 1.93 -15.84]
varray: [1.112, 1.139, 1.142, 1.154]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 1.13675
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [67837.0, 68373.0, 68909.0, 69446.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -12.2861672027
actual_array: [ 18.1 -119. ]
varray: [-10.958, -10.385, -10.561, -10.108]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -10.503
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [67837.0, 68373.0, 68909.0, 69446.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 58.9861
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 3.67815
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 27 - Sample 27 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 27
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.23875
varray: [2.213, 2.268, 2.241, 2.233]
inst: mean_all
operation: mean_all(residuals)
tarray: [70034.0, 70570.0, 71106.0, 71642.0]
source_col: residuals
SAMP_residuals_STD | 0.0227797424627
varray: [2.213, 2.268, 2.241, 2.233]
inst: std_all
operation: std_all(residuals)
tarray: [70034.0, 70570.0, 71106.0, 71642.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.67425
varray: [-1.568, -1.496, -1.857, -1.776]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [70034.0, 70570.0, 71106.0, 71642.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.170110503301
varray: [-1.568, -1.496, -1.857, -1.776]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [70034.0, 70570.0, 71106.0, 71642.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 68.311
varray: [67.842, 68.481, 68.608, 68.313]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [70034.0, 70570.0, 71106.0, 71642.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.335198448684
varray: [67.842, 68.481, 68.608, 68.313]
inst: std_all
operation: std_all(baseline_shift)
tarray: [70034.0, 70570.0, 71106.0, 71642.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.3055
varray: [25.715, 26.704, 26.349, 26.454]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [70034.0, 70570.0, 71106.0, 71642.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.420886762285
varray: [25.715, 26.704, 26.349, 26.454]
inst: std_all
operation: std_all(slope_shift)
tarray: [70034.0, 70570.0, 71106.0, 71642.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0446056797579
varray: [0.0, -0.052, -0.077, -0.105]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70034.0, 70570.0, 71106.0, 71642.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.149992222021
varray: [-14.443, -14.696, -14.793, -14.698]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70034.0, 70570.0, 71106.0, 71642.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.40175
varray: [63.198, 64.599, 62.81, 63.0]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [70034.0, 70570.0, 71106.0, 71642.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.813734754081
varray: [63.198, 64.599, 62.81, 63.0]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [70034.0, 70570.0, 71106.0, 71642.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.23875
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.67425
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 68.311
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.3055
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.094
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.094
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 4.38279519758
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.40965
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [0.0, -0.052, -0.077, -0.105]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [70034.0, 70570.0, 71106.0, 71642.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-14.443, -14.696, -14.793, -14.698]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [70034.0, 70570.0, 71106.0, 71642.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.0585
varray: [0.0, -0.052, -0.077, -0.105]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [70034.0, 70570.0, 71106.0, 71642.0]
source_col: d(18_16)Mean
dD_mean | -14.6575
varray: [-14.443, -14.696, -14.793, -14.698]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [70034.0, 70570.0, 71106.0, 71642.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0446056797579
varray: [0.0, -0.052, -0.077, -0.105]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70034.0, 70570.0, 71106.0, 71642.0]
source_col: d(18_16)Mean
dD_stdev | 0.149992222021
varray: [-14.443, -14.696, -14.793, -14.698]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70034.0, 70570.0, 71106.0, 71642.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.0585
varray: [0.0, -0.052, -0.077, -0.105]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [70034.0, 70570.0, 71106.0, 71642.0]
source_col: d(18_16)Mean
measured_dD_mean | -14.6575
varray: [-14.443, -14.696, -14.793, -14.698]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [70034.0, 70570.0, 71106.0, 71642.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0446056797579
varray: [0.0, -0.052, -0.077, -0.105]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70034.0, 70570.0, 71106.0, 71642.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.149992222021
varray: [-14.443, -14.696, -14.793, -14.698]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70034.0, 70570.0, 71106.0, 71642.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.929692959575
actual_array: [ 1.93 -15.84]
varray: [0.0, -0.052, -0.077, -0.105]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.0585
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [70034.0, 70570.0, 71106.0, 71642.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -16.8157611963
actual_array: [ 18.1 -119. ]
varray: [-14.443, -14.696, -14.793, -14.698]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -14.6575
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [70034.0, 70570.0, 71106.0, 71642.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.0386
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 7.03365
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 28 - Sample 28 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 28
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.4285
varray: [2.46, 2.361, 2.465, 2.428]
inst: mean_all
operation: mean_all(residuals)
tarray: [72230.0, 72766.0, 73302.0, 73839.0]
source_col: residuals
SAMP_residuals_STD | 0.0478922401508
varray: [2.46, 2.361, 2.465, 2.428]
inst: std_all
operation: std_all(residuals)
tarray: [72230.0, 72766.0, 73302.0, 73839.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.26375
varray: [-2.197, -2.247, -2.209, -2.402]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [72230.0, 72766.0, 73302.0, 73839.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0945987138743
varray: [-2.197, -2.247, -2.209, -2.402]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [72230.0, 72766.0, 73302.0, 73839.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 71.41625
varray: [71.776, 70.493, 71.902, 71.494]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [72230.0, 72766.0, 73302.0, 73839.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.638698872709
varray: [71.776, 70.493, 71.902, 71.494]
inst: std_all
operation: std_all(baseline_shift)
tarray: [72230.0, 72766.0, 73302.0, 73839.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 25.1795
varray: [25.819, 24.387, 25.46, 25.052]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [72230.0, 72766.0, 73302.0, 73839.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.614261887689
varray: [25.819, 24.387, 25.46, 25.052]
inst: std_all
operation: std_all(slope_shift)
tarray: [72230.0, 72766.0, 73302.0, 73839.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0260895892391
varray: [0.529, 0.479, 0.498, 0.47]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72230.0, 72766.0, 73302.0, 73839.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.34531278092
varray: [-13.595, -14.383, -14.088, -13.789]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72230.0, 72766.0, 73302.0, 73839.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 65.80775
varray: [66.572, 64.002, 66.792, 65.865]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [72230.0, 72766.0, 73302.0, 73839.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.26712755869
varray: [66.572, 64.002, 66.792, 65.865]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [72230.0, 72766.0, 73302.0, 73839.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.4285
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.26375
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 71.41625
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 25.1795
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.28375
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.28375
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 13.2299801842
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.99635
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [0.529, 0.479, 0.498, 0.47]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [72230.0, 72766.0, 73302.0, 73839.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-13.595, -14.383, -14.088, -13.789]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [72230.0, 72766.0, 73302.0, 73839.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.494
varray: [0.529, 0.479, 0.498, 0.47]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72230.0, 72766.0, 73302.0, 73839.0]
source_col: d(18_16)Mean
dD_mean | -13.96375
varray: [-13.595, -14.383, -14.088, -13.789]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72230.0, 72766.0, 73302.0, 73839.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0260895892391
varray: [0.529, 0.479, 0.498, 0.47]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72230.0, 72766.0, 73302.0, 73839.0]
source_col: d(18_16)Mean
dD_stdev | 0.34531278092
varray: [-13.595, -14.383, -14.088, -13.789]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72230.0, 72766.0, 73302.0, 73839.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.494
varray: [0.529, 0.479, 0.498, 0.47]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72230.0, 72766.0, 73302.0, 73839.0]
source_col: d(18_16)Mean
measured_dD_mean | -13.96375
varray: [-13.595, -14.383, -14.088, -13.789]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72230.0, 72766.0, 73302.0, 73839.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0260895892391
varray: [0.529, 0.479, 0.498, 0.47]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72230.0, 72766.0, 73302.0, 73839.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.34531278092
varray: [-13.595, -14.383, -14.088, -13.789]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72230.0, 72766.0, 73302.0, 73839.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.360597271009
actual_array: [ 1.93 -15.84]
varray: [0.529, 0.479, 0.498, 0.47]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 0.494
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [72230.0, 72766.0, 73302.0, 73839.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -16.0593751491
actual_array: [ 18.1 -119. ]
varray: [-13.595, -14.383, -14.088, -13.789]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -13.96375
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [72230.0, 72766.0, 73302.0, 73839.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 58.84885
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 10.1389
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 29 - Sample 29 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 29
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.204
varray: [2.215, 2.185, 2.22, 2.196]
inst: mean_all
operation: mean_all(residuals)
tarray: [74422.0, 74959.0, 75495.0, 76030.0]
source_col: residuals
SAMP_residuals_STD | 0.0163503312912
varray: [2.215, 2.185, 2.22, 2.196]
inst: std_all
operation: std_all(residuals)
tarray: [74422.0, 74959.0, 75495.0, 76030.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.13375
varray: [-1.265, -1.235, -0.915, -1.12]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [74422.0, 74959.0, 75495.0, 76030.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.15865975125
varray: [-1.265, -1.235, -0.915, -1.12]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [74422.0, 74959.0, 75495.0, 76030.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 61.1855
varray: [61.603, 61.133, 61.013, 60.993]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [74422.0, 74959.0, 75495.0, 76030.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.285116935075
varray: [61.603, 61.133, 61.013, 60.993]
inst: std_all
operation: std_all(baseline_shift)
tarray: [74422.0, 74959.0, 75495.0, 76030.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.258
varray: [28.44, 28.141, 28.334, 28.117]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [74422.0, 74959.0, 75495.0, 76030.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.155424150848
varray: [28.44, 28.141, 28.334, 28.117]
inst: std_all
operation: std_all(slope_shift)
tarray: [74422.0, 74959.0, 75495.0, 76030.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0258907963055
varray: [0.708, 0.756, 0.764, 0.758]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74422.0, 74959.0, 75495.0, 76030.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.831462316244
varray: [-4.015, -2.771, -2.422, -2.123]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74422.0, 74959.0, 75495.0, 76030.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.22575
varray: [63.78, 63.955, 65.237, 63.931]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [74422.0, 74959.0, 75495.0, 76030.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.678602178501
varray: [63.78, 63.955, 65.237, 63.931]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [74422.0, 74959.0, 75495.0, 76030.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.204
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.13375
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 61.1855
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.258
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.05925
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.05925
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.7625597389
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.41435
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [0.708, 0.756, 0.764, 0.758]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [74422.0, 74959.0, 75495.0, 76030.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-4.015, -2.771, -2.422, -2.123]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [74422.0, 74959.0, 75495.0, 76030.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.7465
varray: [0.708, 0.756, 0.764, 0.758]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [74422.0, 74959.0, 75495.0, 76030.0]
source_col: d(18_16)Mean
dD_mean | -2.83275
varray: [-4.015, -2.771, -2.422, -2.123]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [74422.0, 74959.0, 75495.0, 76030.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0258907963055
varray: [0.708, 0.756, 0.764, 0.758]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74422.0, 74959.0, 75495.0, 76030.0]
source_col: d(18_16)Mean
dD_stdev | 0.831462316244
varray: [-4.015, -2.771, -2.422, -2.123]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74422.0, 74959.0, 75495.0, 76030.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.7465
varray: [0.708, 0.756, 0.764, 0.758]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [74422.0, 74959.0, 75495.0, 76030.0]
source_col: d(18_16)Mean
measured_dD_mean | -2.83275
varray: [-4.015, -2.771, -2.422, -2.123]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [74422.0, 74959.0, 75495.0, 76030.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0258907963055
varray: [0.708, 0.756, 0.764, 0.758]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74422.0, 74959.0, 75495.0, 76030.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.831462316244
varray: [-4.015, -2.771, -2.422, -2.123]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74422.0, 74959.0, 75495.0, 76030.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.100512816054
actual_array: [ 1.93 -15.84]
varray: [0.708, 0.756, 0.764, 0.758]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 0.7465
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [74422.0, 74959.0, 75495.0, 76030.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -3.9233995219
actual_array: [ 18.1 -119. ]
varray: [-4.015, -2.771, -2.422, -2.123]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -2.83275
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [74422.0, 74959.0, 75495.0, 76030.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.07335
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.09185
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 30 - Sample 30 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 30
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1395
varray: [2.171, 2.106, 2.14, 2.141]
inst: mean_all
operation: mean_all(residuals)
tarray: [76616.0, 77152.0, 77687.0, 78222.0]
source_col: residuals
SAMP_residuals_STD | 0.0265643871879
varray: [2.171, 2.106, 2.14, 2.141]
inst: std_all
operation: std_all(residuals)
tarray: [76616.0, 77152.0, 77687.0, 78222.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.96975
varray: [-0.958, -0.966, -0.928, -1.027]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [76616.0, 77152.0, 77687.0, 78222.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.041524089394
varray: [-0.958, -0.966, -0.928, -1.027]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [76616.0, 77152.0, 77687.0, 78222.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 60.1055
varray: [60.407, 59.861, 60.039, 60.115]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [76616.0, 77152.0, 77687.0, 78222.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.227445964279
varray: [60.407, 59.861, 60.039, 60.115]
inst: std_all
operation: std_all(baseline_shift)
tarray: [76616.0, 77152.0, 77687.0, 78222.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.043
varray: [27.503, 26.607, 26.923, 27.139]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [76616.0, 77152.0, 77687.0, 78222.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.37652445693
varray: [27.503, 26.607, 26.923, 27.139]
inst: std_all
operation: std_all(slope_shift)
tarray: [76616.0, 77152.0, 77687.0, 78222.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0232862620444
varray: [-0.087, -0.102, -0.075, -0.047]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76616.0, 77152.0, 77687.0, 78222.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0990904468991
varray: [-3.269, -3.172, -3.204, -3.034]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76616.0, 77152.0, 77687.0, 78222.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.327
varray: [63.403, 63.057, 63.601, 63.247]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [76616.0, 77152.0, 77687.0, 78222.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.231049778186
varray: [63.403, 63.057, 63.601, 63.247]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [76616.0, 77152.0, 77687.0, 78222.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1395
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.96975
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 60.1055
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.043
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.00525
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.00525
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.244783774333
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.4844
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-0.087, -0.102, -0.075, -0.047]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [76616.0, 77152.0, 77687.0, 78222.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-3.269, -3.172, -3.204, -3.034]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [76616.0, 77152.0, 77687.0, 78222.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.07775
varray: [-0.087, -0.102, -0.075, -0.047]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [76616.0, 77152.0, 77687.0, 78222.0]
source_col: d(18_16)Mean
dD_mean | -3.16975
varray: [-3.269, -3.172, -3.204, -3.034]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [76616.0, 77152.0, 77687.0, 78222.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0232862620444
varray: [-0.087, -0.102, -0.075, -0.047]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76616.0, 77152.0, 77687.0, 78222.0]
source_col: d(18_16)Mean
dD_stdev | 0.0990904468991
varray: [-3.269, -3.172, -3.204, -3.034]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76616.0, 77152.0, 77687.0, 78222.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.07775
varray: [-0.087, -0.102, -0.075, -0.047]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [76616.0, 77152.0, 77687.0, 78222.0]
source_col: d(18_16)Mean
measured_dD_mean | -3.16975
varray: [-3.269, -3.172, -3.204, -3.034]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [76616.0, 77152.0, 77687.0, 78222.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0232862620444
varray: [-0.087, -0.102, -0.075, -0.047]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76616.0, 77152.0, 77687.0, 78222.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0990904468991
varray: [-3.269, -3.172, -3.204, -3.034]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76616.0, 77152.0, 77687.0, 78222.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.949521180399
actual_array: [ 1.93 -15.84]
varray: [-0.087, -0.102, -0.075, -0.047]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.07775
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [76616.0, 77152.0, 77687.0, 78222.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -4.29082596933
actual_array: [ 18.1 -119. ]
varray: [-3.269, -3.172, -3.204, -3.034]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -3.16975
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [76616.0, 77152.0, 77687.0, 78222.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.13785
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.17185
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 31 - Sample 31 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 31
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.08675
varray: [2.15, 2.07, 2.066, 2.061]
inst: mean_all
operation: mean_all(residuals)
tarray: [78808.0, 79344.0, 79880.0, 80416.0]
source_col: residuals
SAMP_residuals_STD | 0.0423270992785
varray: [2.15, 2.07, 2.066, 2.061]
inst: std_all
operation: std_all(residuals)
tarray: [78808.0, 79344.0, 79880.0, 80416.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.924
varray: [-1.039, -0.896, -0.846, -0.915]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [78808.0, 79344.0, 79880.0, 80416.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0820040649399
varray: [-1.039, -0.896, -0.846, -0.915]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [78808.0, 79344.0, 79880.0, 80416.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 60.81225
varray: [61.611, 60.694, 60.483, 60.461]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [78808.0, 79344.0, 79880.0, 80416.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.542760459749
varray: [61.611, 60.694, 60.483, 60.461]
inst: std_all
operation: std_all(baseline_shift)
tarray: [78808.0, 79344.0, 79880.0, 80416.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.66725
varray: [29.339, 26.879, 27.178, 27.273]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [78808.0, 79344.0, 79880.0, 80416.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.12707360747
varray: [29.339, 26.879, 27.178, 27.273]
inst: std_all
operation: std_all(slope_shift)
tarray: [78808.0, 79344.0, 79880.0, 80416.0]
source_col: slope_shift
SAMP_d18O_STD | 0.176277243757
varray: [-8.59, -8.856, -8.943, -8.981]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78808.0, 79344.0, 79880.0, 80416.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.03928745761
varray: [-58.34, -62.942, -64.408, -65.093]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78808.0, 79344.0, 79880.0, 80416.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.78
varray: [63.173, 63.651, 62.279, 62.017]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [78808.0, 79344.0, 79880.0, 80416.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.762946044401
varray: [63.173, 63.651, 62.279, 62.017]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [78808.0, 79344.0, 79880.0, 80416.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.08675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.924
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 60.81225
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.66725
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.058
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.058
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.70427788787
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.0314
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-8.59, -8.856, -8.943, -8.981]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [78808.0, 79344.0, 79880.0, 80416.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-58.34, -62.942, -64.408, -65.093]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [78808.0, 79344.0, 79880.0, 80416.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.8425
varray: [-8.59, -8.856, -8.943, -8.981]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [78808.0, 79344.0, 79880.0, 80416.0]
source_col: d(18_16)Mean
dD_mean | -62.69575
varray: [-58.34, -62.942, -64.408, -65.093]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [78808.0, 79344.0, 79880.0, 80416.0]
source_col: d(D_H)Mean
d18O_stdev | 0.176277243757
varray: [-8.59, -8.856, -8.943, -8.981]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78808.0, 79344.0, 79880.0, 80416.0]
source_col: d(18_16)Mean
dD_stdev | 3.03928745761
varray: [-58.34, -62.942, -64.408, -65.093]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78808.0, 79344.0, 79880.0, 80416.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.8425
varray: [-8.59, -8.856, -8.943, -8.981]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [78808.0, 79344.0, 79880.0, 80416.0]
source_col: d(18_16)Mean
measured_dD_mean | -62.69575
varray: [-58.34, -62.942, -64.408, -65.093]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [78808.0, 79344.0, 79880.0, 80416.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.176277243757
varray: [-8.59, -8.856, -8.943, -8.981]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78808.0, 79344.0, 79880.0, 80416.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.03928745761
varray: [-58.34, -62.942, -64.408, -65.093]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78808.0, 79344.0, 79880.0, 80416.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.97754187969
actual_array: [ 1.93 -15.84]
varray: [-8.59, -8.856, -8.943, -8.981]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -8.8425
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [78808.0, 79344.0, 79880.0, 80416.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -69.191201963
actual_array: [ 18.1 -119. ]
varray: [-58.34, -62.942, -64.408, -65.093]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -62.69575
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [78808.0, 79344.0, 79880.0, 80416.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.1906
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.4651
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 32 - Sample 32 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 32
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.2195
varray: [2.224, 2.231, 2.219, 2.204]
inst: mean_all
operation: mean_all(residuals)
tarray: [81002.0, 81538.0, 82074.0, 82609.0]
source_col: residuals
SAMP_residuals_STD | 0.0114455231423
varray: [2.224, 2.231, 2.219, 2.204]
inst: std_all
operation: std_all(residuals)
tarray: [81002.0, 81538.0, 82074.0, 82609.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.4345
varray: [-1.4, -1.352, -1.448, -1.538]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [81002.0, 81538.0, 82074.0, 82609.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.079353638858
varray: [-1.4, -1.352, -1.448, -1.538]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [81002.0, 81538.0, 82074.0, 82609.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 210.68825
varray: [209.896, 210.92, 210.965, 210.972]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [81002.0, 81538.0, 82074.0, 82609.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.528669004829
varray: [209.896, 210.92, 210.965, 210.972]
inst: std_all
operation: std_all(baseline_shift)
tarray: [81002.0, 81538.0, 82074.0, 82609.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 22.546
varray: [22.739, 22.153, 22.324, 22.968]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [81002.0, 81538.0, 82074.0, 82609.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.373749470457
varray: [22.739, 22.153, 22.324, 22.968]
inst: std_all
operation: std_all(slope_shift)
tarray: [81002.0, 81538.0, 82074.0, 82609.0]
source_col: slope_shift
SAMP_d18O_STD | 0.203798920507
varray: [-1.445, -1.05, -1.046, -1.019]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [81002.0, 81538.0, 82074.0, 82609.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.48367865608
varray: [-21.482, -17.276, -16.588, -16.008]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [81002.0, 81538.0, 82074.0, 82609.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.381
varray: [60.816, 62.782, 61.848, 60.078]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [81002.0, 81538.0, 82074.0, 82609.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.1829234971
varray: [60.816, 62.782, 61.848, 60.078]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [81002.0, 81538.0, 82074.0, 82609.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.2195
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.4345
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 210.68825
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 22.546
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.07475
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.07475
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 3.48525469169
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.4304
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-1.445, -1.05, -1.046, -1.019]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [81002.0, 81538.0, 82074.0, 82609.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-21.482, -17.276, -16.588, -16.008]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [81002.0, 81538.0, 82074.0, 82609.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.14
varray: [-1.445, -1.05, -1.046, -1.019]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [81002.0, 81538.0, 82074.0, 82609.0]
source_col: d(18_16)Mean
dD_mean | -17.8385
varray: [-21.482, -17.276, -16.588, -16.008]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [81002.0, 81538.0, 82074.0, 82609.0]
source_col: d(D_H)Mean
d18O_stdev | 0.203798920507
varray: [-1.445, -1.05, -1.046, -1.019]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [81002.0, 81538.0, 82074.0, 82609.0]
source_col: d(18_16)Mean
dD_stdev | 2.48367865608
varray: [-21.482, -17.276, -16.588, -16.008]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [81002.0, 81538.0, 82074.0, 82609.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.14
varray: [-1.445, -1.05, -1.046, -1.019]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [81002.0, 81538.0, 82074.0, 82609.0]
source_col: d(18_16)Mean
measured_dD_mean | -17.8385
varray: [-21.482, -17.276, -16.588, -16.008]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [81002.0, 81538.0, 82074.0, 82609.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.203798920507
varray: [-1.445, -1.05, -1.046, -1.019]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [81002.0, 81538.0, 82074.0, 82609.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.48367865608
varray: [-21.482, -17.276, -16.588, -16.008]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [81002.0, 81538.0, 82074.0, 82609.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.04367845674
actual_array: [ 1.93 -15.84]
varray: [-1.445, -1.05, -1.046, -1.019]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.14
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [81002.0, 81538.0, 82074.0, 82609.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -20.2839615799
actual_array: [ 18.1 -119. ]
varray: [-21.482, -17.276, -16.588, -16.008]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -17.8385
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [81002.0, 81538.0, 82074.0, 82609.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.05785
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 149.4109
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 33 - Sample 33 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 33
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.6365
varray: [1.556, 1.521, 1.744, 1.725]
inst: mean_all
operation: mean_all(residuals)
tarray: [83196.0, 83732.0, 84268.0, 84804.0]
source_col: residuals
SAMP_residuals_STD | 0.114322642843
varray: [1.556, 1.521, 1.744, 1.725]
inst: std_all
operation: std_all(residuals)
tarray: [83196.0, 83732.0, 84268.0, 84804.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0025
varray: [-0.95, -0.914, -1.022, -1.124]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [83196.0, 83732.0, 84268.0, 84804.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0926120942426
varray: [-0.95, -0.914, -1.022, -1.124]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [83196.0, 83732.0, 84268.0, 84804.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 114.21375
varray: [110.681, 108.674, 119.127, 118.373]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [83196.0, 83732.0, 84268.0, 84804.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 5.31063520212
varray: [110.681, 108.674, 119.127, 118.373]
inst: std_all
operation: std_all(baseline_shift)
tarray: [83196.0, 83732.0, 84268.0, 84804.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 16.5635
varray: [15.303, 14.686, 18.139, 18.126]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [83196.0, 83732.0, 84268.0, 84804.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.82915946088
varray: [15.303, 14.686, 18.139, 18.126]
inst: std_all
operation: std_all(slope_shift)
tarray: [83196.0, 83732.0, 84268.0, 84804.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0259020591202
varray: [-0.921, -0.901, -0.89, -0.949]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83196.0, 83732.0, 84268.0, 84804.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.334998507459
varray: [-14.673, -14.061, -14.037, -13.937]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83196.0, 83732.0, 84268.0, 84804.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 49.93175
varray: [47.815, 47.806, 52.485, 51.621]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [83196.0, 83732.0, 84268.0, 84804.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.47467807671
varray: [47.815, 47.806, 52.485, 51.621]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [83196.0, 83732.0, 84268.0, 84804.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.6365
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0025
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 114.21375
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 16.5635
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.50825
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.50825
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -23.6974006294
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -13.87965
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-0.921, -0.901, -0.89, -0.949]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [83196.0, 83732.0, 84268.0, 84804.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-14.673, -14.061, -14.037, -13.937]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [83196.0, 83732.0, 84268.0, 84804.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.91525
varray: [-0.921, -0.901, -0.89, -0.949]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [83196.0, 83732.0, 84268.0, 84804.0]
source_col: d(18_16)Mean
dD_mean | -14.177
varray: [-14.673, -14.061, -14.037, -13.937]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [83196.0, 83732.0, 84268.0, 84804.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0259020591202
varray: [-0.921, -0.901, -0.89, -0.949]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83196.0, 83732.0, 84268.0, 84804.0]
source_col: d(18_16)Mean
dD_stdev | 0.334998507459
varray: [-14.673, -14.061, -14.037, -13.937]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83196.0, 83732.0, 84268.0, 84804.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.91525
varray: [-0.921, -0.901, -0.89, -0.949]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [83196.0, 83732.0, 84268.0, 84804.0]
source_col: d(18_16)Mean
measured_dD_mean | -14.177
varray: [-14.673, -14.061, -14.037, -13.937]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [83196.0, 83732.0, 84268.0, 84804.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0259020591202
varray: [-0.921, -0.901, -0.89, -0.949]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83196.0, 83732.0, 84268.0, 84804.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.334998507459
varray: [-14.673, -14.061, -14.037, -13.937]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83196.0, 83732.0, 84268.0, 84804.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.81217754089
actual_array: [ 1.93 -15.84]
varray: [-0.921, -0.901, -0.89, -0.949]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.91525
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [83196.0, 83732.0, 84268.0, 84804.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -16.29187868
actual_array: [ 18.1 -119. ]
varray: [-14.673, -14.061, -14.037, -13.937]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -14.177
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [83196.0, 83732.0, 84268.0, 84804.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.64085
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 52.9364
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 34 - Sample 34 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 34
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.20475
varray: [2.212, 2.202, 2.224, 2.181]
inst: mean_all
operation: mean_all(residuals)
tarray: [85392.0, 85927.0, 86464.0, 87000.0]
source_col: residuals
SAMP_residuals_STD | 0.0182094297915
varray: [2.212, 2.202, 2.224, 2.181]
inst: std_all
operation: std_all(residuals)
tarray: [85392.0, 85927.0, 86464.0, 87000.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.5505
varray: [-1.682, -1.501, -1.421, -1.598]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [85392.0, 85927.0, 86464.0, 87000.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.113679373679
varray: [-1.682, -1.501, -1.421, -1.598]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [85392.0, 85927.0, 86464.0, 87000.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 171.39525
varray: [171.917, 170.964, 171.838, 170.862]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [85392.0, 85927.0, 86464.0, 87000.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.559339714902
varray: [171.917, 170.964, 171.838, 170.862]
inst: std_all
operation: std_all(baseline_shift)
tarray: [85392.0, 85927.0, 86464.0, 87000.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 22.393
varray: [22.768, 22.41, 22.234, 22.16]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [85392.0, 85927.0, 86464.0, 87000.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.271098997908
varray: [22.768, 22.41, 22.234, 22.16]
inst: std_all
operation: std_all(slope_shift)
tarray: [85392.0, 85927.0, 86464.0, 87000.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0196701974232
varray: [-0.733, -0.724, -0.752, -0.768]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85392.0, 85927.0, 86464.0, 87000.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.104399872286
varray: [-16.118, -16.273, -16.37, -16.231]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85392.0, 85927.0, 86464.0, 87000.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 60.97575
varray: [60.852, 60.751, 62.098, 60.202]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [85392.0, 85927.0, 86464.0, 87000.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.800824731137
varray: [60.852, 60.751, 62.098, 60.202]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [85392.0, 85927.0, 86464.0, 87000.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.20475
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.5505
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 171.39525
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 22.393
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.06
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.06
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.79752884952
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.83565
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-0.733, -0.724, -0.752, -0.768]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [85392.0, 85927.0, 86464.0, 87000.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-16.118, -16.273, -16.37, -16.231]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [85392.0, 85927.0, 86464.0, 87000.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.74425
varray: [-0.733, -0.724, -0.752, -0.768]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [85392.0, 85927.0, 86464.0, 87000.0]
source_col: d(18_16)Mean
dD_mean | -16.248
varray: [-16.118, -16.273, -16.37, -16.231]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [85392.0, 85927.0, 86464.0, 87000.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0196701974232
varray: [-0.733, -0.724, -0.752, -0.768]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85392.0, 85927.0, 86464.0, 87000.0]
source_col: d(18_16)Mean
dD_stdev | 0.104399872286
varray: [-16.118, -16.273, -16.37, -16.231]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85392.0, 85927.0, 86464.0, 87000.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.74425
varray: [-0.733, -0.724, -0.752, -0.768]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [85392.0, 85927.0, 86464.0, 87000.0]
source_col: d(18_16)Mean
measured_dD_mean | -16.248
varray: [-16.118, -16.273, -16.37, -16.231]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [85392.0, 85927.0, 86464.0, 87000.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0196701974232
varray: [-0.733, -0.724, -0.752, -0.768]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85392.0, 85927.0, 86464.0, 87000.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.104399872286
varray: [-16.118, -16.273, -16.37, -16.231]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85392.0, 85927.0, 86464.0, 87000.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.63604113774
actual_array: [ 1.93 -15.84]
varray: [-0.733, -0.724, -0.752, -0.768]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.74425
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [85392.0, 85927.0, 86464.0, 87000.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -18.5498613881
actual_array: [ 18.1 -119. ]
varray: [-16.118, -16.273, -16.37, -16.231]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -16.248
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [85392.0, 85927.0, 86464.0, 87000.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.0726
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 110.1179
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 35 - Sample 35 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 35
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1715
varray: [2.176, 2.149, 2.193, 2.168]
inst: mean_all
operation: mean_all(residuals)
tarray: [87588.0, 88124.0, 88660.0, 89197.0]
source_col: residuals
SAMP_residuals_STD | 0.0182665450118
varray: [2.176, 2.149, 2.193, 2.168]
inst: std_all
operation: std_all(residuals)
tarray: [87588.0, 88124.0, 88660.0, 89197.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.016
varray: [-1.071, -1.139, -0.919, -0.935]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [87588.0, 88124.0, 88660.0, 89197.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.106652082336
varray: [-1.071, -1.139, -0.919, -0.935]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [87588.0, 88124.0, 88660.0, 89197.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 61.48025
varray: [62.323, 61.371, 61.404, 60.823]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [87588.0, 88124.0, 88660.0, 89197.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.621813141815
varray: [62.323, 61.371, 61.404, 60.823]
inst: std_all
operation: std_all(baseline_shift)
tarray: [87588.0, 88124.0, 88660.0, 89197.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.66775
varray: [27.725, 27.593, 27.592, 27.761]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [87588.0, 88124.0, 88660.0, 89197.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0881263297772
varray: [27.725, 27.593, 27.592, 27.761]
inst: std_all
operation: std_all(slope_shift)
tarray: [87588.0, 88124.0, 88660.0, 89197.0]
source_col: slope_shift
SAMP_d18O_STD | 0.01808314132
varray: [-0.626, -0.591, -0.62, -0.593]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87588.0, 88124.0, 88660.0, 89197.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.758288643372
varray: [-3.913, -2.511, -2.33, -2.372]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87588.0, 88124.0, 88660.0, 89197.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.09675
varray: [64.259, 62.93, 65.627, 63.571]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [87588.0, 88124.0, 88660.0, 89197.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.15552423168
varray: [64.259, 62.93, 65.627, 63.571]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [87588.0, 88124.0, 88660.0, 89197.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1715
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.016
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 61.48025
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.66775
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.02675
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.02675
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.24723161208
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.28535
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-0.626, -0.591, -0.62, -0.593]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [87588.0, 88124.0, 88660.0, 89197.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-3.913, -2.511, -2.33, -2.372]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [87588.0, 88124.0, 88660.0, 89197.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.6075
varray: [-0.626, -0.591, -0.62, -0.593]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [87588.0, 88124.0, 88660.0, 89197.0]
source_col: d(18_16)Mean
dD_mean | -2.7815
varray: [-3.913, -2.511, -2.33, -2.372]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [87588.0, 88124.0, 88660.0, 89197.0]
source_col: d(D_H)Mean
d18O_stdev | 0.01808314132
varray: [-0.626, -0.591, -0.62, -0.593]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87588.0, 88124.0, 88660.0, 89197.0]
source_col: d(18_16)Mean
dD_stdev | 0.758288643372
varray: [-3.913, -2.511, -2.33, -2.372]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87588.0, 88124.0, 88660.0, 89197.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.6075
varray: [-0.626, -0.591, -0.62, -0.593]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [87588.0, 88124.0, 88660.0, 89197.0]
source_col: d(18_16)Mean
measured_dD_mean | -2.7815
varray: [-3.913, -2.511, -2.33, -2.372]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [87588.0, 88124.0, 88660.0, 89197.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.01808314132
varray: [-0.626, -0.591, -0.62, -0.593]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87588.0, 88124.0, 88660.0, 89197.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.758288643372
varray: [-3.913, -2.511, -2.33, -2.372]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87588.0, 88124.0, 88660.0, 89197.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.49518351708
actual_array: [ 1.93 -15.84]
varray: [-0.626, -0.591, -0.62, -0.593]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.6075
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [87588.0, 88124.0, 88660.0, 89197.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -3.86752235445
actual_array: [ 18.1 -119. ]
varray: [-3.913, -2.511, -2.33, -2.372]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -2.7815
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [87588.0, 88124.0, 88660.0, 89197.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.10585
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.2029
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 36 - Sample 36 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 36
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.138
varray: [2.133, 2.151, 2.119, 2.149]
inst: mean_all
operation: mean_all(residuals)
tarray: [89780.0, 90316.0, 90853.0, 91389.0]
source_col: residuals
SAMP_residuals_STD | 0.0150111069989
varray: [2.133, 2.151, 2.119, 2.149]
inst: std_all
operation: std_all(residuals)
tarray: [89780.0, 90316.0, 90853.0, 91389.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.98775
varray: [-0.908, -1.153, -0.957, -0.933]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [89780.0, 90316.0, 90853.0, 91389.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.111968373511
varray: [-0.908, -1.153, -0.957, -0.933]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [89780.0, 90316.0, 90853.0, 91389.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 60.70325
varray: [60.359, 61.005, 60.514, 60.935]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [89780.0, 90316.0, 90853.0, 91389.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.315745018435
varray: [60.359, 61.005, 60.514, 60.935]
inst: std_all
operation: std_all(baseline_shift)
tarray: [89780.0, 90316.0, 90853.0, 91389.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.11975
varray: [27.142, 27.593, 26.599, 27.145]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [89780.0, 90316.0, 90853.0, 91389.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.406726259295
varray: [27.142, 27.593, 26.599, 27.145]
inst: std_all
operation: std_all(slope_shift)
tarray: [89780.0, 90316.0, 90853.0, 91389.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0302806318736
varray: [-0.226, -0.202, -0.167, -0.162]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89780.0, 90316.0, 90853.0, 91389.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.092105012531
varray: [-4.314, -4.485, -4.372, -4.273]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89780.0, 90316.0, 90853.0, 91389.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.4215
varray: [63.162, 63.729, 63.22, 63.575]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [89780.0, 90316.0, 90853.0, 91389.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.274506223366
varray: [63.162, 63.729, 63.22, 63.575]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [89780.0, 90316.0, 90853.0, 91389.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.138
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.98775
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 60.70325
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.11975
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.00675
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.00675
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.314721995571
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.3899
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-0.226, -0.202, -0.167, -0.162]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [89780.0, 90316.0, 90853.0, 91389.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-4.314, -4.485, -4.372, -4.273]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [89780.0, 90316.0, 90853.0, 91389.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.18925
varray: [-0.226, -0.202, -0.167, -0.162]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [89780.0, 90316.0, 90853.0, 91389.0]
source_col: d(18_16)Mean
dD_mean | -4.361
varray: [-4.314, -4.485, -4.372, -4.273]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [89780.0, 90316.0, 90853.0, 91389.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0302806318736
varray: [-0.226, -0.202, -0.167, -0.162]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89780.0, 90316.0, 90853.0, 91389.0]
source_col: d(18_16)Mean
dD_stdev | 0.092105012531
varray: [-4.314, -4.485, -4.372, -4.273]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89780.0, 90316.0, 90853.0, 91389.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.18925
varray: [-0.226, -0.202, -0.167, -0.162]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [89780.0, 90316.0, 90853.0, 91389.0]
source_col: d(18_16)Mean
measured_dD_mean | -4.361
varray: [-4.314, -4.485, -4.372, -4.273]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [89780.0, 90316.0, 90853.0, 91389.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0302806318736
varray: [-0.226, -0.202, -0.167, -0.162]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89780.0, 90316.0, 90853.0, 91389.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.092105012531
varray: [-4.314, -4.485, -4.372, -4.273]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89780.0, 90316.0, 90853.0, 91389.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.06437035556
actual_array: [ 1.93 -15.84]
varray: [-0.226, -0.202, -0.167, -0.162]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.18925
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [89780.0, 90316.0, 90853.0, 91389.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -5.58962939812
actual_array: [ 18.1 -119. ]
varray: [-4.314, -4.485, -4.372, -4.273]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -4.361
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [89780.0, 90316.0, 90853.0, 91389.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.13935
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.5741
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 37 - Sample 37 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 37
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.15425
varray: [2.128, 2.151, 2.137, 2.201]
inst: mean_all
operation: mean_all(residuals)
tarray: [91974.0, 92509.0, 93045.0, 93581.0]
source_col: residuals
SAMP_residuals_STD | 0.0325717157464
varray: [2.128, 2.151, 2.137, 2.201]
inst: std_all
operation: std_all(residuals)
tarray: [91974.0, 92509.0, 93045.0, 93581.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.04475
varray: [-0.915, -1.028, -1.084, -1.152]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [91974.0, 92509.0, 93045.0, 93581.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.100264234235
varray: [-0.915, -1.028, -1.084, -1.152]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [91974.0, 92509.0, 93045.0, 93581.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 60.739
varray: [60.268, 60.775, 60.639, 61.274]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [91974.0, 92509.0, 93045.0, 93581.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.416077717099
varray: [60.268, 60.775, 60.639, 61.274]
inst: std_all
operation: std_all(baseline_shift)
tarray: [91974.0, 92509.0, 93045.0, 93581.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.5555
varray: [27.061, 27.623, 27.233, 28.305]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [91974.0, 92509.0, 93045.0, 93581.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.552220668453
varray: [27.061, 27.623, 27.233, 28.305]
inst: std_all
operation: std_all(slope_shift)
tarray: [91974.0, 92509.0, 93045.0, 93581.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0181199705666
varray: [-0.543, -0.515, -0.522, -0.554]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [91974.0, 92509.0, 93045.0, 93581.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.128164932801
varray: [-2.639, -2.566, -2.788, -2.842]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [91974.0, 92509.0, 93045.0, 93581.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.73325
varray: [63.318, 63.466, 63.476, 64.673]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [91974.0, 92509.0, 93045.0, 93581.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.630651184623
varray: [63.318, 63.466, 63.476, 64.673]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [91974.0, 92509.0, 93045.0, 93581.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.15425
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.04475
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 60.739
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.5555
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0095
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0095
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.44294206784
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.07815
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-0.543, -0.515, -0.522, -0.554]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [91974.0, 92509.0, 93045.0, 93581.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-2.639, -2.566, -2.788, -2.842]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [91974.0, 92509.0, 93045.0, 93581.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.5335
varray: [-0.543, -0.515, -0.522, -0.554]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [91974.0, 92509.0, 93045.0, 93581.0]
source_col: d(18_16)Mean
dD_mean | -2.70875
varray: [-2.639, -2.566, -2.788, -2.842]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [91974.0, 92509.0, 93045.0, 93581.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0181199705666
varray: [-0.543, -0.515, -0.522, -0.554]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [91974.0, 92509.0, 93045.0, 93581.0]
source_col: d(18_16)Mean
dD_stdev | 0.128164932801
varray: [-2.639, -2.566, -2.788, -2.842]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [91974.0, 92509.0, 93045.0, 93581.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.5335
varray: [-0.543, -0.515, -0.522, -0.554]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [91974.0, 92509.0, 93045.0, 93581.0]
source_col: d(18_16)Mean
measured_dD_mean | -2.70875
varray: [-2.639, -2.566, -2.788, -2.842]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [91974.0, 92509.0, 93045.0, 93581.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0181199705666
varray: [-0.543, -0.515, -0.522, -0.554]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [91974.0, 92509.0, 93045.0, 93581.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.128164932801
varray: [-2.639, -2.566, -2.788, -2.842]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [91974.0, 92509.0, 93045.0, 93581.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.41896074613
actual_array: [ 1.93 -15.84]
varray: [-0.543, -0.515, -0.522, -0.554]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.5335
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [91974.0, 92509.0, 93045.0, 93581.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -3.78820403382
actual_array: [ 18.1 -119. ]
varray: [-2.639, -2.566, -2.788, -2.842]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -2.70875
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [91974.0, 92509.0, 93045.0, 93581.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.1231
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.53835
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 38 - Sample 38 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 38
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1305
varray: [2.112, 2.112, 2.166, 2.132]
inst: mean_all
operation: mean_all(residuals)
tarray: [94165.0, 94701.0, 95237.0, 95773.0]
source_col: residuals
SAMP_residuals_STD | 0.0254754784057
varray: [2.112, 2.112, 2.166, 2.132]
inst: std_all
operation: std_all(residuals)
tarray: [94165.0, 94701.0, 95237.0, 95773.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.10975
varray: [-1.102, -1.152, -1.093, -1.092]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [94165.0, 94701.0, 95237.0, 95773.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0285233822211
varray: [-1.102, -1.152, -1.093, -1.092]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [94165.0, 94701.0, 95237.0, 95773.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 60.3225
varray: [60.019, 60.049, 60.885, 60.337]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [94165.0, 94701.0, 95237.0, 95773.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.401468554186
varray: [60.019, 60.049, 60.885, 60.337]
inst: std_all
operation: std_all(baseline_shift)
tarray: [94165.0, 94701.0, 95237.0, 95773.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.0765
varray: [26.959, 26.685, 27.51, 27.152]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [94165.0, 94701.0, 95237.0, 95773.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.346747266271
varray: [26.959, 26.685, 27.51, 27.152]
inst: std_all
operation: std_all(slope_shift)
tarray: [94165.0, 94701.0, 95237.0, 95773.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0242263630507
varray: [-0.636, -0.691, -0.646, -0.65]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94165.0, 94701.0, 95237.0, 95773.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.103584104315
varray: [-3.473, -3.646, -3.667, -3.708]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94165.0, 94701.0, 95237.0, 95773.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.901
varray: [62.145, 62.307, 64.131, 63.021]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [94165.0, 94701.0, 95237.0, 95773.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.904004424768
varray: [62.145, 62.307, 64.131, 63.021]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [94165.0, 94701.0, 95237.0, 95773.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1305
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.10975
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 60.3225
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.0765
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.01425
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.01425
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.66441310176
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.9104
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-0.636, -0.691, -0.646, -0.65]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [94165.0, 94701.0, 95237.0, 95773.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-3.473, -3.646, -3.667, -3.708]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [94165.0, 94701.0, 95237.0, 95773.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.65575
varray: [-0.636, -0.691, -0.646, -0.65]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [94165.0, 94701.0, 95237.0, 95773.0]
source_col: d(18_16)Mean
dD_mean | -3.6235
varray: [-3.473, -3.646, -3.667, -3.708]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [94165.0, 94701.0, 95237.0, 95773.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0242263630507
varray: [-0.636, -0.691, -0.646, -0.65]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94165.0, 94701.0, 95237.0, 95773.0]
source_col: d(18_16)Mean
dD_stdev | 0.103584104315
varray: [-3.473, -3.646, -3.667, -3.708]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94165.0, 94701.0, 95237.0, 95773.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.65575
varray: [-0.636, -0.691, -0.646, -0.65]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [94165.0, 94701.0, 95237.0, 95773.0]
source_col: d(18_16)Mean
measured_dD_mean | -3.6235
varray: [-3.473, -3.646, -3.667, -3.708]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [94165.0, 94701.0, 95237.0, 95773.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0242263630507
varray: [-0.636, -0.691, -0.646, -0.65]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94165.0, 94701.0, 95237.0, 95773.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.103584104315
varray: [-3.473, -3.646, -3.667, -3.708]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94165.0, 94701.0, 95237.0, 95773.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.54488282382
actual_array: [ 1.93 -15.84]
varray: [-0.636, -0.691, -0.646, -0.65]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.65575
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [94165.0, 94701.0, 95237.0, 95773.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -4.78554332993
actual_array: [ 18.1 -119. ]
varray: [-3.473, -3.646, -3.667, -3.708]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -3.6235
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [94165.0, 94701.0, 95237.0, 95773.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.14685
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.95485
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 39 - Sample 39 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 39
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0725
varray: [2.062, 2.066, 2.08, 2.082]
inst: mean_all
operation: mean_all(residuals)
tarray: [96358.0, 96894.0, 97430.0, 97966.0]
source_col: residuals
SAMP_residuals_STD | 0.00998331942125
varray: [2.062, 2.066, 2.08, 2.082]
inst: std_all
operation: std_all(residuals)
tarray: [96358.0, 96894.0, 97430.0, 97966.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.042
varray: [-1.057, -1.065, -1.052, -0.994]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [96358.0, 96894.0, 97430.0, 97966.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0324448249597
varray: [-1.057, -1.065, -1.052, -0.994]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [96358.0, 96894.0, 97430.0, 97966.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 59.77425
varray: [59.685, 59.826, 59.736, 59.85]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [96358.0, 96894.0, 97430.0, 97966.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0771248987033
varray: [59.685, 59.826, 59.736, 59.85]
inst: std_all
operation: std_all(baseline_shift)
tarray: [96358.0, 96894.0, 97430.0, 97966.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.12775
varray: [25.965, 26.234, 26.313, 25.999]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [96358.0, 96894.0, 97430.0, 97966.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.171921251353
varray: [25.965, 26.234, 26.313, 25.999]
inst: std_all
operation: std_all(slope_shift)
tarray: [96358.0, 96894.0, 97430.0, 97966.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0472678537698
varray: [-0.994, -1.039, -1.078, -1.102]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96358.0, 96894.0, 97430.0, 97966.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.254198577232
varray: [-7.859, -8.289, -8.203, -8.464]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96358.0, 96894.0, 97430.0, 97966.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.743
varray: [61.484, 61.842, 61.61, 62.036]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [96358.0, 96894.0, 97430.0, 97966.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.245234581574
varray: [61.484, 61.842, 61.61, 62.036]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [96358.0, 96894.0, 97430.0, 97966.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0725
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.042
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 59.77425
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.12775
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.07225
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.07225
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.36869098963
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.0684
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-0.994, -1.039, -1.078, -1.102]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [96358.0, 96894.0, 97430.0, 97966.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-7.859, -8.289, -8.203, -8.464]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [96358.0, 96894.0, 97430.0, 97966.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.05325
varray: [-0.994, -1.039, -1.078, -1.102]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [96358.0, 96894.0, 97430.0, 97966.0]
source_col: d(18_16)Mean
dD_mean | -8.20375
varray: [-7.859, -8.289, -8.203, -8.464]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [96358.0, 96894.0, 97430.0, 97966.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0472678537698
varray: [-0.994, -1.039, -1.078, -1.102]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96358.0, 96894.0, 97430.0, 97966.0]
source_col: d(18_16)Mean
dD_stdev | 0.254198577232
varray: [-7.859, -8.289, -8.203, -8.464]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96358.0, 96894.0, 97430.0, 97966.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.05325
varray: [-0.994, -1.039, -1.078, -1.102]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [96358.0, 96894.0, 97430.0, 97966.0]
source_col: d(18_16)Mean
measured_dD_mean | -8.20375
varray: [-7.859, -8.289, -8.203, -8.464]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [96358.0, 96894.0, 97430.0, 97966.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0472678537698
varray: [-0.994, -1.039, -1.078, -1.102]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96358.0, 96894.0, 97430.0, 97966.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.254198577232
varray: [-7.859, -8.289, -8.203, -8.464]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96358.0, 96894.0, 97430.0, 97966.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.95432270836
actual_array: [ 1.93 -15.84]
varray: [-0.994, -1.039, -1.078, -1.102]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.05325
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [96358.0, 96894.0, 97430.0, 97966.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -9.77932667074
actual_array: [ 18.1 -119. ]
varray: [-7.859, -8.289, -8.203, -8.464]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -8.20375
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [96358.0, 96894.0, 97430.0, 97966.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.20485
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.5031
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 40 - Sample 40 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 40
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.06475
varray: [2.088, 2.047, 2.06, 2.064]
inst: mean_all
operation: mean_all(residuals)
tarray: [98550.0, 99085.0, 99622.0, 100158.0]
source_col: residuals
SAMP_residuals_STD | 0.0171148084029
varray: [2.088, 2.047, 2.06, 2.064]
inst: std_all
operation: std_all(residuals)
tarray: [98550.0, 99085.0, 99622.0, 100158.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.95425
varray: [-1.009, -0.962, -0.888, -0.958]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [98550.0, 99085.0, 99622.0, 100158.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0498689950437
varray: [-1.009, -0.962, -0.888, -0.958]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [98550.0, 99085.0, 99622.0, 100158.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 60.34375
varray: [60.637, 60.223, 60.123, 60.392]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [98550.0, 99085.0, 99622.0, 100158.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.224821966602
varray: [60.637, 60.223, 60.123, 60.392]
inst: std_all
operation: std_all(baseline_shift)
tarray: [98550.0, 99085.0, 99622.0, 100158.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.1475
varray: [27.572, 27.105, 26.924, 26.989]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [98550.0, 99085.0, 99622.0, 100158.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.292734806039
varray: [27.572, 27.105, 26.924, 26.989]
inst: std_all
operation: std_all(slope_shift)
tarray: [98550.0, 99085.0, 99622.0, 100158.0]
source_col: slope_shift
SAMP_d18O_STD | 0.191484333563
varray: [-8.604, -8.871, -9.003, -9.017]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [98550.0, 99085.0, 99622.0, 100158.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.01743869035
varray: [-58.517, -63.039, -64.643, -65.15]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [98550.0, 99085.0, 99622.0, 100158.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.04875
varray: [61.764, 61.412, 62.42, 62.599]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [98550.0, 99085.0, 99622.0, 100158.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.555918084493
varray: [61.764, 61.412, 62.42, 62.599]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [98550.0, 99085.0, 99622.0, 100158.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.06475
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.95425
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 60.34375
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.1475
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.08
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.08
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.73003846602
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.76265
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-8.604, -8.871, -9.003, -9.017]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [98550.0, 99085.0, 99622.0, 100158.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-58.517, -63.039, -64.643, -65.15]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [98550.0, 99085.0, 99622.0, 100158.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.87375
varray: [-8.604, -8.871, -9.003, -9.017]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [98550.0, 99085.0, 99622.0, 100158.0]
source_col: d(18_16)Mean
dD_mean | -62.83725
varray: [-58.517, -63.039, -64.643, -65.15]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [98550.0, 99085.0, 99622.0, 100158.0]
source_col: d(D_H)Mean
d18O_stdev | 0.191484333563
varray: [-8.604, -8.871, -9.003, -9.017]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [98550.0, 99085.0, 99622.0, 100158.0]
source_col: d(18_16)Mean
dD_stdev | 3.01743869035
varray: [-58.517, -63.039, -64.643, -65.15]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [98550.0, 99085.0, 99622.0, 100158.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.87375
varray: [-8.604, -8.871, -9.003, -9.017]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [98550.0, 99085.0, 99622.0, 100158.0]
source_col: d(18_16)Mean
measured_dD_mean | -62.83725
varray: [-58.517, -63.039, -64.643, -65.15]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [98550.0, 99085.0, 99622.0, 100158.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.191484333563
varray: [-8.604, -8.871, -9.003, -9.017]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [98550.0, 99085.0, 99622.0, 100158.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.01743869035
varray: [-58.517, -63.039, -64.643, -65.15]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [98550.0, 99085.0, 99622.0, 100158.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.0097305499
actual_array: [ 1.93 -15.84]
varray: [-8.604, -8.871, -9.003, -9.017]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -8.87375
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [98550.0, 99085.0, 99622.0, 100158.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -69.3454774595
actual_array: [ 18.1 -119. ]
varray: [-58.517, -63.039, -64.643, -65.15]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -62.83725
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [98550.0, 99085.0, 99622.0, 100158.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.2126
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.9336
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 41 - Sample 41 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 41
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.108
varray: [2.105, 2.123, 2.11, 2.094]
inst: mean_all
operation: mean_all(residuals)
tarray: [100742.0, 101278.0, 101815.0, 102351.0]
source_col: residuals
SAMP_residuals_STD | 0.0120277457018
varray: [2.105, 2.123, 2.11, 2.094]
inst: std_all
operation: std_all(residuals)
tarray: [100742.0, 101278.0, 101815.0, 102351.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.06575
varray: [-1.103, -1.104, -1.077, -0.979]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [100742.0, 101278.0, 101815.0, 102351.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0591685445711
varray: [-1.103, -1.104, -1.077, -0.979]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [100742.0, 101278.0, 101815.0, 102351.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 60.18325
varray: [60.281, 60.441, 60.24, 59.771]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [100742.0, 101278.0, 101815.0, 102351.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.288190186046
varray: [60.281, 60.441, 60.24, 59.771]
inst: std_all
operation: std_all(baseline_shift)
tarray: [100742.0, 101278.0, 101815.0, 102351.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.02425
varray: [27.082, 27.295, 26.889, 26.831]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [100742.0, 101278.0, 101815.0, 102351.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.209983134243
varray: [27.082, 27.295, 26.889, 26.831]
inst: std_all
operation: std_all(slope_shift)
tarray: [100742.0, 101278.0, 101815.0, 102351.0]
source_col: slope_shift
SAMP_d18O_STD | 0.185464057614
varray: [-1.579, -1.285, -1.244, -1.149]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100742.0, 101278.0, 101815.0, 102351.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.97786292219
varray: [-13.443, -8.82, -7.393, -6.919]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100742.0, 101278.0, 101815.0, 102351.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.31675
varray: [62.38, 62.868, 62.2, 61.819]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [100742.0, 101278.0, 101815.0, 102351.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.43560790856
varray: [62.38, 62.868, 62.2, 61.819]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [100742.0, 101278.0, 101815.0, 102351.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.108
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.06575
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 60.18325
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.02425
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.03675
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.03675
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.71348642033
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.49465
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-1.579, -1.285, -1.244, -1.149]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [100742.0, 101278.0, 101815.0, 102351.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-13.443, -8.82, -7.393, -6.919]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [100742.0, 101278.0, 101815.0, 102351.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.31425
varray: [-1.579, -1.285, -1.244, -1.149]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [100742.0, 101278.0, 101815.0, 102351.0]
source_col: d(18_16)Mean
dD_mean | -9.14375
varray: [-13.443, -8.82, -7.393, -6.919]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [100742.0, 101278.0, 101815.0, 102351.0]
source_col: d(D_H)Mean
d18O_stdev | 0.185464057614
varray: [-1.579, -1.285, -1.244, -1.149]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100742.0, 101278.0, 101815.0, 102351.0]
source_col: d(18_16)Mean
dD_stdev | 2.97786292219
varray: [-13.443, -8.82, -7.393, -6.919]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100742.0, 101278.0, 101815.0, 102351.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.31425
varray: [-1.579, -1.285, -1.244, -1.149]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [100742.0, 101278.0, 101815.0, 102351.0]
source_col: d(18_16)Mean
measured_dD_mean | -9.14375
varray: [-13.443, -8.82, -7.393, -6.919]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [100742.0, 101278.0, 101815.0, 102351.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.185464057614
varray: [-1.579, -1.285, -1.244, -1.149]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100742.0, 101278.0, 101815.0, 102351.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.97786292219
varray: [-13.443, -8.82, -7.393, -6.919]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100742.0, 101278.0, 101815.0, 102351.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.2231624816
actual_array: [ 1.93 -15.84]
varray: [-1.579, -1.285, -1.244, -1.149]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.31425
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [100742.0, 101278.0, 101815.0, 102351.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -10.8041956933
actual_array: [ 18.1 -119. ]
varray: [-13.443, -8.82, -7.393, -6.919]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -9.14375
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [100742.0, 101278.0, 101815.0, 102351.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.16935
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.0941
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 42 - Sample 42 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 42
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.17625
varray: [2.192, 2.161, 2.157, 2.195]
inst: mean_all
operation: mean_all(residuals)
tarray: [102937.0, 103472.0, 104008.0, 104543.0]
source_col: residuals
SAMP_residuals_STD | 0.0200229035523
varray: [2.192, 2.161, 2.157, 2.195]
inst: std_all
operation: std_all(residuals)
tarray: [102937.0, 103472.0, 104008.0, 104543.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.11875
varray: [-1.061, -1.117, -1.121, -1.176]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [102937.0, 103472.0, 104008.0, 104543.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0469778316514
varray: [-1.061, -1.117, -1.121, -1.176]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [102937.0, 103472.0, 104008.0, 104543.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 60.77325
varray: [60.989, 60.683, 60.476, 60.945]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [102937.0, 103472.0, 104008.0, 104543.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.239825457364
varray: [60.989, 60.683, 60.476, 60.945]
inst: std_all
operation: std_all(baseline_shift)
tarray: [102937.0, 103472.0, 104008.0, 104543.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.791
varray: [27.896, 27.627, 27.801, 27.84]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [102937.0, 103472.0, 104008.0, 104543.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.116077560278
varray: [27.896, 27.627, 27.801, 27.84]
inst: std_all
operation: std_all(slope_shift)
tarray: [102937.0, 103472.0, 104008.0, 104543.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0478008019459
varray: [-0.405, -0.317, -0.358, -0.297]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [102937.0, 103472.0, 104008.0, 104543.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.481146114328
varray: [-3.541, -2.886, -2.925, -2.365]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [102937.0, 103472.0, 104008.0, 104543.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.038
varray: [65.004, 64.03, 62.496, 64.622]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [102937.0, 103472.0, 104008.0, 104543.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.10333434038
varray: [65.004, 64.03, 62.496, 64.622]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [102937.0, 103472.0, 104008.0, 104543.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.17625
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.11875
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 60.77325
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.791
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0315
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0315
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.468702646
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.2266
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-0.405, -0.317, -0.358, -0.297]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [102937.0, 103472.0, 104008.0, 104543.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-3.541, -2.886, -2.925, -2.365]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [102937.0, 103472.0, 104008.0, 104543.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.34425
varray: [-0.405, -0.317, -0.358, -0.297]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [102937.0, 103472.0, 104008.0, 104543.0]
source_col: d(18_16)Mean
dD_mean | -2.92925
varray: [-3.541, -2.886, -2.925, -2.365]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [102937.0, 103472.0, 104008.0, 104543.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0478008019459
varray: [-0.405, -0.317, -0.358, -0.297]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [102937.0, 103472.0, 104008.0, 104543.0]
source_col: d(18_16)Mean
dD_stdev | 0.481146114328
varray: [-3.541, -2.886, -2.925, -2.365]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [102937.0, 103472.0, 104008.0, 104543.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.34425
varray: [-0.405, -0.317, -0.358, -0.297]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [102937.0, 103472.0, 104008.0, 104543.0]
source_col: d(18_16)Mean
measured_dD_mean | -2.92925
varray: [-3.541, -2.886, -2.925, -2.365]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [102937.0, 103472.0, 104008.0, 104543.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0478008019459
varray: [-0.405, -0.317, -0.358, -0.297]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [102937.0, 103472.0, 104008.0, 104543.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.481146114328
varray: [-3.541, -2.886, -2.925, -2.365]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [102937.0, 103472.0, 104008.0, 104543.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.22402615959
actual_array: [ 1.93 -15.84]
varray: [-0.405, -0.317, -0.358, -0.297]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.34425
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [102937.0, 103472.0, 104008.0, 104543.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -4.02861213963
actual_array: [ 18.1 -119. ]
varray: [-3.541, -2.886, -2.925, -2.365]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -2.92925
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [102937.0, 103472.0, 104008.0, 104543.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.1011
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.5041
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 43 - Sample 43 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 43
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1115
varray: [2.077, 2.109, 2.137, 2.123]
inst: mean_all
operation: mean_all(residuals)
tarray: [105127.0, 105662.0, 106197.0, 106733.0]
source_col: residuals
SAMP_residuals_STD | 0.0256839768468
varray: [2.077, 2.109, 2.137, 2.123]
inst: std_all
operation: std_all(residuals)
tarray: [105127.0, 105662.0, 106197.0, 106733.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.05525
varray: [-1.176, -0.988, -1.014, -1.043]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [105127.0, 105662.0, 106197.0, 106733.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0835758138858
varray: [-1.176, -0.988, -1.014, -1.043]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [105127.0, 105662.0, 106197.0, 106733.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 60.0305
varray: [59.873, 59.883, 60.276, 60.09]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [105127.0, 105662.0, 106197.0, 106733.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.191809801627
varray: [59.873, 59.883, 60.276, 60.09]
inst: std_all
operation: std_all(baseline_shift)
tarray: [105127.0, 105662.0, 106197.0, 106733.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.7395
varray: [26.026, 26.816, 27.38, 26.736]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [105127.0, 105662.0, 106197.0, 106733.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.555333833533
varray: [26.026, 26.816, 27.38, 26.736]
inst: std_all
operation: std_all(slope_shift)
tarray: [105127.0, 105662.0, 106197.0, 106733.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0297377425729
varray: [-1.148, -1.143, -1.205, -1.146]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [105127.0, 105662.0, 106197.0, 106733.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0703278038901
varray: [-5.045, -4.928, -5.093, -5.046]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [105127.0, 105662.0, 106197.0, 106733.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.59125
varray: [61.762, 63.09, 63.311, 62.202]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [105127.0, 105662.0, 106197.0, 106733.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.731656283146
varray: [61.762, 63.09, 63.311, 62.202]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [105127.0, 105662.0, 106197.0, 106733.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1115
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.05525
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 60.0305
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.7395
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.03325
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.03325
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.55029723744
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.22015
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-1.148, -1.143, -1.205, -1.146]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [105127.0, 105662.0, 106197.0, 106733.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-5.045, -4.928, -5.093, -5.046]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [105127.0, 105662.0, 106197.0, 106733.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.1605
varray: [-1.148, -1.143, -1.205, -1.146]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [105127.0, 105662.0, 106197.0, 106733.0]
source_col: d(18_16)Mean
dD_mean | -5.028
varray: [-5.045, -4.928, -5.093, -5.046]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [105127.0, 105662.0, 106197.0, 106733.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0297377425729
varray: [-1.148, -1.143, -1.205, -1.146]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [105127.0, 105662.0, 106197.0, 106733.0]
source_col: d(18_16)Mean
dD_stdev | 0.0703278038901
varray: [-5.045, -4.928, -5.093, -5.046]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [105127.0, 105662.0, 106197.0, 106733.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.1605
varray: [-1.148, -1.143, -1.205, -1.146]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [105127.0, 105662.0, 106197.0, 106733.0]
source_col: d(18_16)Mean
measured_dD_mean | -5.028
varray: [-5.045, -4.928, -5.093, -5.046]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [105127.0, 105662.0, 106197.0, 106733.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0297377425729
varray: [-1.148, -1.143, -1.205, -1.146]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [105127.0, 105662.0, 106197.0, 106733.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0703278038901
varray: [-5.045, -4.928, -5.093, -5.046]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [105127.0, 105662.0, 106197.0, 106733.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.06479422437
actual_array: [ 1.93 -15.84]
varray: [-1.148, -1.143, -1.205, -1.146]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.1605
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [105127.0, 105662.0, 106197.0, 106733.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -6.31685028963
actual_array: [ 18.1 -119. ]
varray: [-5.045, -4.928, -5.093, -5.046]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -5.028
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [105127.0, 105662.0, 106197.0, 106733.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.16585
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.24685
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 44 - Sample 44 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Sample Variables
Sample 44
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.093
varray: [2.09, 2.097, 2.089, 2.096]
inst: mean_all
operation: mean_all(residuals)
tarray: [107318.0, 107853.0, 108389.0, 108925.0]
source_col: residuals
SAMP_residuals_STD | 0.00408248290464
varray: [2.09, 2.097, 2.089, 2.096]
inst: std_all
operation: std_all(residuals)
tarray: [107318.0, 107853.0, 108389.0, 108925.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.97825
varray: [-1.028, -0.96, -1.065, -0.86]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [107318.0, 107853.0, 108389.0, 108925.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0900310131751
varray: [-1.028, -0.96, -1.065, -0.86]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [107318.0, 107853.0, 108389.0, 108925.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 60.868
varray: [60.753, 60.997, 60.925, 60.797]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [107318.0, 107853.0, 108389.0, 108925.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.112777066227
varray: [60.753, 60.997, 60.925, 60.797]
inst: std_all
operation: std_all(baseline_shift)
tarray: [107318.0, 107853.0, 108389.0, 108925.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.78125
varray: [27.546, 27.884, 27.93, 27.765]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [107318.0, 107853.0, 108389.0, 108925.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.171552470108
varray: [27.546, 27.884, 27.93, 27.765]
inst: std_all
operation: std_all(slope_shift)
tarray: [107318.0, 107853.0, 108389.0, 108925.0]
source_col: slope_shift
SAMP_d18O_STD | 0.178199326598
varray: [-8.61, -8.931, -8.963, -8.994]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [107318.0, 107853.0, 108389.0, 108925.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.11874408483
varray: [-58.175, -63.406, -64.429, -64.986]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [107318.0, 107853.0, 108389.0, 108925.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.8425
varray: [62.683, 62.902, 62.175, 63.61]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [107318.0, 107853.0, 108389.0, 108925.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.595425058257
varray: [62.683, 62.902, 62.175, 63.61]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [107318.0, 107853.0, 108389.0, 108925.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.093
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.97825
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 60.868
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.78125
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.05175
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.05175
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.41286863271
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.9689
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.534
varray: [-8.61, -8.931, -8.963, -8.994]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [107318.0, 107853.0, 108389.0, 108925.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3799
varray: [-58.175, -63.406, -64.429, -64.986]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [107318.0, 107853.0, 108389.0, 108925.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.8745
varray: [-8.61, -8.931, -8.963, -8.994]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [107318.0, 107853.0, 108389.0, 108925.0]
source_col: d(18_16)Mean
dD_mean | -62.749
varray: [-58.175, -63.406, -64.429, -64.986]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [107318.0, 107853.0, 108389.0, 108925.0]
source_col: d(D_H)Mean
d18O_stdev | 0.178199326598
varray: [-8.61, -8.931, -8.963, -8.994]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [107318.0, 107853.0, 108389.0, 108925.0]
source_col: d(18_16)Mean
dD_stdev | 3.11874408483
varray: [-58.175, -63.406, -64.429, -64.986]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [107318.0, 107853.0, 108389.0, 108925.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.8745
varray: [-8.61, -8.931, -8.963, -8.994]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [107318.0, 107853.0, 108389.0, 108925.0]
source_col: d(18_16)Mean
measured_dD_mean | -62.749
varray: [-58.175, -63.406, -64.429, -64.986]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [107318.0, 107853.0, 108389.0, 108925.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.178199326598
varray: [-8.61, -8.931, -8.963, -8.994]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [107318.0, 107853.0, 108389.0, 108925.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.11874408483
varray: [-58.175, -63.406, -64.429, -64.986]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [107318.0, 107853.0, 108389.0, 108925.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.0105030779
actual_array: [ 1.93 -15.84]
varray: [-8.61, -8.931, -8.963, -8.994]
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -8.8745
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: [107318.0, 107853.0, 108389.0, 108925.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -69.2492597028
actual_array: [ 18.1 -119. ]
varray: [-58.175, -63.406, -64.429, -64.986]
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -62.749
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: [107318.0, 107853.0, 108389.0, 108925.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 59.18435
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.40935
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 1 - Sample 1 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 1 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -10.178
operation: col(d(18_16)Mean)
inst: col
dD_actual | -76.906
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.353156888
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -10.178
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -84.6844413536
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -76.906
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 1 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.394
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.711
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5456075308
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -9.394
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.9301183807
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -70.711
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 1 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.175
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.643
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3200293303
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -9.175
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.6754065312
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -68.643
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 1 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.137
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.221
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2808879073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -9.137
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.2153057573
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -68.221
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 1 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.117
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.061
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2602871584
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -9.117
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.0408599662
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -68.061
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 1 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.099
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.621
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2417464844
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -9.099
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.5611340408
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -67.621
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 1 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.135
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.659
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2788278325
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -9.135
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.6025649162
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -67.659
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 1 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.142
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.757
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2860380946
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -9.142
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.7094129632
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -67.757
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 1 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.108
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.325
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2510168214
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -9.108
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.2384093273
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -67.325
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 1 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.113
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.465
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2561670087
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -9.113
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3910493945
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -67.465
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 2 - Sample 2 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 2 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.872
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.819
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7178
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3669
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.202
operation: col(d(18_16)Mean)
inst: col
dD_actual | 8.397
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.39870668568
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 2.202
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 8.32024186699
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: 8.397
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 2 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.872
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.819
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7178
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3669
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.573
operation: col(d(18_16)Mean)
inst: col
dD_actual | 15.173
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.78085057791
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 2.573
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 15.7080211186
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: 15.173
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 2 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.872
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.819
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7178
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3669
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.675
operation: col(d(18_16)Mean)
inst: col
dD_actual | 16.9
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.88591439734
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 2.675
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 17.5909453759
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: 16.9
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 2 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.872
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.819
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7178
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3669
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.758
operation: col(d(18_16)Mean)
inst: col
dD_actual | 17.787
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.9714075053
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 2.758
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 18.5580292302
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: 17.787
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 2 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.872
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.819
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7178
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3669
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.763
operation: col(d(18_16)Mean)
inst: col
dD_actual | 18.49
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.97655769253
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 2.763
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.3245004247
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: 18.49
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 2 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.872
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.819
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7178
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3669
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.82
operation: col(d(18_16)Mean)
inst: col
dD_actual | 18.808
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.03526982692
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 2.82
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.6712114344
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: 18.808
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 2 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.872
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.819
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7178
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3669
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.831
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.204
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.04660023882
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 2.831
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.1029647673
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: 19.204
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 2 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.872
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.819
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7178
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3669
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.807
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.363
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.02187934013
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 2.807
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.2763202722
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: 19.363
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 2 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.872
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.819
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7178
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3669
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.877
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.728
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.0939819613
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 2.877
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.674274733
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: 19.728
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 2 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.872
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.819
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7178
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3669
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.872
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.819
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.08883177408
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 2.872
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.7734907767
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: 19.819
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 3 - Sample 3 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 3 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.758
operation: col(d(18_16)Mean)
inst: col
dD_actual | -94.826
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.0406909424
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -13.758
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -104.222369953
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -94.826
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 3 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.383
operation: col(d(18_16)Mean)
inst: col
dD_actual | -105.397
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.6844643457
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -14.383
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -115.747785311
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -105.397
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 3 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.504
operation: col(d(18_16)Mean)
inst: col
dD_actual | -107.942
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.8090988766
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -14.504
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -118.522563676
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -107.942
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 3 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.623
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.355
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9316733326
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -14.623
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.063138068
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -109.355
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 3 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.595
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.043
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9028322842
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -14.595
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.81325497
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -110.043
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 3 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.656
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.713
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9656645683
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -14.656
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.54374672
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -110.713
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 3 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.68
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.903
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.990385467
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -14.68
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.750901096
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -110.903
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 3 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.679
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.333
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9893554296
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -14.679
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.21972416
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -111.333
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 3 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.721
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.644
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0326170023
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -14.721
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.558803166
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -111.644
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 3 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.741
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.643
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0532177512
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -14.741
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.55771288
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -111.643
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 4 - Sample 4 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 4 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.463
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.588
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.6166801145
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -9.463
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.9765855672
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -72.588
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 4 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.218
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.083
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3643209404
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -9.218
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.1551324567
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -69.083
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 4 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.203
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.329
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3488703787
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -9.203
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.3330566663
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -68.329
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 4 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.149
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.776
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2932483567
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -9.149
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.7301284009
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -67.776
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 4 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.176
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.679
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3210593677
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -9.176
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.62437064
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -67.679
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 4 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.185
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.517
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3303297047
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -9.185
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.4477442766
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -67.517
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 4 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.12
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.375
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2633772708
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -9.12
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.292923637
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -67.375
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 4 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.146
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.526
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2901582444
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -9.146
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.4575568523
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -67.526
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 4 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.123
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.242
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2664673831
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -9.123
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.1479155732
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -67.242
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 4 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.11
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.294
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2530768963
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -9.11
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.2046104553
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -67.294
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 5 - Sample 5 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 5 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.501
operation: col(d(18_16)Mean)
inst: col
dD_actual | -24.141
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.41552197452
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.501
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.1554903186
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -24.141
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 5 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.329
operation: col(d(18_16)Mean)
inst: col
dD_actual | -20.517
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.23835553392
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.329
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -23.204293151
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -20.517
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 5 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.288
operation: col(d(18_16)Mean)
inst: col
dD_actual | -19.645
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.19612399866
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.288
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -22.2535635897
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -19.645
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 5 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.188
operation: col(d(18_16)Mean)
inst: col
dD_actual | -19.112
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.09312025412
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.188
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -21.6724410482
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -19.112
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 6 - Sample 6 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 6 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.481
operation: col(d(18_16)Mean)
inst: col
dD_actual | -8.309
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.373987757799
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 0.481
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -9.89407929267
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -8.309
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 6 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.51
operation: col(d(18_16)Mean)
inst: col
dD_actual | -7.675
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.344116671884
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 0.51
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -9.20283784558
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -7.675
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 6 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.51
operation: col(d(18_16)Mean)
inst: col
dD_actual | -7.193
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.344116671884
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 0.51
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.67731989999
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -7.193
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 6 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.565
operation: col(d(18_16)Mean)
inst: col
dD_actual | -6.505
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.287464612388
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 0.565
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -7.92720299841
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -6.505
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 7 - Sample 7 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 7 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.197
operation: col(d(18_16)Mean)
inst: col
dD_actual | -12.022
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.07235314576
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.197
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -13.9423119316
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -12.022
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 7 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.273
operation: col(d(18_16)Mean)
inst: col
dD_actual | -12.911
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.15063599161
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.273
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -14.9115763582
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -12.911
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 7 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.279
operation: col(d(18_16)Mean)
inst: col
dD_actual | -12.856
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.15681621628
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.279
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -14.8516106175
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -12.856
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 7 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.258
operation: col(d(18_16)Mean)
inst: col
dD_actual | -12.661
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.13518542993
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.258
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -14.6390048097
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -12.661
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 8 - Sample 8 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 8 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.287
operation: col(d(18_16)Mean)
inst: col
dD_actual | -15.399
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.16505651584
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.287
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.6242084093
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -15.399
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 8 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.231
operation: col(d(18_16)Mean)
inst: col
dD_actual | -15.417
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.1073744189
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.231
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.6438335608
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -15.417
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 8 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.265
operation: col(d(18_16)Mean)
inst: col
dD_actual | -15.45
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.14239569204
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.265
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.6798130052
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -15.45
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 8 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.241
operation: col(d(18_16)Mean)
inst: col
dD_actual | -15.495
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.11767479335
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.241
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.7288758839
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -15.495
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 9 - Sample 9 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 9 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.507
operation: col(d(18_16)Mean)
inst: col
dD_actual | -22.072
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.42170219919
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.507
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -24.8996881829
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -22.072
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 9 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.578
operation: col(d(18_16)Mean)
inst: col
dD_actual | -22.982
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.49483485781
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.578
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.8918486196
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -22.982
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 9 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.561
operation: col(d(18_16)Mean)
inst: col
dD_actual | -23.338
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.47732422124
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.561
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -26.2799905047
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -23.338
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 9 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.604
operation: col(d(18_16)Mean)
inst: col
dD_actual | -23.366
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.52161583139
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.604
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -26.3105185182
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -23.366
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 10 - Sample 10 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 10 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.914
operation: col(d(18_16)Mean)
inst: col
dD_actual | -22.621
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.84092743946
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.914
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.4982553035
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -22.621
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 10 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.963
operation: col(d(18_16)Mean)
inst: col
dD_actual | -22.53
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.89139927428
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.963
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.3990392598
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -22.53
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 10 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.927
operation: col(d(18_16)Mean)
inst: col
dD_actual | -22.21
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.85431792625
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.927
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.0501476777
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -22.21
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 10 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.959
operation: col(d(18_16)Mean)
inst: col
dD_actual | -22.396
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.8872791245
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.959
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.2529409098
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -22.396
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 11 - Sample 11 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 11 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.39
operation: col(d(18_16)Mean)
inst: col
dD_actual | -29.411
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.33122526345
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -2.39
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -32.9012985619
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -29.411
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 11 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.502
operation: col(d(18_16)Mean)
inst: col
dD_actual | -29.924
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.44658945733
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -2.502
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -33.4606153795
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -29.924
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 11 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.483
operation: col(d(18_16)Mean)
inst: col
dD_actual | -30.056
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.42701874587
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -2.483
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -33.6045331571
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -30.056
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 11 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.515
operation: col(d(18_16)Mean)
inst: col
dD_actual | -30.155
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.45997994412
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -2.515
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -33.7124714903
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -30.155
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 12 - Sample 12 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 12 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.695
operation: col(d(18_16)Mean)
inst: col
dD_actual | -5.64
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.876477700878
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 1.695
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -6.98410544046
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -5.64
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 12 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.905
operation: col(d(18_16)Mean)
inst: col
dD_actual | -3.408
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.0927855644
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 1.905
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -4.55058665509
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -3.408
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 12 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.939
operation: col(d(18_16)Mean)
inst: col
dD_actual | -2.909
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.12780683755
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 1.939
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -4.0065338442
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -2.909
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 12 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.93
operation: col(d(18_16)Mean)
inst: col
dD_actual | -2.675
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.11853650054
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 1.93
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -3.75140687477
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -2.675
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 13 - Sample 13 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 13 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.48
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.554
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.60415330574
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -8.48
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.6755091183
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -58.554
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 13 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.875
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.55
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0110180967
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -8.875
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.1225789444
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -63.55
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 13 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.922
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.861
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0594298566
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -8.922
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.5519441449
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -64.861
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 13 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.028
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.426
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1686138258
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -9.028
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.1679558446
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -65.426
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 14 - Sample 14 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 14 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.022
operation: col(d(18_16)Mean)
inst: col
dD_actual | -11.737
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.846774945223
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 0.022
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -13.6315803663
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -11.737
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 14 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.288
operation: col(d(18_16)Mean)
inst: col
dD_actual | -7.239
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.572784984755
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 0.288
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.72747306492
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -7.239
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 14 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.388
operation: col(d(18_16)Mean)
inst: col
dD_actual | -5.886
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.469781240218
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 0.388
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -7.25231584422
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -5.886
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 14 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.476
operation: col(d(18_16)Mean)
inst: col
dD_actual | -5.027
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.379137945026
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 0.476
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -6.31576000344
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -5.027
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 15 - Sample 15 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 15 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.197
operation: col(d(18_16)Mean)
inst: col
dD_actual | -8.665
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.07235314576
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.197
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -10.2822211778
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -8.665
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 15 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.206
operation: col(d(18_16)Mean)
inst: col
dD_actual | -8.876
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.08162348277
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.206
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -10.5122715648
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -8.876
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 15 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.214
operation: col(d(18_16)Mean)
inst: col
dD_actual | -8.922
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.08986378233
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.214
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -10.5624247297
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -8.922
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 15 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.151
operation: col(d(18_16)Mean)
inst: col
dD_actual | -8.929
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.02497142327
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.151
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -10.5700567331
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -8.929
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 16 - Sample 16 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 16 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.912
operation: col(d(18_16)Mean)
inst: col
dD_actual | -9.563
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.13003327189
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 2.912
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -11.2612981802
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -9.563
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 16 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.999
operation: col(d(18_16)Mean)
inst: col
dD_actual | -9.316
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.21964652964
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 2.999
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -10.9919974902
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -9.316
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 16 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.075
operation: col(d(18_16)Mean)
inst: col
dD_actual | -9.358
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.29792937549
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 3.075
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -11.0377895103
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -9.358
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 16 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.056
operation: col(d(18_16)Mean)
inst: col
dD_actual | -9.299
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.27835866402
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 3.056
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -10.9734626249
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -9.299
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 17 - Sample 17 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 17 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.1
operation: col(d(18_16)Mean)
inst: col
dD_actual | -21.889
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.972439513558
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.1
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -24.7001658094
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -21.889
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 17 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.252
operation: col(d(18_16)Mean)
inst: col
dD_actual | -23.393
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.12900520525
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.252
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -26.3399562454
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -23.393
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 17 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.255
operation: col(d(18_16)Mean)
inst: col
dD_actual | -23.831
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.13209531759
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.255
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -26.8175015985
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -23.831
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 17 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.31
operation: col(d(18_16)Mean)
inst: col
dD_actual | -24.089
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.18874737709
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.31
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.0987954365
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -24.089
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 18 - Sample 18 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 18 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.002
operation: col(d(18_16)Mean)
inst: col
dD_actual | -20.816
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.90153328928
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.002
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -23.5302887231
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -20.816
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 18 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.001
operation: col(d(18_16)Mean)
inst: col
dD_actual | -20.42
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.90050325183
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.001
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -23.0985353902
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -20.42
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 18 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.042
operation: col(d(18_16)Mean)
inst: col
dD_actual | -20.108
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.94273478709
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.042
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -22.7583660976
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -20.108
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 18 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.015
operation: col(d(18_16)Mean)
inst: col
dD_actual | -20.357
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.91492377607
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.015
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -23.0298473599
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -20.357
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 19 - Sample 19 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 19 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.89
operation: col(d(18_16)Mean)
inst: col
dD_actual | -19.84
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.7861690954
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.89
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -22.4661693976
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -19.84
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 19 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.92
operation: col(d(18_16)Mean)
inst: col
dD_actual | -20.085
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.81707021876
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.92
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -22.7332895151
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -20.085
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 19 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.903
operation: col(d(18_16)Mean)
inst: col
dD_actual | -19.898
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.79955958219
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.903
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -22.5294059968
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -19.898
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 19 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.912
operation: col(d(18_16)Mean)
inst: col
dD_actual | -20.067
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.8088299192
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.912
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -22.7136643636
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -20.067
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 20 - Sample 20 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 20 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.872
operation: col(d(18_16)Mean)
inst: col
dD_actual | -22.0
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.76762842138
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.872
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -24.8211875769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -22.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 20 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.919
operation: col(d(18_16)Mean)
inst: col
dD_actual | -22.397
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.81604018131
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.919
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.254031196
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -22.397
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 20 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.864
operation: col(d(18_16)Mean)
inst: col
dD_actual | -22.381
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.75938812182
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.864
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.2365866169
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -22.381
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 20 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.926
operation: col(d(18_16)Mean)
inst: col
dD_actual | -22.617
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.82325044343
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.926
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.4938941587
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -22.617
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 21 - Sample 21 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 21 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.204
operation: col(d(18_16)Mean)
inst: col
dD_actual | -14.371
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.07956340788
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.204
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.5033942017
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -14.371
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 21 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.145
operation: col(d(18_16)Mean)
inst: col
dD_actual | -13.5
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.0187911986
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.145
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -15.5537549266
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -13.5
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 21 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.151
operation: col(d(18_16)Mean)
inst: col
dD_actual | -13.255
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.02497142327
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.151
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -15.286634809
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -13.255
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 21 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.17
operation: col(d(18_16)Mean)
inst: col
dD_actual | -13.237
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.04454213473
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.17
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -15.2670096575
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -13.237
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 22 - Sample 22 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 22 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.622
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.422
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.75041862298
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -8.622
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -65.6218775349
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -59.422
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 22 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.875
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.799
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0110180967
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -8.875
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.3940602067
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -63.799
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 22 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.98
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.874
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1191720284
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -8.98
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.5661178654
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -64.874
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 22 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.002
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.328
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1418328522
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -9.002
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.0611077976
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -65.328
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 23 - Sample 23 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 23 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.779
operation: col(d(18_16)Mean)
inst: col
dD_actual | -22.247
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.70187238433
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.779
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.0904882669
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -22.247
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 23 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.51
operation: col(d(18_16)Mean)
inst: col
dD_actual | -18.58
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.42479231153
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.51
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -21.0924087929
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -18.58
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 23 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.426
operation: col(d(18_16)Mean)
inst: col
dD_actual | -17.764
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.33826916612
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.426
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -20.2027352585
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -17.764
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 23 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.424
operation: col(d(18_16)Mean)
inst: col
dD_actual | -17.367
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.33620909123
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.424
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -19.7698916394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -17.367
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 24 - Sample 24 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 24 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.275
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.754
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.21277095723
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -2.275
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.9141219498
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -25.754
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 24 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.326
operation: col(d(18_16)Mean)
inst: col
dD_actual | -26.673
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.26530286695
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -2.326
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -29.9160949623
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -26.673
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 24 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.35
operation: col(d(18_16)Mean)
inst: col
dD_actual | -26.66
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.29002376564
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -2.35
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -29.9019212417
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -26.66
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 24 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.302
operation: col(d(18_16)Mean)
inst: col
dD_actual | -26.707
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.24058196826
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -2.302
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -29.9531646929
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -26.707
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 25 - Sample 25 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 25 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.02
operation: col(d(18_16)Mean)
inst: col
dD_actual | -15.089
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.181202425254
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 1.02
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.2862196891
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -15.089
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 25 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.169
operation: col(d(18_16)Mean)
inst: col
dD_actual | -14.069
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.334678004614
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 1.169
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.174127771
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -14.069
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 25 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.168
operation: col(d(18_16)Mean)
inst: col
dD_actual | -13.77
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.333647967169
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 1.168
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -15.848132199
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -13.77
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 25 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.171
operation: col(d(18_16)Mean)
inst: col
dD_actual | -13.825
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.336738079505
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 1.171
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -15.9080979397
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -13.825
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 26 - Sample 26 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 26 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.112
operation: col(d(18_16)Mean)
inst: col
dD_actual | -10.958
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.275965870228
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 1.112
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -12.782247421
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -10.958
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 26 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.139
operation: col(d(18_16)Mean)
inst: col
dD_actual | -10.385
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.303776881253
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 1.139
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -12.1575134318
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -10.385
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 26 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.142
operation: col(d(18_16)Mean)
inst: col
dD_actual | -10.561
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.306866993589
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 1.142
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -12.3494038019
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -10.561
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 26 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.154
operation: col(d(18_16)Mean)
inst: col
dD_actual | -10.108
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.319227442933
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 1.154
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -11.855504156
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -10.108
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 27 - Sample 27 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 27 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
dD_actual | -14.443
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.869435769021
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.5818948077
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -14.443
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 27 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.052
operation: col(d(18_16)Mean)
inst: col
dD_actual | -14.696
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.92299771618
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.052
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.8577372148
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -14.696
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 27 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.077
operation: col(d(18_16)Mean)
inst: col
dD_actual | -14.793
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.948748652315
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.077
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9634949756
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -14.793
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 27 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.105
operation: col(d(18_16)Mean)
inst: col
dD_actual | -14.698
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.977589700785
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.105
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.8599177872
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -14.698
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 28 - Sample 28 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 28 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.529
operation: col(d(18_16)Mean)
inst: col
dD_actual | -13.595
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.324545960422
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 0.529
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -15.657332115
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -13.595
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 28 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.479
operation: col(d(18_16)Mean)
inst: col
dD_actual | -14.383
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.37604783269
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 0.479
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.516477636
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -14.383
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 28 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.498
operation: col(d(18_16)Mean)
inst: col
dD_actual | -14.088
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.356477121228
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 0.498
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.1948432087
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -14.088
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 28 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.47
operation: col(d(18_16)Mean)
inst: col
dD_actual | -13.789
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.385318169698
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 0.47
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -15.8688476367
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -13.789
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 29 - Sample 29 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 29 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.708
operation: col(d(18_16)Mean)
inst: col
dD_actual | -4.015
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.140169257701
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 0.708
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -5.21239037494
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -4.015
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 29 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.756
operation: col(d(18_16)Mean)
inst: col
dD_actual | -2.771
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.090727460323
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 0.756
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -3.85607434941
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -2.771
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 29 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.764
operation: col(d(18_16)Mean)
inst: col
dD_actual | -2.422
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.08248716076
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 0.764
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -3.47556446764
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -2.422
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 29 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.758
operation: col(d(18_16)Mean)
inst: col
dD_actual | -2.123
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.0886673854322
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: 0.758
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -3.14956889559
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -2.123
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 30 - Sample 30 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 30 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.087
operation: col(d(18_16)Mean)
inst: col
dD_actual | -3.269
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.959049026768
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.087
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -4.3990368741
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -3.269
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 30 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.102
operation: col(d(18_16)Mean)
inst: col
dD_actual | -3.172
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.974499588449
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.102
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -4.29327911327
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -3.172
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 30 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.075
operation: col(d(18_16)Mean)
inst: col
dD_actual | -3.204
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.946688577424
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.075
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -4.32816827148
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -3.204
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 30 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.047
operation: col(d(18_16)Mean)
inst: col
dD_actual | -3.034
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.917847528954
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.047
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -4.14281961847
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -3.034
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 31 - Sample 31 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 31 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.59
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.34
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.71745742473
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -8.59
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.4421878728
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -58.34
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 31 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.856
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.942
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.9914473852
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -8.856
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.4596849383
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -62.942
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 31 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.943
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.408
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0810606429
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -8.943
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.0580444989
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -64.408
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 31 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.981
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.093
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1202020659
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -8.981
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.8048905419
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -65.093
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 32 - Sample 32 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 32 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.445
operation: col(d(18_16)Mean)
inst: col
dD_actual | -21.482
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.35783987758
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.445
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -24.2564193284
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -21.482
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 32 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.05
operation: col(d(18_16)Mean)
inst: col
dD_actual | -17.276
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.95097508666
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.05
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -19.6706755957
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -17.276
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 32 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.046
operation: col(d(18_16)Mean)
inst: col
dD_actual | -16.588
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.94685493688
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.046
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.9205586941
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -16.588
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 32 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.019
operation: col(d(18_16)Mean)
inst: col
dD_actual | -16.008
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.91904392585
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.019
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.2881927015
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -16.008
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 33 - Sample 33 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 33 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.921
operation: col(d(18_16)Mean)
inst: col
dD_actual | -14.673
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.81810025621
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.921
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.8326606323
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -14.673
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 33 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.901
operation: col(d(18_16)Mean)
inst: col
dD_actual | -14.061
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.7974995073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.901
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.1654054815
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -14.061
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 33 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.89
operation: col(d(18_16)Mean)
inst: col
dD_actual | -14.037
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.7861690954
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.89
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.1392386128
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -14.037
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 33 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.949
operation: col(d(18_16)Mean)
inst: col
dD_actual | -13.937
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.84694130468
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.949
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.0302099934
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -13.937
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 34 - Sample 34 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 34 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.733
operation: col(d(18_16)Mean)
inst: col
dD_actual | -16.118
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.62445321648
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.733
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.4081241829
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -16.118
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 34 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.724
operation: col(d(18_16)Mean)
inst: col
dD_actual | -16.273
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.61518287947
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.724
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.577118543
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -16.273
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 34 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.752
operation: col(d(18_16)Mean)
inst: col
dD_actual | -16.37
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.64402392794
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.752
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.6828763038
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -16.37
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 34 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.768
operation: col(d(18_16)Mean)
inst: col
dD_actual | -16.231
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.66050452706
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.768
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.5313265228
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -16.231
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 35 - Sample 35 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 35 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.626
operation: col(d(18_16)Mean)
inst: col
dD_actual | -3.913
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.51423920982
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.626
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -5.10118118314
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -3.913
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 35 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.591
operation: col(d(18_16)Mean)
inst: col
dD_actual | -2.511
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.47818789923
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.591
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -3.57259993892
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -2.511
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 35 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.62
operation: col(d(18_16)Mean)
inst: col
dD_actual | -2.33
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.50805898515
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.62
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -3.37525813778
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -2.33
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 35 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.593
operation: col(d(18_16)Mean)
inst: col
dD_actual | -2.372
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.48024797412
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.593
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -3.42105015794
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -2.372
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 36 - Sample 36 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 36 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.226
operation: col(d(18_16)Mean)
inst: col
dD_actual | -4.314
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.10222423167
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.226
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -5.538385947
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -4.314
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 36 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.202
operation: col(d(18_16)Mean)
inst: col
dD_actual | -4.485
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.07750333299
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.202
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -5.7248248862
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -4.485
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 36 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.167
operation: col(d(18_16)Mean)
inst: col
dD_actual | -4.372
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.0414520224
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.167
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -5.60162254626
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -4.372
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 36 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.162
operation: col(d(18_16)Mean)
inst: col
dD_actual | -4.273
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.03630183517
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.162
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -5.49368421304
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -4.273
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 37 - Sample 37 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 37 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.543
operation: col(d(18_16)Mean)
inst: col
dD_actual | -2.639
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.42874610186
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.543
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -3.71215657178
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -2.639
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 37 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.515
operation: col(d(18_16)Mean)
inst: col
dD_actual | -2.566
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.39990505339
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.515
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -3.6325656796
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -2.566
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 37 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.522
operation: col(d(18_16)Mean)
inst: col
dD_actual | -2.788
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.4071153155
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.522
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -3.87460921471
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -2.788
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 37 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.554
operation: col(d(18_16)Mean)
inst: col
dD_actual | -2.842
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.44007651376
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.554
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -3.93348466919
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -2.842
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 38 - Sample 38 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 38 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.636
operation: col(d(18_16)Mean)
inst: col
dD_actual | -3.473
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.52453958428
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.636
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -4.62145525771
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -3.473
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 38 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.691
operation: col(d(18_16)Mean)
inst: col
dD_actual | -3.646
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.58119164377
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.691
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -4.8100747693
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -3.646
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 38 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.646
operation: col(d(18_16)Mean)
inst: col
dD_actual | -3.667
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.53483995873
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.646
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -4.83297077938
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -3.667
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 38 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.65
operation: col(d(18_16)Mean)
inst: col
dD_actual | -3.708
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.53896010851
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.65
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -4.87767251334
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -3.708
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 39 - Sample 39 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 39 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.994
operation: col(d(18_16)Mean)
inst: col
dD_actual | -7.859
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.89329298972
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.994
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -9.4034505053
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -7.859
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 39 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.039
operation: col(d(18_16)Mean)
inst: col
dD_actual | -8.289
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.93964467476
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.039
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -9.87227356879
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -8.289
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 39 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.078
operation: col(d(18_16)Mean)
inst: col
dD_actual | -8.203
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.97981613513
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.078
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -9.77850895609
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -8.203
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 39 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.102
operation: col(d(18_16)Mean)
inst: col
dD_actual | -8.464
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.00453703382
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.102
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -10.0630736528
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -8.464
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 40 - Sample 40 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 40 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.604
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.517
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.73187794897
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -8.604
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.6351685291
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -58.517
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 40 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.871
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.039
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0068979469
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -8.871
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.5654426991
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -63.039
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 40 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.003
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.643
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1428628897
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -9.003
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.3142617546
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -64.643
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 40 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.017
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.15
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1572834139
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -9.017
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.867036855
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -65.15
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 41 - Sample 41 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 41 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.579
operation: col(d(18_16)Mean)
inst: col
dD_actual | -13.443
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.49586489526
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.579
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -15.4916086135
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -13.443
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 41 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.285
operation: col(d(18_16)Mean)
inst: col
dD_actual | -8.82
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.19303388632
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.285
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -10.4512155379
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -8.82
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 41 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.244
operation: col(d(18_16)Mean)
inst: col
dD_actual | -7.393
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.15080235106
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.244
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.89537713882
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -7.393
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 41 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.149
operation: col(d(18_16)Mean)
inst: col
dD_actual | -6.919
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.05294879375
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.149
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.37858148279
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -6.919
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 42 - Sample 42 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 42 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.405
operation: col(d(18_16)Mean)
inst: col
dD_actual | -3.541
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.2866009344
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.405
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -4.69559471891
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -3.541
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 42 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.317
operation: col(d(18_16)Mean)
inst: col
dD_actual | -2.886
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.1959576392
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.317
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -3.98145726174
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -2.886
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 42 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.358
operation: col(d(18_16)Mean)
inst: col
dD_actual | -2.925
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.23818917446
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.358
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -4.02397842331
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -2.925
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 42 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.297
operation: col(d(18_16)Mean)
inst: col
dD_actual | -2.365
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.1753568903
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -0.297
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -3.41341815458
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -2.365
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 43 - Sample 43 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 43 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.148
operation: col(d(18_16)Mean)
inst: col
dD_actual | -5.045
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.0519187563
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.148
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -6.33538515493
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -5.045
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 43 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.143
operation: col(d(18_16)Mean)
inst: col
dD_actual | -4.928
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.04676856908
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.143
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -6.20782167021
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -4.928
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 43 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.205
operation: col(d(18_16)Mean)
inst: col
dD_actual | -5.093
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.11063089069
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.205
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -6.38771889225
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -5.093
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 43 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.146
operation: col(d(18_16)Mean)
inst: col
dD_actual | -5.046
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.04985868141
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -1.146
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -6.33647544112
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -5.046
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean

## 44 - Sample 44 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 11:50:32
Detail Variables
Detail 44 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.61
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.175
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.73805817364
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -8.61
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.2622906507
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -58.175
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 44 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.931
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.406
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0687001936
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -8.931
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.9655777324
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -63.406
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 44 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.963
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.429
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1016613919
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -8.963
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.080940509
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -64.429
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean
Detail 44 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7178
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3669
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.741
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.643
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3799
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.994
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.986
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1335925527
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.844081710748
measure_array: [ 2.7178 -14.534 ]
inst: fit_type1
actual_mean: -8.994
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.97083849184
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.6882299192
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.765754266958
measure_array: [ 17.3669 -108.3799]
inst: fit_type1
actual_mean: -64.986
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.917190371991
tarray: []
source_col: d(D_H)Mean