## Summary
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
221003_22-233_22-254_HIDS2052
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -10.2005 | -74.4738 |  | 9.8999 |  | 0.17253 | 1.8417
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.25805 | 3.8943
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.37092 | 5.7273
4 |  | EV | -10.3395 | -74.8965 |  |  |  | 0.12266 | 1.9029
5 |  | Mill Creek Porter Fork Upper-01 | -17.2911 | -128.0719 |  |  |  | 0.17912 | 2.9735
6 |  | Porter Fork-02 | -17.4731 | -130.2768 |  |  |  | 0.026961 | 0.27531
7 |  | Mill Creek Porter Fork Lower-03 | -17.0174 | -129.2209 |  | 4.7848 |  | 0.47245 | 1.565
8 |  | Mill Creek Upstream of Church Fork-04 | -17.4628 | -130.8026 |  |  |  | 0.070798 | 0.31887
9 |  | Church Fork-05 | -16.3149 | -122.939 |  |  |  | 0.065059 | 0.61399
10 |  | Mill Creek Down Stream of Church Fork-06 | -17.4012 | -130.1062 |  |  |  | 0.023301 | 0.46268
11 |  | Gate House-07 | -17.3595 | -129.9877 |  |  |  | 0.054948 | 0.19476
12 |  | TALL4-D-5 | 1.9177 | -26.1706 | -3.4843 | 371.5286 | 4.003 | 0.45618 | 5.8283 | -22.207 | -0.6123
13 |  | EV | -10.1366 | -72.4843 |  |  |  | 0.18962 | 2.0721
14 |  | TALL4-A-1 | 0.25063 | -26.9565 | -2.2775 | 72.2303 |  | 0.2495 | 2.6387 | -8.0758 | -4.4355
15 |  | TALL4-D-1 | -0.54581 | -19.726 | -1.4963 | 232.0078 |  | 0.022 | 0.54942 | -11.5813 | -5.5568
16 |  | TALL4-B-4 | -1.5585 | -16.7009 |  | 481.3936 |  | 0.041857 | 0.2909 | -16.8858 | -7.1528
17 |  | TALL4-3040-2 | -3.5016 | -22.5853 |  |  |  | 0.059757 | 0.32366
18 |  | TALL4-deep-2 | -4.0968 | -30.8234 |  |  |  | 0.026209 | 0.38034
19 |  | TALL4-1020-2 | -2.5045 | -16.0508 |  |  |  | 0.092124 | 0.78179
20 |  | TALL4-2030-1 | -3.0854 | -23.0534 |  |  |  | 0.015513 | 0.35914
21 |  | TALL4-C-1 | -1.7407 | -18.5599 |  | 47.2351 |  | 0.0093274 | 0.16638 |  | 1.8595
22 |  | EV | -9.9591 | -70.4502 |  |  |  | 0.17116 | 3.0587
23 |  | TALL4-3040-1 | -2.8792 | -29.2585 |  |  |  | 0.20155 | 2.2939
24 |  | TALL4-C-5 | -1.9093 | -27.6048 |  | 84.7943 |  | 0.020982 | 0.22741 |  | -3.586
25 |  | TALL4-C-3 | -2.0675 | -25.3383 |  |  |  | 0.036142 | 0.19408
26 |  | TALL4-A-2 | -0.97499 | -19.8875 |  | 56.5608 |  | 0.054341 | 0.17892 |  | -2.4123
27 |  | TALL4-D-2 | -2.282 | -30.0076 |  | 12.8906 |  | 0.056323 | 0.7749
28 | * | TALL4-C-4 | -2.0051 | -19.9037 |  | 285.0676 |  | 0.4445 | 1.5835 | -21.5348 | -33.9013
29 |  | TALL4-B-1 | -1.7642 | -23.1333 |  | 944.6916 |  | 0.020248 | 0.26149 | -34.5097 | -12.5393
30 |  | TALL4-A-4 | 0.055889 | -23.4255 |  | 14.9413 |  | 0.064086 | 0.076081
31 |  | EV | -10.0202 | -70.8897 |  |  |  | 0.26735 | 2.8418
32 |  | TALL4-D-3 | -1.5108 | -27.0229 |  | 8.5863 |  | 0.22546 | 2.5142
33 |  | TALL4-A-5 | -3.962 | -39.3666 |  |  |  | 0.062431 | 0.75948
34 |  | TALL4-3040-3 | -4.0217 | -30.495 |  |  |  | 0.095424 | 0.41828
35 |  | TALL4-010-2 | -1.349 | -15.2716 |  |  |  | 0.14507 | 1.2206
36 |  | TALL4-2030-3 | -3.9985 | -26.8387 |  |  |  | 0.081602 | 0.72077
37 |  | TALL4-C-2 | -2.9506 | -26.982 |  |  |  | 0.027749 | 0.16667
38 |  | TALL4-B-2 | -2.2338 | -19.6484 |  | 780.7351 |  | 0.016338 | 0.4027 | -27.7305 | -11.0158
39 |  | TALL4-1020-3 | -2.0965 | -19.2208 |  | 7.7006 |  | 0.054811 | 0.25515
40 |  | EV | -10.0232 | -70.4657 |  |  |  | 0.18373 | 2.9051
41 |  | TALL4-010-3 | -2.3462 | -18.0783 |  |  |  | 0.23443 | 3.2471
42 |  | TALL4-B-3 | 0.55597 | -14.8856 | -2.4428 | 1187.1301 |  | 0.014175 | 0.12802 | -46.5615 | -14.5555
43 |  | TALL4-deep-3 | -3.6432 | -30.5597 |  | 5.2161 |  | 0.066935 | 0.79475
44 |  | TALL4-010-1 | -0.5935 | -9.7836 |  |  |  | 0.069883 | 0.99278
45 |  | TALL4-B-5 | -1.8606 | -20.9321 |  |  |  | 0.058324 | 0.76721
46 |  | TALL4-D-4 | -2.2558 | -33.8153 |  | 895.9916 |  | 0.065148 | 0.98502 | -32.6715 | -10.1823
47 |  | TALL4-A-3 | 0.0039579 | -24.0619 |  | 5.9943 |  | 0.099514 | 1.0397
48 |  | TALL4-2030-2 | -2.0788 | -17.1783 |  |  |  | 0.025974 | 0.88854
49 |  | EV | -10.03 | -70.5442 |  |  |  | 0.22021 | 2.9944
50 |  | TALL4-1020-1 | -1.5687 | -17.4434 |  |  |  | 0.23243 | 2.9624
51 |  | EV | -10.018 | -70.4257 |  |  |  | 0.22313 | 3.2237

## Detail
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
1 |  | EV | -7.995 | -65.765 | -10.6554 | -79.1802 | 0.562 | 1317.118 | 2022/10/03 12:00:45
1 |  | EV | -7.635 | -61.473 | -10.2959 | -75.1667 | 0.782 | 1311.236 | 2022/10/03 12:09:42
1 |  | EV | -7.553 | -60.565 | -10.214 | -74.3176 | 0.433 | 1309.752 | 2022/10/03 12:18:39
1 |  | EV | -7.494 | -60.185 | -10.1551 | -73.9623 | 0.496 | 1309.495 | 2022/10/03 12:27:35
1 |  | EV | -7.464 | -59.953 | -10.1251 | -73.7453 | 0.255 | 1307.261 | 2022/10/03 12:36:32
1 |  | EV | -7.417 | -59.925 | -10.0782 | -73.7191 | 0.756 | 1305.221 | 2022/10/03 12:45:29
1 |  | EV | -7.498 | -60.02 | -10.159 | -73.808 | 0.418 | 1301.843 | 2022/10/03 12:54:26
1 |  | EV | -7.429 | -59.898 | -10.0901 | -73.6939 | 0.325 | 1299.231 | 2022/10/03 13:03:22
1 |  | EV | -7.447 | -59.796 | -10.1081 | -73.5985 | 0.454 | 1297.083 | 2022/10/03 13:12:19
1 |  | EV | -7.463 | -59.74 | -10.1241 | -73.5462 | 0.641 | 1295.568 | 2022/10/03 13:21:16
Sample 1 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -10.2005 | -74.4738 |  | 9.8999 |  | 0.17253 | 1.8417
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
2 |  | PZ | 3.939 | 27.899 | 1.2628 | 8.4062 | 0.048 | 1296.994 | 2022/10/03 13:31:04
2 |  | PZ | 4.421 | 36.041 | 1.7441 | 16.0198 | 0.269 | 1295.029 | 2022/10/03 13:40:01
2 |  | PZ | 4.625 | 38.497 | 1.9479 | 18.3165 | 0.559 | 1296.664 | 2022/10/03 13:48:58
2 |  | PZ | 4.658 | 39.122 | 1.9808 | 18.9009 | 0.29 | 1297.916 | 2022/10/03 13:57:55
2 |  | PZ | 4.724 | 39.579 | 2.0467 | 19.3283 | 0.262 | 1296.061 | 2022/10/03 14:06:52
2 |  | PZ | 4.653 | 39.6 | 1.9758 | 19.3479 | 0.704 | 1294.379 | 2022/10/03 14:15:49
2 |  | PZ | 4.781 | 40.205 | 2.1037 | 19.9137 | 0.49 | 1296.587 | 2022/10/03 14:24:46
2 |  | PZ | 4.725 | 40.48 | 2.0477 | 20.1708 | 0.388 | 1296.434 | 2022/10/03 14:33:43
2 |  | PZ | 4.792 | 40.805 | 2.1147 | 20.4747 | 0.106 | 1297.634 | 2022/10/03 14:42:39
2 |  | PZ | 4.753 | 40.427 | 2.0757 | 20.1212 | 0.435 | 1294.01 | 2022/10/03 14:51:36
Sample 2 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.25805 | 3.8943
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
3 |  | UT2 | -12.197 | -92.886 | -14.8518 | -104.5414 | 0.528 | 1296.484 | 2022/10/03 15:01:25
3 |  | UT2 | -13.002 | -105.698 | -15.6557 | -116.522 | 0.427 | 1294.971 | 2022/10/03 15:10:21
3 |  | UT2 | -13.187 | -108.406 | -15.8405 | -119.0543 | 0.251 | 1294.189 | 2022/10/03 15:19:18
3 |  | UT2 | -13.345 | -110.117 | -15.9983 | -120.6543 | 0.792 | 1293.121 | 2022/10/03 15:28:14
3 |  | UT2 | -13.214 | -109.983 | -15.8675 | -120.529 | 0.368 | 1297.242 | 2022/10/03 15:37:11
3 |  | UT2 | -13.363 | -110.818 | -16.0163 | -121.3098 | 0.244 | 1293.196 | 2022/10/03 15:46:08
3 |  | UT2 | -13.338 | -111.128 | -15.9913 | -121.5997 | 0.387 | 1295.33 | 2022/10/03 15:55:04
3 |  | UT2 | -13.4 | -111.395 | -16.0532 | -121.8494 | 0.334 | 1294.452 | 2022/10/03 16:04:01
3 |  | UT2 | -13.416 | -111.414 | -16.0692 | -121.8671 | 0.569 | 1294.923 | 2022/10/03 16:12:58
3 |  | UT2 | -13.403 | -111.634 | -16.0562 | -122.0729 | 0.405 | 1294.003 | 2022/10/03 16:21:55
Sample 3 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.37092 | 5.7273
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
4 |  | EV | -8.005 | -66.361 | -10.6654 | -79.7375 | 0.313 | 1293.335 | 2022/10/03 16:31:44
4 |  | EV | -7.739 | -61.962 | -10.3997 | -75.624 | 0.601 | 1293.796 | 2022/10/03 16:40:41
4 |  | EV | -7.69 | -61.034 | -10.3508 | -74.7562 | 0.636 | 1293.394 | 2022/10/03 16:49:38
4 |  | EV | -7.661 | -60.792 | -10.3218 | -74.5299 | 0.449 | 1294.472 | 2022/10/03 16:58:34
4 |  | EV | -7.628 | -60.472 | -10.2889 | -74.2307 | 0.297 | 1294.288 | 2022/10/03 17:07:31
4 |  | EV | -7.602 | -60.476 | -10.2629 | -74.2344 | 0.448 | 1293.449 | 2022/10/03 17:16:28
4 |  | EV | -7.632 | -60.23 | -10.2929 | -74.0044 | 0.411 | 1293.886 | 2022/10/03 17:25:25
4 |  | EV | -7.616 | -60.175 | -10.2769 | -73.9529 | 0.52 | 1294.647 | 2022/10/03 17:34:22
4 |  | EV | -7.609 | -60.101 | -10.2699 | -73.8837 | 0.56 | 1294.148 | 2022/10/03 17:43:19
4 |  | EV | -7.605 | -60.237 | -10.2659 | -74.0109 | 0.4 | 1293.491 | 2022/10/03 17:52:16
Sample 4 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
4 |  | EV | -10.3395 | -74.8965 |  |  |  | 0.12266 | 1.9029
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
5 |  | Mill Creek Porter Fork Upper-01 | -14.375 | -113.704 | -17.0269 | -124.0086 | 0.625 | 1291.298 | 2022/10/03 18:02:06
5 |  | Mill Creek Porter Fork Upper-01 | -14.684 | -118.569 | -17.3355 | -128.5579 | 0.491 | 1294.377 | 2022/10/03 18:11:03
5 |  | Mill Creek Porter Fork Upper-01 | -14.757 | -119.794 | -17.4084 | -129.7034 | 0.133 | 1295.473 | 2022/10/03 18:20:00
5 |  | Mill Creek Porter Fork Upper-01 | -14.742 | -120.13 | -17.3934 | -130.0176 | 0.332 | 1295.667 | 2022/10/03 18:28:56
Sample 5 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
5 |  | Mill Creek Porter Fork Upper-01 | -17.2911 | -128.0719 |  |  |  | 0.17912 | 2.9735
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
6 |  | Porter Fork-02 | -14.788 | -120.079 | -17.4394 | -129.9699 | 0.152 | 1292.304 | 2022/10/03 18:38:46
6 |  | Porter Fork-02 | -14.843 | -120.316 | -17.4943 | -130.1915 | 0.491 | 1292.407 | 2022/10/03 18:47:43
6 |  | Porter Fork-02 | -14.844 | -120.509 | -17.4953 | -130.372 | 0.433 | 1292.236 | 2022/10/03 18:56:40
6 |  | Porter Fork-02 | -14.812 | -120.725 | -17.4633 | -130.574 | 0.194 | 1292.75 | 2022/10/03 19:05:37
Sample 6 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
6 |  | Porter Fork-02 | -17.4731 | -130.2768 |  |  |  | 0.026961 | 0.27531
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
7 |  | Mill Creek Porter Fork Lower-03 | -14.742 | -120.455 | -17.3934 | -130.3215 | 0.531 | 1296.738 | 2022/10/03 19:15:28
7 |  | Mill Creek Porter Fork Lower-03 | -14.426 | -119.345 | -17.0779 | -129.2835 | 0.628 | 1304.216 | 2022/10/03 19:24:24
7 |  | Mill Creek Porter Fork Lower-03 | -13.684 | -117.046 | -16.3368 | -127.1337 | 0.553 | 1305.382 | 2022/10/03 19:33:21
7 |  | Mill Creek Porter Fork Lower-03 | -14.61 | -120.266 | -17.2616 | -130.1448 | 0.237 | 1294.727 | 2022/10/03 19:42:18
Sample 7 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
7 |  | Mill Creek Porter Fork Lower-03 | -17.0174 | -129.2209 |  | 4.7848 |  | 0.47245 | 1.565
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
8 |  | Mill Creek Upstream of Church Fork-04 | -14.76 | -120.746 | -17.4114 | -130.5936 | 0.114 | 1296.01 | 2022/10/03 19:52:06
8 |  | Mill Creek Upstream of Church Fork-04 | -14.746 | -120.691 | -17.3974 | -130.5422 | 0.416 | 1296.139 | 2022/10/03 20:01:02
8 |  | Mill Creek Upstream of Church Fork-04 | -14.845 | -121.06 | -17.4963 | -130.8872 | 0.381 | 1292.588 | 2022/10/03 20:09:59
8 |  | Mill Creek Upstream of Church Fork-04 | -14.895 | -121.381 | -17.5462 | -131.1874 | 0.155 | 1295.392 | 2022/10/03 20:18:56
Sample 8 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
8 |  | Mill Creek Upstream of Church Fork-04 | -17.4628 | -130.8026 |  |  |  | 0.070798 | 0.31887
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
9 |  | Church Fork-05 | -13.757 | -113.459 | -16.4097 | -123.7795 | 0.311 | 1292.015 | 2022/10/03 20:28:44
9 |  | Church Fork-05 | -13.65 | -112.405 | -16.3029 | -122.7938 | 0.654 | 1294.538 | 2022/10/03 20:37:41
9 |  | Church Fork-05 | -13.627 | -112.295 | -16.2799 | -122.691 | -0.016 | 1295.05 | 2022/10/03 20:46:38
9 |  | Church Fork-05 | -13.614 | -112.082 | -16.2669 | -122.4918 | 0.487 | 1294.181 | 2022/10/03 20:55:35
Sample 9 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
9 |  | Church Fork-05 | -16.3149 | -122.939 |  |  |  | 0.065059 | 0.61399
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
10 |  | Mill Creek Down Stream of Church Fork-06 | -14.732 | -119.584 | -17.3835 | -129.507 | 0.474 | 1291.582 | 2022/10/03 21:05:23
10 |  | Mill Creek Down Stream of Church Fork-06 | -14.73 | -120.273 | -17.3815 | -130.1513 | 0.785 | 1291.968 | 2022/10/03 21:14:19
10 |  | Mill Creek Down Stream of Church Fork-06 | -14.779 | -120.357 | -17.4304 | -130.2299 | 0.333 | 1294.827 | 2022/10/03 21:23:16
10 |  | Mill Creek Down Stream of Church Fork-06 | -14.758 | -120.685 | -17.4094 | -130.5366 | 0.435 | 1292.898 | 2022/10/03 21:32:13
Sample 10 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
10 |  | Mill Creek Down Stream of Church Fork-06 | -17.4012 | -130.1062 |  |  |  | 0.023301 | 0.46268
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
11 |  | Gate House-07 | -14.658 | -120.02 | -17.3095 | -129.9147 | 0.225 | 1294.771 | 2022/10/03 21:42:01
11 |  | Gate House-07 | -14.725 | -120.195 | -17.3765 | -130.0784 | 0.447 | 1292.782 | 2022/10/03 21:50:58
11 |  | Gate House-07 | -14.778 | -120.31 | -17.4294 | -130.1859 | 0.421 | 1292.302 | 2022/10/03 21:59:55
11 |  | Gate House-07 | -14.671 | -119.867 | -17.3225 | -129.7717 | 0.605 | 1296.314 | 2022/10/03 22:08:52
Sample 11 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
11 |  | Gate House-07 | -17.3595 | -129.9877 |  |  |  | 0.054948 | 0.19476
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
12 |  | TALL4-D-5 | 3.951 | -17.514 | 1.2748 | -34.0601 | -3.445 | 1660.834 | 2022/10/03 22:18:41
12 |  | TALL4-D-5 | 4.605 | -8.337 | 1.9279 | -25.4786 | -3.557 | 1662.643 | 2022/10/03 22:27:38
12 |  | TALL4-D-5 | 4.844 | -5.756 | 2.1666 | -23.065 | -3.513 | 1673.854 | 2022/10/03 22:36:35
12 |  | TALL4-D-5 | 4.979 | -4.701 | 2.3014 | -22.0785 | -3.422 | 1670.707 | 2022/10/03 22:45:32
Sample 12 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
12 |  | TALL4-D-5 | 1.9177 | -26.1706 | -3.4843 | 371.5286 | 4.003 | 0.45618 | 5.8283 | -22.207 | -0.6123
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
13 |  | EV | -7.193 | -55.529 | -9.8544 | -69.6084 | 0.493 | 1300.898 | 2022/10/03 22:55:21
13 |  | EV | -7.569 | -59.379 | -10.23 | -73.2086 | 0.257 | 1294.741 | 2022/10/03 23:04:17
13 |  | EV | -7.543 | -59.459 | -10.204 | -73.2834 | 0.455 | 1297.629 | 2022/10/03 23:13:14
13 |  | EV | -7.597 | -60.051 | -10.2579 | -73.837 | 0.343 | 1294.314 | 2022/10/03 23:22:11
Sample 13 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
13 |  | EV | -10.1366 | -72.4843 |  |  |  | 0.18962 | 2.0721
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
14 |  | TALL4-A-1 | 2.562 | -13.722 | -0.11239 | -30.5142 | -2.141 | 1365.926 | 2022/10/03 23:32:01
14 |  | TALL4-A-1 | 2.971 | -9.62 | 0.29607 | -26.6783 | -2.313 | 1368.14 | 2022/10/03 23:40:58
14 |  | TALL4-A-1 | 3.053 | -8.46 | 0.37796 | -25.5936 | -2.285 | 1369.149 | 2022/10/03 23:49:55
14 |  | TALL4-A-1 | 3.116 | -7.868 | 0.44088 | -25.04 | -2.371 | 1367.63 | 2022/10/03 23:58:51
Sample 14 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
14 |  | TALL4-A-1 | 0.25063 | -26.9565 | -2.2775 | 72.2303 |  | 0.2495 | 2.6387 | -8.0758 | -4.4355
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
15 |  | TALL4-D-1 | 2.107 | -2.966 | -0.56678 | -20.4561 | -1.63 | 1527.561 | 2022/10/04 00:08:38
15 |  | TALL4-D-1 | 2.147 | -2.109 | -0.52684 | -19.6547 | -1.548 | 1532.62 | 2022/10/04 00:17:35
15 |  | TALL4-D-1 | 2.147 | -1.978 | -0.52684 | -19.5322 | -1.546 | 1527.242 | 2022/10/04 00:26:31
15 |  | TALL4-D-1 | 2.111 | -1.688 | -0.56279 | -19.261 | -1.261 | 1522.532 | 2022/10/04 00:35:28
Sample 15 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
15 |  | TALL4-D-1 | -0.54581 | -19.726 | -1.4963 | 232.0078 |  | 0.022 | 0.54942 | -11.5813 | -5.5568
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
16 |  | TALL4-B-4 | 1.122 | 0.706 | -1.5505 | -17.0224 | -0.437 | 1784.819 | 2022/10/04 00:45:15
16 |  | TALL4-B-4 | 1.156 | 1.296 | -1.5165 | -16.4706 | -0.638 | 1772.166 | 2022/10/04 00:54:11
16 |  | TALL4-B-4 | 1.056 | 0.911 | -1.6164 | -16.8307 | -0.364 | 1775.586 | 2022/10/04 01:03:08
16 |  | TALL4-B-4 | 1.122 | 1.286 | -1.5505 | -16.48 | -0.412 | 1774.927 | 2022/10/04 01:12:05
Sample 16 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
16 |  | TALL4-B-4 | -1.5585 | -16.7009 |  | 481.3936 |  | 0.041857 | 0.2909 | -16.8858 | -7.1528
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
17 |  | TALL4-3040-2 | -0.791 | -4.835 | -3.4609 | -22.2038 | 0.296 | 1304.044 | 2022/10/04 01:21:53
17 |  | TALL4-3040-2 | -0.918 | -5.616 | -3.5878 | -22.9341 | 0.257 | 1297.529 | 2022/10/04 01:30:49
17 |  | TALL4-3040-2 | -0.792 | -5.324 | -3.4619 | -22.6611 | 0.226 | 1298.779 | 2022/10/04 01:39:45
17 |  | TALL4-3040-2 | -0.826 | -5.197 | -3.4959 | -22.5423 | -0.043 | 1298.311 | 2022/10/04 01:48:42
Sample 17 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
17 |  | TALL4-3040-2 | -3.5016 | -22.5853 |  |  |  | 0.059757 | 0.32366
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
18 |  | TALL4-deep-2 | -1.414 | -13.56 | -4.0831 | -30.3627 | 0.476 | 1294.005 | 2022/10/04 01:58:30
18 |  | TALL4-deep-2 | -1.398 | -13.956 | -4.0671 | -30.733 | 0.208 | 1296.394 | 2022/10/04 02:07:26
18 |  | TALL4-deep-2 | -1.445 | -14.284 | -4.1141 | -31.0397 | 0.225 | 1295.494 | 2022/10/04 02:16:23
18 |  | TALL4-deep-2 | -1.454 | -14.411 | -4.1231 | -31.1584 | 0.38 | 1294.28 | 2022/10/04 02:25:20
Sample 18 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
18 |  | TALL4-deep-2 | -4.0968 | -30.8234 |  |  |  | 0.026209 | 0.38034
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
19 |  | TALL4-1020-2 | 0.043 | 0.601 | -2.628 | -17.1205 | 0.588 | 1295.658 | 2022/10/04 02:35:08
19 |  | TALL4-1020-2 | 0.264 | 1.905 | -2.4073 | -15.9012 | 0.394 | 1299.06 | 2022/10/04 02:44:04
19 |  | TALL4-1020-2 | 0.193 | 2.15 | -2.4782 | -15.672 | 0.553 | 1297.827 | 2022/10/04 02:53:00
19 |  | TALL4-1020-2 | 0.167 | 2.324 | -2.5042 | -15.5093 | 0.394 | 1298.171 | 2022/10/04 03:01:57
Sample 19 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
19 |  | TALL4-1020-2 | -2.5045 | -16.0508 |  |  |  | 0.092124 | 0.78179
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
20 |  | TALL4-2030-1 | -0.396 | -5.383 | -3.0665 | -22.7163 | 0.146 | 1294.299 | 2022/10/04 03:11:46
20 |  | TALL4-2030-1 | -0.43 | -5.521 | -3.1004 | -22.8453 | 0.302 | 1297.773 | 2022/10/04 03:20:42
20 |  | TALL4-2030-1 | -0.425 | -5.898 | -3.0954 | -23.1978 | 0.404 | 1295.422 | 2022/10/04 03:29:39
20 |  | TALL4-2030-1 | -0.409 | -6.172 | -3.0794 | -23.4541 | 0.292 | 1293.848 | 2022/10/04 03:38:36
Sample 20 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
20 |  | TALL4-2030-1 | -3.0854 | -23.0534 |  |  |  | 0.015513 | 0.35914
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
21 |  | TALL4-C-1 | 0.926 | -1.17 | -1.7462 | -18.7766 | -0.061 | 1342.325 | 2022/10/04 03:48:25
21 |  | TALL4-C-1 | 0.933 | -0.831 | -1.7392 | -18.4596 | -0.165 | 1344.207 | 2022/10/04 03:57:21
21 |  | TALL4-C-1 | 0.923 | -0.947 | -1.7492 | -18.5681 | 0.105 | 1341.457 | 2022/10/04 04:06:18
21 |  | TALL4-C-1 | 0.944 | -0.805 | -1.7282 | -18.4353 | -0.255 | 1342.875 | 2022/10/04 04:15:14
Sample 21 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
21 |  | TALL4-C-1 | -1.7407 | -18.5599 |  | 47.2351 |  | 0.0093274 | 0.16638 |  | 1.8595
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
22 |  | EV | -7.047 | -51.913 | -9.7086 | -66.227 | 0.432 | 1293.085 | 2022/10/04 04:24:59
22 |  | EV | -7.333 | -57.301 | -9.9943 | -71.2654 | 0.477 | 1293.921 | 2022/10/04 04:33:56
22 |  | EV | -7.422 | -57.885 | -10.0831 | -71.8115 | 0.64 | 1293.696 | 2022/10/04 04:42:53
22 |  | EV | -7.389 | -58.618 | -10.0502 | -72.497 | 0.495 | 1293.604 | 2022/10/04 04:51:49
Sample 22 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
22 |  | EV | -9.9591 | -70.4502 |  |  |  | 0.17116 | 3.0587
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
23 |  | TALL4-3040-1 | -0.502 | -15.701 | -3.1723 | -32.3647 | 0.36 | 1296.381 | 2022/10/04 05:01:35
23 |  | TALL4-3040-1 | -0.179 | -12.088 | -2.8498 | -28.9862 | 0.35 | 1297.695 | 2022/10/04 05:10:32
23 |  | TALL4-3040-1 | -0.075 | -11.057 | -2.7459 | -28.0221 | 0.327 | 1298.186 | 2022/10/04 05:19:28
23 |  | TALL4-3040-1 | -0.078 | -10.671 | -2.7489 | -27.6611 | 0.253 | 1298.308 | 2022/10/04 05:28:25
Sample 23 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
23 |  | TALL4-3040-1 | -2.8792 | -29.2585 |  |  |  | 0.20155 | 2.2939
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
24 |  | TALL4-C-5 | 0.761 | -10.435 | -1.911 | -27.4404 | -0.304 | 1382.702 | 2022/10/04 05:38:12
24 |  | TALL4-C-5 | 0.735 | -10.418 | -1.937 | -27.4245 | -0.195 | 1379.838 | 2022/10/04 05:47:09
24 |  | TALL4-C-5 | 0.77 | -10.697 | -1.902 | -27.6854 | -0.202 | 1381.146 | 2022/10/04 05:56:05
24 |  | TALL4-C-5 | 0.785 | -10.893 | -1.887 | -27.8687 | -0.155 | 1377.415 | 2022/10/04 06:05:02
Sample 24 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
24 |  | TALL4-C-5 | -1.9093 | -27.6048 |  | 84.7943 |  | 0.020982 | 0.22741 |  | -3.586
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
25 |  | TALL4-C-3 | 0.619 | -8.177 | -2.0528 | -25.329 | 0.258 | 1300.405 | 2022/10/04 06:14:49
25 |  | TALL4-C-3 | 0.565 | -8.442 | -2.1067 | -25.5768 | 0.109 | 1291.598 | 2022/10/04 06:23:46
25 |  | TALL4-C-3 | 0.586 | -7.97 | -2.0858 | -25.1354 | 0.534 | 1299.511 | 2022/10/04 06:32:42
25 |  | TALL4-C-3 | 0.647 | -8.159 | -2.0248 | -25.3121 | 0.037 | 1299.909 | 2022/10/04 06:41:39
Sample 25 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
25 |  | TALL4-C-3 | -2.0675 | -25.3383 |  |  |  | 0.036142 | 0.19408
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
26 |  | TALL4-A-2 | 1.774 | -2.4 | -0.89934 | -19.9268 | -0.793 | 1356.608 | 2022/10/04 06:51:27
26 |  | TALL4-A-2 | 1.683 | -2.303 | -0.99022 | -19.8361 | -0.658 | 1352.85 | 2022/10/04 07:00:24
26 |  | TALL4-A-2 | 1.691 | -2.151 | -0.98223 | -19.694 | -0.529 | 1352.272 | 2022/10/04 07:09:21
26 |  | TALL4-A-2 | 1.645 | -2.578 | -1.0282 | -20.0933 | -0.752 | 1346.437 | 2022/10/04 07:18:17
Sample 26 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
26 |  | TALL4-A-2 | -0.97499 | -19.8875 |  | 56.5608 |  | 0.054341 | 0.17892 |  | -2.4123
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
27 |  | TALL4-D-2 | 0.458 | -12.03 | -2.2136 | -28.9319 | -0.559 | 1308.872 | 2022/10/04 07:28:06
27 |  | TALL4-D-2 | 0.357 | -13.412 | -2.3145 | -30.2243 | -0.124 | 1308.951 | 2022/10/04 07:37:02
27 |  | TALL4-D-2 | 0.411 | -13.602 | -2.2605 | -30.4019 | -0.597 | 1309.144 | 2022/10/04 07:45:59
27 |  | TALL4-D-2 | 0.332 | -13.677 | -2.3394 | -30.4721 | -0.471 | 1306.519 | 2022/10/04 07:54:55
Sample 27 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
27 |  | TALL4-D-2 | -2.282 | -30.0076 |  | 12.8906 |  | 0.056323 | 0.7749
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
28 |  | TALL4-C-4 |  |  | -2.671 | -17.6825 |  |  | 2022/10/04 08:04:43
28 |  | TALL4-C-4 | 0.889 | -3.167 | -1.7832 | -20.644 | -0.508 | 2107.398 | 2022/10/04 08:13:39
28 |  | TALL4-C-4 | 0.889 | -3.167 | -1.7832 | -20.644 | -0.508 | 2107.398 | 2022/10/04 08:22:36
28 |  | TALL4-C-4 | 0.889 | -3.167 | -1.7832 | -20.644 | -0.508 | 2107.398 | 2022/10/04 08:31:34
Sample 28 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
28 | * | TALL4-C-4 | -2.0051 | -19.9037 |  | 285.0676 |  | 0.4445 | 1.5835 | -21.5348 | -33.9013
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
29 |  | TALL4-B-1 | 0.926 | -5.473 | -1.7462 | -22.8004 | -1.01 | 2238.908 | 2022/10/04 08:41:19
29 |  | TALL4-B-1 | 0.897 | -5.966 | -1.7752 | -23.2614 | -0.775 | 2242.039 | 2022/10/04 08:50:16
29 |  | TALL4-B-1 | 0.885 | -5.805 | -1.7872 | -23.1109 | -0.873 | 2242.677 | 2022/10/04 08:59:13
29 |  | TALL4-B-1 | 0.924 | -6.072 | -1.7482 | -23.3605 | -0.996 | 2237.066 | 2022/10/04 09:08:10
Sample 29 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
29 |  | TALL4-B-1 | -1.7642 | -23.1333 |  | 944.6916 |  | 0.020248 | 0.26149 | -34.5097 | -12.5393
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
30 |  | TALL4-A-4 | 2.647 | -6.037 | -0.0275 | -23.3278 | -0.79 | 1318.599 | 2022/10/04 09:17:55
30 |  | TALL4-A-4 | 2.732 | -6.208 | 0.057387 | -23.4877 | -0.715 | 1309.669 | 2022/10/04 09:26:52
30 |  | TALL4-A-4 | 2.74 | -6.135 | 0.065376 | -23.4195 | -0.728 | 1307.318 | 2022/10/04 09:35:49
30 |  | TALL4-A-4 | 2.803 | -6.186 | 0.12829 | -23.4671 | -0.994 | 1306.103 | 2022/10/04 09:44:45
Sample 30 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
30 |  | TALL4-A-4 | 0.055889 | -23.4255 |  | 14.9413 |  | 0.064086 | 0.076081
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
31 |  | EV | -6.96 | -52.752 | -9.6218 | -67.0116 | 0.609 | 1293.515 | 2022/10/04 09:54:31
31 |  | EV | -7.51 | -57.563 | -10.171 | -71.5104 | 0.783 | 1293.574 | 2022/10/04 10:03:28
31 |  | EV | -7.454 | -58.131 | -10.1151 | -72.0416 | 0.274 | 1294.568 | 2022/10/04 10:12:25
31 |  | EV | -7.512 | -59.151 | -10.173 | -72.9954 | 0.493 | 1294.27 | 2022/10/04 10:21:22
Sample 31 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
31 |  | EV | -10.0202 | -70.8897 |  |  |  | 0.26735 | 2.8418
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
32 |  | TALL4-D-3 | 0.837 | -13.665 | -1.8351 | -30.4609 | 0.014 | 1306.069 | 2022/10/04 10:31:08
32 |  | TALL4-D-3 | 1.199 | -9.517 | -1.4736 | -26.582 | -0.278 | 1303.894 | 2022/10/04 10:40:05
32 |  | TALL4-D-3 | 1.259 | -8.597 | -1.4137 | -25.7217 | 0.205 | 1303.112 | 2022/10/04 10:49:02
32 |  | TALL4-D-3 | 1.352 | -8.175 | -1.3208 | -25.3271 | 0.045 | 1303.194 | 2022/10/04 10:57:59
Sample 32 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
32 |  | TALL4-D-3 | -1.5108 | -27.0229 |  | 8.5863 |  | 0.22546 | 2.5142
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
33 |  | TALL4-A-5 | -1.201 | -22.053 | -3.8704 | -38.3046 | 0.302 | 1294.661 | 2022/10/04 11:07:45
33 |  | TALL4-A-5 | -1.335 | -23.571 | -4.0042 | -39.7241 | 0.275 | 1292.463 | 2022/10/04 11:16:42
33 |  | TALL4-A-5 | -1.306 | -23.493 | -3.9753 | -39.6511 | 0.231 | 1295.727 | 2022/10/04 11:25:38
33 |  | TALL4-A-5 | -1.329 | -23.638 | -3.9982 | -39.7867 | 0.165 | 1294.692 | 2022/10/04 11:34:35
Sample 33 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
33 |  | TALL4-A-5 | -3.962 | -39.3666 |  |  |  | 0.062431 | 0.75948
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
34 |  | TALL4-3040-3 | -1.224 | -13.244 | -3.8934 | -30.0672 | 0.396 | 1307.017 | 2022/10/04 11:44:21
34 |  | TALL4-3040-3 | -1.379 | -13.724 | -4.0482 | -30.516 | 0.669 | 1292.584 | 2022/10/04 11:53:17
34 |  | TALL4-3040-3 | -1.453 | -14.251 | -4.1221 | -31.0088 | 0.304 | 1288.221 | 2022/10/04 12:02:14
34 |  | TALL4-3040-3 | -1.354 | -13.587 | -4.0232 | -30.3879 | 0.512 | 1293.031 | 2022/10/04 12:11:10
Sample 34 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
34 |  | TALL4-3040-3 | -4.0217 | -30.495 |  |  |  | 0.095424 | 0.41828
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
35 |  | TALL4-010-2 | 1.116 | 0.786 | -1.5565 | -16.9475 | 0.47 | 1285.172 | 2022/10/04 12:20:58
35 |  | TALL4-010-2 | 1.442 | 3.445 | -1.2309 | -14.4611 | 0.223 | 1293.334 | 2022/10/04 12:29:55
35 |  | TALL4-010-2 | 1.337 | 2.844 | -1.3358 | -15.0231 | 0.247 | 1286.777 | 2022/10/04 12:38:52
35 |  | TALL4-010-2 | 1.4 | 3.238 | -1.2728 | -14.6546 | 0.67 | 1288.47 | 2022/10/04 12:47:49
Sample 35 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
35 |  | TALL4-010-2 | -1.349 | -15.2716 |  |  |  | 0.14507 | 1.2206
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
36 |  | TALL4-2030-3 | -1.225 | -8.736 | -3.8944 | -25.8517 | 0.464 | 1285.666 | 2022/10/04 12:57:32
36 |  | TALL4-2030-3 | -1.305 | -9.957 | -3.9743 | -26.9935 | 0.264 | 1288.076 | 2022/10/04 13:06:29
36 |  | TALL4-2030-3 | -1.382 | -10.134 | -4.0512 | -27.159 | 0.45 | 1287.612 | 2022/10/04 13:15:25
36 |  | TALL4-2030-3 | -1.405 | -10.339 | -4.0741 | -27.3507 | 0.556 | 1287.343 | 2022/10/04 13:24:22
Sample 36 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
36 |  | TALL4-2030-3 | -3.9985 | -26.8387 |  |  |  | 0.081602 | 0.72077
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
37 |  | TALL4-C-2 | -0.321 | -10.021 | -2.9916 | -27.0533 | 0.222 | 1294.823 | 2022/10/04 13:34:07
37 |  | TALL4-C-2 | -0.262 | -9.715 | -2.9326 | -26.7672 | 0.589 | 1295.014 | 2022/10/04 13:43:04
37 |  | TALL4-C-2 | -0.264 | -9.941 | -2.9346 | -26.9785 | 0.148 | 1294.764 | 2022/10/04 13:52:01
37 |  | TALL4-C-2 | -0.273 | -10.102 | -2.9436 | -27.129 | 0.198 | 1294.487 | 2022/10/04 14:00:58
Sample 37 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
37 |  | TALL4-C-2 | -2.9506 | -26.982 |  |  |  | 0.027749 | 0.16667
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
38 |  | TALL4-B-2 | 0.426 | -2.621 | -2.2456 | -20.1335 | -0.192 | 2068.841 | 2022/10/04 14:10:42
38 |  | TALL4-B-2 | 0.46 | -2.221 | -2.2116 | -19.7594 | -0.471 | 2075.904 | 2022/10/04 14:19:40
38 |  | TALL4-B-2 | 0.425 | -1.777 | -2.2466 | -19.3442 | -0.48 | 2078.787 | 2022/10/04 14:28:36
38 |  | TALL4-B-2 | 0.44 | -1.79 | -2.2316 | -19.3564 | -0.263 | 2081.332 | 2022/10/04 14:37:33
Sample 38 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
38 |  | TALL4-B-2 | -2.2338 | -19.6484 |  | 780.7351 |  | 0.016338 | 0.4027 | -27.7305 | -11.0158
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
39 |  | TALL4-1020-3 | 0.53 | -1.939 | -2.1417 | -19.4957 | 0.268 | 1307.607 | 2022/10/04 14:47:19
39 |  | TALL4-1020-3 | 0.533 | -1.665 | -2.1387 | -19.2395 | 0.319 | 1302.118 | 2022/10/04 14:56:16
39 |  | TALL4-1020-3 | 0.645 | -1.316 | -2.0268 | -18.9131 | 0.392 | 1301.419 | 2022/10/04 15:05:13
39 |  | TALL4-1020-3 | 0.593 | -1.66 | -2.0788 | -19.2348 | 0.168 | 1301.582 | 2022/10/04 15:14:09
Sample 39 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
39 |  | TALL4-1020-3 | -2.0965 | -19.2208 |  | 7.7006 |  | 0.054811 | 0.25515
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
40 |  | EV | -7.095 | -52.246 | -9.7566 | -66.5384 | 0.292 | 1292.837 | 2022/10/04 15:23:55
40 |  | EV | -7.389 | -56.759 | -10.0502 | -70.7586 | 0.318 | 1293.918 | 2022/10/04 15:32:52
40 |  | EV | -7.467 | -58.265 | -10.1281 | -72.1669 | 0.291 | 1293.161 | 2022/10/04 15:41:48
40 |  | EV | -7.497 | -58.513 | -10.158 | -72.3988 | 0.287 | 1293.886 | 2022/10/04 15:50:45
Sample 40 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
40 |  | EV | -10.0232 | -70.4657 |  |  |  | 0.18373 | 2.9051
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
41 |  | TALL4-010-3 | -0.02 | -5.214 | -2.691 | -22.5582 | 0.522 | 1294.3 | 2022/10/04 16:00:31
41 |  | TALL4-010-3 | 0.387 | 0.345 | -2.2845 | -17.3599 | 0.455 | 1294.426 | 2022/10/04 16:09:28
41 |  | TALL4-010-3 | 0.496 | 1.58 | -2.1756 | -16.2051 | 0.444 | 1295.442 | 2022/10/04 16:18:25
41 |  | TALL4-010-3 | 0.438 | 1.596 | -2.2336 | -16.1901 | 0.508 | 1294.543 | 2022/10/04 16:27:21
Sample 41 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
41 |  | TALL4-010-3 | -2.3462 | -18.0783 |  |  |  | 0.23443 | 3.2471
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
42 |  | TALL4-B-3 | 3.227 | 3.019 | 0.55173 | -14.8594 | -2.041 | 2487.185 | 2022/10/04 16:37:07
42 |  | TALL4-B-3 | 3.216 | 2.803 | 0.54075 | -15.0614 | -2.194 | 2470.136 | 2022/10/04 16:46:03
42 |  | TALL4-B-3 | 3.25 | 3.082 | 0.5747 | -14.8005 | -2.688 | 2485.549 | 2022/10/04 16:55:00
42 |  | TALL4-B-3 | 3.232 | 3.06 | 0.55672 | -14.8211 | -2.848 | 2487.574 | 2022/10/04 17:03:56
Sample 42 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
42 |  | TALL4-B-3 | 0.55597 | -14.8856 | -2.4428 | 1187.1301 |  | 0.014175 | 0.12802 | -46.5615 | -14.5555
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
43 |  | TALL4-deep-3 | -0.908 | -12.754 | -3.5778 | -29.609 | 0.231 | 1310.491 | 2022/10/04 17:13:42
43 |  | TALL4-deep-3 | -0.931 | -13.562 | -3.6008 | -30.3645 | 0.24 | 1302.164 | 2022/10/04 17:22:39
43 |  | TALL4-deep-3 | -1 | -14.201 | -3.6697 | -30.9621 | 0.104 | 1297.768 | 2022/10/04 17:31:35
43 |  | TALL4-deep-3 | -1.055 | -14.566 | -3.7246 | -31.3034 | 0.471 | 1292.365 | 2022/10/04 17:40:32
Sample 43 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
43 |  | TALL4-deep-3 | -3.6432 | -30.5597 |  | 5.2161 |  | 0.066935 | 0.79475
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
44 |  | TALL4-010-1 | 1.981 | 6.986 | -0.69262 | -11.1498 | 0.362 | 1295.807 | 2022/10/04 17:50:18
44 |  | TALL4-010-1 | 2.1 | 8.691 | -0.57377 | -9.5555 | 0.277 | 1294.752 | 2022/10/04 17:59:15
44 |  | TALL4-010-1 | 2.095 | 8.95 | -0.57877 | -9.3133 | 0.269 | 1295.773 | 2022/10/04 18:08:11
44 |  | TALL4-010-1 | 2.145 | 9.161 | -0.52883 | -9.116 | 0.328 | 1293.685 | 2022/10/04 18:17:08
Sample 44 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
44 |  | TALL4-010-1 | -0.5935 | -9.7836 |  |  |  | 0.069883 | 0.99278
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
45 |  | TALL4-B-5 | 0.892 | -2.363 | -1.7802 | -19.8922 | 0.007 | 1299.48 | 2022/10/04 18:26:54
45 |  | TALL4-B-5 | 0.795 | -3.728 | -1.877 | -21.1686 | -0.013 | 1296.153 | 2022/10/04 18:35:51
45 |  | TALL4-B-5 | 0.806 | -3.685 | -1.8661 | -21.1284 | 0.152 | 1298.822 | 2022/10/04 18:44:48
45 |  | TALL4-B-5 | 0.753 | -4.124 | -1.919 | -21.5389 | 0.124 | 1295.897 | 2022/10/04 18:53:45
Sample 45 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
45 |  | TALL4-B-5 | -1.8606 | -20.9321 |  |  |  | 0.058324 | 0.76721
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
46 |  | TALL4-D-4 | 0.466 | -15.791 | -2.2056 | -32.4489 | -1.031 | 2189.929 | 2022/10/04 19:03:31
46 |  | TALL4-D-4 | 0.478 | -17.534 | -2.1936 | -34.0788 | -0.956 | 2192.944 | 2022/10/04 19:12:28
46 |  | TALL4-D-4 | 0.358 | -17.825 | -2.3135 | -34.3509 | -0.716 | 2188.912 | 2022/10/04 19:21:24
46 |  | TALL4-D-4 | 0.361 | -17.859 | -2.3105 | -34.3827 | -1.025 | 2194.105 | 2022/10/04 19:30:21
Sample 46 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
46 |  | TALL4-D-4 | -2.2558 | -33.8153 |  | 895.9916 |  | 0.065148 | 0.98502 | -32.6715 | -10.1823
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
47 |  | TALL4-A-3 | 2.54 | -8.366 | -0.13436 | -25.5057 | -1.059 | 1301.747 | 2022/10/04 19:40:08
47 |  | TALL4-A-3 | 2.695 | -6.372 | 0.020436 | -23.6411 | -0.91 | 1303.79 | 2022/10/04 19:49:05
47 |  | TALL4-A-3 | 2.702 | -6.445 | 0.027427 | -23.7093 | -0.838 | 1299.291 | 2022/10/04 19:58:01
47 |  | TALL4-A-3 | 2.777 | -6.105 | 0.10233 | -23.3914 | -0.728 | 1301.073 | 2022/10/04 20:06:58
Sample 47 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
47 |  | TALL4-A-3 | 0.0039579 | -24.0619 |  | 5.9943 |  | 0.099514 | 1.0397
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
48 |  | TALL4-2030-2 | 0.581 | -0.75 | -2.0908 | -18.3839 | 0.306 | 1289.604 | 2022/10/04 20:16:45
48 |  | TALL4-2030-2 | 0.573 | 0.853 | -2.0987 | -16.8849 | 0.372 | 1294.886 | 2022/10/04 20:25:42
48 |  | TALL4-2030-2 | 0.587 | 0.77 | -2.0848 | -16.9625 | 0.496 | 1295.765 | 2022/10/04 20:34:38
48 |  | TALL4-2030-2 | 0.631 | 1.284 | -2.0408 | -16.4819 | 0.516 | 1294.267 | 2022/10/04 20:43:34
Sample 48 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
48 |  | TALL4-2030-2 | -2.0788 | -17.1783 |  |  |  | 0.025974 | 0.88854
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
49 |  | EV | -7.047 | -52.158 | -9.7086 | -66.4561 | 0.491 | 1291.151 | 2022/10/04 20:53:21
49 |  | EV | -7.414 | -57.07 | -10.0752 | -71.0494 | 0.443 | 1293.374 | 2022/10/04 21:02:18
49 |  | EV | -7.478 | -58.167 | -10.1391 | -72.0752 | 0.568 | 1293.156 | 2022/10/04 21:11:15
49 |  | EV | -7.536 | -58.724 | -10.197 | -72.5961 | 0.299 | 1294.443 | 2022/10/04 21:20:11
Sample 49 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
49 |  | EV | -10.03 | -70.5442 |  |  |  | 0.22021 | 2.9944
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
50 |  | TALL4-1020-1 | 0.771 | -4.059 | -1.901 | -21.4782 | 0.373 | 1291.316 | 2022/10/04 21:29:56
50 |  | TALL4-1020-1 | 1.125 | 0.738 | -1.5475 | -16.9924 | 0.242 | 1291.182 | 2022/10/04 21:38:53
50 |  | TALL4-1020-1 | 1.225 | 1.914 | -1.4476 | -15.8927 | 0.334 | 1290.065 | 2022/10/04 21:47:49
50 |  | TALL4-1020-1 | 1.294 | 2.43 | -1.3787 | -15.4102 | 0.638 | 1289.368 | 2022/10/04 21:56:46
Sample 50 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
50 |  | TALL4-1020-1 | -1.5687 | -17.4434 |  |  |  | 0.23243 | 2.9624
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
51 |  | EV | -7.03 | -51.715 | -9.6917 | -66.0419 | 0.446 | 1293.216 | 2022/10/04 22:06:32
51 |  | EV | -7.398 | -56.879 | -10.0592 | -70.8708 | 0.219 | 1293.955 | 2022/10/04 22:15:29
51 |  | EV | -7.508 | -58.282 | -10.169 | -72.1828 | 0.275 | 1293.737 | 2022/10/04 22:24:25
51 |  | EV | -7.491 | -58.736 | -10.1521 | -72.6073 | 0.407 | 1293.734 | 2022/10/04 22:33:22
Sample 51 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
51 |  | EV | -10.018 | -70.4257 |  |  |  | 0.22313 | 3.2237

## Source
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Line | Analysis | Time Code | Port | Inj Nr | d(18_16)Mean | d(D_H)Mean | H2O_Mean | Ignore | Good | Identifier 1 | Identifier 2 | Gas Configuration | Timestamp Mean | d(18_16)_SD | d(D_H)_SD | H2O_SD | d(18_16)_Sl | d(D_H)_Sl | H2O_Sl | baseline_shift | slope_shift | residuals | baseline_curvature | interval | ch4_ppm | h16od_adjust | h16od_shift | n2_flag | DAS Temp | Tray | Sample | Job | Method | Error Code
1 | P-1292 | 2022/10/03 12:00:45 | 2-01 | 1 | -7.995 | -65.765 | 19722 | -1 | 1 | EV | standard | H2O | 1664798752 | 0.316 | 0.735 | 98.76 | 0.001 | 0 | 2.336 | 1317.118 | 2.186 | 3.172 | 120.929 | 0.87 | 0.562 | 0 | 0 | 1 | 43.188 | 1 | 1 | 1 | GJB_Rinse | 0
2 | P-1292 | 2022/10/03 12:09:42 | 2-01 | 2 | -7.635 | -61.473 | 18749 | -1 | 1 | EV | standard | H2O | 1664799289 | 0.304 | 0.72 | 90.119 | 0 | 0 | 2.131 | 1311.236 | 1.864 | 3.047 | 116.402 | 0.865 | 0.782 | 0 | 0 | 1 | 42.438 | 1 | 1 | 1 | GJB_Rinse | 0
3 | P-1292 | 2022/10/03 12:18:39 | 2-01 | 3 | -7.553 | -60.565 | 18640 | -1 | 1 | EV | standard | H2O | 1664799825 | 0.316 | 0.799 | 86.01 | 0.001 | 0 | 2.031 | 1309.752 | 1.448 | 3.013 | 115.564 | 0.866 | 0.433 | 0 | 0 | 1 | 42.188 | 1 | 1 | 1 | GJB_Rinse | 0
4 | P-1292 | 2022/10/03 12:27:35 | 2-01 | 4 | -7.494 | -60.185 | 18757 | 0 | 1 | EV | standard | H2O | 1664800362 | 0.307 | 0.712 | 81.076 | 0.001 | -0.001 | 1.923 | 1309.495 | 1.562 | 2.993 | 115.301 | 0.866 | 0.496 | 0 | 0 | 1 | 42.188 | 1 | 1 | 1 | GJB_Rinse | 0
5 | P-1292 | 2022/10/03 12:36:32 | 2-01 | 5 | -7.464 | -59.953 | 18589 | 0 | 1 | EV | standard | H2O | 1664800898 | 0.315 | 0.715 | 85.196 | 0 | 0.001 | 2.025 | 1307.261 | 1.581 | 2.962 | 114.134 | 0.866 | 0.255 | 0 | 0 | 1 | 42.188 | 1 | 1 | 1 | GJB_Rinse | 0
6 | P-1292 | 2022/10/03 12:45:29 | 2-01 | 6 | -7.417 | -59.925 | 18687 | 0 | 1 | EV | standard | H2O | 1664801435 | 0.29 | 0.73 | 78.455 | 0.001 | 0 | 1.864 | 1305.221 | 1.999 | 3.019 | 115.583 | 0.865 | 0.756 | 0 | 0 | 1 | 42.312 | 1 | 1 | 1 | GJB_Rinse | 0
7 | P-1292 | 2022/10/03 12:54:26 | 2-01 | 7 | -7.498 | -60.02 | 18297 | 0 | 1 | EV | standard | H2O | 1664801973 | 0.316 | 0.769 | 80.08 | 0.002 | 0.002 | 1.91 | 1301.843 | 1.442 | 2.949 | 113.707 | 0.862 | 0.418 | 0 | 0 | 1 | 42.25 | 1 | 1 | 1 | GJB_Rinse | 0
8 | P-1292 | 2022/10/03 13:03:22 | 2-01 | 8 | -7.429 | -59.898 | 18141 | 0 | 1 | EV | standard | H2O | 1664802509 | 0.304 | 0.7 | 80.513 | 0 | -0.001 | 1.914 | 1299.231 | 1.84 | 2.89 | 111.731 | 0.87 | 0.325 | 0 | 0 | 1 | 42.125 | 1 | 1 | 1 | GJB_Rinse | 0
9 | P-1292 | 2022/10/03 13:12:19 | 2-01 | 9 | -7.447 | -59.796 | 18061 | 0 | 1 | EV | standard | H2O | 1664803046 | 0.298 | 0.701 | 77.176 | 0.002 | 0.001 | 1.827 | 1297.083 | 1.338 | 2.886 | 111.991 | 0.863 | 0.454 | 0 | 0 | 1 | 42.062 | 1 | 1 | 1 | GJB_Rinse | 0
10 | P-1292 | 2022/10/03 13:21:16 | 2-01 | 10 | -7.463 | -59.74 | 17951 | 0 | 1 | EV | standard | H2O | 1664803582 | 0.337 | 0.819 | 76.541 | 0.001 | 0.004 | 1.827 | 1295.568 | 1.681 | 2.913 | 112.179 | 0.862 | 0.641 | 0 | 0 | 1 | 42.125 | 1 | 1 | 1 | GJB_Rinse | 0
11 | P-1293 | 2022/10/03 13:31:04 | 2-02 | 1 | 3.939 | 27.899 | 18238 | -1 | 1 | PZ | standard | H2O | 1664804171 | 0.319 | 0.83 | 77.377 | 0 | -0.008 | 1.833 | 1296.994 | 1.929 | 2.881 | 110.905 | 0.863 | 0.048 | 0 | 0 | 1 | 41.625 | 1 | 2 | 1 | GJB_Rinse | 0
12 | P-1293 | 2022/10/03 13:40:01 | 2-02 | 2 | 4.421 | 36.041 | 17930 | -1 | 1 | PZ | standard | H2O | 1664804708 | 0.349 | 0.817 | 76.791 | 0.001 | -0.007 | 1.822 | 1295.029 | 1.13 | 2.915 | 111.355 | 0.861 | 0.269 | 0 | 0 | 1 | 41.938 | 1 | 2 | 1 | GJB_Rinse | 0
13 | P-1293 | 2022/10/03 13:48:58 | 2-02 | 3 | 4.625 | 38.497 | 18188 | -1 | 1 | PZ | standard | H2O | 1664805244 | 0.311 | 0.73 | 79.551 | 0 | -0.003 | 1.883 | 1296.664 | 1.563 | 2.963 | 112.65 | 0.864 | 0.559 | 0 | 0 | 1 | 42.062 | 1 | 2 | 1 | GJB_Rinse | 0
14 | P-1293 | 2022/10/03 13:57:55 | 2-02 | 4 | 4.658 | 39.122 | 18346 | 0 | 1 | PZ | standard | H2O | 1664805781 | 0.304 | 0.747 | 79.773 | 0.001 | -0.003 | 1.885 | 1297.916 | 1.157 | 2.996 | 113.853 | 0.864 | 0.29 | 0 | 0 | 1 | 42.188 | 1 | 2 | 1 | GJB_Rinse | 0
15 | P-1293 | 2022/10/03 14:06:52 | 2-02 | 5 | 4.724 | 39.579 | 18067 | 0 | 1 | PZ | standard | H2O | 1664806318 | 0.314 | 0.779 | 76.045 | 0 | -0.005 | 1.811 | 1296.061 | 1.217 | 2.933 | 111.685 | 0.864 | 0.262 | 0 | 0 | 1 | 42.25 | 1 | 2 | 1 | GJB_Rinse | 0
16 | P-1293 | 2022/10/03 14:15:49 | 2-02 | 6 | 4.653 | 39.6 | 17918 | 0 | 1 | PZ | standard | H2O | 1664806855 | 0.306 | 0.722 | 75.934 | 0.001 | -0.002 | 1.815 | 1294.379 | 1.215 | 2.96 | 112.977 | 0.861 | 0.704 | 0 | 0 | 1 | 42.25 | 1 | 2 | 1 | GJB_Rinse | 0
17 | P-1293 | 2022/10/03 14:24:46 | 2-02 | 7 | 4.781 | 40.205 | 18187 | 0 | 1 | PZ | standard | H2O | 1664807392 | 0.338 | 0.756 | 78.764 | 0.002 | 0.001 | 1.863 | 1296.587 | 1.236 | 2.997 | 113.49 | 0.864 | 0.49 | 0 | 0 | 1 | 42.312 | 1 | 2 | 1 | GJB_Rinse | 0
18 | P-1293 | 2022/10/03 14:33:43 | 2-02 | 8 | 4.725 | 40.48 | 18140 | 0 | 1 | PZ | standard | H2O | 1664807929 | 0.337 | 0.735 | 78.997 | -0.001 | 0 | 1.867 | 1296.434 | 1.241 | 2.891 | 111.58 | 0.865 | 0.388 | 0 | 0 | 1 | 42.312 | 1 | 2 | 1 | GJB_Rinse | 0
19 | P-1293 | 2022/10/03 14:42:39 | 2-02 | 9 | 4.792 | 40.805 | 18296 | 0 | 1 | PZ | standard | H2O | 1664808466 | 0.311 | 0.76 | 77.469 | 0.001 | 0 | 1.841 | 1297.634 | 1.211 | 2.966 | 112.135 | 0.865 | 0.106 | 0 | 0 | 1 | 42.312 | 1 | 2 | 1 | GJB_Rinse | 0
20 | P-1293 | 2022/10/03 14:51:36 | 2-02 | 10 | 4.753 | 40.427 | 17834 | 0 | 1 | PZ | standard | H2O | 1664809002 | 0.269 | 0.719 | 78.462 | 0 | -0.001 | 1.872 | 1294.01 | 1.549 | 2.89 | 110.428 | 0.873 | 0.435 | 0 | 0 | 1 | 42.438 | 1 | 2 | 1 | GJB_Rinse | 0
21 | P-1294 | 2022/10/03 15:01:25 | 2-03 | 1 | -12.197 | -92.886 | 18209 | -1 | 1 | UT2 | standard | H2O | 1664809591 | 0.292 | 0.861 | 74.642 | 0.002 | 0.014 | 1.774 | 1296.484 | 2.257 | 2.922 | 112.128 | 0.864 | 0.528 | 0 | 0 | 1 | 42.5 | 1 | 3 | 1 | GJB_Rinse | 0
22 | P-1294 | 2022/10/03 15:10:21 | 2-03 | 2 | -13.002 | -105.698 | 17957 | -1 | 1 | UT2 | standard | H2O | 1664810127 | 0.314 | 0.941 | 74.279 | 0.001 | 0.014 | 1.77 | 1294.971 | 1.406 | 2.861 | 111.633 | 0.862 | 0.427 | 0 | 0 | 1 | 42.5 | 1 | 3 | 1 | GJB_Rinse | 0
23 | P-1294 | 2022/10/03 15:19:18 | 2-03 | 3 | -13.187 | -108.406 | 17866 | -1 | 1 | UT2 | standard | H2O | 1664810664 | 0.321 | 0.849 | 73.625 | 0.001 | 0.012 | 1.744 | 1294.189 | 2.533 | 2.801 | 108.345 | 0.863 | 0.251 | 0 | 0 | 1 | 42.438 | 1 | 3 | 1 | GJB_Rinse | 0
24 | P-1294 | 2022/10/03 15:28:14 | 2-03 | 4 | -13.345 | -110.117 | 17750 | 0 | 1 | UT2 | standard | H2O | 1664811201 | 0.296 | 0.811 | 76.654 | 0.001 | 0.009 | 1.826 | 1293.121 | 1.513 | 2.898 | 111.651 | 0.864 | 0.792 | 0 | 0 | 1 | 42.438 | 1 | 3 | 1 | GJB_Rinse | 0
25 | P-1294 | 2022/10/03 15:37:11 | 2-03 | 5 | -13.214 | -109.983 | 18314 | 0 | 1 | UT2 | standard | H2O | 1664811737 | 0.307 | 0.746 | 78.129 | 0.002 | 0.006 | 1.858 | 1297.242 | 2.473 | 2.842 | 112.434 | 0.864 | 0.368 | 0 | 0 | 1 | 42.375 | 1 | 3 | 1 | GJB_Rinse | 0
26 | P-1294 | 2022/10/03 15:46:08 | 2-03 | 6 | -13.363 | -110.818 | 17703 | 0 | 1 | UT2 | standard | H2O | 1664812274 | 0.301 | 0.792 | 78.788 | 0 | 0.005 | 1.872 | 1293.196 | 1.773 | 2.805 | 108.068 | 0.865 | 0.244 | 0 | 0 | 1 | 42.188 | 1 | 3 | 1 | GJB_Rinse | 0
27 | P-1294 | 2022/10/03 15:55:04 | 2-03 | 7 | -13.338 | -111.128 | 18063 | 0 | 1 | UT2 | standard | H2O | 1664812811 | 0.315 | 0.804 | 76.003 | 0.001 | 0.006 | 1.807 | 1295.33 | 2.607 | 2.83 | 110.239 | 0.864 | 0.387 | 0 | 0 | 1 | 42.375 | 1 | 3 | 1 | GJB_Rinse | 0
28 | P-1294 | 2022/10/03 16:04:01 | 2-03 | 8 | -13.4 | -111.395 | 17923 | 0 | 1 | UT2 | standard | H2O | 1664813348 | 0.293 | 0.794 | 75.097 | 0.001 | 0.002 | 1.779 | 1294.452 | 2.125 | 2.833 | 109.716 | 0.867 | 0.334 | 0 | 0 | 1 | 42.312 | 1 | 3 | 1 | GJB_Rinse | 0
29 | P-1294 | 2022/10/03 16:12:58 | 2-03 | 9 | -13.416 | -111.414 | 18015 | 0 | 1 | UT2 | standard | H2O | 1664813885 | 0.318 | 0.731 | 76.751 | 0.001 | 0.007 | 1.828 | 1294.923 | 1.717 | 2.911 | 112.245 | 0.869 | 0.569 | 0 | 0 | 1 | 42.25 | 1 | 3 | 1 | GJB_Rinse | 0
30 | P-1294 | 2022/10/03 16:21:55 | 2-03 | 10 | -13.403 | -111.634 | 17870 | 0 | 1 | UT2 | standard | H2O | 1664814421 | 0.339 | 0.76 | 76.232 | 0.001 | 0.006 | 1.816 | 1294.003 | 2.277 | 2.812 | 109.679 | 0.863 | 0.405 | 0 | 0 | 1 | 42 | 1 | 3 | 1 | GJB_Rinse | 0
31 | P-1295 | 2022/10/03 16:31:44 | 2-04 | 1 | -8.005 | -66.361 | 17770 | -1 | 1 | EV | standard | H2O | 1664815011 | 0.31 | 0.684 | 74.269 | 0 | -0.001 | 1.766 | 1293.335 | 1.743 | 2.825 | 109.41 | 0.869 | 0.313 | 0 | 0 | 1 | 42.188 | 1 | 4 | 1 | GJB_Rinse | 0
32 | P-1295 | 2022/10/03 16:40:41 | 2-04 | 2 | -7.739 | -61.962 | 17872 | -1 | 1 | EV | standard | H2O | 1664815547 | 0.319 | 0.744 | 75.827 | 0.001 | 0.001 | 1.801 | 1293.796 | 2.213 | 2.877 | 110.444 | 0.866 | 0.601 | 0 | 0 | 1 | 42.188 | 1 | 4 | 1 | GJB_Rinse | 0
33 | P-1295 | 2022/10/03 16:49:38 | 2-04 | 3 | -7.69 | -61.034 | 17844 | -1 | 1 | EV | standard | H2O | 1664816084 | 0.327 | 0.71 | 74.812 | 0.001 | 0.002 | 1.771 | 1293.394 | 2.644 | 2.835 | 109.689 | 0.868 | 0.636 | 0 | 0 | 1 | 42.188 | 1 | 4 | 1 | GJB_Rinse | 0
34 | P-1295 | 2022/10/03 16:58:34 | 2-04 | 4 | -7.661 | -60.792 | 17932 | 0 | 1 | EV | standard | H2O | 1664816621 | 0.337 | 0.69 | 74.572 | 0 | 0.002 | 1.77 | 1294.472 | 1.367 | 2.867 | 111.625 | 0.867 | 0.449 | 0 | 0 | 1 | 42.312 | 1 | 4 | 1 | GJB_Rinse | 0
35 | P-1295 | 2022/10/03 17:07:31 | 2-04 | 5 | -7.628 | -60.472 | 17899 | 0 | 1 | EV | standard | H2O | 1664817158 | 0.334 | 0.714 | 75.606 | 0 | 0 | 1.801 | 1294.288 | 1.705 | 2.854 | 110.256 | 0.869 | 0.297 | 0 | 0 | 1 | 42.312 | 1 | 4 | 1 | GJB_Rinse | 0
36 | P-1295 | 2022/10/03 17:16:28 | 2-04 | 6 | -7.602 | -60.476 | 17804 | 0 | 1 | EV | standard | H2O | 1664817695 | 0.291 | 0.737 | 75.918 | 0.001 | 0.001 | 1.802 | 1293.449 | 2.103 | 2.854 | 109.458 | 0.866 | 0.448 | 0 | 0 | 1 | 42.312 | 1 | 4 | 1 | GJB_Rinse | 0
37 | P-1295 | 2022/10/03 17:25:25 | 2-04 | 7 | -7.632 | -60.23 | 17852 | 0 | 1 | EV | standard | H2O | 1664818232 | 0.319 | 0.788 | 73.743 | 0.001 | 0.001 | 1.756 | 1293.886 | 1.478 | 2.851 | 110.908 | 0.863 | 0.411 | 0 | 0 | 1 | 42.312 | 1 | 4 | 1 | GJB_Rinse | 0
38 | P-1295 | 2022/10/03 17:34:22 | 2-04 | 8 | -7.616 | -60.175 | 17951 | 0 | 1 | EV | standard | H2O | 1664818768 | 0.287 | 0.737 | 76.974 | 0 | 0.005 | 1.824 | 1294.647 | 0.975 | 2.893 | 112.281 | 0.867 | 0.52 | 0 | 0 | 1 | 42.375 | 1 | 4 | 1 | GJB_Rinse | 0
39 | P-1295 | 2022/10/03 17:43:19 | 2-04 | 9 | -7.609 | -60.101 | 17922 | 0 | 1 | EV | standard | H2O | 1664819306 | 0.337 | 0.772 | 77.726 | 0.001 | 0.003 | 1.844 | 1294.148 | 2.088 | 2.845 | 110.802 | 0.867 | 0.56 | 0 | 0 | 1 | 42.375 | 1 | 4 | 1 | GJB_Rinse | 0
40 | P-1295 | 2022/10/03 17:52:16 | 2-04 | 10 | -7.605 | -60.237 | 17836 | 0 | 1 | EV | standard | H2O | 1664819843 | 0.296 | 0.75 | 75.358 | 0.001 | 0.004 | 1.794 | 1293.491 | 2.437 | 2.848 | 109.172 | 0.869 | 0.4 | 0 | 0 | 1 | 42.312 | 1 | 4 | 1 | GJB_Rinse | 0
41 | P-1296 | 2022/10/03 18:02:06 | 2-05 | 1 | -14.375 | -113.704 | 17534 | -1 | 1 | Mill Creek Porter Fork Upper-01 | 22-254 | H2O | 1664820432 | 0.332 | 0.812 | 76.74 | 0.002 | 0.008 | 1.837 | 1291.298 | 2.194 | 2.799 | 109.036 | 0.868 | 0.625 | 0 | 0 | 1 | 42.438 | 1 | 5 | 1 | GJB_Rinse | 0
42 | P-1296 | 2022/10/03 18:11:03 | 2-05 | 2 | -14.684 | -118.569 | 17930 | -1 | 1 | Mill Creek Porter Fork Upper-01 | 22-254 | H2O | 1664820969 | 0.336 | 0.832 | 76.076 | 0.002 | 0.008 | 1.807 | 1294.377 | 2.107 | 2.853 | 110.82 | 0.871 | 0.491 | 0 | 0 | 1 | 42.438 | 1 | 5 | 1 | GJB_Rinse | 0
43 | P-1296 | 2022/10/03 18:20:00 | 2-05 | 3 | -14.757 | -119.794 | 18057 | -1 | 1 | Mill Creek Porter Fork Upper-01 | 22-254 | H2O | 1664821506 | 0.285 | 0.804 | 77.566 | 0.001 | 0.008 | 1.848 | 1295.473 | 2.481 | 2.862 | 109.214 | 0.869 | 0.133 | 0 | 0 | 1 | 42.438 | 1 | 5 | 1 | GJB_Rinse | 0
44 | P-1296 | 2022/10/03 18:28:56 | 2-05 | 4 | -14.742 | -120.13 | 18093 | 0 | 1 | Mill Creek Porter Fork Upper-01 | 22-254 | H2O | 1664822043 | 0.307 | 0.807 | 75.126 | 0.001 | 0.006 | 1.796 | 1295.667 | 2.029 | 2.849 | 111.315 | 0.871 | 0.332 | 0 | 0 | 1 | 42.5 | 1 | 5 | 1 | GJB_Rinse | 0
45 | P-1297 | 2022/10/03 18:38:46 | 2-06 | 1 | -14.788 | -120.079 | 17614 | -1 | 1 | Porter Fork-02 | 22-254 | H2O | 1664822633 | 0.29 | 0.808 | 74.138 | 0.001 | 0.007 | 1.769 | 1292.304 | 2.142 | 2.724 | 107.178 | 0.866 | 0.152 | 0 | 0 | 1 | 42.562 | 1 | 6 | 1 | GJB_Rinse | 0
46 | P-1297 | 2022/10/03 18:47:43 | 2-06 | 2 | -14.843 | -120.316 | 17664 | -1 | 1 | Porter Fork-02 | 22-254 | H2O | 1664823170 | 0.334 | 0.652 | 75.824 | 0 | 0.004 | 1.802 | 1292.407 | 2.125 | 2.807 | 108.709 | 0.865 | 0.491 | 0 | 0 | 1 | 42.562 | 1 | 6 | 1 | GJB_Rinse | 0
47 | P-1297 | 2022/10/03 18:56:40 | 2-06 | 3 | -14.844 | -120.509 | 17652 | -1 | 1 | Porter Fork-02 | 22-254 | H2O | 1664823706 | 0.317 | 0.711 | 76.803 | 0.001 | 0.005 | 1.825 | 1292.236 | 2.231 | 2.831 | 108.908 | 0.865 | 0.433 | 0 | 0 | 1 | 42.438 | 1 | 6 | 1 | GJB_Rinse | 0
48 | P-1297 | 2022/10/03 19:05:37 | 2-06 | 4 | -14.812 | -120.725 | 17662 | 0 | 1 | Porter Fork-02 | 22-254 | H2O | 1664824243 | 0.293 | 0.765 | 73.759 | 0.001 | 0.002 | 1.752 | 1292.75 | 2.022 | 2.714 | 107.494 | 0.866 | 0.194 | 0 | 0 | 1 | 42.5 | 1 | 6 | 1 | GJB_Rinse | 0
49 | P-1298 | 2022/10/03 19:15:28 | 2-07 | 1 | -14.742 | -120.455 | 18241 | -1 | 1 | Mill Creek Porter Fork Lower-03 | 22-254 | H2O | 1664824834 | 0.301 | 0.753 | 75.228 | 0.001 | 0.005 | 1.792 | 1296.738 | 1.911 | 2.899 | 113.092 | 0.868 | 0.531 | 0 | 0 | 1 | 42.438 | 1 | 7 | 1 | GJB_Rinse | 0
50 | P-1298 | 2022/10/03 19:24:24 | 2-07 | 2 | -14.426 | -119.345 | 19301 | -1 | 1 | Mill Creek Porter Fork Lower-03 | 22-254 | H2O | 1664825371 | 0.325 | 0.654 | 79.076 | 0.001 | 0.002 | 1.869 | 1304.216 | 2.238 | 3.113 | 119.77 | 0.864 | 0.628 | 0 | 0 | 1 | 42.312 | 1 | 7 | 1 | GJB_Rinse | 0
51 | P-1298 | 2022/10/03 19:33:21 | 2-07 | 3 | -13.684 | -117.046 | 19441 | -1 | 1 | Mill Creek Porter Fork Lower-03 | 22-254 | H2O | 1664825907 | 0.328 | 0.699 | 80.752 | 0 | 0 | 1.895 | 1305.382 | 2.193 | 3.12 | 120.07 | 0.869 | 0.553 | 0 | 0 | 1 | 42.25 | 1 | 7 | 1 | GJB_Rinse | 0
52 | P-1298 | 2022/10/03 19:42:18 | 2-07 | 4 | -14.61 | -120.266 | 17957 | 0 | 1 | Mill Creek Porter Fork Lower-03 | 22-254 | H2O | 1664826445 | 0.288 | 0.77 | 73.838 | 0.002 | 0.005 | 1.763 | 1294.727 | 2.123 | 2.804 | 109.964 | 0.871 | 0.237 | 0 | 0 | 1 | 42.25 | 1 | 7 | 1 | GJB_Rinse | 0
53 | P-1299 | 2022/10/03 19:52:06 | 2-08 | 1 | -14.76 | -120.746 | 18146 | -1 | 1 | Mill Creek Upstream of Church Fork-04 | 22-254 | H2O | 1664827033 | 0.348 | 0.767 | 75.876 | 0.001 | 0.004 | 1.801 | 1296.01 | 3.015 | 2.799 | 108.656 | 0.871 | 0.114 | 0 | 0 | 1 | 42.312 | 1 | 8 | 1 | GJB_Rinse | 0
54 | P-1299 | 2022/10/03 20:01:02 | 2-08 | 2 | -14.746 | -120.691 | 18162 | -1 | 1 | Mill Creek Upstream of Church Fork-04 | 22-254 | H2O | 1664827569 | 0.296 | 0.672 | 76.54 | 0.001 | 0.003 | 1.823 | 1296.139 | 2.019 | 2.893 | 111.581 | 0.864 | 0.416 | 0 | 0 | 1 | 42.375 | 1 | 8 | 1 | GJB_Rinse | 0
55 | P-1299 | 2022/10/03 20:09:59 | 2-08 | 3 | -14.845 | -121.06 | 17696 | -1 | 1 | Mill Creek Upstream of Church Fork-04 | 22-254 | H2O | 1664828106 | 0.273 | 0.766 | 73.481 | 0 | 0.002 | 1.75 | 1292.588 | 2.29 | 2.835 | 108.877 | 0.867 | 0.381 | 0 | 0 | 1 | 42.375 | 1 | 8 | 1 | GJB_Rinse | 0
56 | P-1299 | 2022/10/03 20:18:56 | 2-08 | 4 | -14.895 | -121.381 | 18083 | 0 | 1 | Mill Creek Upstream of Church Fork-04 | 22-254 | H2O | 1664828643 | 0.341 | 0.788 | 75.51 | 0.002 | 0.005 | 1.805 | 1295.392 | 2.296 | 2.806 | 110.439 | 0.871 | 0.155 | 0 | 0 | 1 | 42.438 | 1 | 8 | 1 | GJB_Rinse | 0
57 | P-1300 | 2022/10/03 20:28:44 | 2-09 | 1 | -13.757 | -113.459 | 17602 | -1 | 1 | Church Fork-05 | 22-254 | H2O | 1664829231 | 0.312 | 0.691 | 77.217 | 0 | 0.005 | 1.84 | 1292.015 | 1.833 | 2.783 | 108.73 | 0.873 | 0.311 | 0 | 0 | 1 | 42.5 | 1 | 9 | 1 | GJB_Rinse | 0
58 | P-1300 | 2022/10/03 20:37:41 | 2-09 | 2 | -13.65 | -112.405 | 18001 | -1 | 1 | Church Fork-05 | 22-254 | H2O | 1664829768 | 0.313 | 0.779 | 77.335 | 0.002 | 0.004 | 1.838 | 1294.538 | 2.332 | 2.917 | 111.898 | 0.871 | 0.654 | 0 | 0 | 1 | 42.5 | 1 | 9 | 1 | GJB_Rinse | 0
59 | P-1300 | 2022/10/03 20:46:38 | 2-09 | 3 | -13.627 | -112.295 | 17998 | -1 | 1 | Church Fork-05 | 22-254 | H2O | 1664830305 | 0.336 | 0.703 | 78.577 | 0.001 | 0.003 | 1.855 | 1295.05 | 2 | 2.76 | 109.796 | 0.871 | -0.016 | 0 | 0 | 1 | 42.375 | 1 | 9 | 1 | GJB_Rinse | 0
60 | P-1300 | 2022/10/03 20:55:35 | 2-09 | 4 | -13.614 | -112.082 | 17921 | 0 | 1 | Church Fork-05 | 22-254 | H2O | 1664830841 | 0.317 | 0.7 | 75.401 | 0.001 | 0.003 | 1.79 | 1294.181 | 1.847 | 2.847 | 111.343 | 0.871 | 0.487 | 0 | 0 | 1 | 42.438 | 1 | 9 | 1 | GJB_Rinse | 0
61 | P-1301 | 2022/10/03 21:05:23 | 2-10 | 1 | -14.732 | -119.584 | 17583 | -1 | 1 | Mill Creek Down Stream of Church Fork-06 | 22-254 | H2O | 1664831429 | 0.301 | 0.746 | 76.223 | -0.001 | 0 | 1.818 | 1291.582 | 2.482 | 2.764 | 108.3 | 0.873 | 0.474 | 0 | 0 | 1 | 42.438 | 1 | 10 | 1 | GJB_Rinse | 0
62 | P-1301 | 2022/10/03 21:14:19 | 2-10 | 2 | -14.73 | -120.273 | 17645 | -1 | 1 | Mill Creek Down Stream of Church Fork-06 | 22-254 | H2O | 1664831966 | 0.322 | 0.756 | 76.173 | 0.002 | 0.002 | 1.816 | 1291.968 | 2.048 | 2.854 | 110.717 | 0.867 | 0.785 | 0 | 0 | 1 | 42.438 | 1 | 10 | 1 | GJB_Rinse | 0
63 | P-1301 | 2022/10/03 21:23:16 | 2-10 | 3 | -14.779 | -120.357 | 18014 | -1 | 1 | Mill Creek Down Stream of Church Fork-06 | 22-254 | H2O | 1664832503 | 0.312 | 0.707 | 76.834 | 0 | 0.002 | 1.841 | 1294.827 | 2.145 | 2.837 | 111.137 | 0.869 | 0.333 | 0 | 0 | 1 | 42.375 | 1 | 10 | 1 | GJB_Rinse | 0
64 | P-1301 | 2022/10/03 21:32:13 | 2-10 | 4 | -14.758 | -120.685 | 17737 | 0 | 1 | Mill Creek Down Stream of Church Fork-06 | 22-254 | H2O | 1664833039 | 0.287 | 0.657 | 73.203 | 0.001 | 0.005 | 1.747 | 1292.898 | 2.161 | 2.852 | 109.628 | 0.866 | 0.435 | 0 | 0 | 1 | 42.438 | 1 | 10 | 1 | GJB_Rinse | 0
65 | P-1302 | 2022/10/03 21:42:01 | 2-11 | 1 | -14.658 | -120.02 | 18013 | -1 | 1 | Gate House-07 | 22-254 | H2O | 1664833629 | 0.315 | 0.786 | 75.798 | 0 | 0.002 | 1.804 | 1294.771 | 2.918 | 2.789 | 109.1 | 0.869 | 0.225 | 0 | 0 | 1 | 42.438 | 1 | 11 | 1 | GJB_Rinse | 0
66 | P-1302 | 2022/10/03 21:50:58 | 2-11 | 2 | -14.725 | -120.195 | 17763 | -1 | 1 | Gate House-07 | 22-254 | H2O | 1664834165 | 0.294 | 0.676 | 75.615 | 0.001 | 0.001 | 1.791 | 1292.782 | 2.864 | 2.798 | 108.556 | 0.869 | 0.447 | 0 | 0 | 1 | 42.375 | 1 | 11 | 1 | GJB_Rinse | 0
67 | P-1302 | 2022/10/03 21:59:55 | 2-11 | 3 | -14.778 | -120.31 | 17663 | -1 | 1 | Gate House-07 | 22-254 | H2O | 1664834702 | 0.317 | 0.79 | 77.172 | 0.002 | 0.007 | 1.834 | 1292.302 | 2.19 | 2.824 | 108.623 | 0.865 | 0.421 | 0 | 0 | 1 | 42.375 | 1 | 11 | 1 | GJB_Rinse | 0
68 | P-1302 | 2022/10/03 22:08:52 | 2-11 | 4 | -14.671 | -119.867 | 18220 | 0 | 1 | Gate House-07 | 22-254 | H2O | 1664835239 | 0.354 | 0.75 | 78.598 | 0.002 | 0.005 | 1.858 | 1296.314 | 1.777 | 2.908 | 113.35 | 0.865 | 0.605 | 0 | 0 | 1 | 42.375 | 1 | 11 | 1 | GJB_Rinse | 0
69 | P-1303 | 2022/10/03 22:18:41 | 2-12 | 1 | 3.951 | -17.514 | 16986 | -1 | 1 | TALL4-D-5 | 22-233 | H2O | 1664835828 | 0.353 | 1.016 | 71.965 | 0 | -0.009 | 1.723 | 1660.834 | -21.551 | 2.999 | 109.989 | 0.874 | -3.445 | 0 | 0 | 1 | 42.375 | 1 | 12 | 1 | GJB_Rinse | 0
70 | P-1303 | 2022/10/03 22:27:38 | 2-12 | 2 | 4.605 | -8.337 | 16892 | -1 | 1 | TALL4-D-5 | 22-233 | H2O | 1664836365 | 0.331 | 0.95 | 73.961 | 0 | -0.004 | 1.772 | 1662.643 | -22.052 | 2.965 | 109.384 | 0.879 | -3.557 | 0 | 0 | 1 | 42.375 | 1 | 12 | 1 | GJB_Rinse | 0
71 | P-1303 | 2022/10/03 22:36:35 | 2-12 | 3 | 4.844 | -5.756 | 17223 | -1 | 1 | TALL4-D-5 | 22-233 | H2O | 1664836902 | 0.367 | 1.05 | 74.76 | -0.001 | -0.006 | 1.781 | 1673.854 | -23.231 | 3.051 | 112.871 | 0.868 | -3.513 | 0 | 0 | 1 | 42.375 | 1 | 12 | 1 | GJB_Rinse | 0
72 | P-1303 | 2022/10/03 22:45:32 | 2-12 | 4 | 4.979 | -4.701 | 17118 | 0 | 1 | TALL4-D-5 | 22-233 | H2O | 1664837438 | 0.356 | 1.033 | 76.625 | 0 | -0.003 | 1.827 | 1670.707 | -21.994 | 3.03 | 110.746 | 0.867 | -3.422 | 0 | 0 | 1 | 42.312 | 1 | 12 | 1 | GJB_Rinse | 0
73 | P-1304 | 2022/10/03 22:55:21 | 2-13 | 1 | -7.193 | -55.529 | 17919 | -1 | 1 | EV | standard | H2O | 1664838028 | 0.307 | 0.763 | 74.901 | 0.002 | 0.003 | 1.777 | 1300.898 | 1.577 | 2.894 | 111.049 | 0.867 | 0.493 | 0 | 0 | 1 | 42.375 | 1 | 13 | 1 | GJB_Rinse | 0
74 | P-1304 | 2022/10/03 23:04:17 | 2-13 | 2 | -7.569 | -59.379 | 17667 | -1 | 1 | EV | standard | H2O | 1664838564 | 0.302 | 0.689 | 74.678 | 0.001 | 0.001 | 1.772 | 1294.741 | 1.239 | 2.793 | 109.441 | 0.868 | 0.257 | 0 | 0 | 1 | 42.562 | 1 | 13 | 1 | GJB_Rinse | 0
75 | P-1304 | 2022/10/03 23:13:14 | 2-13 | 3 | -7.543 | -59.459 | 18227 | -1 | 1 | EV | standard | H2O | 1664839101 | 0.287 | 0.747 | 78.859 | 0.002 | 0.003 | 1.871 | 1297.629 | 2.089 | 2.926 | 112.525 | 0.877 | 0.455 | 0 | 0 | 1 | 42.562 | 1 | 13 | 1 | GJB_Rinse | 0
76 | P-1304 | 2022/10/03 23:22:11 | 2-13 | 4 | -7.597 | -60.051 | 17809 | 0 | 1 | EV | standard | H2O | 1664839637 | 0.301 | 0.732 | 77.202 | 0.002 | 0.002 | 1.837 | 1294.314 | 1.712 | 2.854 | 110.09 | 0.873 | 0.343 | 0 | 0 | 1 | 42.5 | 1 | 13 | 1 | GJB_Rinse | 0
77 | P-1305 | 2022/10/03 23:32:01 | 2-14 | 1 | 2.562 | -13.722 | 16556 | -1 | 1 | TALL4-A-1 | 22-233 | H2O | 1664840227 | 0.324 | 0.885 | 72.934 | 0 | -0.004 | 1.757 | 1365.926 | -7.713 | 2.756 | 107.5 | 0.869 | -2.141 | 0 | 0 | 1 | 42.562 | 1 | 14 | 1 | GJB_Rinse | 0
78 | P-1305 | 2022/10/03 23:40:58 | 2-14 | 2 | 2.971 | -9.62 | 16576 | -1 | 1 | TALL4-A-1 | 22-233 | H2O | 1664840764 | 0.325 | 0.747 | 71.227 | 0 | 0 | 1.709 | 1368.14 | -8.242 | 2.724 | 107.105 | 0.868 | -2.313 | 0 | 0 | 1 | 42.438 | 1 | 14 | 1 | GJB_Rinse | 0
79 | P-1305 | 2022/10/03 23:49:55 | 2-14 | 3 | 3.053 | -8.46 | 16657 | -1 | 1 | TALL4-A-1 | 22-233 | H2O | 1664841301 | 0.308 | 0.84 | 73.198 | 0 | -0.001 | 1.748 | 1369.149 | -8.083 | 2.705 | 106.964 | 0.866 | -2.285 | 0 | 0 | 1 | 42.375 | 1 | 14 | 1 | GJB_Rinse | 0
80 | P-1305 | 2022/10/03 23:58:51 | 2-14 | 4 | 3.116 | -7.868 | 16515 | 0 | 1 | TALL4-A-1 | 22-233 | H2O | 1664841838 | 0.33 | 0.813 | 72.102 | 0 | -0.004 | 1.723 | 1367.63 | -8.265 | 2.674 | 106.128 | 0.866 | -2.371 | 0 | 0 | 1 | 42.375 | 1 | 14 | 1 | GJB_Rinse | 0
81 | P-1306 | 2022/10/04 00:08:38 | 2-15 | 1 | 2.107 | -2.966 | 16965 | -1 | 1 | TALL4-D-1 | 22-233 | H2O | 1664842425 | 0.388 | 0.879 | 74.468 | 0 | -0.002 | 1.788 | 1527.561 | -11.618 | 2.701 | 105.788 | 0.871 | -1.63 | 0 | 0 | 1 | 42.438 | 1 | 15 | 1 | GJB_Rinse | 0
82 | P-1306 | 2022/10/04 00:17:35 | 2-15 | 2 | 2.147 | -2.109 | 17104 | -1 | 1 | TALL4-D-1 | 22-233 | H2O | 1664842961 | 0.357 | 0.892 | 71.572 | 0.002 | -0.002 | 1.714 | 1532.62 | -12.038 | 2.711 | 107.338 | 0.869 | -1.548 | 0 | 0 | 1 | 42.5 | 1 | 15 | 1 | GJB_Rinse | 0
83 | P-1306 | 2022/10/04 00:26:31 | 2-15 | 3 | 2.147 | -1.978 | 16848 | -1 | 1 | TALL4-D-1 | 22-233 | H2O | 1664843497 | 0.361 | 0.773 | 74.119 | 0 | 0 | 1.776 | 1527.242 | -11.175 | 2.696 | 104.978 | 0.868 | -1.546 | 0 | 0 | 1 | 42.438 | 1 | 15 | 1 | GJB_Rinse | 0
84 | P-1306 | 2022/10/04 00:35:28 | 2-15 | 4 | 2.111 | -1.688 | 16593 | 0 | 1 | TALL4-D-1 | 22-233 | H2O | 1664844034 | 0.342 | 0.922 | 72.655 | 0 | -0.001 | 1.737 | 1522.532 | -11.494 | 2.666 | 105.108 | 0.865 | -1.261 | 0 | 0 | 1 | 42.375 | 1 | 15 | 1 | GJB_Rinse | 0
85 | P-1307 | 2022/10/04 00:45:15 | 2-16 | 1 | 1.122 | 0.706 | 17701 | -1 | 1 | TALL4-B-4 | 22-233 | H2O | 1664844621 | 0.391 | 0.982 | 74.226 | 0.001 | -0.004 | 1.767 | 1784.819 | -16.738 | 2.818 | 105.445 | 0.873 | -0.437 | 0 | 0 | 1 | 42.375 | 1 | 16 | 1 | GJB_Rinse | 0
86 | P-1307 | 2022/10/04 00:54:11 | 2-16 | 2 | 1.156 | 1.296 | 17275 | -1 | 1 | TALL4-B-4 | 22-233 | H2O | 1664845158 | 0.382 | 0.957 | 75.567 | -0.001 | -0.003 | 1.791 | 1772.166 | -17.007 | 2.724 | 103.188 | 0.871 | -0.638 | 0 | 0 | 1 | 42.375 | 1 | 16 | 1 | GJB_Rinse | 0
87 | P-1307 | 2022/10/04 01:03:08 | 2-16 | 3 | 1.056 | 0.911 | 17352 | -1 | 1 | TALL4-B-4 | 22-233 | H2O | 1664845695 | 0.388 | 1.081 | 75.431 | 0.002 | 0.002 | 1.801 | 1775.586 | -16.939 | 2.787 | 104.838 | 0.876 | -0.364 | 0 | 0 | 1 | 42.438 | 1 | 16 | 1 | GJB_Rinse | 0
88 | P-1307 | 2022/10/04 01:12:05 | 2-16 | 4 | 1.122 | 1.286 | 17334 | 0 | 1 | TALL4-B-4 | 22-233 | H2O | 1664846231 | 0.389 | 1.065 | 75.924 | 0 | 0.001 | 1.812 | 1774.927 | -16.859 | 2.757 | 103.357 | 0.867 | -0.412 | 0 | 0 | 1 | 42.375 | 1 | 16 | 1 | GJB_Rinse | 0
89 | P-1308 | 2022/10/04 01:21:53 | 2-17 | 1 | -0.791 | -4.835 | 18086 | -1 | 1 | TALL4-3040-2 | 22-233 | H2O | 1664846819 | 0.292 | 0.689 | 76.067 | 0.001 | -0.001 | 1.817 | 1304.044 | 1.368 | 2.881 | 111.065 | 0.872 | 0.296 | 0 | 0 | 1 | 42.375 | 1 | 17 | 1 | GJB_Rinse | 0
90 | P-1308 | 2022/10/04 01:30:49 | 2-17 | 2 | -0.918 | -5.616 | 17906 | -1 | 1 | TALL4-3040-2 | 22-233 | H2O | 1664847355 | 0.327 | 0.781 | 75.709 | -0.002 | -0.001 | 1.797 | 1297.529 | 1.631 | 2.834 | 109.431 | 0.871 | 0.257 | 0 | 0 | 1 | 42.312 | 1 | 17 | 1 | GJB_Rinse | 0
91 | P-1308 | 2022/10/04 01:39:45 | 2-17 | 3 | -0.792 | -5.324 | 18247 | -1 | 1 | TALL4-3040-2 | 22-233 | H2O | 1664847893 | 0.327 | 0.774 | 77.634 | 0.002 | 0 | 1.842 | 1298.779 | 1.775 | 2.909 | 112.114 | 0.867 | 0.226 | 0 | 0 | 1 | 42.375 | 1 | 17 | 1 | GJB_Rinse | 0
92 | P-1308 | 2022/10/04 01:48:42 | 2-17 | 4 | -0.826 | -5.197 | 18199 | 0 | 1 | TALL4-3040-2 | 22-233 | H2O | 1664848429 | 0.318 | 0.737 | 75.962 | 0 | -0.001 | 1.802 | 1298.311 | 1.455 | 2.873 | 110.919 | 0.877 | -0.043 | 0 | 0 | 1 | 42.438 | 1 | 17 | 1 | GJB_Rinse | 0
93 | P-1309 | 2022/10/04 01:58:30 | 2-18 | 1 | -1.414 | -13.56 | 17765 | -1 | 1 | TALL4-deep-2 | 22-233 | H2O | 1664849017 | 0.344 | 0.762 | 75.382 | 0.001 | 0.002 | 1.797 | 1294.005 | 1.218 | 2.894 | 111.18 | 0.873 | 0.476 | 0 | 0 | 1 | 42.438 | 1 | 18 | 1 | GJB_Rinse | 0
94 | P-1309 | 2022/10/04 02:07:26 | 2-18 | 2 | -1.398 | -13.956 | 18018 | -1 | 1 | TALL4-deep-2 | 22-233 | H2O | 1664849553 | 0.312 | 0.761 | 73.52 | 0.001 | 0.004 | 1.763 | 1296.394 | 1.013 | 2.844 | 111.31 | 0.873 | 0.208 | 0 | 0 | 1 | 42.438 | 1 | 18 | 1 | GJB_Rinse | 0
95 | P-1309 | 2022/10/04 02:16:23 | 2-18 | 3 | -1.445 | -14.284 | 17924 | -1 | 1 | TALL4-deep-2 | 22-233 | H2O | 1664850090 | 0.295 | 0.717 | 75.417 | 0 | 0.001 | 1.796 | 1295.494 | 1.175 | 2.882 | 110.864 | 0.864 | 0.225 | 0 | 0 | 1 | 42.375 | 1 | 18 | 1 | GJB_Rinse | 0
96 | P-1309 | 2022/10/04 02:25:20 | 2-18 | 4 | -1.454 | -14.411 | 17818 | 0 | 1 | TALL4-deep-2 | 22-233 | H2O | 1664850626 | 0.32 | 0.771 | 75.264 | 0.002 | 0.003 | 1.792 | 1294.28 | 1.648 | 2.863 | 110.255 | 0.869 | 0.38 | 0 | 0 | 1 | 42.25 | 1 | 18 | 1 | GJB_Rinse | 0
97 | P-1310 | 2022/10/04 02:35:08 | 2-19 | 1 | 0.043 | 0.601 | 17693 | -1 | 1 | TALL4-1020-2 | 22-233 | H2O | 1664851214 | 0.316 | 0.747 | 78.432 | 0 | 0.001 | 1.868 | 1295.658 | 1.639 | 2.875 | 110.084 | 0.863 | 0.588 | 0 | 0 | 1 | 42.312 | 1 | 19 | 1 | GJB_Rinse | 0
98 | P-1310 | 2022/10/04 02:44:04 | 2-19 | 2 | 0.264 | 1.905 | 18091 | -1 | 1 | TALL4-1020-2 | 22-233 | H2O | 1664851750 | 0.298 | 0.759 | 77.809 | 0.001 | 0.002 | 1.851 | 1299.06 | 1.239 | 2.892 | 111.614 | 0.869 | 0.394 | 0 | 0 | 1 | 42.375 | 1 | 19 | 1 | GJB_Rinse | 0
99 | P-1310 | 2022/10/04 02:53:00 | 2-19 | 3 | 0.193 | 2.15 | 17918 | -1 | 1 | TALL4-1020-2 | 22-233 | H2O | 1664852287 | 0.327 | 0.75 | 74.081 | 0.002 | 0 | 1.754 | 1297.827 | 1.031 | 2.909 | 112.217 | 0.863 | 0.553 | 0 | 0 | 1 | 42.312 | 1 | 19 | 1 | GJB_Rinse | 0
100 | P-1310 | 2022/10/04 03:01:57 | 2-19 | 4 | 0.167 | 2.324 | 17985 | 0 | 1 | TALL4-1020-2 | 22-233 | H2O | 1664852824 | 0.305 | 0.695 | 73.893 | 0.001 | -0.002 | 1.764 | 1298.171 | 1.25 | 2.913 | 111.708 | 0.872 | 0.394 | 0 | 0 | 1 | 42.438 | 1 | 19 | 1 | GJB_Rinse | 0
101 | P-1311 | 2022/10/04 03:11:46 | 2-20 | 1 | -0.396 | -5.383 | 17799 | -1 | 1 | TALL4-2030-1 | 22-233 | H2O | 1664853412 | 0.321 | 0.742 | 71.475 | 0.001 | 0 | 1.694 | 1294.299 | 1.686 | 2.792 | 108.89 | 0.867 | 0.146 | 0 | 0 | 1 | 42.375 | 1 | 20 | 1 | GJB_Rinse | 0
102 | P-1311 | 2022/10/04 03:20:42 | 2-20 | 2 | -0.43 | -5.521 | 18287 | -1 | 1 | TALL4-2030-1 | 22-233 | H2O | 1664853949 | 0.323 | 0.721 | 77.392 | 0.002 | 0 | 1.842 | 1297.773 | 1.533 | 2.901 | 112.313 | 0.864 | 0.302 | 0 | 0 | 1 | 42.375 | 1 | 20 | 1 | GJB_Rinse | 0
103 | P-1311 | 2022/10/04 03:29:39 | 2-20 | 3 | -0.425 | -5.898 | 17992 | -1 | 1 | TALL4-2030-1 | 22-233 | H2O | 1664854485 | 0.346 | 0.747 | 77.445 | 0.001 | 0.001 | 1.842 | 1295.422 | 1.842 | 2.928 | 111.177 | 0.87 | 0.404 | 0 | 0 | 1 | 42.438 | 1 | 20 | 1 | GJB_Rinse | 0
104 | P-1311 | 2022/10/04 03:38:36 | 2-20 | 4 | -0.409 | -6.172 | 17726 | 0 | 1 | TALL4-2030-1 | 22-233 | H2O | 1664855022 | 0.29 | 0.789 | 76.824 | 0 | -0.001 | 1.83 | 1293.848 | 0.879 | 2.863 | 110.513 | 0.868 | 0.292 | 0 | 0 | 1 | 42.562 | 1 | 20 | 1 | GJB_Rinse | 0
105 | P-1312 | 2022/10/04 03:48:25 | 2-21 | 1 | 0.926 | -1.17 | 18097 | -1 | 1 | TALL4-C-1 | 22-233 | H2O | 1664855611 | 0.333 | 0.722 | 78.081 | 0 | 0.003 | 1.843 | 1342.325 | -1.934 | 2.878 | 113.502 | 0.871 | -0.061 | 0 | 0 | 1 | 42.438 | 1 | 21 | 1 | GJB_Rinse | 0
106 | P-1312 | 2022/10/04 03:57:21 | 2-21 | 2 | 0.933 | -0.831 | 18234 | -1 | 1 | TALL4-C-1 | 22-233 | H2O | 1664856148 | 0.32 | 0.793 | 74.396 | 0 | 0.001 | 1.765 | 1344.207 | -1.416 | 2.907 | 112.757 | 0.866 | -0.165 | 0 | 0 | 1 | 42.375 | 1 | 21 | 1 | GJB_Rinse | 0
107 | P-1312 | 2022/10/04 04:06:18 | 2-21 | 3 | 0.923 | -0.947 | 17959 | -1 | 1 | TALL4-C-1 | 22-233 | H2O | 1664856684 | 0.319 | 0.729 | 76.937 | 0 | 0.001 | 1.837 | 1341.457 | -1.84 | 2.924 | 113.342 | 0.866 | 0.105 | 0 | 0 | 1 | 42.312 | 1 | 21 | 1 | GJB_Rinse | 0
108 | P-1312 | 2022/10/04 04:15:14 | 2-21 | 4 | 0.944 | -0.805 | 18093 | 0 | 1 | TALL4-C-1 | 22-233 | H2O | 1664857221 | 0.33 | 0.715 | 74.8 | 0.001 | 0 | 1.783 | 1342.875 | -1.67 | 2.909 | 113.276 | 0.867 | -0.255 | 0 | 0 | 1 | 42.312 | 1 | 21 | 1 | GJB_Rinse | 0
109 | P-1313 | 2022/10/04 04:24:59 | 2-22 | 1 | -7.047 | -51.913 | 17633 | -1 | 1 | EV | standard | H2O | 1664857806 | 0.299 | 0.828 | 77.785 | 0.001 | 0.005 | 1.852 | 1293.085 | 1.971 | 2.824 | 108.692 | 0.864 | 0.432 | 0 | 0 | 1 | 42.312 | 1 | 22 | 1 | GJB_Rinse | 0
110 | P-1313 | 2022/10/04 04:33:56 | 2-22 | 2 | -7.333 | -57.301 | 17815 | -1 | 1 | EV | standard | H2O | 1664858342 | 0.339 | 0.756 | 74.129 | 0.001 | 0.004 | 1.76 | 1293.921 | 1.861 | 2.846 | 109.946 | 0.865 | 0.477 | 0 | 0 | 1 | 42.375 | 1 | 22 | 1 | GJB_Rinse | 0
111 | P-1313 | 2022/10/04 04:42:53 | 2-22 | 3 | -7.422 | -57.885 | 17817 | -1 | 1 | EV | standard | H2O | 1664858879 | 0.326 | 0.746 | 75.157 | 0 | 0.005 | 1.788 | 1293.696 | 1.314 | 2.874 | 112.13 | 0.87 | 0.64 | 0 | 0 | 1 | 42.438 | 1 | 22 | 1 | GJB_Rinse | 0
112 | P-1313 | 2022/10/04 04:51:49 | 2-22 | 4 | -7.389 | -58.618 | 17821 | 0 | 1 | EV | standard | H2O | 1664859416 | 0.348 | 0.74 | 72.134 | 0.001 | 0.005 | 1.729 | 1293.604 | 1.906 | 2.876 | 110.503 | 0.869 | 0.495 | 0 | 0 | 1 | 42.438 | 1 | 22 | 1 | GJB_Rinse | 0
113 | P-1314 | 2022/10/04 05:01:35 | 2-23 | 1 | -0.502 | -15.701 | 17722 | -1 | 1 | TALL4-3040-1 | 22-233 | H2O | 1664860002 | 0.329 | 0.726 | 76.242 | 0.002 | 0 | 1.817 | 1296.381 | 0.885 | 2.871 | 109.658 | 0.868 | 0.36 | 0 | 0 | 1 | 42.375 | 1 | 23 | 1 | GJB_Rinse | 0
114 | P-1314 | 2022/10/04 05:10:32 | 2-23 | 2 | -0.179 | -12.088 | 17921 | -1 | 1 | TALL4-3040-1 | 22-233 | H2O | 1664860539 | 0.322 | 0.785 | 75.671 | 0 | 0.001 | 1.814 | 1297.695 | 1.385 | 2.881 | 110.936 | 0.869 | 0.35 | 0 | 0 | 1 | 42.375 | 1 | 23 | 1 | GJB_Rinse | 0
115 | P-1314 | 2022/10/04 05:19:28 | 2-23 | 3 | -0.075 | -11.057 | 17990 | -1 | 1 | TALL4-3040-1 | 22-233 | H2O | 1664861074 | 0.292 | 0.737 | 74.72 | 0.001 | -0.003 | 1.799 | 1298.186 | 1.258 | 2.899 | 111.337 | 0.87 | 0.327 | 0 | 0 | 1 | 42.438 | 1 | 23 | 1 | GJB_Rinse | 0
116 | P-1314 | 2022/10/04 05:28:25 | 2-23 | 4 | -0.078 | -10.671 | 17983 | 0 | 1 | TALL4-3040-1 | 22-233 | H2O | 1664861612 | 0.302 | 0.716 | 77.148 | 0.001 | 0.002 | 1.837 | 1298.308 | 0.911 | 2.883 | 111.945 | 0.869 | 0.253 | 0 | 0 | 1 | 42.375 | 1 | 23 | 1 | GJB_Rinse | 0
117 | P-1315 | 2022/10/04 05:38:12 | 2-24 | 1 | 0.761 | -10.435 | 17694 | -1 | 1 | TALL4-C-5 | 22-233 | H2O | 1664862199 | 0.312 | 0.793 | 74.8 | 0.001 | 0.001 | 1.782 | 1382.702 | -2.827 | 2.768 | 107.959 | 0.863 | -0.304 | 0 | 0 | 1 | 42.312 | 1 | 24 | 1 | GJB_Rinse | 0
118 | P-1315 | 2022/10/04 05:47:09 | 2-24 | 2 | 0.735 | -10.418 | 17399 | -1 | 1 | TALL4-C-5 | 22-233 | H2O | 1664862735 | 0.307 | 0.806 | 74.613 | 0.001 | 0.002 | 1.778 | 1379.838 | -2.873 | 2.789 | 108.346 | 0.863 | -0.195 | 0 | 0 | 1 | 42.188 | 1 | 24 | 1 | GJB_Rinse | 0
119 | P-1315 | 2022/10/04 05:56:05 | 2-24 | 3 | 0.77 | -10.697 | 17479 | -1 | 1 | TALL4-C-5 | 22-233 | H2O | 1664863272 | 0.316 | 0.751 | 72.652 | 0.001 | 0 | 1.737 | 1381.146 | -2.676 | 2.776 | 108.358 | 0.871 | -0.202 | 0 | 0 | 1 | 42.125 | 1 | 24 | 1 | GJB_Rinse | 0
120 | P-1315 | 2022/10/04 06:05:02 | 2-24 | 4 | 0.785 | -10.893 | 17150 | 0 | 1 | TALL4-C-5 | 22-233 | H2O | 1664863808 | 0.313 | 0.816 | 76.082 | 0 | 0.003 | 1.816 | 1377.415 | -2.921 | 2.739 | 106.432 | 0.866 | -0.155 | 0 | 0 | 1 | 42.25 | 1 | 24 | 1 | GJB_Rinse | 0
121 | P-1316 | 2022/10/04 06:14:49 | 2-25 | 1 | 0.619 | -8.177 | 17743 | -1 | 1 | TALL4-C-3 | 22-233 | H2O | 1664864396 | 0.322 | 0.689 | 74.977 | 0.002 | -0.002 | 1.796 | 1300.405 | 0.706 | 2.825 | 109.96 | 0.864 | 0.258 | 0 | 0 | 1 | 42.312 | 1 | 25 | 1 | GJB_Rinse | 0
122 | P-1316 | 2022/10/04 06:23:46 | 2-25 | 2 | 0.565 | -8.442 | 16751 | -1 | 1 | TALL4-C-3 | 22-233 | H2O | 1664864932 | 0.321 | 0.775 | 74.842 | 0.001 | -0.001 | 1.789 | 1291.598 | 0.431 | 2.67 | 104.241 | 0.866 | 0.109 | 0 | 0 | 1 | 42.25 | 1 | 25 | 1 | GJB_Rinse | 0
123 | P-1316 | 2022/10/04 06:32:42 | 2-25 | 3 | 0.586 | -7.97 | 17867 | -1 | 1 | TALL4-C-3 | 22-233 | H2O | 1664865469 | 0.309 | 0.825 | 75.282 | 0 | -0.001 | 1.801 | 1299.511 | -0.008 | 2.95 | 113.68 | 0.87 | 0.534 | 0 | 0 | 1 | 42.312 | 1 | 25 | 1 | GJB_Rinse | 0
124 | P-1316 | 2022/10/04 06:41:39 | 2-25 | 4 | 0.647 | -8.159 | 17898 | 0 | 1 | TALL4-C-3 | 22-233 | H2O | 1664866006 | 0.295 | 0.747 | 74.413 | 0.001 | 0.002 | 1.768 | 1299.909 | 0.394 | 2.835 | 111.017 | 0.866 | 0.037 | 0 | 0 | 1 | 42.375 | 1 | 25 | 1 | GJB_Rinse | 0
125 | P-1317 | 2022/10/04 06:51:27 | 2-26 | 1 | 1.774 | -2.4 | 17760 | -1 | 1 | TALL4-A-2 | 22-233 | H2O | 1664866593 | 0.307 | 0.786 | 75.879 | 0.001 | -0.001 | 1.797 | 1356.608 | -4.091 | 2.806 | 112.056 | 0.868 | -0.793 | 0 | 0 | 1 | 42.375 | 1 | 26 | 1 | GJB_Rinse | 0
126 | P-1317 | 2022/10/04 07:00:24 | 2-26 | 2 | 1.683 | -2.303 | 17343 | -1 | 1 | TALL4-A-2 | 22-233 | H2O | 1664867130 | 0.327 | 0.769 | 72.673 | 0 | 0.001 | 1.727 | 1352.85 | -3.594 | 2.761 | 109.134 | 0.865 | -0.658 | 0 | 0 | 1 | 42.375 | 1 | 26 | 1 | GJB_Rinse | 0
127 | P-1317 | 2022/10/04 07:09:21 | 2-26 | 3 | 1.691 | -2.151 | 17291 | -1 | 1 | TALL4-A-2 | 22-233 | H2O | 1664867667 | 0.305 | 0.823 | 74.304 | 0.001 | 0 | 1.791 | 1352.272 | -3.88 | 2.809 | 110.221 | 0.863 | -0.529 | 0 | 0 | 1 | 42.438 | 1 | 26 | 1 | GJB_Rinse | 0
128 | P-1317 | 2022/10/04 07:18:17 | 2-26 | 4 | 1.645 | -2.578 | 16730 | 0 | 1 | TALL4-A-2 | 22-233 | H2O | 1664868204 | 0.323 | 0.794 | 72.805 | 0 | 0 | 1.757 | 1346.437 | -3.16 | 2.684 | 104.379 | 0.863 | -0.752 | 0 | 0 | 1 | 42.188 | 1 | 26 | 1 | GJB_Rinse | 0
129 | P-1318 | 2022/10/04 07:28:06 | 2-27 | 1 | 0.458 | -12.03 | 17558 | -1 | 1 | TALL4-D-2 | 22-233 | H2O | 1664868792 | 0.314 | 0.769 | 75.034 | -0.001 | 0 | 1.79 | 1308.872 | 0.131 | 2.736 | 107.039 | 0.871 | -0.559 | 0 | 0 | 1 | 42.25 | 1 | 27 | 1 | GJB_Rinse | 0
130 | P-1318 | 2022/10/04 07:37:02 | 2-27 | 2 | 0.357 | -13.412 | 17666 | -1 | 1 | TALL4-D-2 | 22-233 | H2O | 1664869329 | 0.337 | 0.773 | 73.333 | 0 | 0.001 | 1.756 | 1308.951 | -0.843 | 2.808 | 111.261 | 0.873 | -0.124 | 0 | 0 | 1 | 42.312 | 1 | 27 | 1 | GJB_Rinse | 0
131 | P-1318 | 2022/10/04 07:45:59 | 2-27 | 3 | 0.411 | -13.602 | 17672 | -1 | 1 | TALL4-D-2 | 22-233 | H2O | 1664869865 | 0.293 | 0.702 | 72.781 | 0.001 | 0.002 | 1.736 | 1309.144 | -0.965 | 2.763 | 110.267 | 0.867 | -0.597 | 0 | 0 | 1 | 42.25 | 1 | 27 | 1 | GJB_Rinse | 0
132 | P-1318 | 2022/10/04 07:54:55 | 2-27 | 4 | 0.332 | -13.677 | 17348 | 0 | 1 | TALL4-D-2 | 22-233 | H2O | 1664870402 | 0.325 | 0.778 | 75.767 | 0 | 0.002 | 1.81 | 1306.519 | -0.536 | 2.701 | 106.797 | 0.865 | -0.471 | 0 | 0 | 1 | 42.312 | 1 | 27 | 1 | GJB_Rinse | 0
133 | P-1319 | 2022/10/04 08:04:43 | 2-28 | 1 | [5 empty] | TALL4-C-4 | 22-233 | [18 empty] | 1 | 28 | 1 | GJB_Rinse | 0
134 | P-1319 | 2022/10/04 08:13:39 | 2-28 | 2 | 0.889 | -3.167 | 17596 | 0 | 1 | TALL4-C-4 | 22-233 | H2O | 1664872600 | 0.43 | 1.018 | 73.917 | 0.001 | 0.002 | 1.752 | 2107.398 | -28.713 | 2.814 | 103.278 | 0.876 | -0.508 | 0 | 0 | 1 | 42.688 | 1 | 28 | 1 | GJB_Rinse | 0
135 | P-1319 | 2022/10/04 08:22:36 | 2-28 | 3 | 0.889 | -3.167 | 17596 | 0 | 1 | TALL4-C-4 | 22-233 | H2O | 1664872600 | 0.43 | 1.018 | 73.917 | 0.001 | 0.002 | 1.752 | 2107.398 | -28.713 | 2.814 | 103.278 | 0.876 | -0.508 | 0 | 0 | 1 | 42.688 | 1 | 28 | 1 | GJB_Rinse | 0
136 | P-1319 | 2022/10/04 08:31:34 | 2-28 | 4 | 0.889 | -3.167 | 17596 | 0 | 1 | TALL4-C-4 | 22-233 | H2O | 1664872600 | 0.43 | 1.018 | 73.917 | 0.001 | 0.002 | 1.752 | 2107.398 | -28.713 | 2.814 | 103.278 | 0.876 | -0.508 | 0 | 0 | 1 | 42.688 | 1 | 28 | 1 | GJB_Rinse | 0
137 | P-1320 | 2022/10/04 08:41:19 | 2-29 | 1 | 0.926 | -5.473 | 16888 | -1 | 1 | TALL4-B-1 | 22-233 | H2O | 1664873186 | 0.427 | 1.322 | 73.128 | 0.002 | -0.002 | 1.749 | 2238.908 | -34.559 | 2.656 | 98.038 | 0.89 | -1.01 | 0 | 0 | 1 | 42.75 | 1 | 29 | 1 | GJB_Rinse | 0
138 | P-1320 | 2022/10/04 08:50:16 | 2-29 | 2 | 0.897 | -5.966 | 16921 | -1 | 1 | TALL4-B-1 | 22-233 | H2O | 1664873723 | 0.438 | 1.174 | 76.179 | 0.002 | 0 | 1.825 | 2242.039 | -34.808 | 2.757 | 100.061 | 0.873 | -0.775 | 0 | 0 | 1 | 42.625 | 1 | 29 | 1 | GJB_Rinse | 0
139 | P-1320 | 2022/10/04 08:59:13 | 2-29 | 3 | 0.885 | -5.805 | 16941 | -1 | 1 | TALL4-B-1 | 22-233 | H2O | 1664874260 | 0.411 | 1.399 | 74.282 | -0.001 | 0 | 1.776 | 2242.677 | -35.138 | 2.766 | 99.606 | 0.869 | -0.873 | 0 | 0 | 1 | 42.688 | 1 | 29 | 1 | GJB_Rinse | 0
140 | P-1320 | 2022/10/04 09:08:10 | 2-29 | 4 | 0.924 | -6.072 | 16844 | 0 | 1 | TALL4-B-1 | 22-233 | H2O | 1664874797 | 0.436 | 1.25 | 71.774 | 0.001 | -0.001 | 1.711 | 2237.066 | -33.534 | 2.728 | 97.577 | 0.877 | -0.996 | 0 | 0 | 1 | 42.5 | 1 | 29 | 1 | GJB_Rinse | 0
141 | P-1321 | 2022/10/04 09:17:55 | 2-30 | 1 | 2.647 | -6.037 | 17733 | -1 | 1 | TALL4-A-4 | 22-233 | H2O | 1664875382 | 0.313 | 0.748 | 74.041 | 0 | 0 | 1.765 | 1318.599 | -2.197 | 2.83 | 111.266 | 0.862 | -0.79 | 0 | 0 | 1 | 42.375 | 1 | 30 | 1 | GJB_Rinse | 0
142 | P-1321 | 2022/10/04 09:26:52 | 2-30 | 2 | 2.732 | -6.208 | 17782 | -1 | 1 | TALL4-A-4 | 22-233 | H2O | 1664875919 | 0.311 | 0.842 | 76.751 | 0.001 | 0.003 | 1.816 | 1309.669 | -2.441 | 2.865 | 112.883 | 0.864 | -0.715 | 0 | 0 | 1 | 42.312 | 1 | 30 | 1 | GJB_Rinse | 0
143 | P-1321 | 2022/10/04 09:35:49 | 2-30 | 3 | 2.74 | -6.135 | 17753 | -1 | 1 | TALL4-A-4 | 22-233 | H2O | 1664876455 | 0.342 | 0.728 | 75.763 | 0.001 | -0.001 | 1.805 | 1307.318 | -1.951 | 2.799 | 112.571 | 0.868 | -0.728 | 0 | 0 | 1 | 42.25 | 1 | 30 | 1 | GJB_Rinse | 0
144 | P-1321 | 2022/10/04 09:44:45 | 2-30 | 4 | 2.803 | -6.186 | 17669 | 0 | 1 | TALL4-A-4 | 22-233 | H2O | 1664876992 | 0.315 | 0.775 | 75.052 | 0.001 | 0.003 | 1.8 | 1306.103 | -2.305 | 2.765 | 110.625 | 0.863 | -0.994 | 0 | 0 | 1 | 42.188 | 1 | 30 | 1 | GJB_Rinse | 0
145 | P-1322 | 2022/10/04 09:54:31 | 2-31 | 1 | -6.96 | -52.752 | 17640 | -1 | 1 | EV | standard | H2O | 1664877578 | 0.33 | 0.754 | 75.328 | 0.001 | 0.005 | 1.793 | 1293.515 | 1.901 | 2.827 | 109.587 | 0.863 | 0.609 | 0 | 0 | 1 | 42.188 | 1 | 31 | 1 | GJB_Rinse | 0
146 | P-1322 | 2022/10/04 10:03:28 | 2-31 | 2 | -7.51 | -57.563 | 17660 | -1 | 1 | EV | standard | H2O | 1664878115 | 0.341 | 0.886 | 80.392 | 0.001 | 0.011 | 1.919 | 1293.574 | 1.915 | 2.842 | 110.603 | 0.862 | 0.783 | 0 | 0 | 1 | 42.125 | 1 | 31 | 1 | GJB_Rinse | 0
147 | P-1322 | 2022/10/04 10:12:25 | 2-31 | 3 | -7.454 | -58.131 | 17778 | -1 | 1 | EV | standard | H2O | 1664878651 | 0.316 | 0.704 | 74.374 | 0 | 0.002 | 1.771 | 1294.568 | 2.3 | 2.805 | 109.03 | 0.869 | 0.274 | 0 | 0 | 1 | 42.188 | 1 | 31 | 1 | GJB_Rinse | 0
148 | P-1322 | 2022/10/04 10:21:22 | 2-31 | 4 | -7.512 | -59.151 | 17763 | 0 | 1 | EV | standard | H2O | 1664879188 | 0.291 | 0.759 | 74.438 | 0.001 | 0.003 | 1.768 | 1294.27 | 1.835 | 2.841 | 109.794 | 0.865 | 0.493 | 0 | 0 | 1 | 42.188 | 1 | 31 | 1 | GJB_Rinse | 0
149 | P-1323 | 2022/10/04 10:31:08 | 2-32 | 1 | 0.837 | -13.665 | 17499 | -1 | 1 | TALL4-D-3 | 22-233 | H2O | 1664879775 | 0.287 | 0.762 | 76.45 | 0.001 | -0.001 | 1.819 | 1306.069 | 0.133 | 2.798 | 108.573 | 0.866 | 0.014 | 0 | 0 | 1 | 42.25 | 1 | 32 | 1 | GJB_Rinse | 0
150 | P-1323 | 2022/10/04 10:40:05 | 2-32 | 2 | 1.199 | -9.517 | 17131 | -1 | 1 | TALL4-D-3 | 22-233 | H2O | 1664880312 | 0.296 | 0.743 | 73.034 | 0.001 | -0.002 | 1.75 | 1303.894 | 0.137 | 2.677 | 105.207 | 0.864 | -0.278 | 0 | 0 | 1 | 42.25 | 1 | 32 | 1 | GJB_Rinse | 0
151 | P-1323 | 2022/10/04 10:49:02 | 2-32 | 3 | 1.259 | -8.597 | 17107 | -1 | 1 | TALL4-D-3 | 22-233 | H2O | 1664880848 | 0.341 | 0.767 | 74.013 | 0.001 | 0 | 1.773 | 1303.112 | 0.318 | 2.737 | 107.221 | 0.863 | 0.205 | 0 | 0 | 1 | 42.125 | 1 | 32 | 1 | GJB_Rinse | 0
152 | P-1323 | 2022/10/04 10:57:59 | 2-32 | 4 | 1.352 | -8.175 | 17094 | 0 | 1 | TALL4-D-3 | 22-233 | H2O | 1664881385 | 0.294 | 0.79 | 73.692 | 0 | 0 | 1.771 | 1303.194 | 0.116 | 2.71 | 106.765 | 0.866 | 0.045 | 0 | 0 | 1 | 42.062 | 1 | 32 | 1 | GJB_Rinse | 0
153 | P-1324 | 2022/10/04 11:07:45 | 2-33 | 1 | -1.201 | -22.053 | 17646 | -1 | 1 | TALL4-A-5 | 22-233 | H2O | 1664881972 | 0.309 | 0.692 | 74.354 | 0.001 | 0.003 | 1.772 | 1294.661 | 0.896 | 2.847 | 110.11 | 0.862 | 0.302 | 0 | 0 | 1 | 42.188 | 1 | 33 | 1 | GJB_Rinse | 0
154 | P-1324 | 2022/10/04 11:16:42 | 2-33 | 2 | -1.335 | -23.571 | 17327 | -1 | 1 | TALL4-A-5 | 22-233 | H2O | 1664882508 | 0.325 | 0.708 | 74.429 | 0.001 | 0.002 | 1.782 | 1292.463 | 0.607 | 2.759 | 107.684 | 0.864 | 0.275 | 0 | 0 | 1 | 42.188 | 1 | 33 | 1 | GJB_Rinse | 0
155 | P-1324 | 2022/10/04 11:25:38 | 2-33 | 3 | -1.306 | -23.493 | 17800 | -1 | 1 | TALL4-A-5 | 22-233 | H2O | 1664883045 | 0.338 | 0.714 | 77.211 | 0.001 | 0.004 | 1.836 | 1295.727 | 1.627 | 2.85 | 109.682 | 0.865 | 0.231 | 0 | 0 | 1 | 42.25 | 1 | 33 | 1 | GJB_Rinse | 0
156 | P-1324 | 2022/10/04 11:34:35 | 2-33 | 4 | -1.329 | -23.638 | 17640 | 0 | 1 | TALL4-A-5 | 22-233 | H2O | 1664883581 | 0.319 | 0.749 | 74.369 | 0.001 | 0.004 | 1.762 | 1294.692 | 0.946 | 2.768 | 108.99 | 0.863 | 0.165 | 0 | 0 | 1 | 42.25 | 1 | 33 | 1 | GJB_Rinse | 0
157 | P-1325 | 2022/10/04 11:44:21 | 2-34 | 1 | -1.224 | -13.244 | 19615 | -1 | 1 | TALL4-3040-3 | 22-233 | H2O | 1664884168 | 0.285 | 0.698 | 77.32 | 0.001 | 0 | 1.829 | 1307.017 | 1.403 | 3.176 | 121.958 | 0.864 | 0.396 | 0 | 0 | 1 | 42.188 | 1 | 34 | 1 | GJB_Rinse | 0
158 | P-1325 | 2022/10/04 11:53:17 | 2-34 | 2 | -1.379 | -13.724 | 17623 | -1 | 1 | TALL4-3040-3 | 22-233 | H2O | 1664884704 | 0.319 | 0.792 | 76.806 | 0.001 | 0.002 | 1.832 | 1292.584 | 1.36 | 2.851 | 109.904 | 0.862 | 0.669 | 0 | 0 | 1 | 42.125 | 1 | 34 | 1 | GJB_Rinse | 0
159 | P-1325 | 2022/10/04 12:02:14 | 2-34 | 3 | -1.453 | -14.251 | 16986 | -1 | 1 | TALL4-3040-3 | 22-233 | H2O | 1664885240 | 0.325 | 0.834 | 73.979 | 0.001 | 0.001 | 1.774 | 1288.221 | 1.492 | 2.709 | 104.738 | 0.863 | 0.304 | 0 | 0 | 1 | 42.062 | 1 | 34 | 1 | GJB_Rinse | 0
160 | P-1325 | 2022/10/04 12:11:10 | 2-34 | 4 | -1.354 | -13.587 | 17678 | 0 | 1 | TALL4-3040-3 | 22-233 | H2O | 1664885776 | 0.316 | 0.764 | 74.979 | 0 | 0.001 | 1.789 | 1293.031 | 1.46 | 2.872 | 109.928 | 0.862 | 0.512 | 0 | 0 | 1 | 42.125 | 1 | 34 | 1 | GJB_Rinse | 0
161 | P-1326 | 2022/10/04 12:20:58 | 2-35 | 1 | 1.116 | 0.786 | 16602 | -1 | 1 | TALL4-010-2 | 22-233 | H2O | 1664886364 | 0.304 | 0.718 | 72.053 | 0 | -0.001 | 1.735 | 1285.172 | 1.632 | 2.656 | 102.849 | 0.86 | 0.47 | 0 | 0 | 1 | 42.188 | 1 | 35 | 1 | GJB_Rinse | 0
162 | P-1326 | 2022/10/04 12:29:55 | 2-35 | 2 | 1.442 | 3.445 | 17705 | -1 | 1 | TALL4-010-2 | 22-233 | H2O | 1664886901 | 0.32 | 0.719 | 77.572 | 0.001 | 0.001 | 1.836 | 1293.334 | 1.278 | 2.872 | 109.528 | 0.868 | 0.223 | 0 | 0 | 1 | 42.25 | 1 | 35 | 1 | GJB_Rinse | 0
163 | P-1326 | 2022/10/04 12:38:52 | 2-35 | 3 | 1.337 | 2.844 | 16803 | -1 | 1 | TALL4-010-2 | 22-233 | H2O | 1664887438 | 0.311 | 0.722 | 74.729 | 0.001 | 0 | 1.786 | 1286.777 | 1.385 | 2.675 | 103.864 | 0.861 | 0.247 | 0 | 0 | 1 | 42.25 | 1 | 35 | 1 | GJB_Rinse | 0
164 | P-1326 | 2022/10/04 12:47:49 | 2-35 | 4 | 1.4 | 3.238 | 17053 | 0 | 1 | TALL4-010-2 | 22-233 | H2O | 1664887975 | 0.31 | 0.807 | 73.502 | 0.001 | 0 | 1.75 | 1288.47 | 1.385 | 2.764 | 106.647 | 0.863 | 0.67 | 0 | 0 | 1 | 42.188 | 1 | 35 | 1 | GJB_Rinse | 0
165 | P-1327 | 2022/10/04 12:57:32 | 2-36 | 1 | -1.225 | -8.736 | 16700 | -1 | 1 | TALL4-2030-3 | 22-233 | H2O | 1664888558 | 0.327 | 0.788 | 73.326 | 0.002 | -0.001 | 1.76 | 1285.666 | 1.639 | 2.663 | 104.156 | 0.862 | 0.464 | 0 | 0 | 1 | 42.25 | 1 | 36 | 1 | GJB_Rinse | 0
166 | P-1327 | 2022/10/04 13:06:29 | 2-36 | 2 | -1.305 | -9.957 | 16987 | -1 | 1 | TALL4-2030-3 | 22-233 | H2O | 1664889095 | 0.311 | 0.747 | 74.003 | 0 | 0.002 | 1.767 | 1288.076 | 1.344 | 2.695 | 104.696 | 0.861 | 0.264 | 0 | 0 | 1 | 42.312 | 1 | 36 | 1 | GJB_Rinse | 0
167 | P-1327 | 2022/10/04 13:15:25 | 2-36 | 3 | -1.382 | -10.134 | 16981 | -1 | 1 | TALL4-2030-3 | 22-233 | H2O | 1664889632 | 0.33 | 0.698 | 71.622 | 0.001 | 0 | 1.717 | 1287.612 | 2.348 | 2.72 | 104.045 | 0.863 | 0.45 | 0 | 0 | 1 | 42.25 | 1 | 36 | 1 | GJB_Rinse | 0
168 | P-1327 | 2022/10/04 13:24:22 | 2-36 | 4 | -1.405 | -10.339 | 16945 | 0 | 1 | TALL4-2030-3 | 22-233 | H2O | 1664890169 | 0.302 | 0.746 | 72.67 | 0.001 | 0 | 1.735 | 1287.343 | 1.789 | 2.684 | 105.161 | 0.861 | 0.556 | 0 | 0 | 1 | 42.25 | 1 | 36 | 1 | GJB_Rinse | 0
169 | P-1328 | 2022/10/04 13:34:07 | 2-37 | 1 | -0.321 | -10.021 | 17804 | -1 | 1 | TALL4-C-2 | 22-233 | H2O | 1664890754 | 0.303 | 0.755 | 73.027 | 0.001 | 0.002 | 1.74 | 1294.823 | 1.407 | 2.821 | 109.957 | 0.862 | 0.222 | 0 | 0 | 1 | 42.25 | 1 | 37 | 1 | GJB_Rinse | 0
170 | P-1328 | 2022/10/04 13:43:04 | 2-37 | 2 | -0.262 | -9.715 | 17841 | -1 | 1 | TALL4-C-2 | 22-233 | H2O | 1664891291 | 0.313 | 0.822 | 75.077 | 0.002 | 0.004 | 1.777 | 1295.014 | 1.239 | 2.891 | 111.401 | 0.864 | 0.589 | 0 | 0 | 1 | 42.312 | 1 | 37 | 1 | GJB_Rinse | 0
171 | P-1328 | 2022/10/04 13:52:01 | 2-37 | 3 | -0.264 | -9.941 | 17770 | -1 | 1 | TALL4-C-2 | 22-233 | H2O | 1664891828 | 0.312 | 0.786 | 76.679 | -0.001 | 0 | 1.842 | 1294.764 | 1.506 | 2.803 | 108.893 | 0.862 | 0.148 | 0 | 0 | 1 | 42.438 | 1 | 37 | 1 | GJB_Rinse | 0
172 | P-1328 | 2022/10/04 14:00:58 | 2-37 | 4 | -0.273 | -10.102 | 17734 | 0 | 1 | TALL4-C-2 | 22-233 | H2O | 1664892365 | 0.322 | 0.799 | 77.102 | 0.001 | 0.005 | 1.839 | 1294.487 | 1.372 | 2.83 | 108.945 | 0.862 | 0.198 | 0 | 0 | 1 | 42.125 | 1 | 37 | 1 | GJB_Rinse | 0
173 | P-1329 | 2022/10/04 14:10:42 | 2-38 | 1 | 0.426 | -2.621 | 16900 | -1 | 1 | TALL4-B-2 | 22-233 | H2O | 1664892949 | 0.4 | 1.209 | 77.738 | 0.001 | 0.002 | 1.86 | 2068.841 | -27.617 | 2.732 | 101.148 | 0.868 | -0.192 | 0 | 0 | 1 | 42.188 | 1 | 38 | 1 | GJB_Rinse | 0
174 | P-1329 | 2022/10/04 14:19:40 | 2-38 | 2 | 0.46 | -2.221 | 16868 | -1 | 1 | TALL4-B-2 | 22-233 | H2O | 1664893486 | 0.438 | 1.122 | 75.045 | 0.002 | 0.001 | 1.799 | 2075.904 | -27.593 | 2.733 | 99.472 | 0.868 | -0.471 | 0 | 0 | 1 | 42.188 | 1 | 38 | 1 | GJB_Rinse | 0
175 | P-1329 | 2022/10/04 14:28:36 | 2-38 | 3 | 0.425 | -1.777 | 16886 | -1 | 1 | TALL4-B-2 | 22-233 | H2O | 1664894023 | 0.388 | 1.128 | 73.017 | 0.001 | 0.001 | 1.746 | 2078.787 | -27.667 | 2.74 | 99.837 | 0.87 | -0.48 | 0 | 0 | 1 | 42.25 | 1 | 38 | 1 | GJB_Rinse | 0
176 | P-1329 | 2022/10/04 14:37:33 | 2-38 | 4 | 0.44 | -1.79 | 16901 | 0 | 1 | TALL4-B-2 | 22-233 | H2O | 1664894560 | 0.394 | 1.109 | 73.642 | 0.001 | 0 | 1.76 | 2081.332 | -28.045 | 2.699 | 100.919 | 0.869 | -0.263 | 0 | 0 | 1 | 42.25 | 1 | 38 | 1 | GJB_Rinse | 0
177 | P-1330 | 2022/10/04 14:47:19 | 2-39 | 1 | 0.53 | -1.939 | 17319 | -1 | 1 | TALL4-1020-3 | 22-233 | H2O | 1664895145 | 0.322 | 0.827 | 73.718 | 0.002 | 0.001 | 1.758 | 1307.607 | 0.492 | 2.785 | 107.468 | 0.862 | 0.268 | 0 | 0 | 1 | 42.312 | 1 | 39 | 1 | GJB_Rinse | 0
178 | P-1330 | 2022/10/04 14:56:16 | 2-39 | 2 | 0.533 | -1.665 | 17691 | -1 | 1 | TALL4-1020-3 | 22-233 | H2O | 1664895682 | 0.33 | 0.742 | 74.382 | 0.001 | 0 | 1.771 | 1302.118 | 0.988 | 2.833 | 109.116 | 0.863 | 0.319 | 0 | 0 | 1 | 42.25 | 1 | 39 | 1 | GJB_Rinse | 0
179 | P-1330 | 2022/10/04 15:05:13 | 2-39 | 3 | 0.645 | -1.316 | 17858 | -1 | 1 | TALL4-1020-3 | 22-233 | H2O | 1664896219 | 0.303 | 0.776 | 75.772 | 0.002 | 0.003 | 1.804 | 1301.419 | 1.005 | 2.847 | 110.687 | 0.863 | 0.392 | 0 | 0 | 1 | 42.25 | 1 | 39 | 1 | GJB_Rinse | 0
180 | P-1330 | 2022/10/04 15:14:09 | 2-39 | 4 | 0.593 | -1.66 | 17971 | 0 | 1 | TALL4-1020-3 | 22-233 | H2O | 1664896756 | 0.316 | 0.698 | 74.055 | 0 | -0.002 | 1.76 | 1301.582 | 1.229 | 2.882 | 110.142 | 0.866 | 0.168 | 0 | 0 | 1 | 42.188 | 1 | 39 | 1 | GJB_Rinse | 0
181 | P-1331 | 2022/10/04 15:23:55 | 2-40 | 1 | -7.095 | -52.246 | 17588 | -1 | 1 | EV | standard | H2O | 1664897342 | 0.333 | 0.855 | 75.866 | 0.002 | 0.008 | 1.813 | 1292.837 | 1.939 | 2.778 | 107.914 | 0.86 | 0.292 | 0 | 0 | 1 | 42.25 | 1 | 40 | 1 | GJB_Rinse | 0
182 | P-1331 | 2022/10/04 15:32:52 | 2-40 | 2 | -7.389 | -56.759 | 17721 | -1 | 1 | EV | standard | H2O | 1664897879 | 0.292 | 0.721 | 75.615 | 0 | 0.006 | 1.802 | 1293.918 | 1.98 | 2.819 | 108.406 | 0.863 | 0.318 | 0 | 0 | 1 | 42.688 | 1 | 40 | 1 | GJB_Rinse | 0
183 | P-1331 | 2022/10/04 15:41:48 | 2-40 | 3 | -7.467 | -58.265 | 17663 | -1 | 1 | EV | standard | H2O | 1664898415 | 0.32 | 0.779 | 76.841 | 0.001 | 0.007 | 1.831 | 1293.161 | 2.123 | 2.821 | 108.551 | 0.863 | 0.291 | 0 | 0 | 1 | 42.438 | 1 | 40 | 1 | GJB_Rinse | 0
184 | P-1331 | 2022/10/04 15:50:45 | 2-40 | 4 | -7.497 | -58.513 | 17761 | 0 | 1 | EV | standard | H2O | 1664898952 | 0.296 | 0.712 | 74.082 | 0.001 | 0.006 | 1.76 | 1293.886 | 2.163 | 2.863 | 108.92 | 0.866 | 0.287 | 0 | 0 | 1 | 42.375 | 1 | 40 | 1 | GJB_Rinse | 0
185 | P-1332 | 2022/10/04 16:00:31 | 2-41 | 1 | -0.02 | -5.214 | 17863 | -1 | 1 | TALL4-010-3 | 22-233 | H2O | 1664899538 | 0.299 | 0.747 | 75.09 | 0 | -0.001 | 1.787 | 1294.3 | 2.075 | 2.874 | 109.77 | 0.864 | 0.522 | 0 | 0 | 1 | 42.438 | 1 | 41 | 1 | GJB_Rinse | 0
186 | P-1332 | 2022/10/04 16:09:28 | 2-41 | 2 | 0.387 | 0.345 | 17876 | -1 | 1 | TALL4-010-3 | 22-233 | H2O | 1664900074 | 0.308 | 0.778 | 77.602 | 0 | -0.002 | 1.846 | 1294.426 | 1.983 | 2.884 | 110.294 | 0.865 | 0.455 | 0 | 0 | 1 | 42.438 | 1 | 41 | 1 | GJB_Rinse | 0
187 | P-1332 | 2022/10/04 16:18:25 | 2-41 | 3 | 0.496 | 1.58 | 17964 | -1 | 1 | TALL4-010-3 | 22-233 | H2O | 1664900611 | 0.333 | 0.693 | 73.55 | 0.001 | -0.002 | 1.748 | 1295.442 | 1.369 | 2.91 | 110.836 | 0.865 | 0.444 | 0 | 0 | 1 | 42.375 | 1 | 41 | 1 | GJB_Rinse | 0
188 | P-1332 | 2022/10/04 16:27:21 | 2-41 | 4 | 0.438 | 1.596 | 17828 | 0 | 1 | TALL4-010-3 | 22-233 | H2O | 1664901148 | 0.331 | 0.806 | 77.713 | -0.001 | 0 | 1.85 | 1294.543 | 0.716 | 2.878 | 111.716 | 0.869 | 0.508 | 0 | 0 | 1 | 42.312 | 1 | 41 | 1 | GJB_Rinse | 0
189 | P-1333 | 2022/10/04 16:37:07 | 2-42 | 1 | 3.227 | 3.019 | 16661 | -1 | 1 | TALL4-B-3 | 22-233 | H2O | 1664901733 | 0.475 | 1.411 | 74.226 | 0.002 | -0.004 | 1.776 | 2487.185 | -46.871 | 2.759 | 99.332 | 0.874 | -2.041 | 0 | 0 | 1 | 42.188 | 1 | 42 | 1 | GJB_Rinse | 0
190 | P-1333 | 2022/10/04 16:46:03 | 2-42 | 2 | 3.216 | 2.803 | 16272 | -1 | 1 | TALL4-B-3 | 22-233 | H2O | 1664902270 | 0.457 | 1.576 | 73.149 | -0.001 | -0.003 | 1.758 | 2470.136 | -45.9 | 2.761 | 97.007 | 0.875 | -2.194 | 0 | 0 | 1 | 42.188 | 1 | 42 | 1 | GJB_Rinse | 0
191 | P-1333 | 2022/10/04 16:55:00 | 2-42 | 3 | 3.25 | 3.082 | 16452 | -1 | 1 | TALL4-B-3 | 22-233 | H2O | 1664902806 | 0.443 | 1.435 | 73.284 | 0.002 | 0.002 | 1.756 | 2485.549 | -46.421 | 2.674 | 95.518 | 0.877 | -2.688 | 0 | 0 | 1 | 42.25 | 1 | 42 | 1 | GJB_Rinse | 0
192 | P-1333 | 2022/10/04 17:03:56 | 2-42 | 4 | 3.232 | 3.06 | 16426 | 0 | 1 | TALL4-B-3 | 22-233 | H2O | 1664903342 | 0.478 | 1.559 | 72.713 | 0.001 | 0.005 | 1.745 | 2487.574 | -47.054 | 2.657 | 95.36 | 0.877 | -2.848 | 0 | 0 | 1 | 42.312 | 1 | 42 | 1 | GJB_Rinse | 0
193 | P-1334 | 2022/10/04 17:13:42 | 2-43 | 1 | -0.908 | -12.754 | 17561 | -1 | 1 | TALL4-deep-3 | 22-233 | H2O | 1664903928 | 0.322 | 0.793 | 73.482 | 0 | 0.001 | 1.757 | 1310.491 | 0.882 | 2.813 | 108.391 | 0.865 | 0.231 | 0 | 0 | 1 | 42.375 | 1 | 43 | 1 | GJB_Rinse | 0
194 | P-1334 | 2022/10/04 17:22:39 | 2-43 | 2 | -0.931 | -13.562 | 18053 | -1 | 1 | TALL4-deep-3 | 22-233 | H2O | 1664904465 | 0.29 | 0.715 | 76.036 | 0 | 0.002 | 1.803 | 1302.164 | 1.546 | 2.909 | 110.793 | 0.867 | 0.24 | 0 | 0 | 1 | 42.438 | 1 | 43 | 1 | GJB_Rinse | 0
195 | P-1334 | 2022/10/04 17:31:35 | 2-43 | 3 | -1 | -14.201 | 17749 | -1 | 1 | TALL4-deep-3 | 22-233 | H2O | 1664905001 | 0.296 | 0.705 | 72.312 | 0 | -0.001 | 1.714 | 1297.768 | 0.854 | 2.818 | 109.425 | 0.867 | 0.104 | 0 | 0 | 1 | 42.438 | 1 | 43 | 1 | GJB_Rinse | 0
196 | P-1334 | 2022/10/04 17:40:32 | 2-43 | 4 | -1.055 | -14.566 | 17329 | 0 | 1 | TALL4-deep-3 | 22-233 | H2O | 1664905539 | 0.32 | 0.818 | 74.147 | 0 | 0.003 | 1.769 | 1292.365 | 1.328 | 2.831 | 108.57 | 0.861 | 0.471 | 0 | 0 | 1 | 42.5 | 1 | 43 | 1 | GJB_Rinse | 0
197 | P-1335 | 2022/10/04 17:50:18 | 2-44 | 1 | 1.981 | 6.986 | 17880 | -1 | 1 | TALL4-010-1 | 22-233 | H2O | 1664906125 | 0.311 | 0.785 | 74.935 | 0.002 | -0.002 | 1.779 | 1295.807 | 0.922 | 2.877 | 111.839 | 0.867 | 0.362 | 0 | 0 | 1 | 42.438 | 1 | 44 | 1 | GJB_Rinse | 0
198 | P-1335 | 2022/10/04 17:59:15 | 2-44 | 2 | 2.1 | 8.691 | 17730 | -1 | 1 | TALL4-010-1 | 22-233 | H2O | 1664906662 | 0.334 | 0.775 | 78.125 | 0.002 | 0.001 | 1.863 | 1294.752 | 1.292 | 2.856 | 109.098 | 0.867 | 0.277 | 0 | 0 | 1 | 42.5 | 1 | 44 | 1 | GJB_Rinse | 0
199 | P-1335 | 2022/10/04 18:08:11 | 2-44 | 3 | 2.095 | 8.95 | 17886 | -1 | 1 | TALL4-010-1 | 22-233 | H2O | 1664907198 | 0.326 | 0.775 | 75.082 | 0 | 0.001 | 1.784 | 1295.773 | 1.363 | 2.876 | 110.023 | 0.865 | 0.269 | 0 | 0 | 1 | 42.5 | 1 | 44 | 1 | GJB_Rinse | 0
200 | P-1335 | 2022/10/04 18:17:08 | 2-44 | 4 | 2.145 | 9.161 | 17650 | 0 | 1 | TALL4-010-1 | 22-233 | H2O | 1664907734 | 0.277 | 0.833 | 75.573 | 0 | 0.002 | 1.811 | 1293.685 | 1.195 | 2.851 | 109.624 | 0.864 | 0.328 | 0 | 0 | 1 | 42.438 | 1 | 44 | 1 | GJB_Rinse | 0
201 | P-1336 | 2022/10/04 18:26:54 | 2-45 | 1 | 0.892 | -2.363 | 17750 | -1 | 1 | TALL4-B-5 | 22-233 | H2O | 1664908320 | 0.318 | 0.729 | 74.322 | -0.001 | 0 | 1.768 | 1299.48 | 0.22 | 2.851 | 110.467 | 0.869 | 0.007 | 0 | 0 | 1 | 42.375 | 1 | 45 | 1 | GJB_Rinse | 0
202 | P-1336 | 2022/10/04 18:35:51 | 2-45 | 2 | 0.795 | -3.728 | 17283 | -1 | 1 | TALL4-B-5 | 22-233 | H2O | 1664908857 | 0.318 | 0.806 | 75.409 | 0.001 | 0.002 | 1.797 | 1296.153 | 0.588 | 2.715 | 106.602 | 0.862 | -0.013 | 0 | 0 | 1 | 42.5 | 1 | 45 | 1 | GJB_Rinse | 0
203 | P-1336 | 2022/10/04 18:44:48 | 2-45 | 3 | 0.806 | -3.685 | 17643 | -1 | 1 | TALL4-B-5 | 22-233 | H2O | 1664909395 | 0.329 | 0.785 | 74.194 | 0 | 0.001 | 1.772 | 1298.822 | 0.021 | 2.871 | 110.79 | 0.861 | 0.152 | 0 | 0 | 1 | 42.438 | 1 | 45 | 1 | GJB_Rinse | 0
204 | P-1336 | 2022/10/04 18:53:45 | 2-45 | 4 | 0.753 | -4.124 | 17301 | 0 | 1 | TALL4-B-5 | 22-233 | H2O | 1664909932 | 0.297 | 0.719 | 73.839 | 0 | -0.001 | 1.769 | 1295.897 | 0.887 | 2.786 | 107.534 | 0.863 | 0.124 | 0 | 0 | 1 | 42.312 | 1 | 45 | 1 | GJB_Rinse | 0
205 | P-1337 | 2022/10/04 19:03:31 | 2-46 | 1 | 0.466 | -15.791 | 17317 | -1 | 1 | TALL4-D-4 | 22-233 | H2O | 1664910517 | 0.425 | 1.244 | 73.377 | 0.002 | 0.005 | 1.756 | 2189.929 | -32.583 | 2.759 | 101.059 | 0.874 | -1.031 | 0 | 0 | 1 | 42.375 | 1 | 46 | 1 | GJB_Rinse | 0
206 | P-1337 | 2022/10/04 19:12:28 | 2-46 | 2 | 0.478 | -17.534 | 17186 | -1 | 1 | TALL4-D-4 | 22-233 | H2O | 1664911054 | 0.455 | 1.225 | 75.392 | 0.001 | 0.002 | 1.797 | 2192.944 | -32.65 | 2.784 | 100.678 | 0.877 | -0.956 | 0 | 0 | 1 | 42.5 | 1 | 46 | 1 | GJB_Rinse | 0
207 | P-1337 | 2022/10/04 19:21:24 | 2-46 | 3 | 0.358 | -17.825 | 17111 | -1 | 1 | TALL4-D-4 | 22-233 | H2O | 1664911591 | 0.43 | 1.317 | 74.531 | 0.001 | 0.005 | 1.777 | 2188.912 | -32.471 | 2.752 | 102.085 | 0.872 | -0.716 | 0 | 0 | 1 | 42.5 | 1 | 46 | 1 | GJB_Rinse | 0
208 | P-1337 | 2022/10/04 19:30:21 | 2-46 | 4 | 0.361 | -17.859 | 17159 | 0 | 1 | TALL4-D-4 | 22-233 | H2O | 1664912127 | 0.419 | 1.295 | 72.913 | 0.002 | 0.004 | 1.747 | 2194.105 | -32.982 | 2.787 | 100.888 | 0.879 | -1.025 | 0 | 0 | 1 | 42.5 | 1 | 46 | 1 | GJB_Rinse | 0
209 | P-1338 | 2022/10/04 19:40:08 | 2-47 | 1 | 2.54 | -8.366 | 15814 | -1 | 1 | TALL4-A-3 | 22-233 | H2O | 1664912714 | 0.303 | 0.814 | 71.96 | 0 | -0.002 | 1.746 | 1301.747 | -3.104 | 2.507 | 100.938 | 0.864 | -1.059 | 0 | 0 | 1 | 42.438 | 1 | 47 | 1 | GJB_Rinse | 0
210 | P-1338 | 2022/10/04 19:49:05 | 2-47 | 2 | 2.695 | -6.372 | 17245 | -1 | 1 | TALL4-A-3 | 22-233 | H2O | 1664913251 | 0.299 | 0.773 | 73.495 | 0 | 0.003 | 1.762 | 1303.79 | -3.079 | 2.782 | 110.486 | 0.868 | -0.91 | 0 | 0 | 1 | 42.5 | 1 | 47 | 1 | GJB_Rinse | 0
211 | P-1338 | 2022/10/04 19:58:01 | 2-47 | 3 | 2.702 | -6.445 | 16946 | -1 | 1 | TALL4-A-3 | 22-233 | H2O | 1664913788 | 0.279 | 0.78 | 71.856 | 0 | 0.002 | 1.723 | 1299.291 | -2.816 | 2.689 | 108.893 | 0.863 | -0.838 | 0 | 0 | 1 | 42.5 | 1 | 47 | 1 | GJB_Rinse | 0
212 | P-1338 | 2022/10/04 20:06:58 | 2-47 | 4 | 2.777 | -6.105 | 17339 | 0 | 1 | TALL4-A-3 | 22-233 | H2O | 1664914325 | 0.304 | 0.716 | 75.104 | 0.001 | 0.001 | 1.777 | 1301.073 | -2.162 | 2.815 | 111.072 | 0.869 | -0.728 | 0 | 0 | 1 | 42.438 | 1 | 47 | 1 | GJB_Rinse | 0
213 | P-1339 | 2022/10/04 20:16:45 | 2-48 | 1 | 0.581 | -0.75 | 17025 | -1 | 1 | TALL4-2030-2 | 22-233 | H2O | 1664914911 | 0.331 | 0.725 | 71.454 | 0.001 | 0 | 1.706 | 1289.604 | 1.36 | 2.666 | 104.964 | 0.867 | 0.306 | 0 | 0 | 1 | 42.438 | 1 | 48 | 1 | GJB_Rinse | 0
214 | P-1339 | 2022/10/04 20:25:42 | 2-48 | 2 | 0.573 | 0.853 | 17777 | -1 | 1 | TALL4-2030-2 | 22-233 | H2O | 1664915448 | 0.301 | 0.8 | 74.048 | 0 | 0.001 | 1.766 | 1294.886 | 1.267 | 2.882 | 110.539 | 0.872 | 0.372 | 0 | 0 | 1 | 42.438 | 1 | 48 | 1 | GJB_Rinse | 0
215 | P-1339 | 2022/10/04 20:34:38 | 2-48 | 3 | 0.587 | 0.77 | 17942 | -1 | 1 | TALL4-2030-2 | 22-233 | H2O | 1664915985 | 0.329 | 0.705 | 75.753 | -0.001 | -0.001 | 1.795 | 1295.765 | 1.703 | 2.919 | 111.096 | 0.864 | 0.496 | 0 | 0 | 1 | 42.438 | 1 | 48 | 1 | GJB_Rinse | 0
216 | P-1339 | 2022/10/04 20:43:34 | 2-48 | 4 | 0.631 | 1.284 | 17747 | 0 | 1 | TALL4-2030-2 | 22-233 | H2O | 1664916521 | 0.341 | 0.671 | 77.216 | 0.002 | 0.002 | 1.836 | 1294.267 | 1.225 | 2.862 | 111.106 | 0.865 | 0.516 | 0 | 0 | 1 | 42.438 | 1 | 48 | 1 | GJB_Rinse | 0
217 | P-1340 | 2022/10/04 20:53:21 | 2-49 | 1 | -7.047 | -52.158 | 17448 | -1 | 1 | EV | standard | H2O | 1664917108 | 0.319 | 0.792 | 74.918 | 0.002 | 0.005 | 1.796 | 1291.151 | 2.064 | 2.815 | 108.055 | 0.863 | 0.491 | 0 | 0 | 1 | 42.438 | 1 | 49 | 1 | GJB_Rinse | 0
218 | P-1340 | 2022/10/04 21:02:18 | 2-49 | 2 | -7.414 | -57.07 | 17712 | -1 | 1 | EV | standard | H2O | 1664917645 | 0.316 | 0.838 | 77.432 | 0.001 | 0.004 | 1.831 | 1293.374 | 1.899 | 2.863 | 109.293 | 0.865 | 0.443 | 0 | 0 | 1 | 42.438 | 1 | 49 | 1 | GJB_Rinse | 0
219 | P-1340 | 2022/10/04 21:11:15 | 2-49 | 3 | -7.478 | -58.167 | 17710 | -1 | 1 | EV | standard | H2O | 1664918181 | 0.289 | 0.808 | 73.457 | 0.001 | 0.005 | 1.752 | 1293.156 | 1.718 | 2.866 | 110.493 | 0.868 | 0.568 | 0 | 0 | 1 | 42.375 | 1 | 49 | 1 | GJB_Rinse | 0
220 | P-1340 | 2022/10/04 21:20:11 | 2-49 | 4 | -7.536 | -58.724 | 17827 | 0 | 1 | EV | standard | H2O | 1664918718 | 0.287 | 0.821 | 75.429 | 0.001 | 0.006 | 1.799 | 1294.443 | 1.53 | 2.831 | 109.425 | 0.867 | 0.299 | 0 | 0 | 1 | 42.375 | 1 | 49 | 1 | GJB_Rinse | 0
221 | P-1341 | 2022/10/04 21:29:56 | 2-50 | 1 | 0.771 | -4.059 | 17433 | -1 | 1 | TALL4-1020-1 | 22-233 | H2O | 1664919303 | 0.293 | 0.695 | 76.515 | -0.001 | -0.001 | 1.828 | 1291.316 | 1.428 | 2.796 | 107.514 | 0.866 | 0.373 | 0 | 0 | 1 | 42.375 | 1 | 50 | 1 | GJB_Rinse | 0
222 | P-1341 | 2022/10/04 21:38:53 | 2-50 | 2 | 1.125 | 0.738 | 17397 | -1 | 1 | TALL4-1020-1 | 22-233 | H2O | 1664919839 | 0.287 | 0.712 | 76.677 | 0 | -0.001 | 1.834 | 1291.182 | 1.241 | 2.792 | 107.933 | 0.871 | 0.242 | 0 | 0 | 1 | 42.375 | 1 | 50 | 1 | GJB_Rinse | 0
223 | P-1341 | 2022/10/04 21:47:49 | 2-50 | 3 | 1.225 | 1.914 | 17263 | -1 | 1 | TALL4-1020-1 | 22-233 | H2O | 1664920376 | 0.289 | 0.791 | 73.151 | 0.001 | -0.001 | 1.753 | 1290.065 | 1.335 | 2.818 | 107.544 | 0.868 | 0.334 | 0 | 0 | 1 | 42.375 | 1 | 50 | 1 | GJB_Rinse | 0
224 | P-1341 | 2022/10/04 21:56:46 | 2-50 | 4 | 1.294 | 2.43 | 17206 | 0 | 1 | TALL4-1020-1 | 22-233 | H2O | 1664920913 | 0.325 | 0.728 | 75.054 | 0 | 0.002 | 1.798 | 1289.368 | 1.706 | 2.815 | 107.654 | 0.869 | 0.638 | 0 | 0 | 1 | 42.438 | 1 | 50 | 1 | GJB_Rinse | 0
225 | P-1342 | 2022/10/04 22:06:32 | 2-51 | 1 | -7.03 | -51.715 | 17754 | -1 | 1 | EV | standard | H2O | 1664921499 | 0.336 | 0.779 | 73.666 | 0.001 | 0.004 | 1.762 | 1293.216 | 1.803 | 2.817 | 109.65 | 0.87 | 0.446 | 0 | 0 | 1 | 42.5 | 1 | 51 | 1 | GJB_Rinse | 0
226 | P-1342 | 2022/10/04 22:15:29 | 2-51 | 2 | -7.398 | -56.879 | 17811 | -1 | 1 | EV | standard | H2O | 1664922035 | 0.344 | 0.806 | 74.908 | 0.002 | 0.008 | 1.774 | 1293.955 | 1.779 | 2.777 | 108.499 | 0.868 | 0.219 | 0 | 0 | 1 | 42.5 | 1 | 51 | 1 | GJB_Rinse | 0
227 | P-1342 | 2022/10/04 22:24:25 | 2-51 | 3 | -7.508 | -58.282 | 17791 | -1 | 1 | EV | standard | H2O | 1664922572 | 0.283 | 0.731 | 74.175 | 0.001 | 0.006 | 1.768 | 1293.737 | 1.531 | 2.842 | 109.353 | 0.863 | 0.275 | 0 | 0 | 1 | 42.438 | 1 | 51 | 1 | GJB_Rinse | 0
228 | P-1342 | 2022/10/04 22:33:22 | 2-51 | 4 | -7.491 | -58.736 | 17829 | 0 | 1 | EV | standard | H2O | 1664923109 | 0.303 | 0.707 | 77.75 | 0.001 | 0.003 | 1.835 | 1293.734 | 2.257 | 2.847 | 109.343 | 0.866 | 0.407 | 0 | 0 | 1 | 42.438 | 1 | 51 | 1 | GJB_Rinse | 0

## Instructions
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Variable | Instruction | Comment
StandardsFile | standards file.csv
OutputBase | ChemCorrect
CorrectionType | M
timecode | Time Code
identifier | Identifier 1
sample | Sample
injection | Inj Nr | Inj Nr
line | Line
job | Job
Variable | Instruction | Comment
baseline_shift_MEAN | mean_standards(baseline_shift) | The mean of the organic baseline of the standards.
baseline_shift_STD | std_standards(baseline_shift) | The standard deviation of the organic baseline of the standards.
ch4_ppm_MEAN | mean_standards(ch4_ppm) | The mean methane concentration of the standards.
ch4_ppm_STD | std_standards(ch4_ppm) | The standard deviation of the methane concentrations of the standards.
residuals_MEAN | mean_standards(residuals) | The mean difference between the measured and modeled spectra of the standards.
residuals_STD | std_standards(residuals) | The standard deviation of the difference between measured and modeled spectra of the standards.
slope_shift_MEAN | mean_standards(slope_shift) | The mean organic slope of the standards.
slope_shift_STD | std_standards(slope_shift) | The standard deviation of the organic slope of the standards.
baseline_curvature_MEAN | mean_standards(baseline_curvature) | The mean of the baseline curvature of the standards.
baseline_curvature_STD | std_standards(baseline_curvature) | The standard deviation of the baseline curvature of the standards.
residuals_MEAN_ABS | abs(residuals_MEAN) | The absolute value of the mean difference between the measured and modeled spectra of the standards.
baseline_shift_MEAN_ABS | abs(baseline_shift_MEAN) | The absolute value of the mean of the organic baseline of the standards.
ch4_ppm_MEAN_ABS | abs(ch4_ppm_MEAN) | The absolute value of the mean methane concentration of the standards.
slope_shift_MEAN_ABS | abs(slope_shift_MEAN) | The absolute value of the mean organic slope of the standards.
residuals_15SIGMA | evaluation(residuals_STD*1.5) | One and half times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_15SIGMA | evaluation(residuals_MEAN+residuals_15SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_2SIGMA | evaluation(residuals_STD*2) | Twice the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_2SIGMA | evaluation(residuals_MEAN+residuals_2SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_3SIGMA | evaluation(residuals_STD*3) | Three times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_3SIGMA | evaluation(residuals_MEAN+residuals_3SIGMA) | The mean difference between measured and modeled spectra of the standards plus three standard deviation.
residuals_10PCT | evaluation(residuals_MEAN*0.1) | 30% of the difference between measured and modeled spectra of the standards.
ch4_ppm_3SIGMA | evaluation(ch4_ppm_STD*3) | Three times the standard deviation of the methanol amplitudes of the standards.
ch4_ppm_PLUS_3SIGMA | evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA) | The mean methanol amplitude of the standards plus three standard deviation.
HALF | float(0.5) | A fixed value of 0.5 or half.
baseline_shift_HALFSIGMA | evaluation(baseline_shift_STD*HALF) | Half the standard deviation of the organic baseline of the standards.
baseline_shift_PLUS_HALFSIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA) | The absolute value of the mean organic baseline of the standards plus one-half standard deviation.
baseline_shift_PLUS_1SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD) | The absolute value of the mean organic baseline of the standards plus one standard deviation.
baseline_shift_3SIGMA | evaluation(baseline_shift_STD*3) | Three times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_3SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA) | The absolute value of the mean organic baseline of the standards plus three standard deviation.
baseline_shift_18SIGMA | evaluation(baseline_shift_STD*18) | Eighteen times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_18SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA) | The absolute value of the mean organic baseline of the standards plus eighteen standard deviation.
slope_shift_5SIGMA | evaluation(slope_shift_STD*5) | Five times the standard deviation of the organic slope of the standards.
slope_shift_PLUS_5SIGMA | evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA) | The absolute value of the mean organic slope of the standards plus five standard deviation.
baseline_curvature_3SIGMA | evaluation(baseline_curvature_STD*3) | Three times the standard deviation of the baseline curvature of the standards.
baseline_curvature_PLUS_3SIGMA | evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
TRUE | evaluation(1>0)
FALSE | evaluation(1<0)
PROC_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the uncalibrated d18O value in order to calibrate it.
PROC_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the uncalibrated dD value in order to calibrate it.
Variable | Instruction | Comment
IS_STANDARD | is_standard_sample() | Standard recognition.
IS_NOT_STANDARD | bool_xor(IS_STANDARD,TRUE)
SAMP_residuals_MEAN | mean_all(residuals) | The mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_residuals_STD | std_all(residuals) | One standard deviation of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN | mean_all(ch4_ppm) | The mean methane concentration of a sample.
SAMP_ch4_ppm_STD | std_all(ch4_ppm) | The standard deviation of the methane concentrations of a sample.
SAMP_baseline_shift_MEAN | mean_all(baseline_shift) | The mean of the organic baseline of a sample.
SAMP_baseline_shift_STD | std_all(baseline_shift) | The standard deviation of the organic baseline of a sample.
SAMP_slope_shift_MEAN | mean_all(slope_shift) | The mean organic slope of a sample.
SAMP_slope_shift_STD | std_all(slope_shift) | The standard deviation of the organic slope of a sample.
SAMP_d18O_STD | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
SAMP_dD_STD | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_baseline_curvature_MEAN | mean_all(baseline_curvature) | The mean of the baseline curvature of the ssamples
SAMP_baseline_curvature_STD | std_all(baseline_curvature) | The standard deviation of the baseline curvature of the samples
SAMP_residuals_MEAN_ABS | abs(SAMP_residuals_MEAN) | Absolute value of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN_ABS | abs(SAMP_ch4_ppm_MEAN) | Absolute value of the mean methane concentration of a sample.
SAMP_baseline_shift_MEAN_ABS | abs(SAMP_baseline_shift_MEAN) | Absolute value of the mean organic baseline of a sample.
SAMP_slope_shift_MEAN_ABS | abs(SAMP_slope_shift_MEAN) | Absolute value of the mean organic slope of a sample.
SAMP_residuals_DIFFERENTIAL | evaluation(SAMP_residuals_MEAN-residuals_MEAN)
SAMP_residuals_DIFFERENTIAL_ABS | abs(SAMP_residuals_DIFFERENTIAL) | Absolute value of the difference between the mean of the sample residual and the mean of all standards residuals
SAMP_residuals_DIFFERENTIAL_PCT | evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
SAMP_baseline_curvature_DIFFERENTIAL | evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
SAMP_STD_DELTA_18_16O_CURR_MEAN | current_standards_mean(d(18_16)Mean) | The d18O mean of the last run standard.
SAMP_STD_DELTA_DH_CURR_MEAN | current_standards_mean(d(D_H)Mean) | The dD mean of the last run standard.
known_dD | known_standard(d(D_H)Mean) | The dD value of the known Standard.
known_d18O | known_standard(d(18_16)Mean) | The d18O value of the known Standard.
d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
measured_d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
measured_dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
measured_d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
measured_dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the mean uncalibrated d18O value in order to calibrate it.
SAMP_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the mean uncalibrated dD value in order to calibrate it.
SAMP_TEST_VALUE | evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
residuals_GT_3SIGMA | evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA) | Threshold of red trigger. If the mean difference between the measured spectra of the sample and the modeled spectra of the standard is greater than that of the mean difference of the standard plus three standard deviations, this will trigger.
residuals_GE_2SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GE_15SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GT_10PCT | evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
ch4_ppm_GE_3SIGMA | evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA) | True if the mean methane concentration of the sample is greater than or equal to that of the standard plus three standard deviations, otherwise false.
baseline_shift_GT_1SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus one standard deviation, otherwise false.
baseline_shift_GT_3SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus three standard deviation, otherwise false. Sets yellow trigger for "other", trigger threshold is approximately sample to sample dD precision specification
baseline_shift_GT_18SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus eighteen standard deviation, otherwise false. Sets red trigger for "other", trigger threshold is approximately 24 hour dD drift specification
slope_shift_GT_5SIGMA | evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA) | True if the mean organic slope of the sample is greater than that of the standard plus five standard deviations, otherwise false.
baseline_curvature_GT_3SIGMA | evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
RED | truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA)) | Red trigger. If the sample is not a standard and the threshold is triggered (measured-modeled > standard+3stdev's), then RED.
NOT_RED | bool_xor(RED,TRUE) | If the sample is determined to be "red" than it can not be "not red."
YELLOW | truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA))) | Truth evaluation used to determine whether or not a given sample is yellow.
NOT_YELLOW | bool_xor(YELLOW,TRUE) | If sample is determined to be "yellow" than it can not be "not yellow."
GREEN | truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW)) | This is the pre-requisite for a sample to be colored green. It must not be a standard, must not be yellow, and must not be red.
NOT_GREEN | bool_xor(GREEN,TRUE) | Pre-requisite for being "not-green."
CYAN | truth_evaluation((IS_STANDARD)) | This is the pre-requisite for a sample to be a standard and therefore cyan colored.
NOT_CYAN | bool_xor(CYAN,TRUE) | Pre-requisite for being not a standard.
METHANE | truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA)) | This is the threshold for amounts of methane being included in the main ChemCorrect spreadsheet. Methane will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured methane concentration must be greater than or equal to that of the standard plus three standard deviations.
OTHER1 | truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA)) | This is the threshold for amounts of "OTHER1" being included in the main ChemCorrect spreadsheet. This column (OTHER1) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic baseline must be greater than that of the standard plus one standard deviation.
OTHER2 | truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA)) | This is the threshold for amounts of "OTHER2" being included in the main ChemCorrect spreadsheet. This column (OTHER2) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic slope must be greater than that of the standard plus three standard deviations.
OTHER3 | truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | truth_evaluation(RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN)) | Organic baseline differential showing the difference between the measured baseline and the standard baseline.
Variable | Instruction | Comment
DTL_TEST | math_evaluation((1 + 2)) | Test evaluation of 1+2 being equal to 3.
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample | FALSE | FALSE | FALSE | FALSE | Testing Sample
B | identifier | TRUE | Name | RED | YELLOW | GREEN | CYAN | Testing Name
C | SAMP_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d<sup>18</sup>O Mean | RED | YELLOW | GREEN | CYAN | Testing18 16 Mean
D | SAMP_DELTA_DH_CALIBRATE | TRUE | Calibrated d<sup>2</sup>H Mean | RED | YELLOW | GREEN | CYAN | This is a test to see that I can use tool tips
E | SAMP_ch4_ppm_MEAN | METHANE | CH<sub>4</sub> | RED | YELLOW | FALSE | FALSE
F | baseline_shift_DIFF | OTHER1 | C<sub>2+</sub> alcohols | RED | YELLOW | FALSE | FALSE
G | SAMP_residuals_DIFFERENTIAL_PCT | OTHER3 | relative %deviation | RED | YELLOW | FALSE | FALSE
H | SAMP_d18O_STD | TRUE | Uncalibrated d<sup>18</sup>O precision | RED | YELLOW | GREEN | CYAN
I | SAMP_dD_STD | TRUE | Uncalibrated d<sup>2</sup>H precision | RED | YELLOW | GREEN | CYAN
J | SAMP_slope_shift_MEAN | OTHER2 | slope | RED | YELLOW | FALSE | FALSE
K | SAMP_baseline_curvature_DIFFERENTIAL | OTHER4 | curvature | RED | YELLOW | FALSE | FALSE
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample
B | identifier | TRUE | Name
C | col(d(18_16)Mean) | TRUE | Measured d(18_16) Mean
D | col(d(D_H)Mean) | TRUE | Measured d(D_H) Mean
E | DTL_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d(18_16) Mean
F | DTL_DELTA_DH_CALIBRATE | TRUE | Calibrated d(D_H) Mean
G | col(ch4_ppm) | METHANE | CH<sub>4</sub>
H | col(baseline_shift) | OTHER1 | Other
I | col(Time Code) | TRUE | Time Code
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | d18O_mean | TRUE | d(18_16) Mean | RED | YELLOW | GREEN | CYAN
B | dD_mean | TRUE | d(D_H) Mean | RED | YELLOW | GREEN | CYAN
C | XY_X(d18O_mean) XY_Y(dD_mean) | TRUE | d18O x dD | gmwl | best_fit

## Process Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Process Variables
Variable | Value | Operations
baseline_shift_MEAN | 1295.48095
varray: [1296.994, 1295.029, 1296.664, 1297.916, 1296.061, 1294.379, 1296.587, 1296.434, 1297.634, 1294.01, 1296.484, 1294.971, 1294.189, 1293.121, 1297.242, 1293.196, 1295.33, 1294.452, 1294.923, 1294.003]
inst: mean_standards
operation: mean_standards(baseline_shift)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0, 10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: baseline_shift
baseline_shift_STD | 1.46374600453
varray: [1296.994, 1295.029, 1296.664, 1297.916, 1296.061, 1294.379, 1296.587, 1296.434, 1297.634, 1294.01, 1296.484, 1294.971, 1294.189, 1293.121, 1297.242, 1293.196, 1295.33, 1294.452, 1294.923, 1294.003]
inst: std_standards
operation: std_standards(baseline_shift)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0, 10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: baseline_shift
ch4_ppm_MEAN | 0.3928
varray: [0.048, 0.269, 0.559, 0.29, 0.262, 0.704, 0.49, 0.388, 0.106, 0.435, 0.528, 0.427, 0.251, 0.792, 0.368, 0.244, 0.387, 0.334, 0.569, 0.405]
inst: mean_standards
operation: mean_standards(ch4_ppm)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0, 10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: ch4_ppm
ch4_ppm_STD | 0.182990249956
varray: [0.048, 0.269, 0.559, 0.29, 0.262, 0.704, 0.49, 0.388, 0.106, 0.435, 0.528, 0.427, 0.251, 0.792, 0.368, 0.244, 0.387, 0.334, 0.569, 0.405]
inst: std_standards
operation: std_standards(ch4_ppm)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0, 10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: ch4_ppm
residuals_MEAN | 2.89535
varray: [2.881, 2.915, 2.963, 2.996, 2.933, 2.96, 2.997, 2.891, 2.966, 2.89, 2.922, 2.861, 2.801, 2.898, 2.842, 2.805, 2.83, 2.833, 2.911, 2.812]
inst: mean_standards
operation: mean_standards(residuals)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0, 10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: residuals
residuals_STD | 0.0621655472361
varray: [2.881, 2.915, 2.963, 2.996, 2.933, 2.96, 2.997, 2.891, 2.966, 2.89, 2.922, 2.861, 2.801, 2.898, 2.842, 2.805, 2.83, 2.833, 2.911, 2.812]
inst: std_standards
operation: std_standards(residuals)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0, 10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: residuals
slope_shift_MEAN | 1.70645
varray: [1.929, 1.13, 1.563, 1.157, 1.217, 1.215, 1.236, 1.241, 1.211, 1.549, 2.257, 1.406, 2.533, 1.513, 2.473, 1.773, 2.607, 2.125, 1.717, 2.277]
inst: mean_standards
operation: mean_standards(slope_shift)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0, 10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: slope_shift
slope_shift_STD | 0.508425482301
varray: [1.929, 1.13, 1.563, 1.157, 1.217, 1.215, 1.236, 1.241, 1.211, 1.549, 2.257, 1.406, 2.533, 1.513, 2.473, 1.773, 2.607, 2.125, 1.717, 2.277]
inst: std_standards
operation: std_standards(slope_shift)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0, 10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: slope_shift
baseline_curvature_MEAN | 111.3598
varray: [110.905, 111.355, 112.65, 113.853, 111.685, 112.977, 113.49, 111.58, 112.135, 110.428, 112.128, 111.633, 108.345, 111.651, 112.434, 108.068, 110.239, 109.716, 112.245, 109.679]
inst: mean_standards
operation: mean_standards(baseline_curvature)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0, 10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: baseline_curvature
baseline_curvature_STD | 1.56035460207
varray: [110.905, 111.355, 112.65, 113.853, 111.685, 112.977, 113.49, 111.58, 112.135, 110.428, 112.128, 111.633, 108.345, 111.651, 112.434, 108.068, 110.239, 109.716, 112.245, 109.679]
inst: std_standards
operation: std_standards(baseline_curvature)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0, 10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: baseline_curvature
residuals_MEAN_ABS | 2.89535
operation: abs(residuals_MEAN)
inst: abs
source_col: residuals_MEAN
baseline_shift_MEAN_ABS | 1295.48095
operation: abs(baseline_shift_MEAN)
inst: abs
source_col: baseline_shift_MEAN
ch4_ppm_MEAN_ABS | 0.3928
operation: abs(ch4_ppm_MEAN)
inst: abs
source_col: ch4_ppm_MEAN
slope_shift_MEAN_ABS | 1.70645
operation: abs(slope_shift_MEAN)
inst: abs
source_col: slope_shift_MEAN
residuals_15SIGMA | 0.0932483208541
operation: evaluation(residuals_STD*1.5)
inst: evaluation
factors: residuals_STD*1.5
residuals_PLUS_15SIGMA | 2.98859832085
operation: evaluation(residuals_MEAN+residuals_15SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_15SIGMA
residuals_2SIGMA | 0.124331094472
operation: evaluation(residuals_STD*2)
inst: evaluation
factors: residuals_STD*2
residuals_PLUS_2SIGMA | 3.01968109447
operation: evaluation(residuals_MEAN+residuals_2SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_2SIGMA
residuals_3SIGMA | 0.186496641708
operation: evaluation(residuals_STD*3)
inst: evaluation
factors: residuals_STD*3
residuals_PLUS_3SIGMA | 3.08184664171
operation: evaluation(residuals_MEAN+residuals_3SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_3SIGMA
residuals_10PCT | 0.289535
operation: evaluation(residuals_MEAN*0.1)
inst: evaluation
factors: residuals_MEAN*0.1
ch4_ppm_3SIGMA | 0.548970749868
operation: evaluation(ch4_ppm_STD*3)
inst: evaluation
factors: ch4_ppm_STD*3
ch4_ppm_PLUS_3SIGMA | 0.941770749868
operation: evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA)
inst: evaluation
factors: ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA
HALF | 0.5
operation: float(0.5)
inst: float
baseline_shift_HALFSIGMA | 0.731873002267
operation: evaluation(baseline_shift_STD*HALF)
inst: evaluation
factors: baseline_shift_STD*HALF
baseline_shift_PLUS_HALFSIGMA | 1296.212823
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA
baseline_shift_PLUS_1SIGMA | 1296.944696
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_STD
baseline_shift_3SIGMA | 4.3912380136
operation: evaluation(baseline_shift_STD*3)
inst: evaluation
factors: baseline_shift_STD*3
baseline_shift_PLUS_3SIGMA | 1299.87218801
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_3SIGMA
baseline_shift_18SIGMA | 26.3474280816
operation: evaluation(baseline_shift_STD*18)
inst: evaluation
factors: baseline_shift_STD*18
baseline_shift_PLUS_18SIGMA | 1321.82837808
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_18SIGMA
slope_shift_5SIGMA | 2.5421274115
operation: evaluation(slope_shift_STD*5)
inst: evaluation
factors: slope_shift_STD*5
slope_shift_PLUS_5SIGMA | 4.2485774115
operation: evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA)
inst: evaluation
factors: slope_shift_MEAN_ABS+slope_shift_5SIGMA
baseline_curvature_3SIGMA | 4.68106380622
operation: evaluation(baseline_curvature_STD*3)
inst: evaluation
factors: baseline_curvature_STD*3
baseline_curvature_PLUS_3SIGMA | 116.040863806
operation: evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
inst: evaluation
factors: baseline_curvature_3SIGMA+baseline_curvature_MEAN
TRUE | True
operation: evaluation(1>0)
inst: evaluation
factors: 1>0
FALSE | False
operation: evaluation(1<0)
inst: evaluation
factors: 1<0
PROC_DELTA_18_16O_CALIBRATE | -5.81110865409
actual_array: [ 1.93 -15.84]
varray: [-7.995, -7.635, -7.553, -7.494, -7.464, -7.417, -7.498, -7.429, -7.447, -7.463, 3.939, 4.421, 4.625, 4.658, 4.724, 4.653, 4.781, 4.725, 4.792, 4.753, -12.197, -13.002, -13.187, -13.345, -13.214, -13.363, -13.338, -13.4, -13.416, -13.403, -8.005, -7.739, -7.69, -7.661, -7.628, -7.602, -7.632, -7.616, -7.609, -7.605, -14.375, -14.684, -14.757, -14.742, -14.788, -14.843, -14.844, -14.812, -14.742, -14.426, -13.684, -14.61, -14.76, -14.746, -14.845, -14.895, -13.757, -13.65, -13.627, -13.614, -14.732, -14.73, -14.779, -14.758, -14.658, -14.725, -14.778, -14.671, 3.951, 4.605, 4.844, 4.979, -7.193, -7.569, -7.543, -7.597, 2.562, 2.971, 3.053, 3.116, 2.107, 2.147, 2.147, 2.111, 1.122, 1.156, 1.056, 1.122, -0.791, -0.918, -0.792, -0.826, -1.414, -1.398, -1.445, -1.454, 0.043, 0.264, 0.193, 0.167, -0.396, -0.43, -0.425, -0.409, 0.926, 0.933, 0.923, 0.944, -7.047, -7.333, -7.422, -7.389, -0.502, -0.179, -0.075, -0.078, 0.761, 0.735, 0.77, 0.785, 0.619, 0.565, 0.586, 0.647, 1.774, 1.683, 1.691, 1.645, 0.458, 0.357, 0.411, 0.332, 0.0, 0.889, 0.889, 0.889, 0.926, 0.897, 0.885, 0.924, 2.647, 2.732, 2.74, 2.803, -6.96, -7.51, -7.454, -7.512, 0.837, 1.199, 1.259, 1.352, -1.201, -1.335, -1.306, -1.329, -1.224, -1.379, -1.453, -1.354, 1.116, 1.442, 1.337, 1.4, -1.225, -1.305, -1.382, -1.405, -0.321, -0.262, -0.264, -0.273, 0.426, 0.46, 0.425, 0.44, 0.53, 0.533, 0.645, 0.593, -7.095, -7.389, -7.467, -7.497, -0.02, 0.387, 0.496, 0.438, 3.227, 3.216, 3.25, 3.232, -0.908, -0.931, -1.0, -1.055, 1.981, 2.1, 2.095, 2.145, 0.892, 0.795, 0.806, 0.753, 0.466, 0.478, 0.358, 0.361, 2.54, 2.695, 2.702, 2.777, 0.581, 0.573, 0.587, 0.631, -7.047, -7.414, -7.478, -7.536, 0.771, 1.125, 1.225, 1.294, -7.03, -7.398, -7.508, -7.491]
fit_rsq: 1.0
fit_intercept: 2.6745368036
except: True
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -3.14428947368
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [1.0, 538.0, 1075.0, 1611.0, 2148.0, 2685.0, 3222.0, 3758.0, 4295.0, 4832.0, 5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0, 10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0, 16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21092.0, 21682.0, 22219.0, 22756.0, 23292.0, 23882.0, 24419.0, 24956.0, 25493.0, 26084.0, 26620.0, 27157.0, 27694.0, 28282.0, 28818.0, 29355.0, 29892.0, 30480.0, 31017.0, 31554.0, 32091.0, 32679.0, 33215.0, 33752.0, 34289.0, 34877.0, 35414.0, 35951.0, 36488.0, 37077.0, 37614.0, 38151.0, 38688.0, 39277.0, 39813.0, 40350.0, 40887.0, 41477.0, 42014.0, 42551.0, 43087.0, 43674.0, 44211.0, 44747.0, 45284.0, 45871.0, 46407.0, 46944.0, 47481.0, 48069.0, 48605.0, 49141.0, 49678.0, 50266.0, 50802.0, 51339.0, 51876.0, 52464.0, 53000.0, 53536.0, 54073.0, 54662.0, 55198.0, 55735.0, 56272.0, 56861.0, 57397.0, 57934.0, 58470.0, 59055.0, 59592.0, 60129.0, 60665.0, 61251.0, 61788.0, 62324.0, 62861.0, 63448.0, 63985.0, 64521.0, 65058.0, 65645.0, 66182.0, 66718.0, 67255.0, 67843.0, 68380.0, 68917.0, 69453.0, 70042.0, 70578.0, 71115.0, 71651.0, 72239.0, 72775.0, 73312.0, 73850.0, 74435.0, 74972.0, 75509.0, 76046.0, 76631.0, 77168.0, 77705.0, 78241.0, 78827.0, 79364.0, 79901.0, 80438.0, 81024.0, 81561.0, 82098.0, 82635.0, 83221.0, 83758.0, 84294.0, 84831.0, 85417.0, 85953.0, 86490.0, 87026.0, 87614.0, 88151.0, 88688.0, 89225.0, 89808.0, 90345.0, 90881.0, 91418.0, 92003.0, 92540.0, 93077.0, 93614.0, 94198.0, 94736.0, 95272.0, 95809.0, 96395.0, 96932.0, 97469.0, 98005.0, 98591.0, 99128.0, 99664.0, 100201.0, 100787.0, 101324.0, 101861.0, 102397.0, 102983.0, 103519.0, 104056.0, 104592.0, 105178.0, 105715.0, 106251.0, 106788.0, 107374.0, 107911.0, 108447.0, 108984.0, 109570.0, 110107.0, 110644.0, 111181.0, 111767.0, 112304.0, 112840.0, 113377.0, 113964.0, 114501.0, 115037.0, 115574.0, 116161.0, 116698.0, 117234.0, 117770.0, 118357.0, 118894.0, 119431.0, 119967.0, 120552.0, 121089.0, 121625.0, 122162.0, 122748.0, 123285.0, 123821.0, 124358.0]
source_col: d(18_16)Mean
PROC_DELTA_DH_CALIBRATE | -48.124750468
actual_array: [ 18.1 -119. ]
varray: [-65.765, -61.473, -60.565, -60.185, -59.953, -59.925, -60.02, -59.898, -59.796, -59.74, 27.899, 36.041, 38.497, 39.122, 39.579, 39.6, 40.205, 40.48, 40.805, 40.427, -92.886, -105.698, -108.406, -110.117, -109.983, -110.818, -111.128, -111.395, -111.414, -111.634, -66.361, -61.962, -61.034, -60.792, -60.472, -60.476, -60.23, -60.175, -60.101, -60.237, -113.704, -118.569, -119.794, -120.13, -120.079, -120.316, -120.509, -120.725, -120.455, -119.345, -117.046, -120.266, -120.746, -120.691, -121.06, -121.381, -113.459, -112.405, -112.295, -112.082, -119.584, -120.273, -120.357, -120.685, -120.02, -120.195, -120.31, -119.867, -17.514, -8.337, -5.756, -4.701, -55.529, -59.379, -59.459, -60.051, -13.722, -9.62, -8.46, -7.868, -2.966, -2.109, -1.978, -1.688, 0.706, 1.296, 0.911, 1.286, -4.835, -5.616, -5.324, -5.197, -13.56, -13.956, -14.284, -14.411, 0.601, 1.905, 2.15, 2.324, -5.383, -5.521, -5.898, -6.172, -1.17, -0.831, -0.947, -0.805, -51.913, -57.301, -57.885, -58.618, -15.701, -12.088, -11.057, -10.671, -10.435, -10.418, -10.697, -10.893, -8.177, -8.442, -7.97, -8.159, -2.4, -2.303, -2.151, -2.578, -12.03, -13.412, -13.602, -13.677, 0.0, -3.167, -3.167, -3.167, -5.473, -5.966, -5.805, -6.072, -6.037, -6.208, -6.135, -6.186, -52.752, -57.563, -58.131, -59.151, -13.665, -9.517, -8.597, -8.175, -22.053, -23.571, -23.493, -23.638, -13.244, -13.724, -14.251, -13.587, 0.786, 3.445, 2.844, 3.238, -8.736, -9.957, -10.134, -10.339, -10.021, -9.715, -9.941, -10.102, -2.621, -2.221, -1.777, -1.79, -1.939, -1.665, -1.316, -1.66, -52.246, -56.759, -58.265, -58.513, -5.214, 0.345, 1.58, 1.596, 3.019, 2.803, 3.082, 3.06, -12.754, -13.562, -14.201, -14.566, 6.986, 8.691, 8.95, 9.161, -2.363, -3.728, -3.685, -4.124, -15.791, -17.534, -17.825, -17.859, -8.366, -6.372, -6.445, -6.105, -0.75, 0.853, 0.77, 1.284, -52.158, -57.07, -58.167, -58.724, -4.059, 0.738, 1.914, 2.43, -51.715, -56.879, -58.282, -58.736]
fit_rsq: 1.0
fit_intercept: 18.9095369074
except: True
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -32.5546008772
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [1.0, 538.0, 1075.0, 1611.0, 2148.0, 2685.0, 3222.0, 3758.0, 4295.0, 4832.0, 5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0, 10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0, 16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21092.0, 21682.0, 22219.0, 22756.0, 23292.0, 23882.0, 24419.0, 24956.0, 25493.0, 26084.0, 26620.0, 27157.0, 27694.0, 28282.0, 28818.0, 29355.0, 29892.0, 30480.0, 31017.0, 31554.0, 32091.0, 32679.0, 33215.0, 33752.0, 34289.0, 34877.0, 35414.0, 35951.0, 36488.0, 37077.0, 37614.0, 38151.0, 38688.0, 39277.0, 39813.0, 40350.0, 40887.0, 41477.0, 42014.0, 42551.0, 43087.0, 43674.0, 44211.0, 44747.0, 45284.0, 45871.0, 46407.0, 46944.0, 47481.0, 48069.0, 48605.0, 49141.0, 49678.0, 50266.0, 50802.0, 51339.0, 51876.0, 52464.0, 53000.0, 53536.0, 54073.0, 54662.0, 55198.0, 55735.0, 56272.0, 56861.0, 57397.0, 57934.0, 58470.0, 59055.0, 59592.0, 60129.0, 60665.0, 61251.0, 61788.0, 62324.0, 62861.0, 63448.0, 63985.0, 64521.0, 65058.0, 65645.0, 66182.0, 66718.0, 67255.0, 67843.0, 68380.0, 68917.0, 69453.0, 70042.0, 70578.0, 71115.0, 71651.0, 72239.0, 72775.0, 73312.0, 73850.0, 74435.0, 74972.0, 75509.0, 76046.0, 76631.0, 77168.0, 77705.0, 78241.0, 78827.0, 79364.0, 79901.0, 80438.0, 81024.0, 81561.0, 82098.0, 82635.0, 83221.0, 83758.0, 84294.0, 84831.0, 85417.0, 85953.0, 86490.0, 87026.0, 87614.0, 88151.0, 88688.0, 89225.0, 89808.0, 90345.0, 90881.0, 91418.0, 92003.0, 92540.0, 93077.0, 93614.0, 94198.0, 94736.0, 95272.0, 95809.0, 96395.0, 96932.0, 97469.0, 98005.0, 98591.0, 99128.0, 99664.0, 100201.0, 100787.0, 101324.0, 101861.0, 102397.0, 102983.0, 103519.0, 104056.0, 104592.0, 105178.0, 105715.0, 106251.0, 106788.0, 107374.0, 107911.0, 108447.0, 108984.0, 109570.0, 110107.0, 110644.0, 111181.0, 111767.0, 112304.0, 112840.0, 113377.0, 113964.0, 114501.0, 115037.0, 115574.0, 116161.0, 116698.0, 117234.0, 117770.0, 118357.0, 118894.0, 119431.0, 119967.0, 120552.0, 121089.0, 121625.0, 122162.0, 122748.0, 123285.0, 123821.0, 124358.0]
source_col: d(D_H)Mean

## 1 - Sample 1 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 1
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9844
varray: [3.172, 3.047, 3.013, 2.993, 2.962, 3.019, 2.949, 2.89, 2.886, 2.913]
inst: mean_all
operation: mean_all(residuals)
tarray: [1.0, 538.0, 1075.0, 1611.0, 2148.0, 2685.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: residuals
SAMP_residuals_STD | 0.086054504692
varray: [3.172, 3.047, 3.013, 2.993, 2.962, 3.019, 2.949, 2.89, 2.886, 2.913]
inst: std_all
operation: std_all(residuals)
tarray: [1.0, 538.0, 1075.0, 1611.0, 2148.0, 2685.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.5122
varray: [0.562, 0.782, 0.433, 0.496, 0.255, 0.756, 0.418, 0.325, 0.454, 0.641]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [1.0, 538.0, 1075.0, 1611.0, 2148.0, 2685.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.173618867382
varray: [0.562, 0.782, 0.433, 0.496, 0.255, 0.756, 0.418, 0.325, 0.454, 0.641]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [1.0, 538.0, 1075.0, 1611.0, 2148.0, 2685.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1305.3808
varray: [1317.118, 1311.236, 1309.752, 1309.495, 1307.261, 1305.221, 1301.843, 1299.231, 1297.083, 1295.568]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [1.0, 538.0, 1075.0, 1611.0, 2148.0, 2685.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 6.88896498919
varray: [1317.118, 1311.236, 1309.752, 1309.495, 1307.261, 1305.221, 1301.843, 1299.231, 1297.083, 1295.568]
inst: std_all
operation: std_all(baseline_shift)
tarray: [1.0, 538.0, 1075.0, 1611.0, 2148.0, 2685.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.6941
varray: [2.186, 1.864, 1.448, 1.562, 1.581, 1.999, 1.442, 1.84, 1.338, 1.681]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [1.0, 538.0, 1075.0, 1611.0, 2148.0, 2685.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.272115763764
varray: [2.186, 1.864, 1.448, 1.562, 1.581, 1.999, 1.442, 1.84, 1.338, 1.681]
inst: std_all
operation: std_all(slope_shift)
tarray: [1.0, 538.0, 1075.0, 1611.0, 2148.0, 2685.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: slope_shift
SAMP_d18O_STD | 0.172530351597
varray: [-7.995, -7.635, -7.553, -7.494, -7.464, -7.417, -7.498, -7.429, -7.447, -7.463]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1075.0, 1611.0, 2148.0, 2685.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.84173885228
varray: [-65.765, -61.473, -60.565, -60.185, -59.953, -59.925, -60.02, -59.898, -59.796, -59.74]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1075.0, 1611.0, 2148.0, 2685.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 114.7521
varray: [120.929, 116.402, 115.564, 115.301, 114.134, 115.583, 113.707, 111.731, 111.991, 112.179]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [1.0, 538.0, 1075.0, 1611.0, 2148.0, 2685.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.73603474028
varray: [120.929, 116.402, 115.564, 115.301, 114.134, 115.583, 113.707, 111.731, 111.991, 112.179]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [1.0, 538.0, 1075.0, 1611.0, 2148.0, 2685.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9844
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.5122
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1305.3808
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.6941
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.08905
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.08905
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 3.07562125477
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 3.3923
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 0
varray: [-7.995, -7.635, -7.553, -7.494, -7.464, -7.417, -7.498, -7.429, -7.447, -7.463]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [1.0, 538.0, 1075.0, 1611.0, 2148.0, 2685.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 0
varray: [-65.765, -61.473, -60.565, -60.185, -59.953, -59.925, -60.02, -59.898, -59.796, -59.74]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [1.0, 538.0, 1075.0, 1611.0, 2148.0, 2685.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.5395
varray: [-7.995, -7.635, -7.553, -7.494, -7.464, -7.417, -7.498, -7.429, -7.447, -7.463]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1075.0, 1611.0, 2148.0, 2685.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: d(18_16)Mean
dD_mean | -60.732
varray: [-65.765, -61.473, -60.565, -60.185, -59.953, -59.925, -60.02, -59.898, -59.796, -59.74]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1075.0, 1611.0, 2148.0, 2685.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: d(D_H)Mean
d18O_stdev | 0.172530351597
varray: [-7.995, -7.635, -7.553, -7.494, -7.464, -7.417, -7.498, -7.429, -7.447, -7.463]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1075.0, 1611.0, 2148.0, 2685.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: d(18_16)Mean
dD_stdev | 1.84173885228
varray: [-65.765, -61.473, -60.565, -60.185, -59.953, -59.925, -60.02, -59.898, -59.796, -59.74]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1075.0, 1611.0, 2148.0, 2685.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.5395
varray: [-7.995, -7.635, -7.553, -7.494, -7.464, -7.417, -7.498, -7.429, -7.447, -7.463]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1075.0, 1611.0, 2148.0, 2685.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: d(18_16)Mean
measured_dD_mean | -60.732
varray: [-65.765, -61.473, -60.565, -60.185, -59.953, -59.925, -60.02, -59.898, -59.796, -59.74]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1075.0, 1611.0, 2148.0, 2685.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.172530351597
varray: [-7.995, -7.635, -7.553, -7.494, -7.464, -7.417, -7.498, -7.429, -7.447, -7.463]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1075.0, 1611.0, 2148.0, 2685.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.84173885228
varray: [-65.765, -61.473, -60.565, -60.185, -59.953, -59.925, -60.02, -59.898, -59.796, -59.74]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1075.0, 1611.0, 2148.0, 2685.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.2004897266
actual_array: [ 1.93 -15.84]
varray: [-7.995, -7.635, -7.553, -7.494, -7.464, -7.417, -7.498, -7.429, -7.447, -7.463]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.5395
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [1.0, 538.0, 1075.0, 1611.0, 2148.0, 2685.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -74.4737841834
actual_array: [ 18.1 -119. ]
varray: [-65.765, -61.473, -60.565, -60.185, -59.953, -59.925, -60.02, -59.898, -59.796, -59.74]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -60.732
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [1.0, 538.0, 1075.0, 1611.0, 2148.0, 2685.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.49655
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 9.89985
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 2 - Sample 2 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 2
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9392
varray: [2.881, 2.915, 2.963, 2.996, 2.933, 2.96, 2.997, 2.891, 2.966, 2.89]
inst: mean_all
operation: mean_all(residuals)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: residuals
SAMP_residuals_STD | 0.043511939613
varray: [2.881, 2.915, 2.963, 2.996, 2.933, 2.96, 2.997, 2.891, 2.966, 2.89]
inst: std_all
operation: std_all(residuals)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3551
varray: [0.048, 0.269, 0.559, 0.29, 0.262, 0.704, 0.49, 0.388, 0.106, 0.435]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.201563582469
varray: [0.048, 0.269, 0.559, 0.29, 0.262, 0.704, 0.49, 0.388, 0.106, 0.435]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1296.1708
varray: [1296.994, 1295.029, 1296.664, 1297.916, 1296.061, 1294.379, 1296.587, 1296.434, 1297.634, 1294.01]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.31425710321
varray: [1296.994, 1295.029, 1296.664, 1297.916, 1296.061, 1294.379, 1296.587, 1296.434, 1297.634, 1294.01]
inst: std_all
operation: std_all(baseline_shift)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.3448
varray: [1.929, 1.13, 1.563, 1.157, 1.217, 1.215, 1.236, 1.241, 1.211, 1.549]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.255103813291
varray: [1.929, 1.13, 1.563, 1.157, 1.217, 1.215, 1.236, 1.241, 1.211, 1.549]
inst: std_all
operation: std_all(slope_shift)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: slope_shift
SAMP_d18O_STD | 0.258046270614
varray: [3.939, 4.421, 4.625, 4.658, 4.724, 4.653, 4.781, 4.725, 4.792, 4.753]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.89433908956
varray: [27.899, 36.041, 38.497, 39.122, 39.579, 39.6, 40.205, 40.48, 40.805, 40.427]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 112.1058
varray: [110.905, 111.355, 112.65, 113.853, 111.685, 112.977, 113.49, 111.58, 112.135, 110.428]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.12046050057
varray: [110.905, 111.355, 112.65, 113.853, 111.685, 112.977, 113.49, 111.58, 112.135, 110.428]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9392
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3551
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1296.1708
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.3448
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.04385
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.04385
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.51449738374
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.746
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 4.6071
varray: [3.939, 4.421, 4.625, 4.658, 4.724, 4.653, 4.781, 4.725, 4.792, 4.753]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 38.2655
varray: [27.899, 36.041, 38.497, 39.122, 39.579, 39.6, 40.205, 40.48, 40.805, 40.427]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: d(D_H)Mean
known_dD | 18.1
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 1.93
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 4.6071
varray: [3.939, 4.421, 4.625, 4.658, 4.724, 4.653, 4.781, 4.725, 4.792, 4.753]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: d(18_16)Mean
dD_mean | 38.2655
varray: [27.899, 36.041, 38.497, 39.122, 39.579, 39.6, 40.205, 40.48, 40.805, 40.427]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: d(D_H)Mean
d18O_stdev | 0.258046270614
varray: [3.939, 4.421, 4.625, 4.658, 4.724, 4.653, 4.781, 4.725, 4.792, 4.753]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: d(18_16)Mean
dD_stdev | 3.89433908956
varray: [27.899, 36.041, 38.497, 39.122, 39.579, 39.6, 40.205, 40.48, 40.805, 40.427]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: d(D_H)Mean
measured_d18O_mean | 4.6071
varray: [3.939, 4.421, 4.625, 4.658, 4.724, 4.653, 4.781, 4.725, 4.792, 4.753]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: d(18_16)Mean
measured_dD_mean | 38.2655
varray: [27.899, 36.041, 38.497, 39.122, 39.579, 39.6, 40.205, 40.48, 40.805, 40.427]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.258046270614
varray: [3.939, 4.421, 4.625, 4.658, 4.724, 4.653, 4.781, 4.725, 4.792, 4.753]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.89433908956
varray: [27.899, 36.041, 38.497, 39.122, 39.579, 39.6, 40.205, 40.48, 40.805, 40.427]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.93
actual_array: [ 1.93 -15.84]
varray: [3.939, 4.421, 4.625, 4.658, 4.724, 4.653, 4.781, 4.725, 4.792, 4.753]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 4.6071
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 18.1
actual_array: [ 18.1 -119. ]
varray: [27.899, 36.041, 38.497, 39.122, 39.579, 39.6, 40.205, 40.48, 40.805, 40.427]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 38.2655
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [5420.0, 5957.0, 6494.0, 7031.0, 7568.0, 8105.0, 8642.0, 9179.0, 9715.0, 10252.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.54175
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.68985
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 3 - Sample 3 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 3
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8515
varray: [2.922, 2.861, 2.801, 2.898, 2.842, 2.805, 2.83, 2.833, 2.911, 2.812]
inst: mean_all
operation: mean_all(residuals)
tarray: [10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: residuals
SAMP_residuals_STD | 0.0446349389803
varray: [2.922, 2.861, 2.801, 2.898, 2.842, 2.805, 2.83, 2.833, 2.911, 2.812]
inst: std_all
operation: std_all(residuals)
tarray: [10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4305
varray: [0.528, 0.427, 0.251, 0.792, 0.368, 0.244, 0.387, 0.334, 0.569, 0.405]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.164027944774
varray: [0.528, 0.427, 0.251, 0.792, 0.368, 0.244, 0.387, 0.334, 0.569, 0.405]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1294.7911
varray: [1296.484, 1294.971, 1294.189, 1293.121, 1297.242, 1293.196, 1295.33, 1294.452, 1294.923, 1294.003]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.31846665571
varray: [1296.484, 1294.971, 1294.189, 1293.121, 1297.242, 1293.196, 1295.33, 1294.452, 1294.923, 1294.003]
inst: std_all
operation: std_all(baseline_shift)
tarray: [10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.0681
varray: [2.257, 1.406, 2.533, 1.513, 2.473, 1.773, 2.607, 2.125, 1.717, 2.277]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.435879302866
varray: [2.257, 1.406, 2.533, 1.513, 2.473, 1.773, 2.607, 2.125, 1.717, 2.277]
inst: std_all
operation: std_all(slope_shift)
tarray: [10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: slope_shift
SAMP_d18O_STD | 0.370923547079
varray: [-12.197, -13.002, -13.187, -13.345, -13.214, -13.363, -13.338, -13.4, -13.416, -13.403]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 5.72725991203
varray: [-92.886, -105.698, -108.406, -110.117, -109.983, -110.818, -111.128, -111.395, -111.414, -111.634]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.6138
varray: [112.128, 111.633, 108.345, 111.651, 112.434, 108.068, 110.239, 109.716, 112.245, 109.679]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.62720679557
varray: [112.128, 111.633, 108.345, 111.651, 112.434, 108.068, 110.239, 109.716, 112.245, 109.679]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8515
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4305
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1294.7911
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.0681
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.04385
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.04385
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.51449738374
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.746
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [-12.197, -13.002, -13.187, -13.345, -13.214, -13.363, -13.338, -13.4, -13.416, -13.403]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-92.886, -105.698, -108.406, -110.117, -109.983, -110.818, -111.128, -111.395, -111.414, -111.634]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: d(D_H)Mean
known_dD | -119.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | -15.84
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -13.1865
varray: [-12.197, -13.002, -13.187, -13.345, -13.214, -13.363, -13.338, -13.4, -13.416, -13.403]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: d(18_16)Mean
dD_mean | -108.3479
varray: [-92.886, -105.698, -108.406, -110.117, -109.983, -110.818, -111.128, -111.395, -111.414, -111.634]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: d(D_H)Mean
d18O_stdev | 0.370923547079
varray: [-12.197, -13.002, -13.187, -13.345, -13.214, -13.363, -13.338, -13.4, -13.416, -13.403]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: d(18_16)Mean
dD_stdev | 5.72725991203
varray: [-92.886, -105.698, -108.406, -110.117, -109.983, -110.818, -111.128, -111.395, -111.414, -111.634]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: d(D_H)Mean
measured_d18O_mean | -13.1865
varray: [-12.197, -13.002, -13.187, -13.345, -13.214, -13.363, -13.338, -13.4, -13.416, -13.403]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: d(18_16)Mean
measured_dD_mean | -108.3479
varray: [-92.886, -105.698, -108.406, -110.117, -109.983, -110.818, -111.128, -111.395, -111.414, -111.634]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.370923547079
varray: [-12.197, -13.002, -13.187, -13.345, -13.214, -13.363, -13.338, -13.4, -13.416, -13.403]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: d(18_16)Mean
measured_dD_stdev | 5.72725991203
varray: [-92.886, -105.698, -108.406, -110.117, -109.983, -110.818, -111.128, -111.395, -111.414, -111.634]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -15.84
actual_array: [ 1.93 -15.84]
varray: [-12.197, -13.002, -13.187, -13.345, -13.214, -13.363, -13.338, -13.4, -13.416, -13.403]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -13.1865
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -119.0
actual_array: [ 18.1 -119. ]
varray: [-92.886, -105.698, -108.406, -110.117, -109.983, -110.818, -111.128, -111.395, -111.414, -111.634]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -108.3479
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [10841.0, 11377.0, 11914.0, 12450.0, 12987.0, 13524.0, 14060.0, 14597.0, 15134.0, 15671.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.62945
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.68985
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 4 - Sample 4 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 4
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8549
varray: [2.825, 2.877, 2.835, 2.867, 2.854, 2.854, 2.851, 2.893, 2.845, 2.848]
inst: mean_all
operation: mean_all(residuals)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21092.0]
source_col: residuals
SAMP_residuals_STD | 0.0198295514602
varray: [2.825, 2.877, 2.835, 2.867, 2.854, 2.854, 2.851, 2.893, 2.845, 2.848]
inst: std_all
operation: std_all(residuals)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21092.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4635
varray: [0.313, 0.601, 0.636, 0.449, 0.297, 0.448, 0.411, 0.52, 0.56, 0.4]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21092.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.114880856156
varray: [0.313, 0.601, 0.636, 0.449, 0.297, 0.448, 0.411, 0.52, 0.56, 0.4]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21092.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1293.8906
varray: [1293.335, 1293.796, 1293.394, 1294.472, 1294.288, 1293.449, 1293.886, 1294.647, 1294.148, 1293.491]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21092.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.477646126565
varray: [1293.335, 1293.796, 1293.394, 1294.472, 1294.288, 1293.449, 1293.886, 1294.647, 1294.148, 1293.491]
inst: std_all
operation: std_all(baseline_shift)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21092.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.8753
varray: [1.743, 2.213, 2.644, 1.367, 1.705, 2.103, 1.478, 0.975, 2.088, 2.437]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21092.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.515406861939
varray: [1.743, 2.213, 2.644, 1.367, 1.705, 2.103, 1.478, 0.975, 2.088, 2.437]
inst: std_all
operation: std_all(slope_shift)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21092.0]
source_col: slope_shift
SAMP_d18O_STD | 0.122657653654
varray: [-8.005, -7.739, -7.69, -7.661, -7.628, -7.602, -7.632, -7.616, -7.609, -7.605]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21092.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.90288903162
varray: [-66.361, -61.962, -61.034, -60.792, -60.472, -60.476, -60.23, -60.175, -60.101, -60.237]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21092.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.4045
varray: [109.41, 110.444, 109.689, 111.625, 110.256, 109.458, 110.908, 112.281, 110.802, 109.172]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21092.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.01990699467
varray: [109.41, 110.444, 109.689, 111.625, 110.256, 109.458, 110.908, 112.281, 110.802, 109.172]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21092.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8549
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4635
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1293.8906
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.8753
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.04045
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.04045
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.39706771202
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.9553
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [-8.005, -7.739, -7.69, -7.661, -7.628, -7.602, -7.632, -7.616, -7.609, -7.605]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21092.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-66.361, -61.962, -61.034, -60.792, -60.472, -60.476, -60.23, -60.175, -60.101, -60.237]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21092.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.6787
varray: [-8.005, -7.739, -7.69, -7.661, -7.628, -7.602, -7.632, -7.616, -7.609, -7.605]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21092.0]
source_col: d(18_16)Mean
dD_mean | -61.184
varray: [-66.361, -61.962, -61.034, -60.792, -60.472, -60.476, -60.23, -60.175, -60.101, -60.237]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21092.0]
source_col: d(D_H)Mean
d18O_stdev | 0.122657653654
varray: [-8.005, -7.739, -7.69, -7.661, -7.628, -7.602, -7.632, -7.616, -7.609, -7.605]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21092.0]
source_col: d(18_16)Mean
dD_stdev | 1.90288903162
varray: [-66.361, -61.962, -61.034, -60.792, -60.472, -60.476, -60.23, -60.175, -60.101, -60.237]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21092.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.6787
varray: [-8.005, -7.739, -7.69, -7.661, -7.628, -7.602, -7.632, -7.616, -7.609, -7.605]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21092.0]
source_col: d(18_16)Mean
measured_dD_mean | -61.184
varray: [-66.361, -61.962, -61.034, -60.792, -60.472, -60.476, -60.23, -60.175, -60.101, -60.237]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21092.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.122657653654
varray: [-8.005, -7.739, -7.69, -7.661, -7.628, -7.602, -7.632, -7.616, -7.609, -7.605]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21092.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.90288903162
varray: [-66.361, -61.962, -61.034, -60.792, -60.472, -60.476, -60.23, -60.175, -60.101, -60.237]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21092.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.339505103
actual_array: [ 1.93 -15.84]
varray: [-8.005, -7.739, -7.69, -7.661, -7.628, -7.602, -7.632, -7.616, -7.609, -7.605]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.6787
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21092.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -74.8964549625
actual_array: [ 18.1 -119. ]
varray: [-66.361, -61.962, -61.034, -60.792, -60.472, -60.476, -60.23, -60.175, -60.101, -60.237]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -61.184
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18944.0, 19481.0, 20018.0, 20555.0, 21092.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.62605
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.59035
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 5 - Sample 5 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 5
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.84075
varray: [2.799, 2.853, 2.862, 2.849]
inst: mean_all
operation: mean_all(residuals)
tarray: [21682.0, 22219.0, 22756.0, 23292.0]
source_col: residuals
SAMP_residuals_STD | 0.028359301825
varray: [2.799, 2.853, 2.862, 2.849]
inst: std_all
operation: std_all(residuals)
tarray: [21682.0, 22219.0, 22756.0, 23292.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.39525
varray: [0.625, 0.491, 0.133, 0.332]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [21682.0, 22219.0, 22756.0, 23292.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.211918813071
varray: [0.625, 0.491, 0.133, 0.332]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [21682.0, 22219.0, 22756.0, 23292.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1294.20375
varray: [1291.298, 1294.377, 1295.473, 1295.667]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [21682.0, 22219.0, 22756.0, 23292.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.0187037714
varray: [1291.298, 1294.377, 1295.473, 1295.667]
inst: std_all
operation: std_all(baseline_shift)
tarray: [21682.0, 22219.0, 22756.0, 23292.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.20275
varray: [2.194, 2.107, 2.481, 2.029]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [21682.0, 22219.0, 22756.0, 23292.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.197363243792
varray: [2.194, 2.107, 2.481, 2.029]
inst: std_all
operation: std_all(slope_shift)
tarray: [21682.0, 22219.0, 22756.0, 23292.0]
source_col: slope_shift
SAMP_d18O_STD | 0.179121001933
varray: [-14.375, -14.684, -14.757, -14.742]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21682.0, 22219.0, 22756.0, 23292.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.97349461913
varray: [-113.704, -118.569, -119.794, -120.13]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21682.0, 22219.0, 22756.0, 23292.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.09625
varray: [109.036, 110.82, 109.214, 111.315]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [21682.0, 22219.0, 22756.0, 23292.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.14187867861
varray: [109.036, 110.82, 109.214, 111.315]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [21682.0, 22219.0, 22756.0, 23292.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.84075
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.39525
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1294.20375
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.20275
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0546
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0546
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.88578237519
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.26355
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [-14.375, -14.684, -14.757, -14.742]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [21682.0, 22219.0, 22756.0, 23292.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-113.704, -118.569, -119.794, -120.13]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [21682.0, 22219.0, 22756.0, 23292.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -14.6395
varray: [-14.375, -14.684, -14.757, -14.742]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21682.0, 22219.0, 22756.0, 23292.0]
source_col: d(18_16)Mean
dD_mean | -118.04925
varray: [-113.704, -118.569, -119.794, -120.13]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21682.0, 22219.0, 22756.0, 23292.0]
source_col: d(D_H)Mean
d18O_stdev | 0.179121001933
varray: [-14.375, -14.684, -14.757, -14.742]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21682.0, 22219.0, 22756.0, 23292.0]
source_col: d(18_16)Mean
dD_stdev | 2.97349461913
varray: [-113.704, -118.569, -119.794, -120.13]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21682.0, 22219.0, 22756.0, 23292.0]
source_col: d(D_H)Mean
measured_d18O_mean | -14.6395
varray: [-14.375, -14.684, -14.757, -14.742]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21682.0, 22219.0, 22756.0, 23292.0]
source_col: d(18_16)Mean
measured_dD_mean | -118.04925
varray: [-113.704, -118.569, -119.794, -120.13]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21682.0, 22219.0, 22756.0, 23292.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.179121001933
varray: [-14.375, -14.684, -14.757, -14.742]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21682.0, 22219.0, 22756.0, 23292.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.97349461913
varray: [-113.704, -118.569, -119.794, -120.13]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21682.0, 22219.0, 22756.0, 23292.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -17.2910728577
actual_array: [ 1.93 -15.84]
varray: [-14.375, -14.684, -14.757, -14.742]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -14.6395
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [21682.0, 22219.0, 22756.0, 23292.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -128.071852129
actual_array: [ 18.1 -119. ]
varray: [-113.704, -118.569, -119.794, -120.13]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -118.04925
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [21682.0, 22219.0, 22756.0, 23292.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.6402
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.2772
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 6 - Sample 6 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 6
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.769
varray: [2.724, 2.807, 2.831, 2.714]
inst: mean_all
operation: mean_all(residuals)
tarray: [23882.0, 24419.0, 24956.0, 25493.0]
source_col: residuals
SAMP_residuals_STD | 0.0587026404858
varray: [2.724, 2.807, 2.831, 2.714]
inst: std_all
operation: std_all(residuals)
tarray: [23882.0, 24419.0, 24956.0, 25493.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3175
varray: [0.152, 0.491, 0.433, 0.194]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [23882.0, 24419.0, 24956.0, 25493.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.169395985785
varray: [0.152, 0.491, 0.433, 0.194]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [23882.0, 24419.0, 24956.0, 25493.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1292.42425
varray: [1292.304, 1292.407, 1292.236, 1292.75]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [23882.0, 24419.0, 24956.0, 25493.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.228260633195
varray: [1292.304, 1292.407, 1292.236, 1292.75]
inst: std_all
operation: std_all(baseline_shift)
tarray: [23882.0, 24419.0, 24956.0, 25493.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.13
varray: [2.142, 2.125, 2.231, 2.022]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [23882.0, 24419.0, 24956.0, 25493.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0857010307211
varray: [2.142, 2.125, 2.231, 2.022]
inst: std_all
operation: std_all(slope_shift)
tarray: [23882.0, 24419.0, 24956.0, 25493.0]
source_col: slope_shift
SAMP_d18O_STD | 0.02696139215
varray: [-14.788, -14.843, -14.844, -14.812]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23882.0, 24419.0, 24956.0, 25493.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.275313609059
varray: [-120.079, -120.316, -120.509, -120.725]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23882.0, 24419.0, 24956.0, 25493.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.07225
varray: [107.178, 108.709, 108.908, 107.494]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [23882.0, 24419.0, 24956.0, 25493.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.863709972541
varray: [107.178, 108.709, 108.908, 107.494]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [23882.0, 24419.0, 24956.0, 25493.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.769
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3175
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1292.42425
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.13
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.12635
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.12635
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.36389382976
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.28755
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [-14.788, -14.843, -14.844, -14.812]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [23882.0, 24419.0, 24956.0, 25493.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-120.079, -120.316, -120.509, -120.725]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [23882.0, 24419.0, 24956.0, 25493.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -14.82175
varray: [-14.788, -14.843, -14.844, -14.812]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23882.0, 24419.0, 24956.0, 25493.0]
source_col: d(18_16)Mean
dD_mean | -120.40725
varray: [-120.079, -120.316, -120.509, -120.725]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23882.0, 24419.0, 24956.0, 25493.0]
source_col: d(D_H)Mean
d18O_stdev | 0.02696139215
varray: [-14.788, -14.843, -14.844, -14.812]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23882.0, 24419.0, 24956.0, 25493.0]
source_col: d(18_16)Mean
dD_stdev | 0.275313609059
varray: [-120.079, -120.316, -120.509, -120.725]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23882.0, 24419.0, 24956.0, 25493.0]
source_col: d(D_H)Mean
measured_d18O_mean | -14.82175
varray: [-14.788, -14.843, -14.844, -14.812]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23882.0, 24419.0, 24956.0, 25493.0]
source_col: d(18_16)Mean
measured_dD_mean | -120.40725
varray: [-120.079, -120.316, -120.509, -120.725]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23882.0, 24419.0, 24956.0, 25493.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.02696139215
varray: [-14.788, -14.843, -14.844, -14.812]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23882.0, 24419.0, 24956.0, 25493.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.275313609059
varray: [-120.079, -120.316, -120.509, -120.725]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23882.0, 24419.0, 24956.0, 25493.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -17.4730811359
actual_array: [ 1.93 -15.84]
varray: [-14.788, -14.843, -14.844, -14.812]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -14.82175
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [23882.0, 24419.0, 24956.0, 25493.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -130.276847034
actual_array: [ 18.1 -119. ]
varray: [-120.079, -120.316, -120.509, -120.725]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -120.40725
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [23882.0, 24419.0, 24956.0, 25493.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.71195
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -3.0567
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 7 - Sample 7 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 7
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.984
varray: [2.899, 3.113, 3.12, 2.804]
inst: mean_all
operation: mean_all(residuals)
tarray: [26084.0, 26620.0, 27157.0, 27694.0]
source_col: residuals
SAMP_residuals_STD | 0.157862809638
varray: [2.899, 3.113, 3.12, 2.804]
inst: std_all
operation: std_all(residuals)
tarray: [26084.0, 26620.0, 27157.0, 27694.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.48725
varray: [0.531, 0.628, 0.553, 0.237]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [26084.0, 26620.0, 27157.0, 27694.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.171923190214
varray: [0.531, 0.628, 0.553, 0.237]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [26084.0, 26620.0, 27157.0, 27694.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1300.26575
varray: [1296.738, 1304.216, 1305.382, 1294.727]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [26084.0, 26620.0, 27157.0, 27694.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 5.31987658848
varray: [1296.738, 1304.216, 1305.382, 1294.727]
inst: std_all
operation: std_all(baseline_shift)
tarray: [26084.0, 26620.0, 27157.0, 27694.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.11625
varray: [1.911, 2.238, 2.193, 2.123]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [26084.0, 26620.0, 27157.0, 27694.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.144783459
varray: [1.911, 2.238, 2.193, 2.123]
inst: std_all
operation: std_all(slope_shift)
tarray: [26084.0, 26620.0, 27157.0, 27694.0]
source_col: slope_shift
SAMP_d18O_STD | 0.472452819514
varray: [-14.742, -14.426, -13.684, -14.61]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26084.0, 26620.0, 27157.0, 27694.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.56501182104
varray: [-120.455, -119.345, -117.046, -120.266]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26084.0, 26620.0, 27157.0, 27694.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 115.724
varray: [113.092, 119.77, 120.07, 109.964]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [26084.0, 26620.0, 27157.0, 27694.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 5.0120806059
varray: [113.092, 119.77, 120.07, 109.964]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [26084.0, 26620.0, 27157.0, 27694.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.984
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.48725
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1300.26575
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.11625
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.08865
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.08865
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 3.06180599927
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 4.3642
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [-14.742, -14.426, -13.684, -14.61]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [26084.0, 26620.0, 27157.0, 27694.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-120.455, -119.345, -117.046, -120.266]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [26084.0, 26620.0, 27157.0, 27694.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -14.3655
varray: [-14.742, -14.426, -13.684, -14.61]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26084.0, 26620.0, 27157.0, 27694.0]
source_col: d(18_16)Mean
dD_mean | -119.278
varray: [-120.455, -119.345, -117.046, -120.266]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26084.0, 26620.0, 27157.0, 27694.0]
source_col: d(D_H)Mean
d18O_stdev | 0.472452819514
varray: [-14.742, -14.426, -13.684, -14.61]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26084.0, 26620.0, 27157.0, 27694.0]
source_col: d(18_16)Mean
dD_stdev | 1.56501182104
varray: [-120.455, -119.345, -117.046, -120.266]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26084.0, 26620.0, 27157.0, 27694.0]
source_col: d(D_H)Mean
measured_d18O_mean | -14.3655
varray: [-14.742, -14.426, -13.684, -14.61]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26084.0, 26620.0, 27157.0, 27694.0]
source_col: d(18_16)Mean
measured_dD_mean | -119.278
varray: [-120.455, -119.345, -117.046, -120.266]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26084.0, 26620.0, 27157.0, 27694.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.472452819514
varray: [-14.742, -14.426, -13.684, -14.61]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26084.0, 26620.0, 27157.0, 27694.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.56501182104
varray: [-120.455, -119.345, -117.046, -120.266]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26084.0, 26620.0, 27157.0, 27694.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -17.0174362692
actual_array: [ 1.93 -15.84]
varray: [-14.742, -14.426, -13.684, -14.61]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -14.3655
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [26084.0, 26620.0, 27157.0, 27694.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -129.220871421
actual_array: [ 18.1 -119. ]
varray: [-120.455, -119.345, -117.046, -120.266]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -119.278
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [26084.0, 26620.0, 27157.0, 27694.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.49695
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 4.7848
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 8 - Sample 8 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 8
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.83325
varray: [2.799, 2.893, 2.835, 2.806]
inst: mean_all
operation: mean_all(residuals)
tarray: [28282.0, 28818.0, 29355.0, 29892.0]
source_col: residuals
SAMP_residuals_STD | 0.0427736289474
varray: [2.799, 2.893, 2.835, 2.806]
inst: std_all
operation: std_all(residuals)
tarray: [28282.0, 28818.0, 29355.0, 29892.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.2665
varray: [0.114, 0.416, 0.381, 0.155]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [28282.0, 28818.0, 29355.0, 29892.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.154001082247
varray: [0.114, 0.416, 0.381, 0.155]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [28282.0, 28818.0, 29355.0, 29892.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1295.03225
varray: [1296.01, 1296.139, 1292.588, 1295.392]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [28282.0, 28818.0, 29355.0, 29892.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.66179308279
varray: [1296.01, 1296.139, 1292.588, 1295.392]
inst: std_all
operation: std_all(baseline_shift)
tarray: [28282.0, 28818.0, 29355.0, 29892.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.405
varray: [3.015, 2.019, 2.29, 2.296]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [28282.0, 28818.0, 29355.0, 29892.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.42669348874
varray: [3.015, 2.019, 2.29, 2.296]
inst: std_all
operation: std_all(slope_shift)
tarray: [28282.0, 28818.0, 29355.0, 29892.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0707978342418
varray: [-14.76, -14.746, -14.845, -14.895]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28282.0, 28818.0, 29355.0, 29892.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.318871447452
varray: [-120.746, -120.691, -121.06, -121.381]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28282.0, 28818.0, 29355.0, 29892.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.88825
varray: [108.656, 111.581, 108.877, 110.439]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [28282.0, 28818.0, 29355.0, 29892.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.37958867662
varray: [108.656, 111.581, 108.877, 110.439]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [28282.0, 28818.0, 29355.0, 29892.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.83325
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.2665
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1295.03225
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.405
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0621
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0621
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.14481841574
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.47155
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [-14.76, -14.746, -14.845, -14.895]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [28282.0, 28818.0, 29355.0, 29892.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-120.746, -120.691, -121.06, -121.381]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [28282.0, 28818.0, 29355.0, 29892.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -14.8115
varray: [-14.76, -14.746, -14.845, -14.895]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28282.0, 28818.0, 29355.0, 29892.0]
source_col: d(18_16)Mean
dD_mean | -120.9695
varray: [-120.746, -120.691, -121.06, -121.381]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28282.0, 28818.0, 29355.0, 29892.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0707978342418
varray: [-14.76, -14.746, -14.845, -14.895]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28282.0, 28818.0, 29355.0, 29892.0]
source_col: d(18_16)Mean
dD_stdev | 0.318871447452
varray: [-120.746, -120.691, -121.06, -121.381]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28282.0, 28818.0, 29355.0, 29892.0]
source_col: d(D_H)Mean
measured_d18O_mean | -14.8115
varray: [-14.76, -14.746, -14.845, -14.895]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28282.0, 28818.0, 29355.0, 29892.0]
source_col: d(18_16)Mean
measured_dD_mean | -120.9695
varray: [-120.746, -120.691, -121.06, -121.381]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28282.0, 28818.0, 29355.0, 29892.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0707978342418
varray: [-14.76, -14.746, -14.845, -14.895]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28282.0, 28818.0, 29355.0, 29892.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.318871447452
varray: [-120.746, -120.691, -121.06, -121.381]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28282.0, 28818.0, 29355.0, 29892.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -17.4628447307
actual_array: [ 1.93 -15.84]
varray: [-14.76, -14.746, -14.845, -14.895]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -14.8115
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [28282.0, 28818.0, 29355.0, 29892.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -130.802613949
actual_array: [ 18.1 -119. ]
varray: [-120.746, -120.691, -121.06, -121.381]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -120.9695
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [28282.0, 28818.0, 29355.0, 29892.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.6477
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.4487
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 9 - Sample 9 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 9
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.82675
varray: [2.783, 2.917, 2.76, 2.847]
inst: mean_all
operation: mean_all(residuals)
tarray: [30480.0, 31017.0, 31554.0, 32091.0]
source_col: residuals
SAMP_residuals_STD | 0.0705330891615
varray: [2.783, 2.917, 2.76, 2.847]
inst: std_all
operation: std_all(residuals)
tarray: [30480.0, 31017.0, 31554.0, 32091.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.359
varray: [0.311, 0.654, -0.016, 0.487]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [30480.0, 31017.0, 31554.0, 32091.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.286553078271
varray: [0.311, 0.654, -0.016, 0.487]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [30480.0, 31017.0, 31554.0, 32091.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1293.946
varray: [1292.015, 1294.538, 1295.05, 1294.181]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [30480.0, 31017.0, 31554.0, 32091.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.33582259301
varray: [1292.015, 1294.538, 1295.05, 1294.181]
inst: std_all
operation: std_all(baseline_shift)
tarray: [30480.0, 31017.0, 31554.0, 32091.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.003
varray: [1.833, 2.332, 2.0, 1.847]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [30480.0, 31017.0, 31554.0, 32091.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.232010057253
varray: [1.833, 2.332, 2.0, 1.847]
inst: std_all
operation: std_all(slope_shift)
tarray: [30480.0, 31017.0, 31554.0, 32091.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0650589476296
varray: [-13.757, -13.65, -13.627, -13.614]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30480.0, 31017.0, 31554.0, 32091.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.613985545432
varray: [-113.459, -112.405, -112.295, -112.082]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30480.0, 31017.0, 31554.0, 32091.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.44175
varray: [108.73, 111.898, 109.796, 111.343]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [30480.0, 31017.0, 31554.0, 32091.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.44683525323
varray: [108.73, 111.898, 109.796, 111.343]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [30480.0, 31017.0, 31554.0, 32091.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.82675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.359
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1293.946
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.003
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0686
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0686
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.36931631754
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.91805
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [-13.757, -13.65, -13.627, -13.614]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [30480.0, 31017.0, 31554.0, 32091.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-113.459, -112.405, -112.295, -112.082]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [30480.0, 31017.0, 31554.0, 32091.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -13.662
varray: [-13.757, -13.65, -13.627, -13.614]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30480.0, 31017.0, 31554.0, 32091.0]
source_col: d(18_16)Mean
dD_mean | -112.56025
varray: [-113.459, -112.405, -112.295, -112.082]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30480.0, 31017.0, 31554.0, 32091.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0650589476296
varray: [-13.757, -13.65, -13.627, -13.614]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30480.0, 31017.0, 31554.0, 32091.0]
source_col: d(18_16)Mean
dD_stdev | 0.613985545432
varray: [-113.459, -112.405, -112.295, -112.082]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30480.0, 31017.0, 31554.0, 32091.0]
source_col: d(D_H)Mean
measured_d18O_mean | -13.662
varray: [-13.757, -13.65, -13.627, -13.614]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30480.0, 31017.0, 31554.0, 32091.0]
source_col: d(18_16)Mean
measured_dD_mean | -112.56025
varray: [-113.459, -112.405, -112.295, -112.082]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30480.0, 31017.0, 31554.0, 32091.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0650589476296
varray: [-13.757, -13.65, -13.627, -13.614]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30480.0, 31017.0, 31554.0, 32091.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.613985545432
varray: [-113.459, -112.405, -112.295, -112.082]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30480.0, 31017.0, 31554.0, 32091.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -16.314869335
actual_array: [ 1.93 -15.84]
varray: [-13.757, -13.65, -13.627, -13.614]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -13.662
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [30480.0, 31017.0, 31554.0, 32091.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -122.939020478
actual_array: [ 18.1 -119. ]
varray: [-113.459, -112.405, -112.295, -112.082]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -112.56025
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [30480.0, 31017.0, 31554.0, 32091.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.6542
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.53495
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 10 - Sample 10 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 10
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.82675
varray: [2.764, 2.854, 2.837, 2.852]
inst: mean_all
operation: mean_all(residuals)
tarray: [32679.0, 33215.0, 33752.0, 34289.0]
source_col: residuals
SAMP_residuals_STD | 0.0425156833808
varray: [2.764, 2.854, 2.837, 2.852]
inst: std_all
operation: std_all(residuals)
tarray: [32679.0, 33215.0, 33752.0, 34289.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.50675
varray: [0.474, 0.785, 0.333, 0.435]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [32679.0, 33215.0, 33752.0, 34289.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.194792838677
varray: [0.474, 0.785, 0.333, 0.435]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [32679.0, 33215.0, 33752.0, 34289.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1292.81875
varray: [1291.582, 1291.968, 1294.827, 1292.898]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [32679.0, 33215.0, 33752.0, 34289.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.44829494579
varray: [1291.582, 1291.968, 1294.827, 1292.898]
inst: std_all
operation: std_all(baseline_shift)
tarray: [32679.0, 33215.0, 33752.0, 34289.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.209
varray: [2.482, 2.048, 2.145, 2.161]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [32679.0, 33215.0, 33752.0, 34289.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.188723784051
varray: [2.482, 2.048, 2.145, 2.161]
inst: std_all
operation: std_all(slope_shift)
tarray: [32679.0, 33215.0, 33752.0, 34289.0]
source_col: slope_shift
SAMP_d18O_STD | 0.023300572239
varray: [-14.732, -14.73, -14.779, -14.758]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32679.0, 33215.0, 33752.0, 34289.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.462676542882
varray: [-119.584, -120.273, -120.357, -120.685]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32679.0, 33215.0, 33752.0, 34289.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.9455
varray: [108.3, 110.717, 111.137, 109.628]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [32679.0, 33215.0, 33752.0, 34289.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.26798541001
varray: [108.3, 110.717, 111.137, 109.628]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [32679.0, 33215.0, 33752.0, 34289.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.82675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.50675
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1292.81875
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.209
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0686
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0686
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.36931631754
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.4143
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [-14.732, -14.73, -14.779, -14.758]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [32679.0, 33215.0, 33752.0, 34289.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-119.584, -120.273, -120.357, -120.685]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [32679.0, 33215.0, 33752.0, 34289.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -14.74975
varray: [-14.732, -14.73, -14.779, -14.758]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32679.0, 33215.0, 33752.0, 34289.0]
source_col: d(18_16)Mean
dD_mean | -120.22475
varray: [-119.584, -120.273, -120.357, -120.685]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32679.0, 33215.0, 33752.0, 34289.0]
source_col: d(D_H)Mean
d18O_stdev | 0.023300572239
varray: [-14.732, -14.73, -14.779, -14.758]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32679.0, 33215.0, 33752.0, 34289.0]
source_col: d(18_16)Mean
dD_stdev | 0.462676542882
varray: [-119.584, -120.273, -120.357, -120.685]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32679.0, 33215.0, 33752.0, 34289.0]
source_col: d(D_H)Mean
measured_d18O_mean | -14.74975
varray: [-14.732, -14.73, -14.779, -14.758]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32679.0, 33215.0, 33752.0, 34289.0]
source_col: d(18_16)Mean
measured_dD_mean | -120.22475
varray: [-119.584, -120.273, -120.357, -120.685]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32679.0, 33215.0, 33752.0, 34289.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.023300572239
varray: [-14.732, -14.73, -14.779, -14.758]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32679.0, 33215.0, 33752.0, 34289.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.462676542882
varray: [-119.584, -120.273, -120.357, -120.685]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32679.0, 33215.0, 33752.0, 34289.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -17.4011766309
actual_array: [ 1.93 -15.84]
varray: [-14.732, -14.73, -14.779, -14.758]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -14.74975
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [32679.0, 33215.0, 33752.0, 34289.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -130.106189032
actual_array: [ 18.1 -119. ]
varray: [-119.584, -120.273, -120.357, -120.685]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -120.22475
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [32679.0, 33215.0, 33752.0, 34289.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.6542
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -2.6622
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 11 - Sample 11 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 11
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.82975
varray: [2.789, 2.798, 2.824, 2.908]
inst: mean_all
operation: mean_all(residuals)
tarray: [34877.0, 35414.0, 35951.0, 36488.0]
source_col: residuals
SAMP_residuals_STD | 0.0542363654141
varray: [2.789, 2.798, 2.824, 2.908]
inst: std_all
operation: std_all(residuals)
tarray: [34877.0, 35414.0, 35951.0, 36488.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4245
varray: [0.225, 0.447, 0.421, 0.605]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [34877.0, 35414.0, 35951.0, 36488.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.15588350351
varray: [0.225, 0.447, 0.421, 0.605]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [34877.0, 35414.0, 35951.0, 36488.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1294.04225
varray: [1294.771, 1292.782, 1292.302, 1296.314]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [34877.0, 35414.0, 35951.0, 36488.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.85370122997
varray: [1294.771, 1292.782, 1292.302, 1296.314]
inst: std_all
operation: std_all(baseline_shift)
tarray: [34877.0, 35414.0, 35951.0, 36488.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.43725
varray: [2.918, 2.864, 2.19, 1.777]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [34877.0, 35414.0, 35951.0, 36488.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.550847453173
varray: [2.918, 2.864, 2.19, 1.777]
inst: std_all
operation: std_all(slope_shift)
tarray: [34877.0, 35414.0, 35951.0, 36488.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0549484607003
varray: [-14.658, -14.725, -14.778, -14.671]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34877.0, 35414.0, 35951.0, 36488.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.194763103967
varray: [-120.02, -120.195, -120.31, -119.867]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34877.0, 35414.0, 35951.0, 36488.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.90725
varray: [109.1, 108.556, 108.623, 113.35]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [34877.0, 35414.0, 35951.0, 36488.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.30791065324
varray: [109.1, 108.556, 108.623, 113.35]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [34877.0, 35414.0, 35951.0, 36488.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.82975
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4245
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1294.04225
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.43725
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0656
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0656
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.26570190132
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.45255
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [-14.658, -14.725, -14.778, -14.671]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [34877.0, 35414.0, 35951.0, 36488.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-120.02, -120.195, -120.31, -119.867]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [34877.0, 35414.0, 35951.0, 36488.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -14.708
varray: [-14.658, -14.725, -14.778, -14.671]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [34877.0, 35414.0, 35951.0, 36488.0]
source_col: d(18_16)Mean
dD_mean | -120.098
varray: [-120.02, -120.195, -120.31, -119.867]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [34877.0, 35414.0, 35951.0, 36488.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0549484607003
varray: [-14.658, -14.725, -14.778, -14.671]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34877.0, 35414.0, 35951.0, 36488.0]
source_col: d(18_16)Mean
dD_stdev | 0.194763103967
varray: [-120.02, -120.195, -120.31, -119.867]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34877.0, 35414.0, 35951.0, 36488.0]
source_col: d(D_H)Mean
measured_d18O_mean | -14.708
varray: [-14.658, -14.725, -14.778, -14.671]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [34877.0, 35414.0, 35951.0, 36488.0]
source_col: d(18_16)Mean
measured_dD_mean | -120.098
varray: [-120.02, -120.195, -120.31, -119.867]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [34877.0, 35414.0, 35951.0, 36488.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0549484607003
varray: [-14.658, -14.725, -14.778, -14.671]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34877.0, 35414.0, 35951.0, 36488.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.194763103967
varray: [-120.02, -120.195, -120.31, -119.867]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34877.0, 35414.0, 35951.0, 36488.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -17.3594820048
actual_array: [ 1.93 -15.84]
varray: [-14.658, -14.725, -14.778, -14.671]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -14.708
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [34877.0, 35414.0, 35951.0, 36488.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -129.987663542
actual_array: [ 18.1 -119. ]
varray: [-120.02, -120.195, -120.31, -119.867]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -120.098
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [34877.0, 35414.0, 35951.0, 36488.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.6512
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.4387
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 12 - Sample 12 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 12
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.01125
varray: [2.999, 2.965, 3.051, 3.03]
inst: mean_all
operation: mean_all(residuals)
tarray: [37077.0, 37614.0, 38151.0, 38688.0]
source_col: residuals
SAMP_residuals_STD | 0.0375088878356
varray: [2.999, 2.965, 3.051, 3.03]
inst: std_all
operation: std_all(residuals)
tarray: [37077.0, 37614.0, 38151.0, 38688.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -3.48425
varray: [-3.445, -3.557, -3.513, -3.422]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [37077.0, 37614.0, 38151.0, 38688.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0620073920325
varray: [-3.445, -3.557, -3.513, -3.422]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [37077.0, 37614.0, 38151.0, 38688.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1667.0095
varray: [1660.834, 1662.643, 1673.854, 1670.707]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [37077.0, 37614.0, 38151.0, 38688.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 6.26423203593
varray: [1660.834, 1662.643, 1673.854, 1670.707]
inst: std_all
operation: std_all(baseline_shift)
tarray: [37077.0, 37614.0, 38151.0, 38688.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -22.207
varray: [-21.551, -22.052, -23.231, -21.994]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [37077.0, 37614.0, 38151.0, 38688.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.718402394205
varray: [-21.551, -22.052, -23.231, -21.994]
inst: std_all
operation: std_all(slope_shift)
tarray: [37077.0, 37614.0, 38151.0, 38688.0]
source_col: slope_shift
SAMP_d18O_STD | 0.456177140301
varray: [3.951, 4.605, 4.844, 4.979]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37077.0, 37614.0, 38151.0, 38688.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 5.82834985223
varray: [-17.514, -8.337, -5.756, -4.701]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37077.0, 37614.0, 38151.0, 38688.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.7475
varray: [109.989, 109.384, 112.871, 110.746]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [37077.0, 37614.0, 38151.0, 38688.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.52137098259
varray: [109.989, 109.384, 112.871, 110.746]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [37077.0, 37614.0, 38151.0, 38688.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.01125
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 3.48425
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1667.0095
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 22.207
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.1159
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1159
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 4.00297027993
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.6123
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [3.951, 4.605, 4.844, 4.979]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [37077.0, 37614.0, 38151.0, 38688.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-17.514, -8.337, -5.756, -4.701]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [37077.0, 37614.0, 38151.0, 38688.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 4.59475
varray: [3.951, 4.605, 4.844, 4.979]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37077.0, 37614.0, 38151.0, 38688.0]
source_col: d(18_16)Mean
dD_mean | -9.077
varray: [-17.514, -8.337, -5.756, -4.701]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37077.0, 37614.0, 38151.0, 38688.0]
source_col: d(D_H)Mean
d18O_stdev | 0.456177140301
varray: [3.951, 4.605, 4.844, 4.979]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37077.0, 37614.0, 38151.0, 38688.0]
source_col: d(18_16)Mean
dD_stdev | 5.82834985223
varray: [-17.514, -8.337, -5.756, -4.701]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37077.0, 37614.0, 38151.0, 38688.0]
source_col: d(D_H)Mean
measured_d18O_mean | 4.59475
varray: [3.951, 4.605, 4.844, 4.979]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37077.0, 37614.0, 38151.0, 38688.0]
source_col: d(18_16)Mean
measured_dD_mean | -9.077
varray: [-17.514, -8.337, -5.756, -4.701]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37077.0, 37614.0, 38151.0, 38688.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.456177140301
varray: [3.951, 4.605, 4.844, 4.979]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37077.0, 37614.0, 38151.0, 38688.0]
source_col: d(18_16)Mean
measured_dD_stdev | 5.82834985223
varray: [-17.514, -8.337, -5.756, -4.701]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37077.0, 37614.0, 38151.0, 38688.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.91766638005
actual_array: [ 1.93 -15.84]
varray: [3.951, 4.605, 4.844, 4.979]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 4.59475
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [37077.0, 37614.0, 38151.0, 38688.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -26.170556102
actual_array: [ 18.1 -119. ]
varray: [-17.514, -8.337, -5.756, -4.701]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -9.077
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [37077.0, 37614.0, 38151.0, 38688.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.4697
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 371.52855
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 13 - Sample 13 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 13
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.86675
varray: [2.894, 2.793, 2.926, 2.854]
inst: mean_all
operation: mean_all(residuals)
tarray: [39277.0, 39813.0, 40350.0, 40887.0]
source_col: residuals
SAMP_residuals_STD | 0.0573141925413
varray: [2.894, 2.793, 2.926, 2.854]
inst: std_all
operation: std_all(residuals)
tarray: [39277.0, 39813.0, 40350.0, 40887.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.387
varray: [0.493, 0.257, 0.455, 0.343]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [39277.0, 39813.0, 40350.0, 40887.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.107542239763
varray: [0.493, 0.257, 0.455, 0.343]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [39277.0, 39813.0, 40350.0, 40887.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1296.8955
varray: [1300.898, 1294.741, 1297.629, 1294.314]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [39277.0, 39813.0, 40350.0, 40887.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.04762426599
varray: [1300.898, 1294.741, 1297.629, 1294.314]
inst: std_all
operation: std_all(baseline_shift)
tarray: [39277.0, 39813.0, 40350.0, 40887.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.65425
varray: [1.577, 1.239, 2.089, 1.712]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [39277.0, 39813.0, 40350.0, 40887.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.351540775255
varray: [1.577, 1.239, 2.089, 1.712]
inst: std_all
operation: std_all(slope_shift)
tarray: [39277.0, 39813.0, 40350.0, 40887.0]
source_col: slope_shift
SAMP_d18O_STD | 0.189619795029
varray: [-7.193, -7.569, -7.543, -7.597]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39277.0, 39813.0, 40350.0, 40887.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.07212314628
varray: [-55.529, -59.379, -59.459, -60.051]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39277.0, 39813.0, 40350.0, 40887.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.77625
varray: [111.049, 109.441, 112.525, 110.09]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [39277.0, 39813.0, 40350.0, 40887.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.33994412197
varray: [111.049, 109.441, 112.525, 110.09]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [39277.0, 39813.0, 40350.0, 40887.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.86675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.387
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1296.8955
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.65425
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0286
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0286
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.987790767956
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.58355
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [-7.193, -7.569, -7.543, -7.597]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [39277.0, 39813.0, 40350.0, 40887.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-55.529, -59.379, -59.459, -60.051]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [39277.0, 39813.0, 40350.0, 40887.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.4755
varray: [-7.193, -7.569, -7.543, -7.597]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39277.0, 39813.0, 40350.0, 40887.0]
source_col: d(18_16)Mean
dD_mean | -58.6045
varray: [-55.529, -59.379, -59.459, -60.051]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39277.0, 39813.0, 40350.0, 40887.0]
source_col: d(D_H)Mean
d18O_stdev | 0.189619795029
varray: [-7.193, -7.569, -7.543, -7.597]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39277.0, 39813.0, 40350.0, 40887.0]
source_col: d(18_16)Mean
dD_stdev | 2.07212314628
varray: [-55.529, -59.379, -59.459, -60.051]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39277.0, 39813.0, 40350.0, 40887.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.4755
varray: [-7.193, -7.569, -7.543, -7.597]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39277.0, 39813.0, 40350.0, 40887.0]
source_col: d(18_16)Mean
measured_dD_mean | -58.6045
varray: [-55.529, -59.379, -59.459, -60.051]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39277.0, 39813.0, 40350.0, 40887.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.189619795029
varray: [-7.193, -7.569, -7.543, -7.597]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39277.0, 39813.0, 40350.0, 40887.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.07212314628
varray: [-55.529, -59.379, -59.459, -60.051]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39277.0, 39813.0, 40350.0, 40887.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.1365746111
actual_array: [ 1.93 -15.84]
varray: [-7.193, -7.569, -7.543, -7.597]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.4755
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [39277.0, 39813.0, 40350.0, 40887.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -72.4843326735
actual_array: [ 18.1 -119. ]
varray: [-55.529, -59.379, -59.459, -60.051]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -58.6045
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [39277.0, 39813.0, 40350.0, 40887.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.6142
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1.41455
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 14 - Sample 14 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 14
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.71475
varray: [2.756, 2.724, 2.705, 2.674]
inst: mean_all
operation: mean_all(residuals)
tarray: [41477.0, 42014.0, 42551.0, 43087.0]
source_col: residuals
SAMP_residuals_STD | 0.0343644680836
varray: [2.756, 2.724, 2.705, 2.674]
inst: std_all
operation: std_all(residuals)
tarray: [41477.0, 42014.0, 42551.0, 43087.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.2775
varray: [-2.141, -2.313, -2.285, -2.371]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [41477.0, 42014.0, 42551.0, 43087.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0977940011793
varray: [-2.141, -2.313, -2.285, -2.371]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [41477.0, 42014.0, 42551.0, 43087.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1367.71125
varray: [1365.926, 1368.14, 1369.149, 1367.63]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [41477.0, 42014.0, 42551.0, 43087.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.34717862587
varray: [1365.926, 1368.14, 1369.149, 1367.63]
inst: std_all
operation: std_all(baseline_shift)
tarray: [41477.0, 42014.0, 42551.0, 43087.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -8.07575
varray: [-7.713, -8.242, -8.083, -8.265]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [41477.0, 42014.0, 42551.0, 43087.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.25501290817
varray: [-7.713, -8.242, -8.083, -8.265]
inst: std_all
operation: std_all(slope_shift)
tarray: [41477.0, 42014.0, 42551.0, 43087.0]
source_col: slope_shift
SAMP_d18O_STD | 0.249498830993
varray: [2.562, 2.971, 3.053, 3.116]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41477.0, 42014.0, 42551.0, 43087.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.63865388661
varray: [-13.722, -9.62, -8.46, -7.868]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41477.0, 42014.0, 42551.0, 43087.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 106.92425
varray: [107.5, 107.105, 106.964, 106.128]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [41477.0, 42014.0, 42551.0, 43087.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.57727888985
varray: [107.5, 107.105, 106.964, 106.128]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [41477.0, 42014.0, 42551.0, 43087.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.71475
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.2775
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1367.71125
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 8.07575
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1806
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1806
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.23758785639
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.43555
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [2.562, 2.971, 3.053, 3.116]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [41477.0, 42014.0, 42551.0, 43087.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-13.722, -9.62, -8.46, -7.868]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [41477.0, 42014.0, 42551.0, 43087.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.9255
varray: [2.562, 2.971, 3.053, 3.116]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [41477.0, 42014.0, 42551.0, 43087.0]
source_col: d(18_16)Mean
dD_mean | -9.9175
varray: [-13.722, -9.62, -8.46, -7.868]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [41477.0, 42014.0, 42551.0, 43087.0]
source_col: d(D_H)Mean
d18O_stdev | 0.249498830993
varray: [2.562, 2.971, 3.053, 3.116]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41477.0, 42014.0, 42551.0, 43087.0]
source_col: d(18_16)Mean
dD_stdev | 2.63865388661
varray: [-13.722, -9.62, -8.46, -7.868]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41477.0, 42014.0, 42551.0, 43087.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.9255
varray: [2.562, 2.971, 3.053, 3.116]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [41477.0, 42014.0, 42551.0, 43087.0]
source_col: d(18_16)Mean
measured_dD_mean | -9.9175
varray: [-13.722, -9.62, -8.46, -7.868]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [41477.0, 42014.0, 42551.0, 43087.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.249498830993
varray: [2.562, 2.971, 3.053, 3.116]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41477.0, 42014.0, 42551.0, 43087.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.63865388661
varray: [-13.722, -9.62, -8.46, -7.868]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41477.0, 42014.0, 42551.0, 43087.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 0.250630338998
actual_array: [ 1.93 -15.84]
varray: [2.562, 2.971, 3.053, 3.116]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 2.9255
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [41477.0, 42014.0, 42551.0, 43087.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -26.9565180263
actual_array: [ 18.1 -119. ]
varray: [-13.722, -9.62, -8.46, -7.868]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -9.9175
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [41477.0, 42014.0, 42551.0, 43087.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.7662
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 72.2303
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 15 - Sample 15 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 15
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.6935
varray: [2.701, 2.711, 2.696, 2.666]
inst: mean_all
operation: mean_all(residuals)
tarray: [43674.0, 44211.0, 44747.0, 45284.0]
source_col: residuals
SAMP_residuals_STD | 0.019364916731
varray: [2.701, 2.711, 2.696, 2.666]
inst: std_all
operation: std_all(residuals)
tarray: [43674.0, 44211.0, 44747.0, 45284.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.49625
varray: [-1.63, -1.548, -1.546, -1.261]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [43674.0, 44211.0, 44747.0, 45284.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.161642352123
varray: [-1.63, -1.548, -1.546, -1.261]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [43674.0, 44211.0, 44747.0, 45284.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1527.48875
varray: [1527.561, 1532.62, 1527.242, 1522.532]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [43674.0, 44211.0, 44747.0, 45284.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 4.12169879014
varray: [1527.561, 1532.62, 1527.242, 1522.532]
inst: std_all
operation: std_all(baseline_shift)
tarray: [43674.0, 44211.0, 44747.0, 45284.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -11.58125
varray: [-11.618, -12.038, -11.175, -11.494]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [43674.0, 44211.0, 44747.0, 45284.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.357128711625
varray: [-11.618, -12.038, -11.175, -11.494]
inst: std_all
operation: std_all(slope_shift)
tarray: [43674.0, 44211.0, 44747.0, 45284.0]
source_col: slope_shift
SAMP_d18O_STD | 0.022
varray: [2.107, 2.147, 2.147, 2.111]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43674.0, 44211.0, 44747.0, 45284.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.549422348168
varray: [-2.966, -2.109, -1.978, -1.688]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43674.0, 44211.0, 44747.0, 45284.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 105.803
varray: [105.788, 107.338, 104.978, 105.108]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [43674.0, 44211.0, 44747.0, 45284.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.08322050695
varray: [105.788, 107.338, 104.978, 105.108]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [43674.0, 44211.0, 44747.0, 45284.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.6935
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.49625
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1527.48875
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 11.58125
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.20185
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.20185
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.97152330461
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.5568
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [2.107, 2.147, 2.147, 2.111]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [43674.0, 44211.0, 44747.0, 45284.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-2.966, -2.109, -1.978, -1.688]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [43674.0, 44211.0, 44747.0, 45284.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.128
varray: [2.107, 2.147, 2.147, 2.111]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [43674.0, 44211.0, 44747.0, 45284.0]
source_col: d(18_16)Mean
dD_mean | -2.18525
varray: [-2.966, -2.109, -1.978, -1.688]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [43674.0, 44211.0, 44747.0, 45284.0]
source_col: d(D_H)Mean
d18O_stdev | 0.022
varray: [2.107, 2.147, 2.147, 2.111]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43674.0, 44211.0, 44747.0, 45284.0]
source_col: d(18_16)Mean
dD_stdev | 0.549422348168
varray: [-2.966, -2.109, -1.978, -1.688]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43674.0, 44211.0, 44747.0, 45284.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.128
varray: [2.107, 2.147, 2.147, 2.111]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [43674.0, 44211.0, 44747.0, 45284.0]
source_col: d(18_16)Mean
measured_dD_mean | -2.18525
varray: [-2.966, -2.109, -1.978, -1.688]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [43674.0, 44211.0, 44747.0, 45284.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.022
varray: [2.107, 2.147, 2.147, 2.111]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43674.0, 44211.0, 44747.0, 45284.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.549422348168
varray: [-2.966, -2.109, -1.978, -1.688]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43674.0, 44211.0, 44747.0, 45284.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.54581192114
actual_array: [ 1.93 -15.84]
varray: [2.107, 2.147, 2.147, 2.111]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 2.128
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [43674.0, 44211.0, 44747.0, 45284.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -19.7259956116
actual_array: [ 18.1 -119. ]
varray: [-2.966, -2.109, -1.978, -1.688]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -2.18525
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [43674.0, 44211.0, 44747.0, 45284.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.78745
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 232.0078
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 16 - Sample 16 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 16
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.7715
varray: [2.818, 2.724, 2.787, 2.757]
inst: mean_all
operation: mean_all(residuals)
tarray: [45871.0, 46407.0, 46944.0, 47481.0]
source_col: residuals
SAMP_residuals_STD | 0.040286474157
varray: [2.818, 2.724, 2.787, 2.757]
inst: std_all
operation: std_all(residuals)
tarray: [45871.0, 46407.0, 46944.0, 47481.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.46275
varray: [-0.437, -0.638, -0.364, -0.412]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [45871.0, 46407.0, 46944.0, 47481.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.120696244073
varray: [-0.437, -0.638, -0.364, -0.412]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [45871.0, 46407.0, 46944.0, 47481.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1776.8745
varray: [1784.819, 1772.166, 1775.586, 1774.927]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [45871.0, 46407.0, 46944.0, 47481.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 5.49963759412
varray: [1784.819, 1772.166, 1775.586, 1774.927]
inst: std_all
operation: std_all(baseline_shift)
tarray: [45871.0, 46407.0, 46944.0, 47481.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -16.88575
varray: [-16.738, -17.007, -16.939, -16.859]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [45871.0, 46407.0, 46944.0, 47481.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.115589431466
varray: [-16.738, -17.007, -16.939, -16.859]
inst: std_all
operation: std_all(slope_shift)
tarray: [45871.0, 46407.0, 46944.0, 47481.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0418568990729
varray: [1.122, 1.156, 1.056, 1.122]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45871.0, 46407.0, 46944.0, 47481.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.290900183339
varray: [0.706, 1.296, 0.911, 1.286]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45871.0, 46407.0, 46944.0, 47481.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 104.207
varray: [105.445, 103.188, 104.838, 103.357]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [45871.0, 46407.0, 46944.0, 47481.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.10930398599
varray: [105.445, 103.188, 104.838, 103.357]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [45871.0, 46407.0, 46944.0, 47481.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.7715
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.46275
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1776.8745
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 16.88575
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.12385
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.12385
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.27754848291
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -7.1528
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [1.122, 1.156, 1.056, 1.122]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [45871.0, 46407.0, 46944.0, 47481.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [0.706, 1.296, 0.911, 1.286]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [45871.0, 46407.0, 46944.0, 47481.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.114
varray: [1.122, 1.156, 1.056, 1.122]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45871.0, 46407.0, 46944.0, 47481.0]
source_col: d(18_16)Mean
dD_mean | 1.04975
varray: [0.706, 1.296, 0.911, 1.286]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45871.0, 46407.0, 46944.0, 47481.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0418568990729
varray: [1.122, 1.156, 1.056, 1.122]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45871.0, 46407.0, 46944.0, 47481.0]
source_col: d(18_16)Mean
dD_stdev | 0.290900183339
varray: [0.706, 1.296, 0.911, 1.286]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45871.0, 46407.0, 46944.0, 47481.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.114
varray: [1.122, 1.156, 1.056, 1.122]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45871.0, 46407.0, 46944.0, 47481.0]
source_col: d(18_16)Mean
measured_dD_mean | 1.04975
varray: [0.706, 1.296, 0.911, 1.286]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45871.0, 46407.0, 46944.0, 47481.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0418568990729
varray: [1.122, 1.156, 1.056, 1.122]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45871.0, 46407.0, 46944.0, 47481.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.290900183339
varray: [0.706, 1.296, 0.911, 1.286]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45871.0, 46407.0, 46944.0, 47481.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.55846703309
actual_array: [ 1.93 -15.84]
varray: [1.122, 1.156, 1.056, 1.122]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 1.114
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [45871.0, 46407.0, 46944.0, 47481.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -16.7009071817
actual_array: [ 18.1 -119. ]
varray: [0.706, 1.296, 0.911, 1.286]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 1.04975
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [45871.0, 46407.0, 46944.0, 47481.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.70945
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 481.39355
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 17 - Sample 17 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 17
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.87425
varray: [2.881, 2.834, 2.909, 2.873]
inst: mean_all
operation: mean_all(residuals)
tarray: [48069.0, 48605.0, 49141.0, 49678.0]
source_col: residuals
SAMP_residuals_STD | 0.0309556133843
varray: [2.881, 2.834, 2.909, 2.873]
inst: std_all
operation: std_all(residuals)
tarray: [48069.0, 48605.0, 49141.0, 49678.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.184
varray: [0.296, 0.257, 0.226, -0.043]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [48069.0, 48605.0, 49141.0, 49678.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.154019479288
varray: [0.296, 0.257, 0.226, -0.043]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [48069.0, 48605.0, 49141.0, 49678.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1299.66575
varray: [1304.044, 1297.529, 1298.779, 1298.311]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [48069.0, 48605.0, 49141.0, 49678.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.96403141852
varray: [1304.044, 1297.529, 1298.779, 1298.311]
inst: std_all
operation: std_all(baseline_shift)
tarray: [48069.0, 48605.0, 49141.0, 49678.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.55725
varray: [1.368, 1.631, 1.775, 1.455]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [48069.0, 48605.0, 49141.0, 49678.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.181773439571
varray: [1.368, 1.631, 1.775, 1.455]
inst: std_all
operation: std_all(slope_shift)
tarray: [48069.0, 48605.0, 49141.0, 49678.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0597571474107
varray: [-0.791, -0.918, -0.792, -0.826]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48069.0, 48605.0, 49141.0, 49678.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.323661345648
varray: [-4.835, -5.616, -5.324, -5.197]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48069.0, 48605.0, 49141.0, 49678.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.88225
varray: [111.065, 109.431, 112.114, 110.919]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [48069.0, 48605.0, 49141.0, 49678.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.10424676439
varray: [111.065, 109.431, 112.114, 110.919]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [48069.0, 48605.0, 49141.0, 49678.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.87425
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.184
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1299.66575
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.55725
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0211
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0211
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.728754727408
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.47755
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [-0.791, -0.918, -0.792, -0.826]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [48069.0, 48605.0, 49141.0, 49678.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-4.835, -5.616, -5.324, -5.197]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [48069.0, 48605.0, 49141.0, 49678.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.83175
varray: [-0.791, -0.918, -0.792, -0.826]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48069.0, 48605.0, 49141.0, 49678.0]
source_col: d(18_16)Mean
dD_mean | -5.243
varray: [-4.835, -5.616, -5.324, -5.197]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48069.0, 48605.0, 49141.0, 49678.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0597571474107
varray: [-0.791, -0.918, -0.792, -0.826]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48069.0, 48605.0, 49141.0, 49678.0]
source_col: d(18_16)Mean
dD_stdev | 0.323661345648
varray: [-4.835, -5.616, -5.324, -5.197]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48069.0, 48605.0, 49141.0, 49678.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.83175
varray: [-0.791, -0.918, -0.792, -0.826]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48069.0, 48605.0, 49141.0, 49678.0]
source_col: d(18_16)Mean
measured_dD_mean | -5.243
varray: [-4.835, -5.616, -5.324, -5.197]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48069.0, 48605.0, 49141.0, 49678.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0597571474107
varray: [-0.791, -0.918, -0.792, -0.826]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48069.0, 48605.0, 49141.0, 49678.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.323661345648
varray: [-4.835, -5.616, -5.324, -5.197]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48069.0, 48605.0, 49141.0, 49678.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.50163634678
actual_array: [ 1.93 -15.84]
varray: [-0.791, -0.918, -0.792, -0.826]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -0.83175
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [48069.0, 48605.0, 49141.0, 49678.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -22.5853353786
actual_array: [ 18.1 -119. ]
varray: [-4.835, -5.616, -5.324, -5.197]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -5.243
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [48069.0, 48605.0, 49141.0, 49678.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.6067
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 4.1848
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 18 - Sample 18 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 18
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.87075
varray: [2.894, 2.844, 2.882, 2.863]
inst: mean_all
operation: mean_all(residuals)
tarray: [50266.0, 50802.0, 51339.0, 51876.0]
source_col: residuals
SAMP_residuals_STD | 0.0219298122807
varray: [2.894, 2.844, 2.882, 2.863]
inst: std_all
operation: std_all(residuals)
tarray: [50266.0, 50802.0, 51339.0, 51876.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.32225
varray: [0.476, 0.208, 0.225, 0.38]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [50266.0, 50802.0, 51339.0, 51876.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.128432537414
varray: [0.476, 0.208, 0.225, 0.38]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [50266.0, 50802.0, 51339.0, 51876.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1295.04325
varray: [1294.005, 1296.394, 1295.494, 1294.28]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [50266.0, 50802.0, 51339.0, 51876.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.1087853339
varray: [1294.005, 1296.394, 1295.494, 1294.28]
inst: std_all
operation: std_all(baseline_shift)
tarray: [50266.0, 50802.0, 51339.0, 51876.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.2635
varray: [1.218, 1.013, 1.175, 1.648]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [50266.0, 50802.0, 51339.0, 51876.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.271104530886
varray: [1.218, 1.013, 1.175, 1.648]
inst: std_all
operation: std_all(slope_shift)
tarray: [50266.0, 50802.0, 51339.0, 51876.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0262090951135
varray: [-1.414, -1.398, -1.445, -1.454]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50266.0, 50802.0, 51339.0, 51876.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.380343156461
varray: [-13.56, -13.956, -14.284, -14.411]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50266.0, 50802.0, 51339.0, 51876.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.90225
varray: [111.18, 111.31, 110.864, 110.255]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [50266.0, 50802.0, 51339.0, 51876.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.470390174926
varray: [111.18, 111.31, 110.864, 110.255]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [50266.0, 50802.0, 51339.0, 51876.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.87075
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.32225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1295.04325
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.2635
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0246
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0246
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.849638212997
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.45755
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [-1.414, -1.398, -1.445, -1.454]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [50266.0, 50802.0, 51339.0, 51876.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-13.56, -13.956, -14.284, -14.411]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [50266.0, 50802.0, 51339.0, 51876.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.42775
varray: [-1.414, -1.398, -1.445, -1.454]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50266.0, 50802.0, 51339.0, 51876.0]
source_col: d(18_16)Mean
dD_mean | -14.05275
varray: [-13.56, -13.956, -14.284, -14.411]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50266.0, 50802.0, 51339.0, 51876.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0262090951135
varray: [-1.414, -1.398, -1.445, -1.454]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50266.0, 50802.0, 51339.0, 51876.0]
source_col: d(18_16)Mean
dD_stdev | 0.380343156461
varray: [-13.56, -13.956, -14.284, -14.411]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50266.0, 50802.0, 51339.0, 51876.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.42775
varray: [-1.414, -1.398, -1.445, -1.454]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50266.0, 50802.0, 51339.0, 51876.0]
source_col: d(18_16)Mean
measured_dD_mean | -14.05275
varray: [-13.56, -13.956, -14.284, -14.411]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50266.0, 50802.0, 51339.0, 51876.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0262090951135
varray: [-1.414, -1.398, -1.445, -1.454]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50266.0, 50802.0, 51339.0, 51876.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.380343156461
varray: [-13.56, -13.956, -14.284, -14.411]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50266.0, 50802.0, 51339.0, 51876.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.09684586031
actual_array: [ 1.93 -15.84]
varray: [-1.414, -1.398, -1.445, -1.454]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -1.42775
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [50266.0, 50802.0, 51339.0, 51876.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -30.823441343
actual_array: [ 18.1 -119. ]
varray: [-13.56, -13.956, -14.284, -14.411]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -14.05275
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [50266.0, 50802.0, 51339.0, 51876.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.6102
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.4377
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 19 - Sample 19 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 19
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.89725
varray: [2.875, 2.892, 2.909, 2.913]
inst: mean_all
operation: mean_all(residuals)
tarray: [52464.0, 53000.0, 53536.0, 54073.0]
source_col: residuals
SAMP_residuals_STD | 0.0174045013335
varray: [2.875, 2.892, 2.909, 2.913]
inst: std_all
operation: std_all(residuals)
tarray: [52464.0, 53000.0, 53536.0, 54073.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.48225
varray: [0.588, 0.394, 0.553, 0.394]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [52464.0, 53000.0, 53536.0, 54073.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.102899222543
varray: [0.588, 0.394, 0.553, 0.394]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [52464.0, 53000.0, 53536.0, 54073.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1297.679
varray: [1295.658, 1299.06, 1297.827, 1298.171]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [52464.0, 53000.0, 53536.0, 54073.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.44401869794
varray: [1295.658, 1299.06, 1297.827, 1298.171]
inst: std_all
operation: std_all(baseline_shift)
tarray: [52464.0, 53000.0, 53536.0, 54073.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.28975
varray: [1.639, 1.239, 1.031, 1.25]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [52464.0, 53000.0, 53536.0, 54073.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.253694534168
varray: [1.639, 1.239, 1.031, 1.25]
inst: std_all
operation: std_all(slope_shift)
tarray: [52464.0, 53000.0, 53536.0, 54073.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0921244629111
varray: [0.043, 0.264, 0.193, 0.167]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52464.0, 53000.0, 53536.0, 54073.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.781793237798
varray: [0.601, 1.905, 2.15, 2.324]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52464.0, 53000.0, 53536.0, 54073.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.40575
varray: [110.084, 111.614, 112.217, 111.708]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [52464.0, 53000.0, 53536.0, 54073.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.920121866929
varray: [110.084, 111.614, 112.217, 111.708]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [52464.0, 53000.0, 53536.0, 54073.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.89725
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.48225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1297.679
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.28975
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0019
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0019
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.0656224636054
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.04595
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [0.043, 0.264, 0.193, 0.167]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [52464.0, 53000.0, 53536.0, 54073.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [0.601, 1.905, 2.15, 2.324]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [52464.0, 53000.0, 53536.0, 54073.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.16675
varray: [0.043, 0.264, 0.193, 0.167]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [52464.0, 53000.0, 53536.0, 54073.0]
source_col: d(18_16)Mean
dD_mean | 1.745
varray: [0.601, 1.905, 2.15, 2.324]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [52464.0, 53000.0, 53536.0, 54073.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0921244629111
varray: [0.043, 0.264, 0.193, 0.167]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52464.0, 53000.0, 53536.0, 54073.0]
source_col: d(18_16)Mean
dD_stdev | 0.781793237798
varray: [0.601, 1.905, 2.15, 2.324]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52464.0, 53000.0, 53536.0, 54073.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.16675
varray: [0.043, 0.264, 0.193, 0.167]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [52464.0, 53000.0, 53536.0, 54073.0]
source_col: d(18_16)Mean
measured_dD_mean | 1.745
varray: [0.601, 1.905, 2.15, 2.324]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [52464.0, 53000.0, 53536.0, 54073.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0921244629111
varray: [0.043, 0.264, 0.193, 0.167]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52464.0, 53000.0, 53536.0, 54073.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.781793237798
varray: [0.601, 1.905, 2.15, 2.324]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52464.0, 53000.0, 53536.0, 54073.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.50446067687
actual_array: [ 1.93 -15.84]
varray: [0.043, 0.264, 0.193, 0.167]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.16675
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [52464.0, 53000.0, 53536.0, 54073.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -16.0507703252
actual_array: [ 18.1 -119. ]
varray: [0.601, 1.905, 2.15, 2.324]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 1.745
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [52464.0, 53000.0, 53536.0, 54073.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.5837
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 2.19805
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 20 - Sample 20 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 20
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.871
varray: [2.792, 2.901, 2.928, 2.863]
inst: mean_all
operation: mean_all(residuals)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: residuals
SAMP_residuals_STD | 0.0590310652679
varray: [2.792, 2.901, 2.928, 2.863]
inst: std_all
operation: std_all(residuals)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.286
varray: [0.146, 0.302, 0.404, 0.292]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.10616967552
varray: [0.146, 0.302, 0.404, 0.292]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1295.3355
varray: [1294.299, 1297.773, 1295.422, 1293.848]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.75460128424
varray: [1294.299, 1297.773, 1295.422, 1293.848]
inst: std_all
operation: std_all(baseline_shift)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.485
varray: [1.686, 1.533, 1.842, 0.879]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.423237521966
varray: [1.686, 1.533, 1.842, 0.879]
inst: std_all
operation: std_all(slope_shift)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0155134350376
varray: [-0.396, -0.43, -0.425, -0.409]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.359142033185
varray: [-5.383, -5.521, -5.898, -6.172]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.72325
varray: [108.89, 112.313, 111.177, 110.513]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.4304084207
varray: [108.89, 112.313, 111.177, 110.513]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.871
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.286
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1295.3355
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.485
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.02435
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.02435
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.841003678312
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.63655
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [-0.396, -0.43, -0.425, -0.409]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-5.383, -5.521, -5.898, -6.172]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.415
varray: [-0.396, -0.43, -0.425, -0.409]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: d(18_16)Mean
dD_mean | -5.7435
varray: [-5.383, -5.521, -5.898, -6.172]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0155134350376
varray: [-0.396, -0.43, -0.425, -0.409]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: d(18_16)Mean
dD_stdev | 0.359142033185
varray: [-5.383, -5.521, -5.898, -6.172]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.415
varray: [-0.396, -0.43, -0.425, -0.409]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: d(18_16)Mean
measured_dD_mean | -5.7435
varray: [-5.383, -5.521, -5.898, -6.172]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0155134350376
varray: [-0.396, -0.43, -0.425, -0.409]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.359142033185
varray: [-5.383, -5.521, -5.898, -6.172]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.08543909046
actual_array: [ 1.93 -15.84]
varray: [-0.396, -0.43, -0.425, -0.409]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -0.415
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -23.0533591063
actual_array: [ 18.1 -119. ]
varray: [-5.383, -5.521, -5.898, -6.172]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -5.7435
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [54662.0, 55198.0, 55735.0, 56272.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.60995
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.14545
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 21 - Sample 21 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 21
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9045
varray: [2.878, 2.907, 2.924, 2.909]
inst: mean_all
operation: mean_all(residuals)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: residuals
SAMP_residuals_STD | 0.0192267175219
varray: [2.878, 2.907, 2.924, 2.909]
inst: std_all
operation: std_all(residuals)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.094
varray: [-0.061, -0.165, 0.105, -0.255]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.154544491976
varray: [-0.061, -0.165, 0.105, -0.255]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1342.716
varray: [1342.325, 1344.207, 1341.457, 1342.875]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.15272488768
varray: [1342.325, 1344.207, 1341.457, 1342.875]
inst: std_all
operation: std_all(baseline_shift)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -1.715
varray: [-1.934, -1.416, -1.84, -1.67]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.22731182694
varray: [-1.934, -1.416, -1.84, -1.67]
inst: std_all
operation: std_all(slope_shift)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: slope_shift
SAMP_d18O_STD | 0.00932737905309
varray: [0.926, 0.933, 0.923, 0.944]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.166375829575
varray: [-1.17, -0.831, -0.947, -0.805]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.21925
varray: [113.502, 112.757, 113.342, 113.276]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.322444181216
varray: [113.502, 112.757, 113.342, 113.276]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9045
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.094
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1342.716
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.715
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.00915
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.00915
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.316023969468
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.85945
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [0.926, 0.933, 0.923, 0.944]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-1.17, -0.831, -0.947, -0.805]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.9315
varray: [0.926, 0.933, 0.923, 0.944]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: d(18_16)Mean
dD_mean | -0.93825
varray: [-1.17, -0.831, -0.947, -0.805]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: d(D_H)Mean
d18O_stdev | 0.00932737905309
varray: [0.926, 0.933, 0.923, 0.944]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: d(18_16)Mean
dD_stdev | 0.166375829575
varray: [-1.17, -0.831, -0.947, -0.805]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.9315
varray: [0.926, 0.933, 0.923, 0.944]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: d(18_16)Mean
measured_dD_mean | -0.93825
varray: [-1.17, -0.831, -0.947, -0.805]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.00932737905309
varray: [0.926, 0.933, 0.923, 0.944]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.166375829575
varray: [-1.17, -0.831, -0.947, -0.805]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.74072497977
actual_array: [ 1.93 -15.84]
varray: [0.926, 0.933, 0.923, 0.944]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.9315
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -18.5599105198
actual_array: [ 18.1 -119. ]
varray: [-1.17, -0.831, -0.947, -0.805]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -0.93825
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [56861.0, 57397.0, 57934.0, 58470.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.57645
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 47.23505
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 22 - Sample 22 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 22
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.855
varray: [2.824, 2.846, 2.874, 2.876]
inst: mean_all
operation: mean_all(residuals)
tarray: [59055.0, 59592.0, 60129.0, 60665.0]
source_col: residuals
SAMP_residuals_STD | 0.0247924719757
varray: [2.824, 2.846, 2.874, 2.876]
inst: std_all
operation: std_all(residuals)
tarray: [59055.0, 59592.0, 60129.0, 60665.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.511
varray: [0.432, 0.477, 0.64, 0.495]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [59055.0, 59592.0, 60129.0, 60665.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0899888882029
varray: [0.432, 0.477, 0.64, 0.495]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [59055.0, 59592.0, 60129.0, 60665.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1293.5765
varray: [1293.085, 1293.921, 1293.696, 1293.604]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [59055.0, 59592.0, 60129.0, 60665.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.353689600262
varray: [1293.085, 1293.921, 1293.696, 1293.604]
inst: std_all
operation: std_all(baseline_shift)
tarray: [59055.0, 59592.0, 60129.0, 60665.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.763
varray: [1.971, 1.861, 1.314, 1.906]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [59055.0, 59592.0, 60129.0, 60665.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.30271989253
varray: [1.971, 1.861, 1.314, 1.906]
inst: std_all
operation: std_all(slope_shift)
tarray: [59055.0, 59592.0, 60129.0, 60665.0]
source_col: slope_shift
SAMP_d18O_STD | 0.171155630933
varray: [-7.047, -7.333, -7.422, -7.389]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59055.0, 59592.0, 60129.0, 60665.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.05866510916
varray: [-51.913, -57.301, -57.885, -58.618]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59055.0, 59592.0, 60129.0, 60665.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.31775
varray: [108.692, 109.946, 112.13, 110.503]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [59055.0, 59592.0, 60129.0, 60665.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.42592995971
varray: [108.692, 109.946, 112.13, 110.503]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [59055.0, 59592.0, 60129.0, 60665.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.855
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.511
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1293.5765
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.763
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.04035
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.04035
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.39361389815
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.04205
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [-7.047, -7.333, -7.422, -7.389]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [59055.0, 59592.0, 60129.0, 60665.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-51.913, -57.301, -57.885, -58.618]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [59055.0, 59592.0, 60129.0, 60665.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.29775
varray: [-7.047, -7.333, -7.422, -7.389]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [59055.0, 59592.0, 60129.0, 60665.0]
source_col: d(18_16)Mean
dD_mean | -56.42925
varray: [-51.913, -57.301, -57.885, -58.618]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [59055.0, 59592.0, 60129.0, 60665.0]
source_col: d(D_H)Mean
d18O_stdev | 0.171155630933
varray: [-7.047, -7.333, -7.422, -7.389]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59055.0, 59592.0, 60129.0, 60665.0]
source_col: d(18_16)Mean
dD_stdev | 3.05866510916
varray: [-51.913, -57.301, -57.885, -58.618]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59055.0, 59592.0, 60129.0, 60665.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.29775
varray: [-7.047, -7.333, -7.422, -7.389]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [59055.0, 59592.0, 60129.0, 60665.0]
source_col: d(18_16)Mean
measured_dD_mean | -56.42925
varray: [-51.913, -57.301, -57.885, -58.618]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [59055.0, 59592.0, 60129.0, 60665.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.171155630933
varray: [-7.047, -7.333, -7.422, -7.389]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59055.0, 59592.0, 60129.0, 60665.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.05866510916
varray: [-51.913, -57.301, -57.885, -58.618]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59055.0, 59592.0, 60129.0, 60665.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.9590603644
actual_array: [ 1.93 -15.84]
varray: [-7.047, -7.333, -7.422, -7.389]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.29775
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [59055.0, 59592.0, 60129.0, 60665.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -70.4502295493
actual_array: [ 18.1 -119. ]
varray: [-51.913, -57.301, -57.885, -58.618]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -56.42925
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [59055.0, 59592.0, 60129.0, 60665.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.62595
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.90445
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 23 - Sample 23 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 23
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8835
varray: [2.871, 2.881, 2.899, 2.883]
inst: mean_all
operation: mean_all(residuals)
tarray: [61251.0, 61788.0, 62324.0, 62861.0]
source_col: residuals
SAMP_residuals_STD | 0.0115902257671
varray: [2.871, 2.881, 2.899, 2.883]
inst: std_all
operation: std_all(residuals)
tarray: [61251.0, 61788.0, 62324.0, 62861.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3225
varray: [0.36, 0.35, 0.327, 0.253]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [61251.0, 61788.0, 62324.0, 62861.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0483494226094
varray: [0.36, 0.35, 0.327, 0.253]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [61251.0, 61788.0, 62324.0, 62861.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1297.6425
varray: [1296.381, 1297.695, 1298.186, 1298.308]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [61251.0, 61788.0, 62324.0, 62861.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.881744672038
varray: [1296.381, 1297.695, 1298.186, 1298.308]
inst: std_all
operation: std_all(baseline_shift)
tarray: [61251.0, 61788.0, 62324.0, 62861.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.10975
varray: [0.885, 1.385, 1.258, 0.911]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [61251.0, 61788.0, 62324.0, 62861.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.250169775686
varray: [0.885, 1.385, 1.258, 0.911]
inst: std_all
operation: std_all(slope_shift)
tarray: [61251.0, 61788.0, 62324.0, 62861.0]
source_col: slope_shift
SAMP_d18O_STD | 0.201548174556
varray: [-0.502, -0.179, -0.075, -0.078]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61251.0, 61788.0, 62324.0, 62861.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.29385503393
varray: [-15.701, -12.088, -11.057, -10.671]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61251.0, 61788.0, 62324.0, 62861.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.969
varray: [109.658, 110.936, 111.337, 111.945]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [61251.0, 61788.0, 62324.0, 62861.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.967438197854
varray: [109.658, 110.936, 111.337, 111.945]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [61251.0, 61788.0, 62324.0, 62861.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8835
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1297.6425
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.10975
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.01185
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.01185
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.409276944065
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.3908
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [-0.502, -0.179, -0.075, -0.078]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [61251.0, 61788.0, 62324.0, 62861.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-15.701, -12.088, -11.057, -10.671]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [61251.0, 61788.0, 62324.0, 62861.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.2085
varray: [-0.502, -0.179, -0.075, -0.078]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [61251.0, 61788.0, 62324.0, 62861.0]
source_col: d(18_16)Mean
dD_mean | -12.37925
varray: [-15.701, -12.088, -11.057, -10.671]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [61251.0, 61788.0, 62324.0, 62861.0]
source_col: d(D_H)Mean
d18O_stdev | 0.201548174556
varray: [-0.502, -0.179, -0.075, -0.078]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61251.0, 61788.0, 62324.0, 62861.0]
source_col: d(18_16)Mean
dD_stdev | 2.29385503393
varray: [-15.701, -12.088, -11.057, -10.671]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61251.0, 61788.0, 62324.0, 62861.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.2085
varray: [-0.502, -0.179, -0.075, -0.078]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [61251.0, 61788.0, 62324.0, 62861.0]
source_col: d(18_16)Mean
measured_dD_mean | -12.37925
varray: [-15.701, -12.088, -11.057, -10.671]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [61251.0, 61788.0, 62324.0, 62861.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.201548174556
varray: [-0.502, -0.179, -0.075, -0.078]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61251.0, 61788.0, 62324.0, 62861.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.29385503393
varray: [-15.701, -12.088, -11.057, -10.671]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61251.0, 61788.0, 62324.0, 62861.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.87921297545
actual_array: [ 1.93 -15.84]
varray: [-0.502, -0.179, -0.075, -0.078]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -0.2085
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [61251.0, 61788.0, 62324.0, 62861.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -29.2585308369
actual_array: [ 18.1 -119. ]
varray: [-15.701, -12.088, -11.057, -10.671]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -12.37925
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [61251.0, 61788.0, 62324.0, 62861.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.59745
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 2.16155
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 24 - Sample 24 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 24
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.768
varray: [2.768, 2.789, 2.776, 2.739]
inst: mean_all
operation: mean_all(residuals)
tarray: [63448.0, 63985.0, 64521.0, 65058.0]
source_col: residuals
SAMP_residuals_STD | 0.0211817531538
varray: [2.768, 2.789, 2.776, 2.739]
inst: std_all
operation: std_all(residuals)
tarray: [63448.0, 63985.0, 64521.0, 65058.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.214
varray: [-0.304, -0.195, -0.202, -0.155]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [63448.0, 63985.0, 64521.0, 65058.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0634717785056
varray: [-0.304, -0.195, -0.202, -0.155]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [63448.0, 63985.0, 64521.0, 65058.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1380.27525
varray: [1382.702, 1379.838, 1381.146, 1377.415]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [63448.0, 63985.0, 64521.0, 65058.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.237523836
varray: [1382.702, 1379.838, 1381.146, 1377.415]
inst: std_all
operation: std_all(baseline_shift)
tarray: [63448.0, 63985.0, 64521.0, 65058.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -2.82425
varray: [-2.827, -2.873, -2.676, -2.921]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [63448.0, 63985.0, 64521.0, 65058.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.106023189287
varray: [-2.827, -2.873, -2.676, -2.921]
inst: std_all
operation: std_all(slope_shift)
tarray: [63448.0, 63985.0, 64521.0, 65058.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0209821352584
varray: [0.761, 0.735, 0.77, 0.785]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63448.0, 63985.0, 64521.0, 65058.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.227409139365
varray: [-10.435, -10.418, -10.697, -10.893]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63448.0, 63985.0, 64521.0, 65058.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 107.77375
varray: [107.959, 108.346, 108.358, 106.432]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [63448.0, 63985.0, 64521.0, 65058.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.913496715922
varray: [107.959, 108.346, 108.358, 106.432]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [63448.0, 63985.0, 64521.0, 65058.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.768
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.214
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1380.27525
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.82425
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.12735
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.12735
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.3984319685
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.58605
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [0.761, 0.735, 0.77, 0.785]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [63448.0, 63985.0, 64521.0, 65058.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-10.435, -10.418, -10.697, -10.893]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [63448.0, 63985.0, 64521.0, 65058.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.76275
varray: [0.761, 0.735, 0.77, 0.785]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63448.0, 63985.0, 64521.0, 65058.0]
source_col: d(18_16)Mean
dD_mean | -10.61075
varray: [-10.435, -10.418, -10.697, -10.893]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63448.0, 63985.0, 64521.0, 65058.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0209821352584
varray: [0.761, 0.735, 0.77, 0.785]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63448.0, 63985.0, 64521.0, 65058.0]
source_col: d(18_16)Mean
dD_stdev | 0.227409139365
varray: [-10.435, -10.418, -10.697, -10.893]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63448.0, 63985.0, 64521.0, 65058.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.76275
varray: [0.761, 0.735, 0.77, 0.785]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63448.0, 63985.0, 64521.0, 65058.0]
source_col: d(18_16)Mean
measured_dD_mean | -10.61075
varray: [-10.435, -10.418, -10.697, -10.893]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63448.0, 63985.0, 64521.0, 65058.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0209821352584
varray: [0.761, 0.735, 0.77, 0.785]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63448.0, 63985.0, 64521.0, 65058.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.227409139365
varray: [-10.435, -10.418, -10.697, -10.893]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63448.0, 63985.0, 64521.0, 65058.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.90925116334
actual_array: [ 1.93 -15.84]
varray: [0.761, 0.735, 0.77, 0.785]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.76275
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [63448.0, 63985.0, 64521.0, 65058.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -27.6047846582
actual_array: [ 18.1 -119. ]
varray: [-10.435, -10.418, -10.697, -10.893]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -10.61075
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [63448.0, 63985.0, 64521.0, 65058.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.71295
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 84.7943
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 25 - Sample 25 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 25
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.82
varray: [2.825, 2.67, 2.95, 2.835]
inst: mean_all
operation: mean_all(residuals)
tarray: [65645.0, 66182.0, 66718.0, 67255.0]
source_col: residuals
SAMP_residuals_STD | 0.114963762407
varray: [2.825, 2.67, 2.95, 2.835]
inst: std_all
operation: std_all(residuals)
tarray: [65645.0, 66182.0, 66718.0, 67255.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.2345
varray: [0.258, 0.109, 0.534, 0.037]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [65645.0, 66182.0, 66718.0, 67255.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.219855255414
varray: [0.258, 0.109, 0.534, 0.037]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [65645.0, 66182.0, 66718.0, 67255.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1297.85575
varray: [1300.405, 1291.598, 1299.511, 1299.909]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [65645.0, 66182.0, 66718.0, 67255.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 4.18783152916
varray: [1300.405, 1291.598, 1299.511, 1299.909]
inst: std_all
operation: std_all(baseline_shift)
tarray: [65645.0, 66182.0, 66718.0, 67255.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.38075
varray: [0.706, 0.431, -0.008, 0.394]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [65645.0, 66182.0, 66718.0, 67255.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.294173843161
varray: [0.706, 0.431, -0.008, 0.394]
inst: std_all
operation: std_all(slope_shift)
tarray: [65645.0, 66182.0, 66718.0, 67255.0]
source_col: slope_shift
SAMP_d18O_STD | 0.036142080737
varray: [0.619, 0.565, 0.586, 0.647]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65645.0, 66182.0, 66718.0, 67255.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.194077304186
varray: [-8.177, -8.442, -7.97, -8.159]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65645.0, 66182.0, 66718.0, 67255.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.7245
varray: [109.96, 104.241, 113.68, 111.017]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [65645.0, 66182.0, 66718.0, 67255.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.9766304413
varray: [109.96, 104.241, 113.68, 111.017]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [65645.0, 66182.0, 66718.0, 67255.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.82
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.2345
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1297.85575
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.38075
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.07535
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.07535
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.60244875404
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.6353
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [0.619, 0.565, 0.586, 0.647]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [65645.0, 66182.0, 66718.0, 67255.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-8.177, -8.442, -7.97, -8.159]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [65645.0, 66182.0, 66718.0, 67255.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.60425
varray: [0.619, 0.565, 0.586, 0.647]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65645.0, 66182.0, 66718.0, 67255.0]
source_col: d(18_16)Mean
dD_mean | -8.187
varray: [-8.177, -8.442, -7.97, -8.159]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65645.0, 66182.0, 66718.0, 67255.0]
source_col: d(D_H)Mean
d18O_stdev | 0.036142080737
varray: [0.619, 0.565, 0.586, 0.647]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65645.0, 66182.0, 66718.0, 67255.0]
source_col: d(18_16)Mean
dD_stdev | 0.194077304186
varray: [-8.177, -8.442, -7.97, -8.159]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65645.0, 66182.0, 66718.0, 67255.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.60425
varray: [0.619, 0.565, 0.586, 0.647]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65645.0, 66182.0, 66718.0, 67255.0]
source_col: d(18_16)Mean
measured_dD_mean | -8.187
varray: [-8.177, -8.442, -7.97, -8.159]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65645.0, 66182.0, 66718.0, 67255.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.036142080737
varray: [0.619, 0.565, 0.586, 0.647]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65645.0, 66182.0, 66718.0, 67255.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.194077304186
varray: [-8.177, -8.442, -7.97, -8.159]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65645.0, 66182.0, 66718.0, 67255.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.06754094169
actual_array: [ 1.93 -15.84]
varray: [0.619, 0.565, 0.586, 0.647]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.60425
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [65645.0, 66182.0, 66718.0, 67255.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -25.3383061166
actual_array: [ 18.1 -119. ]
varray: [-8.177, -8.442, -7.97, -8.159]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -8.187
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [65645.0, 66182.0, 66718.0, 67255.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.66095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 2.3748
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 26 - Sample 26 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 26
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.765
varray: [2.806, 2.761, 2.809, 2.684]
inst: mean_all
operation: mean_all(residuals)
tarray: [67843.0, 68380.0, 68917.0, 69453.0]
source_col: residuals
SAMP_residuals_STD | 0.0582923665672
varray: [2.806, 2.761, 2.809, 2.684]
inst: std_all
operation: std_all(residuals)
tarray: [67843.0, 68380.0, 68917.0, 69453.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.683
varray: [-0.793, -0.658, -0.529, -0.752]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [67843.0, 68380.0, 68917.0, 69453.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.117192149908
varray: [-0.793, -0.658, -0.529, -0.752]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [67843.0, 68380.0, 68917.0, 69453.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1352.04175
varray: [1356.608, 1352.85, 1352.272, 1346.437]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [67843.0, 68380.0, 68917.0, 69453.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 4.20198939988
varray: [1356.608, 1352.85, 1352.272, 1346.437]
inst: std_all
operation: std_all(baseline_shift)
tarray: [67843.0, 68380.0, 68917.0, 69453.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -3.68125
varray: [-4.091, -3.594, -3.88, -3.16]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [67843.0, 68380.0, 68917.0, 69453.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.40278644052
varray: [-4.091, -3.594, -3.88, -3.16]
inst: std_all
operation: std_all(slope_shift)
tarray: [67843.0, 68380.0, 68917.0, 69453.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0543407459156
varray: [1.774, 1.683, 1.691, 1.645]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67843.0, 68380.0, 68917.0, 69453.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.178920839107
varray: [-2.4, -2.303, -2.151, -2.578]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67843.0, 68380.0, 68917.0, 69453.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.9475
varray: [112.056, 109.134, 110.221, 104.379]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [67843.0, 68380.0, 68917.0, 69453.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.27569580395
varray: [112.056, 109.134, 110.221, 104.379]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [67843.0, 68380.0, 68917.0, 69453.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.765
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.683
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1352.04175
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 3.68125
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.13035
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.13035
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.50204638472
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.4123
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [1.774, 1.683, 1.691, 1.645]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [67843.0, 68380.0, 68917.0, 69453.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-2.4, -2.303, -2.151, -2.578]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [67843.0, 68380.0, 68917.0, 69453.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.69825
varray: [1.774, 1.683, 1.691, 1.645]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [67843.0, 68380.0, 68917.0, 69453.0]
source_col: d(18_16)Mean
dD_mean | -2.358
varray: [-2.4, -2.303, -2.151, -2.578]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [67843.0, 68380.0, 68917.0, 69453.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0543407459156
varray: [1.774, 1.683, 1.691, 1.645]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67843.0, 68380.0, 68917.0, 69453.0]
source_col: d(18_16)Mean
dD_stdev | 0.178920839107
varray: [-2.4, -2.303, -2.151, -2.578]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67843.0, 68380.0, 68917.0, 69453.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.69825
varray: [1.774, 1.683, 1.691, 1.645]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [67843.0, 68380.0, 68917.0, 69453.0]
source_col: d(18_16)Mean
measured_dD_mean | -2.358
varray: [-2.4, -2.303, -2.151, -2.578]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [67843.0, 68380.0, 68917.0, 69453.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0543407459156
varray: [1.774, 1.683, 1.691, 1.645]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67843.0, 68380.0, 68917.0, 69453.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.178920839107
varray: [-2.4, -2.303, -2.151, -2.578]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67843.0, 68380.0, 68917.0, 69453.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.974991935303
actual_array: [ 1.93 -15.84]
varray: [1.774, 1.683, 1.691, 1.645]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 1.69825
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [67843.0, 68380.0, 68917.0, 69453.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -19.8875362689
actual_array: [ 18.1 -119. ]
varray: [-2.4, -2.303, -2.151, -2.578]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -2.358
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [67843.0, 68380.0, 68917.0, 69453.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.71595
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 56.5608
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 27 - Sample 27 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 27
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.752
varray: [2.736, 2.808, 2.763, 2.701]
inst: mean_all
operation: mean_all(residuals)
tarray: [70042.0, 70578.0, 71115.0, 71651.0]
source_col: residuals
SAMP_residuals_STD | 0.0451442133612
varray: [2.736, 2.808, 2.763, 2.701]
inst: std_all
operation: std_all(residuals)
tarray: [70042.0, 70578.0, 71115.0, 71651.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.43775
varray: [-0.559, -0.124, -0.597, -0.471]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [70042.0, 70578.0, 71115.0, 71651.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.215721077629
varray: [-0.559, -0.124, -0.597, -0.471]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [70042.0, 70578.0, 71115.0, 71651.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1308.3715
varray: [1308.872, 1308.951, 1309.144, 1306.519]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [70042.0, 70578.0, 71115.0, 71651.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.24027322259
varray: [1308.872, 1308.951, 1309.144, 1306.519]
inst: std_all
operation: std_all(baseline_shift)
tarray: [70042.0, 70578.0, 71115.0, 71651.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -0.55325
varray: [0.131, -0.843, -0.965, -0.536]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [70042.0, 70578.0, 71115.0, 71651.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.490574068482
varray: [0.131, -0.843, -0.965, -0.536]
inst: std_all
operation: std_all(slope_shift)
tarray: [70042.0, 70578.0, 71115.0, 71651.0]
source_col: slope_shift
SAMP_d18O_STD | 0.056323470537
varray: [0.458, 0.357, 0.411, 0.332]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70042.0, 70578.0, 71115.0, 71651.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.774901445346
varray: [-12.03, -13.412, -13.602, -13.677]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70042.0, 70578.0, 71115.0, 71651.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.841
varray: [107.039, 111.261, 110.267, 106.797]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [70042.0, 70578.0, 71115.0, 71651.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.25942588578
varray: [107.039, 111.261, 110.267, 106.797]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [70042.0, 70578.0, 71115.0, 71651.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.752
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.43775
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1308.3715
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.55325
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.14335
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.14335
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.95104218834
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.5188
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [0.458, 0.357, 0.411, 0.332]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [70042.0, 70578.0, 71115.0, 71651.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-12.03, -13.412, -13.602, -13.677]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [70042.0, 70578.0, 71115.0, 71651.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.3895
varray: [0.458, 0.357, 0.411, 0.332]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [70042.0, 70578.0, 71115.0, 71651.0]
source_col: d(18_16)Mean
dD_mean | -13.18025
varray: [-12.03, -13.412, -13.602, -13.677]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [70042.0, 70578.0, 71115.0, 71651.0]
source_col: d(D_H)Mean
d18O_stdev | 0.056323470537
varray: [0.458, 0.357, 0.411, 0.332]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70042.0, 70578.0, 71115.0, 71651.0]
source_col: d(18_16)Mean
dD_stdev | 0.774901445346
varray: [-12.03, -13.412, -13.602, -13.677]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70042.0, 70578.0, 71115.0, 71651.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.3895
varray: [0.458, 0.357, 0.411, 0.332]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [70042.0, 70578.0, 71115.0, 71651.0]
source_col: d(18_16)Mean
measured_dD_mean | -13.18025
varray: [-12.03, -13.412, -13.602, -13.677]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [70042.0, 70578.0, 71115.0, 71651.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.056323470537
varray: [0.458, 0.357, 0.411, 0.332]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70042.0, 70578.0, 71115.0, 71651.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.774901445346
varray: [-12.03, -13.412, -13.602, -13.677]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70042.0, 70578.0, 71115.0, 71651.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.28200611456
actual_array: [ 1.93 -15.84]
varray: [0.458, 0.357, 0.411, 0.332]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.3895
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [70042.0, 70578.0, 71115.0, 71651.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -30.0075558237
actual_array: [ 18.1 -119. ]
varray: [-12.03, -13.412, -13.602, -13.677]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -13.18025
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [70042.0, 70578.0, 71115.0, 71651.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.72895
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 12.89055
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 28 - Sample 28 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 28
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1105
varray: [0.0, 2.814, 2.814, 2.814]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [72239.0, 72775.0, 73312.0, 73850.0]
source_col: residuals
SAMP_residuals_STD | 1.407
varray: [0.0, 2.814, 2.814, 2.814]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [72239.0, 72775.0, 73312.0, 73850.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.381
varray: [0.0, -0.508, -0.508, -0.508]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [72239.0, 72775.0, 73312.0, 73850.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.254
varray: [0.0, -0.508, -0.508, -0.508]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [72239.0, 72775.0, 73312.0, 73850.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1580.5485
varray: [0.0, 2107.398, 2107.398, 2107.398]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [72239.0, 72775.0, 73312.0, 73850.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1053.699
varray: [0.0, 2107.398, 2107.398, 2107.398]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [72239.0, 72775.0, 73312.0, 73850.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -21.53475
varray: [0.0, -28.713, -28.713, -28.713]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [72239.0, 72775.0, 73312.0, 73850.0]
source_col: slope_shift
SAMP_slope_shift_STD | 14.3565
varray: [0.0, -28.713, -28.713, -28.713]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [72239.0, 72775.0, 73312.0, 73850.0]
source_col: slope_shift
SAMP_d18O_STD | 0.4445
varray: [0.0, 0.889, 0.889, 0.889]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72239.0, 72775.0, 73312.0, 73850.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.5835
varray: [0.0, -3.167, -3.167, -3.167]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72239.0, 72775.0, 73312.0, 73850.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 77.4585
varray: [0.0, 103.278, 103.278, 103.278]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [72239.0, 72775.0, 73312.0, 73850.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 51.639
varray: [0.0, 103.278, 103.278, 103.278]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [72239.0, 72775.0, 73312.0, 73850.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1105
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.381
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1580.5485
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 21.53475
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.78485
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.78485
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -27.1072581899
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -33.9013
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [0.0, 0.889, 0.889, 0.889]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [72239.0, 72775.0, 73312.0, 73850.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [0.0, -3.167, -3.167, -3.167]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [72239.0, 72775.0, 73312.0, 73850.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.66675
varray: [0.0, 0.889, 0.889, 0.889]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72239.0, 72775.0, 73312.0, 73850.0]
source_col: d(18_16)Mean
dD_mean | -2.37525
varray: [0.0, -3.167, -3.167, -3.167]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72239.0, 72775.0, 73312.0, 73850.0]
source_col: d(D_H)Mean
d18O_stdev | 0.4445
varray: [0.0, 0.889, 0.889, 0.889]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72239.0, 72775.0, 73312.0, 73850.0]
source_col: d(18_16)Mean
dD_stdev | 1.5835
varray: [0.0, -3.167, -3.167, -3.167]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72239.0, 72775.0, 73312.0, 73850.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.66675
varray: [0.0, 0.889, 0.889, 0.889]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72239.0, 72775.0, 73312.0, 73850.0]
source_col: d(18_16)Mean
measured_dD_mean | -2.37525
varray: [0.0, -3.167, -3.167, -3.167]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72239.0, 72775.0, 73312.0, 73850.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.4445
varray: [0.0, 0.889, 0.889, 0.889]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72239.0, 72775.0, 73312.0, 73850.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.5835
varray: [0.0, -3.167, -3.167, -3.167]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72239.0, 72775.0, 73312.0, 73850.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.00512383666
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.889, 0.889, 0.889]
fit_rsq: 1.0
fit_intercept: 2.6745368036
except: True
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.66675
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [72239.0, 72775.0, 73312.0, 73850.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -19.9036669568
actual_array: [ 18.1 -119. ]
varray: [0.0, -3.167, -3.167, -3.167]
fit_rsq: 1.0
fit_intercept: 18.9095369074
except: True
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -2.37525
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [72239.0, 72775.0, 73312.0, 73850.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1293.37045
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 285.06755
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 29 - Sample 29 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 29
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.72675
varray: [2.656, 2.757, 2.766, 2.728]
inst: mean_all
operation: mean_all(residuals)
tarray: [74435.0, 74972.0, 75509.0, 76046.0]
source_col: residuals
SAMP_residuals_STD | 0.0498756787757
varray: [2.656, 2.757, 2.766, 2.728]
inst: std_all
operation: std_all(residuals)
tarray: [74435.0, 74972.0, 75509.0, 76046.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9135
varray: [-1.01, -0.775, -0.873, -0.996]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [74435.0, 74972.0, 75509.0, 76046.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.11096696205
varray: [-1.01, -0.775, -0.873, -0.996]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [74435.0, 74972.0, 75509.0, 76046.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 2240.1725
varray: [2238.908, 2242.039, 2242.677, 2237.066]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [74435.0, 74972.0, 75509.0, 76046.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.64610688875
varray: [2238.908, 2242.039, 2242.677, 2237.066]
inst: std_all
operation: std_all(baseline_shift)
tarray: [74435.0, 74972.0, 75509.0, 76046.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -34.50975
varray: [-34.559, -34.808, -35.138, -33.534]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [74435.0, 74972.0, 75509.0, 76046.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.692378689736
varray: [-34.559, -34.808, -35.138, -33.534]
inst: std_all
operation: std_all(slope_shift)
tarray: [74435.0, 74972.0, 75509.0, 76046.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0202484567313
varray: [0.926, 0.897, 0.885, 0.924]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74435.0, 74972.0, 75509.0, 76046.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.261489324193
varray: [-5.473, -5.966, -5.805, -6.072]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74435.0, 74972.0, 75509.0, 76046.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 98.8205
varray: [98.038, 100.061, 99.606, 97.577]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [74435.0, 74972.0, 75509.0, 76046.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.19922877995
varray: [98.038, 100.061, 99.606, 97.577]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [74435.0, 74972.0, 75509.0, 76046.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.72675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9135
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 2240.1725
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 34.50975
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1686
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1686
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.82313019151
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -12.5393
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [0.926, 0.897, 0.885, 0.924]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [74435.0, 74972.0, 75509.0, 76046.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-5.473, -5.966, -5.805, -6.072]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [74435.0, 74972.0, 75509.0, 76046.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.908
varray: [0.926, 0.897, 0.885, 0.924]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [74435.0, 74972.0, 75509.0, 76046.0]
source_col: d(18_16)Mean
dD_mean | -5.829
varray: [-5.473, -5.966, -5.805, -6.072]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [74435.0, 74972.0, 75509.0, 76046.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0202484567313
varray: [0.926, 0.897, 0.885, 0.924]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74435.0, 74972.0, 75509.0, 76046.0]
source_col: d(18_16)Mean
dD_stdev | 0.261489324193
varray: [-5.473, -5.966, -5.805, -6.072]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74435.0, 74972.0, 75509.0, 76046.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.908
varray: [0.926, 0.897, 0.885, 0.924]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [74435.0, 74972.0, 75509.0, 76046.0]
source_col: d(18_16)Mean
measured_dD_mean | -5.829
varray: [-5.473, -5.966, -5.805, -6.072]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [74435.0, 74972.0, 75509.0, 76046.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0202484567313
varray: [0.926, 0.897, 0.885, 0.924]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74435.0, 74972.0, 75509.0, 76046.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.261489324193
varray: [-5.473, -5.966, -5.805, -6.072]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74435.0, 74972.0, 75509.0, 76046.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.76419381126
actual_array: [ 1.93 -15.84]
varray: [0.926, 0.897, 0.885, 0.924]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.908
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [74435.0, 74972.0, 75509.0, 76046.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -23.1333112117
actual_array: [ 18.1 -119. ]
varray: [-5.473, -5.966, -5.805, -6.072]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -5.829
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [74435.0, 74972.0, 75509.0, 76046.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.7542
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 944.69155
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 30 - Sample 30 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 30
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.81475
varray: [2.83, 2.865, 2.799, 2.765]
inst: mean_all
operation: mean_all(residuals)
tarray: [76631.0, 77168.0, 77705.0, 78241.0]
source_col: residuals
SAMP_residuals_STD | 0.0427424457263
varray: [2.83, 2.865, 2.799, 2.765]
inst: std_all
operation: std_all(residuals)
tarray: [76631.0, 77168.0, 77705.0, 78241.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.80675
varray: [-0.79, -0.715, -0.728, -0.994]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [76631.0, 77168.0, 77705.0, 78241.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.12905134637
varray: [-0.79, -0.715, -0.728, -0.994]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [76631.0, 77168.0, 77705.0, 78241.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1310.42225
varray: [1318.599, 1309.669, 1307.318, 1306.103]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [76631.0, 77168.0, 77705.0, 78241.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 5.64856662496
varray: [1318.599, 1309.669, 1307.318, 1306.103]
inst: std_all
operation: std_all(baseline_shift)
tarray: [76631.0, 77168.0, 77705.0, 78241.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -2.2235
varray: [-2.197, -2.441, -1.951, -2.305]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [76631.0, 77168.0, 77705.0, 78241.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.207289652419
varray: [-2.197, -2.441, -1.951, -2.305]
inst: std_all
operation: std_all(slope_shift)
tarray: [76631.0, 77168.0, 77705.0, 78241.0]
source_col: slope_shift
SAMP_d18O_STD | 0.06408587988
varray: [2.647, 2.732, 2.74, 2.803]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76631.0, 77168.0, 77705.0, 78241.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0760810970829
varray: [-6.037, -6.208, -6.135, -6.186]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76631.0, 77168.0, 77705.0, 78241.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.83625
varray: [111.266, 112.883, 112.571, 110.625]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [76631.0, 77168.0, 77705.0, 78241.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.06893322991
varray: [111.266, 112.883, 112.571, 110.625]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [76631.0, 77168.0, 77705.0, 78241.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.81475
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.80675
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1310.42225
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.2235
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0806
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0806
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.78377398242
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.47645
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [2.647, 2.732, 2.74, 2.803]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [76631.0, 77168.0, 77705.0, 78241.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-6.037, -6.208, -6.135, -6.186]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [76631.0, 77168.0, 77705.0, 78241.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.7305
varray: [2.647, 2.732, 2.74, 2.803]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [76631.0, 77168.0, 77705.0, 78241.0]
source_col: d(18_16)Mean
dD_mean | -6.1415
varray: [-6.037, -6.208, -6.135, -6.186]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [76631.0, 77168.0, 77705.0, 78241.0]
source_col: d(D_H)Mean
d18O_stdev | 0.06408587988
varray: [2.647, 2.732, 2.74, 2.803]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76631.0, 77168.0, 77705.0, 78241.0]
source_col: d(18_16)Mean
dD_stdev | 0.0760810970829
varray: [-6.037, -6.208, -6.135, -6.186]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76631.0, 77168.0, 77705.0, 78241.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.7305
varray: [2.647, 2.732, 2.74, 2.803]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [76631.0, 77168.0, 77705.0, 78241.0]
source_col: d(18_16)Mean
measured_dD_mean | -6.1415
varray: [-6.037, -6.208, -6.135, -6.186]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [76631.0, 77168.0, 77705.0, 78241.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.06408587988
varray: [2.647, 2.732, 2.74, 2.803]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76631.0, 77168.0, 77705.0, 78241.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0760810970829
varray: [-6.037, -6.208, -6.135, -6.186]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76631.0, 77168.0, 77705.0, 78241.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 0.0558889713155
actual_array: [ 1.93 -15.84]
varray: [2.647, 2.732, 2.74, 2.803]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 2.7305
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [76631.0, 77168.0, 77705.0, 78241.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -23.4255338189
actual_array: [ 18.1 -119. ]
varray: [-6.037, -6.208, -6.135, -6.186]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -6.1415
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [76631.0, 77168.0, 77705.0, 78241.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.6662
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 14.9413
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 31 - Sample 31 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 31
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.82875
varray: [2.827, 2.842, 2.805, 2.841]
inst: mean_all
operation: mean_all(residuals)
tarray: [78827.0, 79364.0, 79901.0, 80438.0]
source_col: residuals
SAMP_residuals_STD | 0.0172506038542
varray: [2.827, 2.842, 2.805, 2.841]
inst: std_all
operation: std_all(residuals)
tarray: [78827.0, 79364.0, 79901.0, 80438.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.53975
varray: [0.609, 0.783, 0.274, 0.493]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [78827.0, 79364.0, 79901.0, 80438.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.213521856805
varray: [0.609, 0.783, 0.274, 0.493]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [78827.0, 79364.0, 79901.0, 80438.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1293.98175
varray: [1293.515, 1293.574, 1294.568, 1294.27]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [78827.0, 79364.0, 79901.0, 80438.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.519901513109
varray: [1293.515, 1293.574, 1294.568, 1294.27]
inst: std_all
operation: std_all(baseline_shift)
tarray: [78827.0, 79364.0, 79901.0, 80438.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.98775
varray: [1.901, 1.915, 2.3, 1.835]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [78827.0, 79364.0, 79901.0, 80438.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.211069301415
varray: [1.901, 1.915, 2.3, 1.835]
inst: std_all
operation: std_all(slope_shift)
tarray: [78827.0, 79364.0, 79901.0, 80438.0]
source_col: slope_shift
SAMP_d18O_STD | 0.267354945095
varray: [-6.96, -7.51, -7.454, -7.512]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78827.0, 79364.0, 79901.0, 80438.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.84182058723
varray: [-52.752, -57.563, -58.131, -59.151]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78827.0, 79364.0, 79901.0, 80438.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.7535
varray: [109.587, 110.603, 109.03, 109.794]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [78827.0, 79364.0, 79901.0, 80438.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.651783450746
varray: [109.587, 110.603, 109.03, 109.794]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [78827.0, 79364.0, 79901.0, 80438.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.82875
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.53975
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1293.98175
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.98775
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0666
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0666
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.30024004006
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.6063
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [-6.96, -7.51, -7.454, -7.512]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [78827.0, 79364.0, 79901.0, 80438.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-52.752, -57.563, -58.131, -59.151]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [78827.0, 79364.0, 79901.0, 80438.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.359
varray: [-6.96, -7.51, -7.454, -7.512]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [78827.0, 79364.0, 79901.0, 80438.0]
source_col: d(18_16)Mean
dD_mean | -56.89925
varray: [-52.752, -57.563, -58.131, -59.151]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [78827.0, 79364.0, 79901.0, 80438.0]
source_col: d(D_H)Mean
d18O_stdev | 0.267354945095
varray: [-6.96, -7.51, -7.454, -7.512]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78827.0, 79364.0, 79901.0, 80438.0]
source_col: d(18_16)Mean
dD_stdev | 2.84182058723
varray: [-52.752, -57.563, -58.131, -59.151]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78827.0, 79364.0, 79901.0, 80438.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.359
varray: [-6.96, -7.51, -7.454, -7.512]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [78827.0, 79364.0, 79901.0, 80438.0]
source_col: d(18_16)Mean
measured_dD_mean | -56.89925
varray: [-52.752, -57.563, -58.131, -59.151]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [78827.0, 79364.0, 79901.0, 80438.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.267354945095
varray: [-6.96, -7.51, -7.454, -7.512]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78827.0, 79364.0, 79901.0, 80438.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.84182058723
varray: [-52.752, -57.563, -58.131, -59.151]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78827.0, 79364.0, 79901.0, 80438.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.0202291273
actual_array: [ 1.93 -15.84]
varray: [-6.96, -7.51, -7.454, -7.512]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.359
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [78827.0, 79364.0, 79901.0, 80438.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -70.8897323505
actual_array: [ 18.1 -119. ]
varray: [-52.752, -57.563, -58.131, -59.151]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -56.89925
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [78827.0, 79364.0, 79901.0, 80438.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.6522
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.4992
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 32 - Sample 32 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 32
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.7305
varray: [2.798, 2.677, 2.737, 2.71]
inst: mean_all
operation: mean_all(residuals)
tarray: [81024.0, 81561.0, 82098.0, 82635.0]
source_col: residuals
SAMP_residuals_STD | 0.051254268115
varray: [2.798, 2.677, 2.737, 2.71]
inst: std_all
operation: std_all(residuals)
tarray: [81024.0, 81561.0, 82098.0, 82635.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.0035
varray: [0.014, -0.278, 0.205, 0.045]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [81024.0, 81561.0, 82098.0, 82635.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.201230382067
varray: [0.014, -0.278, 0.205, 0.045]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [81024.0, 81561.0, 82098.0, 82635.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1304.06725
varray: [1306.069, 1303.894, 1303.112, 1303.194]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [81024.0, 81561.0, 82098.0, 82635.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.37986554297
varray: [1306.069, 1303.894, 1303.112, 1303.194]
inst: std_all
operation: std_all(baseline_shift)
tarray: [81024.0, 81561.0, 82098.0, 82635.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.176
varray: [0.133, 0.137, 0.318, 0.116]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [81024.0, 81561.0, 82098.0, 82635.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0951034524435
varray: [0.133, 0.137, 0.318, 0.116]
inst: std_all
operation: std_all(slope_shift)
tarray: [81024.0, 81561.0, 82098.0, 82635.0]
source_col: slope_shift
SAMP_d18O_STD | 0.22546452049
varray: [0.837, 1.199, 1.259, 1.352]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [81024.0, 81561.0, 82098.0, 82635.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.51422744927
varray: [-13.665, -9.517, -8.597, -8.175]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [81024.0, 81561.0, 82098.0, 82635.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 106.9415
varray: [108.573, 105.207, 107.221, 106.765]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [81024.0, 81561.0, 82098.0, 82635.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.38799075405
varray: [108.573, 105.207, 107.221, 106.765]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [81024.0, 81561.0, 82098.0, 82635.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.7305
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0035
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1304.06725
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.176
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.16485
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.16485
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.69361217124
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.4183
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [0.837, 1.199, 1.259, 1.352]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [81024.0, 81561.0, 82098.0, 82635.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-13.665, -9.517, -8.597, -8.175]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [81024.0, 81561.0, 82098.0, 82635.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.16175
varray: [0.837, 1.199, 1.259, 1.352]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [81024.0, 81561.0, 82098.0, 82635.0]
source_col: d(18_16)Mean
dD_mean | -9.9885
varray: [-13.665, -9.517, -8.597, -8.175]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [81024.0, 81561.0, 82098.0, 82635.0]
source_col: d(D_H)Mean
d18O_stdev | 0.22546452049
varray: [0.837, 1.199, 1.259, 1.352]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [81024.0, 81561.0, 82098.0, 82635.0]
source_col: d(18_16)Mean
dD_stdev | 2.51422744927
varray: [-13.665, -9.517, -8.597, -8.175]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [81024.0, 81561.0, 82098.0, 82635.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.16175
varray: [0.837, 1.199, 1.259, 1.352]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [81024.0, 81561.0, 82098.0, 82635.0]
source_col: d(18_16)Mean
measured_dD_mean | -9.9885
varray: [-13.665, -9.517, -8.597, -8.175]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [81024.0, 81561.0, 82098.0, 82635.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.22546452049
varray: [0.837, 1.199, 1.259, 1.352]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [81024.0, 81561.0, 82098.0, 82635.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.51422744927
varray: [-13.665, -9.517, -8.597, -8.175]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [81024.0, 81561.0, 82098.0, 82635.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.51078036485
actual_array: [ 1.93 -15.84]
varray: [0.837, 1.199, 1.259, 1.352]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 1.16175
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [81024.0, 81561.0, 82098.0, 82635.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -27.0229110027
actual_array: [ 18.1 -119. ]
varray: [-13.665, -9.517, -8.597, -8.175]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -9.9885
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [81024.0, 81561.0, 82098.0, 82635.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.75045
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 8.5863
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 33 - Sample 33 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 33
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.806
varray: [2.847, 2.759, 2.85, 2.768]
inst: mean_all
operation: mean_all(residuals)
tarray: [83221.0, 83758.0, 84294.0, 84831.0]
source_col: residuals
SAMP_residuals_STD | 0.0492273636643
varray: [2.847, 2.759, 2.85, 2.768]
inst: std_all
operation: std_all(residuals)
tarray: [83221.0, 83758.0, 84294.0, 84831.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.24325
varray: [0.302, 0.275, 0.231, 0.165]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [83221.0, 83758.0, 84294.0, 84831.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0598129027329
varray: [0.302, 0.275, 0.231, 0.165]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [83221.0, 83758.0, 84294.0, 84831.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1294.38575
varray: [1294.661, 1292.463, 1295.727, 1294.692]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [83221.0, 83758.0, 84294.0, 84831.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.37422350802
varray: [1294.661, 1292.463, 1295.727, 1294.692]
inst: std_all
operation: std_all(baseline_shift)
tarray: [83221.0, 83758.0, 84294.0, 84831.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.019
varray: [0.896, 0.607, 1.627, 0.946]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [83221.0, 83758.0, 84294.0, 84831.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.431997685179
varray: [0.896, 0.607, 1.627, 0.946]
inst: std_all
operation: std_all(slope_shift)
tarray: [83221.0, 83758.0, 84294.0, 84831.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0624306281671
varray: [-1.201, -1.335, -1.306, -1.329]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83221.0, 83758.0, 84294.0, 84831.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.759481566597
varray: [-22.053, -23.571, -23.493, -23.638]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83221.0, 83758.0, 84294.0, 84831.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.1165
varray: [110.11, 107.684, 109.682, 108.99]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [83221.0, 83758.0, 84294.0, 84831.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.06064304394
varray: [110.11, 107.684, 109.682, 108.99]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [83221.0, 83758.0, 84294.0, 84831.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.806
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.24325
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1294.38575
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.019
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.08935
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.08935
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.08598269639
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.2433
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [-1.201, -1.335, -1.306, -1.329]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [83221.0, 83758.0, 84294.0, 84831.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-22.053, -23.571, -23.493, -23.638]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [83221.0, 83758.0, 84294.0, 84831.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.29275
varray: [-1.201, -1.335, -1.306, -1.329]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [83221.0, 83758.0, 84294.0, 84831.0]
source_col: d(18_16)Mean
dD_mean | -23.18875
varray: [-22.053, -23.571, -23.493, -23.638]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [83221.0, 83758.0, 84294.0, 84831.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0624306281671
varray: [-1.201, -1.335, -1.306, -1.329]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83221.0, 83758.0, 84294.0, 84831.0]
source_col: d(18_16)Mean
dD_stdev | 0.759481566597
varray: [-22.053, -23.571, -23.493, -23.638]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83221.0, 83758.0, 84294.0, 84831.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.29275
varray: [-1.201, -1.335, -1.306, -1.329]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [83221.0, 83758.0, 84294.0, 84831.0]
source_col: d(18_16)Mean
measured_dD_mean | -23.18875
varray: [-22.053, -23.571, -23.493, -23.638]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [83221.0, 83758.0, 84294.0, 84831.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0624306281671
varray: [-1.201, -1.335, -1.306, -1.329]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83221.0, 83758.0, 84294.0, 84831.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.759481566597
varray: [-22.053, -23.571, -23.493, -23.638]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83221.0, 83758.0, 84294.0, 84831.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.96202491345
actual_array: [ 1.93 -15.84]
varray: [-1.201, -1.335, -1.306, -1.329]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -1.29275
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [83221.0, 83758.0, 84294.0, 84831.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -39.3666277093
actual_array: [ 18.1 -119. ]
varray: [-22.053, -23.571, -23.493, -23.638]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -23.18875
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [83221.0, 83758.0, 84294.0, 84831.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.67495
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.0952
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 34 - Sample 34 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 34
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.902
varray: [3.176, 2.851, 2.709, 2.872]
inst: mean_all
operation: mean_all(residuals)
tarray: [85417.0, 85953.0, 86490.0, 87026.0]
source_col: residuals
SAMP_residuals_STD | 0.196490881892
varray: [3.176, 2.851, 2.709, 2.872]
inst: std_all
operation: std_all(residuals)
tarray: [85417.0, 85953.0, 86490.0, 87026.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.47025
varray: [0.396, 0.669, 0.304, 0.512]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [85417.0, 85953.0, 86490.0, 87026.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.157476717856
varray: [0.396, 0.669, 0.304, 0.512]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [85417.0, 85953.0, 86490.0, 87026.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1295.21325
varray: [1307.017, 1292.584, 1288.221, 1293.031]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [85417.0, 85953.0, 86490.0, 87026.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 8.16282722162
varray: [1307.017, 1292.584, 1288.221, 1293.031]
inst: std_all
operation: std_all(baseline_shift)
tarray: [85417.0, 85953.0, 86490.0, 87026.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.42875
varray: [1.403, 1.36, 1.492, 1.46]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [85417.0, 85953.0, 86490.0, 87026.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.058784209898
varray: [1.403, 1.36, 1.492, 1.46]
inst: std_all
operation: std_all(slope_shift)
tarray: [85417.0, 85953.0, 86490.0, 87026.0]
source_col: slope_shift
SAMP_d18O_STD | 0.095423616923
varray: [-1.224, -1.379, -1.453, -1.354]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85417.0, 85953.0, 86490.0, 87026.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.418279412196
varray: [-13.244, -13.724, -14.251, -13.587]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85417.0, 85953.0, 86490.0, 87026.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.632
varray: [121.958, 109.904, 104.738, 109.928]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [85417.0, 85953.0, 86490.0, 87026.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 7.30395125942
varray: [121.958, 109.904, 104.738, 109.928]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [85417.0, 85953.0, 86490.0, 87026.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.902
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.47025
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1295.21325
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.42875
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.00665
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.00665
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.229678622619
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.2722
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [-1.224, -1.379, -1.453, -1.354]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [85417.0, 85953.0, 86490.0, 87026.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-13.244, -13.724, -14.251, -13.587]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [85417.0, 85953.0, 86490.0, 87026.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.3525
varray: [-1.224, -1.379, -1.453, -1.354]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [85417.0, 85953.0, 86490.0, 87026.0]
source_col: d(18_16)Mean
dD_mean | -13.7015
varray: [-13.244, -13.724, -14.251, -13.587]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [85417.0, 85953.0, 86490.0, 87026.0]
source_col: d(D_H)Mean
d18O_stdev | 0.095423616923
varray: [-1.224, -1.379, -1.453, -1.354]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85417.0, 85953.0, 86490.0, 87026.0]
source_col: d(18_16)Mean
dD_stdev | 0.418279412196
varray: [-13.244, -13.724, -14.251, -13.587]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85417.0, 85953.0, 86490.0, 87026.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.3525
varray: [-1.224, -1.379, -1.453, -1.354]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [85417.0, 85953.0, 86490.0, 87026.0]
source_col: d(18_16)Mean
measured_dD_mean | -13.7015
varray: [-13.244, -13.724, -14.251, -13.587]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [85417.0, 85953.0, 86490.0, 87026.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.095423616923
varray: [-1.224, -1.379, -1.453, -1.354]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85417.0, 85953.0, 86490.0, 87026.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.418279412196
varray: [-13.244, -13.724, -14.251, -13.587]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85417.0, 85953.0, 86490.0, 87026.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.02169566586
actual_array: [ 1.93 -15.84]
varray: [-1.224, -1.379, -1.453, -1.354]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -1.3525
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [85417.0, 85953.0, 86490.0, 87026.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -30.4949831325
actual_array: [ 18.1 -119. ]
varray: [-13.244, -13.724, -14.251, -13.587]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -13.7015
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [85417.0, 85953.0, 86490.0, 87026.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.57895
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.2677
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 35 - Sample 35 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 35
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.74175
varray: [2.656, 2.872, 2.675, 2.764]
inst: mean_all
operation: mean_all(residuals)
tarray: [87614.0, 88151.0, 88688.0, 89225.0]
source_col: residuals
SAMP_residuals_STD | 0.0987737313257
varray: [2.656, 2.872, 2.675, 2.764]
inst: std_all
operation: std_all(residuals)
tarray: [87614.0, 88151.0, 88688.0, 89225.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4025
varray: [0.47, 0.223, 0.247, 0.67]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [87614.0, 88151.0, 88688.0, 89225.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.210168979633
varray: [0.47, 0.223, 0.247, 0.67]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [87614.0, 88151.0, 88688.0, 89225.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1288.43825
varray: [1285.172, 1293.334, 1286.777, 1288.47]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [87614.0, 88151.0, 88688.0, 89225.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.53069949396
varray: [1285.172, 1293.334, 1286.777, 1288.47]
inst: std_all
operation: std_all(baseline_shift)
tarray: [87614.0, 88151.0, 88688.0, 89225.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.42
varray: [1.632, 1.278, 1.385, 1.385]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [87614.0, 88151.0, 88688.0, 89225.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.150064430607
varray: [1.632, 1.278, 1.385, 1.385]
inst: std_all
operation: std_all(slope_shift)
tarray: [87614.0, 88151.0, 88688.0, 89225.0]
source_col: slope_shift
SAMP_d18O_STD | 0.145066364123
varray: [1.116, 1.442, 1.337, 1.4]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87614.0, 88151.0, 88688.0, 89225.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.22056117558
varray: [0.786, 3.445, 2.844, 3.238]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87614.0, 88151.0, 88688.0, 89225.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 105.722
varray: [102.849, 109.528, 103.864, 106.647]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [87614.0, 88151.0, 88688.0, 89225.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.00264072665
varray: [102.849, 109.528, 103.864, 106.647]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [87614.0, 88151.0, 88688.0, 89225.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.74175
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4025
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1288.43825
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.42
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1536
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1536
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.30505811042
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.6378
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [1.116, 1.442, 1.337, 1.4]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [87614.0, 88151.0, 88688.0, 89225.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [0.786, 3.445, 2.844, 3.238]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [87614.0, 88151.0, 88688.0, 89225.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.32375
varray: [1.116, 1.442, 1.337, 1.4]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [87614.0, 88151.0, 88688.0, 89225.0]
source_col: d(18_16)Mean
dD_mean | 2.57825
varray: [0.786, 3.445, 2.844, 3.238]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [87614.0, 88151.0, 88688.0, 89225.0]
source_col: d(D_H)Mean
d18O_stdev | 0.145066364123
varray: [1.116, 1.442, 1.337, 1.4]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87614.0, 88151.0, 88688.0, 89225.0]
source_col: d(18_16)Mean
dD_stdev | 1.22056117558
varray: [0.786, 3.445, 2.844, 3.238]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87614.0, 88151.0, 88688.0, 89225.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.32375
varray: [1.116, 1.442, 1.337, 1.4]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [87614.0, 88151.0, 88688.0, 89225.0]
source_col: d(18_16)Mean
measured_dD_mean | 2.57825
varray: [0.786, 3.445, 2.844, 3.238]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [87614.0, 88151.0, 88688.0, 89225.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.145066364123
varray: [1.116, 1.442, 1.337, 1.4]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87614.0, 88151.0, 88688.0, 89225.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.22056117558
varray: [0.786, 3.445, 2.844, 3.238]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87614.0, 88151.0, 88688.0, 89225.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.34899522862
actual_array: [ 1.93 -15.84]
varray: [1.116, 1.442, 1.337, 1.4]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 1.32375
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [87614.0, 88151.0, 88688.0, 89225.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -15.2715879654
actual_array: [ 18.1 -119. ]
varray: [0.786, 3.445, 2.844, 3.238]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 2.57825
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [87614.0, 88151.0, 88688.0, 89225.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.7392
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -7.0427
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 36 - Sample 36 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 36
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.6905
varray: [2.663, 2.695, 2.72, 2.684]
inst: mean_all
operation: mean_all(residuals)
tarray: [89808.0, 90345.0, 90881.0, 91418.0]
source_col: residuals
SAMP_residuals_STD | 0.0237276210354
varray: [2.663, 2.695, 2.72, 2.684]
inst: std_all
operation: std_all(residuals)
tarray: [89808.0, 90345.0, 90881.0, 91418.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4335
varray: [0.464, 0.264, 0.45, 0.556]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [89808.0, 90345.0, 90881.0, 91418.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.122391448503
varray: [0.464, 0.264, 0.45, 0.556]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [89808.0, 90345.0, 90881.0, 91418.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1287.17425
varray: [1285.666, 1288.076, 1287.612, 1287.343]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [89808.0, 90345.0, 90881.0, 91418.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.05009090876
varray: [1285.666, 1288.076, 1287.612, 1287.343]
inst: std_all
operation: std_all(baseline_shift)
tarray: [89808.0, 90345.0, 90881.0, 91418.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.78
varray: [1.639, 1.344, 2.348, 1.789]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [89808.0, 90345.0, 90881.0, 91418.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.421379480595
varray: [1.639, 1.344, 2.348, 1.789]
inst: std_all
operation: std_all(slope_shift)
tarray: [89808.0, 90345.0, 90881.0, 91418.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0816021854283
varray: [-1.225, -1.305, -1.382, -1.405]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89808.0, 90345.0, 90881.0, 91418.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.720771114849
varray: [-8.736, -9.957, -10.134, -10.339]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89808.0, 90345.0, 90881.0, 91418.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 104.5145
varray: [104.156, 104.696, 104.045, 105.161]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [89808.0, 90345.0, 90881.0, 91418.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.51635162438
varray: [104.156, 104.696, 104.045, 105.161]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [89808.0, 90345.0, 90881.0, 91418.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.6905
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4335
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1287.17425
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.78
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.20485
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.20485
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -7.07513772083
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -6.8453
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [-1.225, -1.305, -1.382, -1.405]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [89808.0, 90345.0, 90881.0, 91418.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-8.736, -9.957, -10.134, -10.339]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [89808.0, 90345.0, 90881.0, 91418.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.32925
varray: [-1.225, -1.305, -1.382, -1.405]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [89808.0, 90345.0, 90881.0, 91418.0]
source_col: d(18_16)Mean
dD_mean | -9.7915
varray: [-8.736, -9.957, -10.134, -10.339]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [89808.0, 90345.0, 90881.0, 91418.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0816021854283
varray: [-1.225, -1.305, -1.382, -1.405]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89808.0, 90345.0, 90881.0, 91418.0]
source_col: d(18_16)Mean
dD_stdev | 0.720771114849
varray: [-8.736, -9.957, -10.134, -10.339]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89808.0, 90345.0, 90881.0, 91418.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.32925
varray: [-1.225, -1.305, -1.382, -1.405]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [89808.0, 90345.0, 90881.0, 91418.0]
source_col: d(18_16)Mean
measured_dD_mean | -9.7915
varray: [-8.736, -9.957, -10.134, -10.339]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [89808.0, 90345.0, 90881.0, 91418.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0816021854283
varray: [-1.225, -1.305, -1.382, -1.405]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89808.0, 90345.0, 90881.0, 91418.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.720771114849
varray: [-8.736, -9.957, -10.134, -10.339]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89808.0, 90345.0, 90881.0, 91418.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.99847650279
actual_array: [ 1.93 -15.84]
varray: [-1.225, -1.305, -1.382, -1.405]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -1.32925
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [89808.0, 90345.0, 90881.0, 91418.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -26.8386938711
actual_array: [ 18.1 -119. ]
varray: [-8.736, -9.957, -10.134, -10.339]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -9.7915
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [89808.0, 90345.0, 90881.0, 91418.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.79045
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -8.3067
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 37 - Sample 37 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 37
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.83625
varray: [2.821, 2.891, 2.803, 2.83]
inst: mean_all
operation: mean_all(residuals)
tarray: [92003.0, 92540.0, 93077.0, 93614.0]
source_col: residuals
SAMP_residuals_STD | 0.0381870396863
varray: [2.821, 2.891, 2.803, 2.83]
inst: std_all
operation: std_all(residuals)
tarray: [92003.0, 92540.0, 93077.0, 93614.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.28925
varray: [0.222, 0.589, 0.148, 0.198]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [92003.0, 92540.0, 93077.0, 93614.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.202196892492
varray: [0.222, 0.589, 0.148, 0.198]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [92003.0, 92540.0, 93077.0, 93614.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1294.772
varray: [1294.823, 1295.014, 1294.764, 1294.487]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [92003.0, 92540.0, 93077.0, 93614.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.217909767258
varray: [1294.823, 1295.014, 1294.764, 1294.487]
inst: std_all
operation: std_all(baseline_shift)
tarray: [92003.0, 92540.0, 93077.0, 93614.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.381
varray: [1.407, 1.239, 1.506, 1.372]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [92003.0, 92540.0, 93077.0, 93614.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.110372097923
varray: [1.407, 1.239, 1.506, 1.372]
inst: std_all
operation: std_all(slope_shift)
tarray: [92003.0, 92540.0, 93077.0, 93614.0]
source_col: slope_shift
SAMP_d18O_STD | 0.027748873851
varray: [-0.321, -0.262, -0.264, -0.273]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [92003.0, 92540.0, 93077.0, 93614.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.166674083168
varray: [-10.021, -9.715, -9.941, -10.102]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [92003.0, 92540.0, 93077.0, 93614.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.799
varray: [109.957, 111.401, 108.893, 108.945]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [92003.0, 92540.0, 93077.0, 93614.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.17494964431
varray: [109.957, 111.401, 108.893, 108.945]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [92003.0, 92540.0, 93077.0, 93614.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.83625
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.28925
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1294.772
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.381
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0591
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0591
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.04120399952
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.5608
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [-0.321, -0.262, -0.264, -0.273]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [92003.0, 92540.0, 93077.0, 93614.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-10.021, -9.715, -9.941, -10.102]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [92003.0, 92540.0, 93077.0, 93614.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.28
varray: [-0.321, -0.262, -0.264, -0.273]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [92003.0, 92540.0, 93077.0, 93614.0]
source_col: d(18_16)Mean
dD_mean | -9.94475
varray: [-10.021, -9.715, -9.941, -10.102]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [92003.0, 92540.0, 93077.0, 93614.0]
source_col: d(D_H)Mean
d18O_stdev | 0.027748873851
varray: [-0.321, -0.262, -0.264, -0.273]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [92003.0, 92540.0, 93077.0, 93614.0]
source_col: d(18_16)Mean
dD_stdev | 0.166674083168
varray: [-10.021, -9.715, -9.941, -10.102]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [92003.0, 92540.0, 93077.0, 93614.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.28
varray: [-0.321, -0.262, -0.264, -0.273]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [92003.0, 92540.0, 93077.0, 93614.0]
source_col: d(18_16)Mean
measured_dD_mean | -9.94475
varray: [-10.021, -9.715, -9.941, -10.102]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [92003.0, 92540.0, 93077.0, 93614.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.027748873851
varray: [-0.321, -0.262, -0.264, -0.273]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [92003.0, 92540.0, 93077.0, 93614.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.166674083168
varray: [-10.021, -9.715, -9.941, -10.102]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [92003.0, 92540.0, 93077.0, 93614.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.9506181436
actual_array: [ 1.93 -15.84]
varray: [-0.321, -0.262, -0.264, -0.273]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -0.28
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [92003.0, 92540.0, 93077.0, 93614.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -26.9819998377
actual_array: [ 18.1 -119. ]
varray: [-10.021, -9.715, -9.941, -10.102]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -9.94475
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [92003.0, 92540.0, 93077.0, 93614.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.6447
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.70895
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 38 - Sample 38 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 38
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.726
varray: [2.732, 2.733, 2.74, 2.699]
inst: mean_all
operation: mean_all(residuals)
tarray: [94198.0, 94736.0, 95272.0, 95809.0]
source_col: residuals
SAMP_residuals_STD | 0.0183484785927
varray: [2.732, 2.733, 2.74, 2.699]
inst: std_all
operation: std_all(residuals)
tarray: [94198.0, 94736.0, 95272.0, 95809.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.3515
varray: [-0.192, -0.471, -0.48, -0.263]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [94198.0, 94736.0, 95272.0, 95809.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.146133500608
varray: [-0.192, -0.471, -0.48, -0.263]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [94198.0, 94736.0, 95272.0, 95809.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 2076.216
varray: [2068.841, 2075.904, 2078.787, 2081.332]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [94198.0, 94736.0, 95272.0, 95809.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 5.39355992519
varray: [2068.841, 2075.904, 2078.787, 2081.332]
inst: std_all
operation: std_all(baseline_shift)
tarray: [94198.0, 94736.0, 95272.0, 95809.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -27.7305
varray: [-27.617, -27.593, -27.667, -28.045]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [94198.0, 94736.0, 95272.0, 95809.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.211920582609
varray: [-27.617, -27.593, -27.667, -28.045]
inst: std_all
operation: std_all(slope_shift)
tarray: [94198.0, 94736.0, 95272.0, 95809.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0163375844808
varray: [0.426, 0.46, 0.425, 0.44]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94198.0, 94736.0, 95272.0, 95809.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.402695397706
varray: [-2.621, -2.221, -1.777, -1.79]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94198.0, 94736.0, 95272.0, 95809.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 100.344
varray: [101.148, 99.472, 99.837, 100.919]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [94198.0, 94736.0, 95272.0, 95809.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.815367810664
varray: [101.148, 99.472, 99.837, 100.919]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [94198.0, 94736.0, 95272.0, 95809.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.726
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3515
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 2076.216
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.7305
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.16935
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.16935
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.84903379557
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -11.0158
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [0.426, 0.46, 0.425, 0.44]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [94198.0, 94736.0, 95272.0, 95809.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-2.621, -2.221, -1.777, -1.79]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [94198.0, 94736.0, 95272.0, 95809.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.43775
varray: [0.426, 0.46, 0.425, 0.44]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [94198.0, 94736.0, 95272.0, 95809.0]
source_col: d(18_16)Mean
dD_mean | -2.10225
varray: [-2.621, -2.221, -1.777, -1.79]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [94198.0, 94736.0, 95272.0, 95809.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0163375844808
varray: [0.426, 0.46, 0.425, 0.44]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94198.0, 94736.0, 95272.0, 95809.0]
source_col: d(18_16)Mean
dD_stdev | 0.402695397706
varray: [-2.621, -2.221, -1.777, -1.79]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94198.0, 94736.0, 95272.0, 95809.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.43775
varray: [0.426, 0.46, 0.425, 0.44]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [94198.0, 94736.0, 95272.0, 95809.0]
source_col: d(18_16)Mean
measured_dD_mean | -2.10225
varray: [-2.621, -2.221, -1.777, -1.79]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [94198.0, 94736.0, 95272.0, 95809.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0163375844808
varray: [0.426, 0.46, 0.425, 0.44]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94198.0, 94736.0, 95272.0, 95809.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.402695397706
varray: [-2.621, -2.221, -1.777, -1.79]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94198.0, 94736.0, 95272.0, 95809.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.23382010948
actual_array: [ 1.93 -15.84]
varray: [0.426, 0.46, 0.425, 0.44]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.43775
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [94198.0, 94736.0, 95272.0, 95809.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -19.6483812871
actual_array: [ 18.1 -119. ]
varray: [-2.621, -2.221, -1.777, -1.79]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -2.10225
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [94198.0, 94736.0, 95272.0, 95809.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.75495
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 780.73505
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 39 - Sample 39 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 39
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.83675
varray: [2.785, 2.833, 2.847, 2.882]
inst: mean_all
operation: mean_all(residuals)
tarray: [96395.0, 96932.0, 97469.0, 98005.0]
source_col: residuals
SAMP_residuals_STD | 0.0401860257635
varray: [2.785, 2.833, 2.847, 2.882]
inst: std_all
operation: std_all(residuals)
tarray: [96395.0, 96932.0, 97469.0, 98005.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.28675
varray: [0.268, 0.319, 0.392, 0.168]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [96395.0, 96932.0, 97469.0, 98005.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0941111930998
varray: [0.268, 0.319, 0.392, 0.168]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [96395.0, 96932.0, 97469.0, 98005.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1303.1815
varray: [1307.607, 1302.118, 1301.419, 1301.582]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [96395.0, 96932.0, 97469.0, 98005.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.96540548099
varray: [1307.607, 1302.118, 1301.419, 1301.582]
inst: std_all
operation: std_all(baseline_shift)
tarray: [96395.0, 96932.0, 97469.0, 98005.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.9285
varray: [0.492, 0.988, 1.005, 1.229]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [96395.0, 96932.0, 97469.0, 98005.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.311033224378
varray: [0.492, 0.988, 1.005, 1.229]
inst: std_all
operation: std_all(slope_shift)
tarray: [96395.0, 96932.0, 97469.0, 98005.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0548110390341
varray: [0.53, 0.533, 0.645, 0.593]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96395.0, 96932.0, 97469.0, 98005.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.255148322876
varray: [-1.939, -1.665, -1.316, -1.66]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96395.0, 96932.0, 97469.0, 98005.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.35325
varray: [107.468, 109.116, 110.687, 110.142]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [96395.0, 96932.0, 97469.0, 98005.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.41556475773
varray: [107.468, 109.116, 110.687, 110.142]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [96395.0, 96932.0, 97469.0, 98005.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.83675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.28675
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1303.1815
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.9285
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0586
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0586
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.02393493015
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.00655
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [0.53, 0.533, 0.645, 0.593]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [96395.0, 96932.0, 97469.0, 98005.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-1.939, -1.665, -1.316, -1.66]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [96395.0, 96932.0, 97469.0, 98005.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.57525
varray: [0.53, 0.533, 0.645, 0.593]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [96395.0, 96932.0, 97469.0, 98005.0]
source_col: d(18_16)Mean
dD_mean | -1.645
varray: [-1.939, -1.665, -1.316, -1.66]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [96395.0, 96932.0, 97469.0, 98005.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0548110390341
varray: [0.53, 0.533, 0.645, 0.593]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96395.0, 96932.0, 97469.0, 98005.0]
source_col: d(18_16)Mean
dD_stdev | 0.255148322876
varray: [-1.939, -1.665, -1.316, -1.66]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96395.0, 96932.0, 97469.0, 98005.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.57525
varray: [0.53, 0.533, 0.645, 0.593]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [96395.0, 96932.0, 97469.0, 98005.0]
source_col: d(18_16)Mean
measured_dD_mean | -1.645
varray: [-1.939, -1.665, -1.316, -1.66]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [96395.0, 96932.0, 97469.0, 98005.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0548110390341
varray: [0.53, 0.533, 0.645, 0.593]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96395.0, 96932.0, 97469.0, 98005.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.255148322876
varray: [-1.939, -1.665, -1.316, -1.66]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96395.0, 96932.0, 97469.0, 98005.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.09650247842
actual_array: [ 1.93 -15.84]
varray: [0.53, 0.533, 0.645, 0.593]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.57525
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [96395.0, 96932.0, 97469.0, 98005.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -19.2208011682
actual_array: [ 18.1 -119. ]
varray: [-1.939, -1.665, -1.316, -1.66]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -1.645
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [96395.0, 96932.0, 97469.0, 98005.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.6442
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 7.70055
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 40 - Sample 40 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 40
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.82025
varray: [2.778, 2.819, 2.821, 2.863]
inst: mean_all
operation: mean_all(residuals)
tarray: [98591.0, 99128.0, 99664.0, 100201.0]
source_col: residuals
SAMP_residuals_STD | 0.0347119095797
varray: [2.778, 2.819, 2.821, 2.863]
inst: std_all
operation: std_all(residuals)
tarray: [98591.0, 99128.0, 99664.0, 100201.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.297
varray: [0.292, 0.318, 0.291, 0.287]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [98591.0, 99128.0, 99664.0, 100201.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0141656862406
varray: [0.292, 0.318, 0.291, 0.287]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [98591.0, 99128.0, 99664.0, 100201.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1293.4505
varray: [1292.837, 1293.918, 1293.161, 1293.886]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [98591.0, 99128.0, 99664.0, 100201.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.538023853251
varray: [1292.837, 1293.918, 1293.161, 1293.886]
inst: std_all
operation: std_all(baseline_shift)
tarray: [98591.0, 99128.0, 99664.0, 100201.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.05125
varray: [1.939, 1.98, 2.123, 2.163]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [98591.0, 99128.0, 99664.0, 100201.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.108493855433
varray: [1.939, 1.98, 2.123, 2.163]
inst: std_all
operation: std_all(slope_shift)
tarray: [98591.0, 99128.0, 99664.0, 100201.0]
source_col: slope_shift
SAMP_d18O_STD | 0.183728059915
varray: [-7.095, -7.389, -7.467, -7.497]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [98591.0, 99128.0, 99664.0, 100201.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.90512298936
varray: [-52.246, -56.759, -58.265, -58.513]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [98591.0, 99128.0, 99664.0, 100201.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.44775
varray: [107.914, 108.406, 108.551, 108.92]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [98591.0, 99128.0, 99664.0, 100201.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.416458381274
varray: [107.914, 108.406, 108.551, 108.92]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [98591.0, 99128.0, 99664.0, 100201.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.82025
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.297
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1293.4505
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.05125
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0751
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0751
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.59381421935
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.91205
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [-7.095, -7.389, -7.467, -7.497]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [98591.0, 99128.0, 99664.0, 100201.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-52.246, -56.759, -58.265, -58.513]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [98591.0, 99128.0, 99664.0, 100201.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.362
varray: [-7.095, -7.389, -7.467, -7.497]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [98591.0, 99128.0, 99664.0, 100201.0]
source_col: d(18_16)Mean
dD_mean | -56.44575
varray: [-52.246, -56.759, -58.265, -58.513]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [98591.0, 99128.0, 99664.0, 100201.0]
source_col: d(D_H)Mean
d18O_stdev | 0.183728059915
varray: [-7.095, -7.389, -7.467, -7.497]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [98591.0, 99128.0, 99664.0, 100201.0]
source_col: d(18_16)Mean
dD_stdev | 2.90512298936
varray: [-52.246, -56.759, -58.265, -58.513]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [98591.0, 99128.0, 99664.0, 100201.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.362
varray: [-7.095, -7.389, -7.467, -7.497]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [98591.0, 99128.0, 99664.0, 100201.0]
source_col: d(18_16)Mean
measured_dD_mean | -56.44575
varray: [-52.246, -56.759, -58.265, -58.513]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [98591.0, 99128.0, 99664.0, 100201.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.183728059915
varray: [-7.095, -7.389, -7.467, -7.497]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [98591.0, 99128.0, 99664.0, 100201.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.90512298936
varray: [-52.246, -56.759, -58.265, -58.513]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [98591.0, 99128.0, 99664.0, 100201.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.0232251484
actual_array: [ 1.93 -15.84]
varray: [-7.095, -7.389, -7.467, -7.497]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.362
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [98591.0, 99128.0, 99664.0, 100201.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -70.4656589029
actual_array: [ 18.1 -119. ]
varray: [-52.246, -56.759, -58.265, -58.513]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -56.44575
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [98591.0, 99128.0, 99664.0, 100201.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.6607
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -2.03045
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 41 - Sample 41 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 41
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8865
varray: [2.874, 2.884, 2.91, 2.878]
inst: mean_all
operation: mean_all(residuals)
tarray: [100787.0, 101324.0, 101861.0, 102397.0]
source_col: residuals
SAMP_residuals_STD | 0.0161967074843
varray: [2.874, 2.884, 2.91, 2.878]
inst: std_all
operation: std_all(residuals)
tarray: [100787.0, 101324.0, 101861.0, 102397.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.48225
varray: [0.522, 0.455, 0.444, 0.508]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [100787.0, 101324.0, 101861.0, 102397.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0385086570354
varray: [0.522, 0.455, 0.444, 0.508]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [100787.0, 101324.0, 101861.0, 102397.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1294.67775
varray: [1294.3, 1294.426, 1295.442, 1294.543]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [100787.0, 101324.0, 101861.0, 102397.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.519072490121
varray: [1294.3, 1294.426, 1295.442, 1294.543]
inst: std_all
operation: std_all(baseline_shift)
tarray: [100787.0, 101324.0, 101861.0, 102397.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.53575
varray: [2.075, 1.983, 1.369, 0.716]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [100787.0, 101324.0, 101861.0, 102397.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.629978504924
varray: [2.075, 1.983, 1.369, 0.716]
inst: std_all
operation: std_all(slope_shift)
tarray: [100787.0, 101324.0, 101861.0, 102397.0]
source_col: slope_shift
SAMP_d18O_STD | 0.234434603532
varray: [-0.02, 0.387, 0.496, 0.438]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100787.0, 101324.0, 101861.0, 102397.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.2471461701
varray: [-5.214, 0.345, 1.58, 1.596]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100787.0, 101324.0, 101861.0, 102397.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.654
varray: [109.77, 110.294, 110.836, 111.716]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [100787.0, 101324.0, 101861.0, 102397.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.831068388682
varray: [109.77, 110.294, 110.836, 111.716]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [100787.0, 101324.0, 101861.0, 102397.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8865
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.48225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1294.67775
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.53575
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.00885
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.00885
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.305662527846
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.7058
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [-0.02, 0.387, 0.496, 0.438]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [100787.0, 101324.0, 101861.0, 102397.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-5.214, 0.345, 1.58, 1.596]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [100787.0, 101324.0, 101861.0, 102397.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.32525
varray: [-0.02, 0.387, 0.496, 0.438]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [100787.0, 101324.0, 101861.0, 102397.0]
source_col: d(18_16)Mean
dD_mean | -0.42325
varray: [-5.214, 0.345, 1.58, 1.596]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [100787.0, 101324.0, 101861.0, 102397.0]
source_col: d(D_H)Mean
d18O_stdev | 0.234434603532
varray: [-0.02, 0.387, 0.496, 0.438]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100787.0, 101324.0, 101861.0, 102397.0]
source_col: d(18_16)Mean
dD_stdev | 3.2471461701
varray: [-5.214, 0.345, 1.58, 1.596]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100787.0, 101324.0, 101861.0, 102397.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.32525
varray: [-0.02, 0.387, 0.496, 0.438]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [100787.0, 101324.0, 101861.0, 102397.0]
source_col: d(18_16)Mean
measured_dD_mean | -0.42325
varray: [-5.214, 0.345, 1.58, 1.596]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [100787.0, 101324.0, 101861.0, 102397.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.234434603532
varray: [-0.02, 0.387, 0.496, 0.438]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100787.0, 101324.0, 101861.0, 102397.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.2471461701
varray: [-5.214, 0.345, 1.58, 1.596]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100787.0, 101324.0, 101861.0, 102397.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.34617089853
actual_array: [ 1.93 -15.84]
varray: [-0.02, 0.387, 0.496, 0.438]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.32525
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [100787.0, 101324.0, 101861.0, 102397.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -18.0783276631
actual_array: [ 18.1 -119. ]
varray: [-5.214, 0.345, 1.58, 1.596]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -0.42325
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [100787.0, 101324.0, 101861.0, 102397.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.59445
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.8032
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 42 - Sample 42 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 42
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.71275
varray: [2.759, 2.761, 2.674, 2.657]
inst: mean_all
operation: mean_all(residuals)
tarray: [102983.0, 103519.0, 104056.0, 104592.0]
source_col: residuals
SAMP_residuals_STD | 0.0550053027747
varray: [2.759, 2.761, 2.674, 2.657]
inst: std_all
operation: std_all(residuals)
tarray: [102983.0, 103519.0, 104056.0, 104592.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.44275
varray: [-2.041, -2.194, -2.688, -2.848]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [102983.0, 103519.0, 104056.0, 104592.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.386287781324
varray: [-2.041, -2.194, -2.688, -2.848]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [102983.0, 103519.0, 104056.0, 104592.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 2482.611
varray: [2487.185, 2470.136, 2485.549, 2487.574]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [102983.0, 103519.0, 104056.0, 104592.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 8.36282077611
varray: [2487.185, 2470.136, 2485.549, 2487.574]
inst: std_all
operation: std_all(baseline_shift)
tarray: [102983.0, 103519.0, 104056.0, 104592.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -46.5615
varray: [-46.871, -45.9, -46.421, -47.054]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [102983.0, 103519.0, 104056.0, 104592.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.514998058249
varray: [-46.871, -45.9, -46.421, -47.054]
inst: std_all
operation: std_all(slope_shift)
tarray: [102983.0, 103519.0, 104056.0, 104592.0]
source_col: slope_shift
SAMP_d18O_STD | 0.014174507634
varray: [3.227, 3.216, 3.25, 3.232]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [102983.0, 103519.0, 104056.0, 104592.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.128023435355
varray: [3.019, 2.803, 3.082, 3.06]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [102983.0, 103519.0, 104056.0, 104592.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 96.80425
varray: [99.332, 97.007, 95.518, 95.36]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [102983.0, 103519.0, 104056.0, 104592.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.84127897126
varray: [99.332, 97.007, 95.518, 95.36]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [102983.0, 103519.0, 104056.0, 104592.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.71275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.44275
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 2482.611
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 46.5615
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1826
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1826
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.30666413387
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -14.55555
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [3.227, 3.216, 3.25, 3.232]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [102983.0, 103519.0, 104056.0, 104592.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [3.019, 2.803, 3.082, 3.06]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [102983.0, 103519.0, 104056.0, 104592.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 3.23125
varray: [3.227, 3.216, 3.25, 3.232]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [102983.0, 103519.0, 104056.0, 104592.0]
source_col: d(18_16)Mean
dD_mean | 2.991
varray: [3.019, 2.803, 3.082, 3.06]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [102983.0, 103519.0, 104056.0, 104592.0]
source_col: d(D_H)Mean
d18O_stdev | 0.014174507634
varray: [3.227, 3.216, 3.25, 3.232]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [102983.0, 103519.0, 104056.0, 104592.0]
source_col: d(18_16)Mean
dD_stdev | 0.128023435355
varray: [3.019, 2.803, 3.082, 3.06]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [102983.0, 103519.0, 104056.0, 104592.0]
source_col: d(D_H)Mean
measured_d18O_mean | 3.23125
varray: [3.227, 3.216, 3.25, 3.232]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [102983.0, 103519.0, 104056.0, 104592.0]
source_col: d(18_16)Mean
measured_dD_mean | 2.991
varray: [3.019, 2.803, 3.082, 3.06]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [102983.0, 103519.0, 104056.0, 104592.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.014174507634
varray: [3.227, 3.216, 3.25, 3.232]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [102983.0, 103519.0, 104056.0, 104592.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.128023435355
varray: [3.019, 2.803, 3.082, 3.06]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [102983.0, 103519.0, 104056.0, 104592.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 0.555974816788
actual_array: [ 1.93 -15.84]
varray: [3.227, 3.216, 3.25, 3.232]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 3.23125
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [102983.0, 103519.0, 104056.0, 104592.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -14.8856203458
actual_array: [ 18.1 -119. ]
varray: [3.019, 2.803, 3.082, 3.06]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 2.991
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [102983.0, 103519.0, 104056.0, 104592.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.7682
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1187.13005
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 43 - Sample 43 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 43
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.84275
varray: [2.813, 2.909, 2.818, 2.831]
inst: mean_all
operation: mean_all(residuals)
tarray: [105178.0, 105715.0, 106251.0, 106788.0]
source_col: residuals
SAMP_residuals_STD | 0.0448135024295
varray: [2.813, 2.909, 2.818, 2.831]
inst: std_all
operation: std_all(residuals)
tarray: [105178.0, 105715.0, 106251.0, 106788.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.2615
varray: [0.231, 0.24, 0.104, 0.471]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [105178.0, 105715.0, 106251.0, 106788.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.15284959928
varray: [0.231, 0.24, 0.104, 0.471]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [105178.0, 105715.0, 106251.0, 106788.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1300.697
varray: [1310.491, 1302.164, 1297.768, 1292.365]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [105178.0, 105715.0, 106251.0, 106788.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 7.66106585274
varray: [1310.491, 1302.164, 1297.768, 1292.365]
inst: std_all
operation: std_all(baseline_shift)
tarray: [105178.0, 105715.0, 106251.0, 106788.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.1525
varray: [0.882, 1.546, 0.854, 1.328]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [105178.0, 105715.0, 106251.0, 106788.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.340546130013
varray: [0.882, 1.546, 0.854, 1.328]
inst: std_all
operation: std_all(slope_shift)
tarray: [105178.0, 105715.0, 106251.0, 106788.0]
source_col: slope_shift
SAMP_d18O_STD | 0.066935292136
varray: [-0.908, -0.931, -1.0, -1.055]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [105178.0, 105715.0, 106251.0, 106788.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.794752529617
varray: [-12.754, -13.562, -14.201, -14.566]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [105178.0, 105715.0, 106251.0, 106788.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.29475
varray: [108.391, 110.793, 109.425, 108.57]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [105178.0, 105715.0, 106251.0, 106788.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.09601471249
varray: [108.391, 110.793, 109.425, 108.57]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [105178.0, 105715.0, 106251.0, 106788.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.84275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.2615
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1300.697
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.1525
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0526
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0526
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.81670609771
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.06505
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [-0.908, -0.931, -1.0, -1.055]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [105178.0, 105715.0, 106251.0, 106788.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-12.754, -13.562, -14.201, -14.566]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [105178.0, 105715.0, 106251.0, 106788.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.9735
varray: [-0.908, -0.931, -1.0, -1.055]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [105178.0, 105715.0, 106251.0, 106788.0]
source_col: d(18_16)Mean
dD_mean | -13.77075
varray: [-12.754, -13.562, -14.201, -14.566]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [105178.0, 105715.0, 106251.0, 106788.0]
source_col: d(D_H)Mean
d18O_stdev | 0.066935292136
varray: [-0.908, -0.931, -1.0, -1.055]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [105178.0, 105715.0, 106251.0, 106788.0]
source_col: d(18_16)Mean
dD_stdev | 0.794752529617
varray: [-12.754, -13.562, -14.201, -14.566]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [105178.0, 105715.0, 106251.0, 106788.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.9735
varray: [-0.908, -0.931, -1.0, -1.055]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [105178.0, 105715.0, 106251.0, 106788.0]
source_col: d(18_16)Mean
measured_dD_mean | -13.77075
varray: [-12.754, -13.562, -14.201, -14.566]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [105178.0, 105715.0, 106251.0, 106788.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.066935292136
varray: [-0.908, -0.931, -1.0, -1.055]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [105178.0, 105715.0, 106251.0, 106788.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.794752529617
varray: [-12.754, -13.562, -14.201, -14.566]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [105178.0, 105715.0, 106251.0, 106788.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.64319834098
actual_array: [ 1.93 -15.84]
varray: [-0.908, -0.931, -1.0, -1.055]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -0.9735
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [105178.0, 105715.0, 106251.0, 106788.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -30.5597396623
actual_array: [ 18.1 -119. ]
varray: [-12.754, -13.562, -14.201, -14.566]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -13.77075
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [105178.0, 105715.0, 106251.0, 106788.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.6382
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 5.21605
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 44 - Sample 44 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 44
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.865
varray: [2.877, 2.856, 2.876, 2.851]
inst: mean_all
operation: mean_all(residuals)
tarray: [107374.0, 107911.0, 108447.0, 108984.0]
source_col: residuals
SAMP_residuals_STD | 0.0134412301024
varray: [2.877, 2.856, 2.876, 2.851]
inst: std_all
operation: std_all(residuals)
tarray: [107374.0, 107911.0, 108447.0, 108984.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.309
varray: [0.362, 0.277, 0.269, 0.328]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [107374.0, 107911.0, 108447.0, 108984.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0439469377014
varray: [0.362, 0.277, 0.269, 0.328]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [107374.0, 107911.0, 108447.0, 108984.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1295.00425
varray: [1295.807, 1294.752, 1295.773, 1293.685]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [107374.0, 107911.0, 108447.0, 108984.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.00655100053
varray: [1295.807, 1294.752, 1295.773, 1293.685]
inst: std_all
operation: std_all(baseline_shift)
tarray: [107374.0, 107911.0, 108447.0, 108984.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.193
varray: [0.922, 1.292, 1.363, 1.195]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [107374.0, 107911.0, 108447.0, 108984.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.193344252565
varray: [0.922, 1.292, 1.363, 1.195]
inst: std_all
operation: std_all(slope_shift)
tarray: [107374.0, 107911.0, 108447.0, 108984.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0698826397135
varray: [1.981, 2.1, 2.095, 2.145]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [107374.0, 107911.0, 108447.0, 108984.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.992784300171
varray: [6.986, 8.691, 8.95, 9.161]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [107374.0, 107911.0, 108447.0, 108984.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.146
varray: [111.839, 109.098, 110.023, 109.624]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [107374.0, 107911.0, 108447.0, 108984.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.19054133346
varray: [111.839, 109.098, 110.023, 109.624]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [107374.0, 107911.0, 108447.0, 108984.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.865
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.309
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1295.00425
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.193
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.03035
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.03035
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.04823251075
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.2138
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [1.981, 2.1, 2.095, 2.145]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [107374.0, 107911.0, 108447.0, 108984.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [6.986, 8.691, 8.95, 9.161]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [107374.0, 107911.0, 108447.0, 108984.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.08025
varray: [1.981, 2.1, 2.095, 2.145]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [107374.0, 107911.0, 108447.0, 108984.0]
source_col: d(18_16)Mean
dD_mean | 8.447
varray: [6.986, 8.691, 8.95, 9.161]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [107374.0, 107911.0, 108447.0, 108984.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0698826397135
varray: [1.981, 2.1, 2.095, 2.145]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [107374.0, 107911.0, 108447.0, 108984.0]
source_col: d(18_16)Mean
dD_stdev | 0.992784300171
varray: [6.986, 8.691, 8.95, 9.161]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [107374.0, 107911.0, 108447.0, 108984.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.08025
varray: [1.981, 2.1, 2.095, 2.145]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [107374.0, 107911.0, 108447.0, 108984.0]
source_col: d(18_16)Mean
measured_dD_mean | 8.447
varray: [6.986, 8.691, 8.95, 9.161]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [107374.0, 107911.0, 108447.0, 108984.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0698826397135
varray: [1.981, 2.1, 2.095, 2.145]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [107374.0, 107911.0, 108447.0, 108984.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.992784300171
varray: [6.986, 8.691, 8.95, 9.161]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [107374.0, 107911.0, 108447.0, 108984.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.59349858938
actual_array: [ 1.93 -15.84]
varray: [1.981, 2.1, 2.095, 2.145]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 2.08025
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [107374.0, 107911.0, 108447.0, 108984.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -9.78364740194
actual_array: [ 18.1 -119. ]
varray: [6.986, 8.691, 8.95, 9.161]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 8.447
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [107374.0, 107911.0, 108447.0, 108984.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.61595
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.4767
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 45 - Sample 45 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 45
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.80575
varray: [2.851, 2.715, 2.871, 2.786]
inst: mean_all
operation: mean_all(residuals)
tarray: [109570.0, 110107.0, 110644.0, 111181.0]
source_col: residuals
SAMP_residuals_STD | 0.0705472654797
varray: [2.851, 2.715, 2.871, 2.786]
inst: std_all
operation: std_all(residuals)
tarray: [109570.0, 110107.0, 110644.0, 111181.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0675
varray: [0.007, -0.013, 0.152, 0.124]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [109570.0, 110107.0, 110644.0, 111181.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0826095232605
varray: [0.007, -0.013, 0.152, 0.124]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [109570.0, 110107.0, 110644.0, 111181.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1297.588
varray: [1299.48, 1296.153, 1298.822, 1295.897]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [109570.0, 110107.0, 110644.0, 111181.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.827669372
varray: [1299.48, 1296.153, 1298.822, 1295.897]
inst: std_all
operation: std_all(baseline_shift)
tarray: [109570.0, 110107.0, 110644.0, 111181.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.429
varray: [0.22, 0.588, 0.021, 0.887]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [109570.0, 110107.0, 110644.0, 111181.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.38522287921
varray: [0.22, 0.588, 0.021, 0.887]
inst: std_all
operation: std_all(slope_shift)
tarray: [109570.0, 110107.0, 110644.0, 111181.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0583238087462
varray: [0.892, 0.795, 0.806, 0.753]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [109570.0, 110107.0, 110644.0, 111181.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.767214441991
varray: [-2.363, -3.728, -3.685, -4.124]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [109570.0, 110107.0, 110644.0, 111181.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.84825
varray: [110.467, 106.602, 110.79, 107.534]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [109570.0, 110107.0, 110644.0, 111181.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.0947264539
varray: [110.467, 106.602, 110.79, 107.534]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [109570.0, 110107.0, 110644.0, 111181.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.80575
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0675
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1297.588
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.429
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0896
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0896
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.09461723108
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.51155
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [0.892, 0.795, 0.806, 0.753]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [109570.0, 110107.0, 110644.0, 111181.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-2.363, -3.728, -3.685, -4.124]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [109570.0, 110107.0, 110644.0, 111181.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.8115
varray: [0.892, 0.795, 0.806, 0.753]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [109570.0, 110107.0, 110644.0, 111181.0]
source_col: d(18_16)Mean
dD_mean | -3.475
varray: [-2.363, -3.728, -3.685, -4.124]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [109570.0, 110107.0, 110644.0, 111181.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0583238087462
varray: [0.892, 0.795, 0.806, 0.753]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [109570.0, 110107.0, 110644.0, 111181.0]
source_col: d(18_16)Mean
dD_stdev | 0.767214441991
varray: [-2.363, -3.728, -3.685, -4.124]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [109570.0, 110107.0, 110644.0, 111181.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.8115
varray: [0.892, 0.795, 0.806, 0.753]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [109570.0, 110107.0, 110644.0, 111181.0]
source_col: d(18_16)Mean
measured_dD_mean | -3.475
varray: [-2.363, -3.728, -3.685, -4.124]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [109570.0, 110107.0, 110644.0, 111181.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0583238087462
varray: [0.892, 0.795, 0.806, 0.753]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [109570.0, 110107.0, 110644.0, 111181.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.767214441991
varray: [-2.363, -3.728, -3.685, -4.124]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [109570.0, 110107.0, 110644.0, 111181.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.86056582142
actual_array: [ 1.93 -15.84]
varray: [0.892, 0.795, 0.806, 0.753]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.8115
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [109570.0, 110107.0, 110644.0, 111181.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -20.9320567561
actual_array: [ 18.1 -119. ]
varray: [-2.363, -3.728, -3.685, -4.124]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -3.475
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [109570.0, 110107.0, 110644.0, 111181.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.6752
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 2.10705
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 46 - Sample 46 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 46
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.7705
varray: [2.759, 2.784, 2.752, 2.787]
inst: mean_all
operation: mean_all(residuals)
tarray: [111767.0, 112304.0, 112840.0, 113377.0]
source_col: residuals
SAMP_residuals_STD | 0.0175973482851
varray: [2.759, 2.784, 2.752, 2.787]
inst: std_all
operation: std_all(residuals)
tarray: [111767.0, 112304.0, 112840.0, 113377.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.932
varray: [-1.031, -0.956, -0.716, -1.025]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [111767.0, 112304.0, 112840.0, 113377.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.147966212359
varray: [-1.031, -0.956, -0.716, -1.025]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [111767.0, 112304.0, 112840.0, 113377.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 2191.4725
varray: [2189.929, 2192.944, 2188.912, 2194.105]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [111767.0, 112304.0, 112840.0, 113377.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.45179668815
varray: [2189.929, 2192.944, 2188.912, 2194.105]
inst: std_all
operation: std_all(baseline_shift)
tarray: [111767.0, 112304.0, 112840.0, 113377.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -32.6715
varray: [-32.583, -32.65, -32.471, -32.982]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [111767.0, 112304.0, 112840.0, 113377.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.219776401524
varray: [-32.583, -32.65, -32.471, -32.982]
inst: std_all
operation: std_all(slope_shift)
tarray: [111767.0, 112304.0, 112840.0, 113377.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0651479086387
varray: [0.466, 0.478, 0.358, 0.361]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [111767.0, 112304.0, 112840.0, 113377.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.985024999683
varray: [-15.791, -17.534, -17.825, -17.859]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [111767.0, 112304.0, 112840.0, 113377.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 101.1775
varray: [101.059, 100.678, 102.085, 100.888]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [111767.0, 112304.0, 112840.0, 113377.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.624742346892
varray: [101.059, 100.678, 102.085, 100.888]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [111767.0, 112304.0, 112840.0, 113377.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.7705
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.932
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 2191.4725
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 32.6715
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.12485
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.12485
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.31208662165
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -10.1823
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [0.466, 0.478, 0.358, 0.361]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [111767.0, 112304.0, 112840.0, 113377.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-15.791, -17.534, -17.825, -17.859]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [111767.0, 112304.0, 112840.0, 113377.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.41575
varray: [0.466, 0.478, 0.358, 0.361]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [111767.0, 112304.0, 112840.0, 113377.0]
source_col: d(18_16)Mean
dD_mean | -17.25225
varray: [-15.791, -17.534, -17.825, -17.859]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [111767.0, 112304.0, 112840.0, 113377.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0651479086387
varray: [0.466, 0.478, 0.358, 0.361]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [111767.0, 112304.0, 112840.0, 113377.0]
source_col: d(18_16)Mean
dD_stdev | 0.985024999683
varray: [-15.791, -17.534, -17.825, -17.859]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [111767.0, 112304.0, 112840.0, 113377.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.41575
varray: [0.466, 0.478, 0.358, 0.361]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [111767.0, 112304.0, 112840.0, 113377.0]
source_col: d(18_16)Mean
measured_dD_mean | -17.25225
varray: [-15.791, -17.534, -17.825, -17.859]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [111767.0, 112304.0, 112840.0, 113377.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0651479086387
varray: [0.466, 0.478, 0.358, 0.361]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [111767.0, 112304.0, 112840.0, 113377.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.985024999683
varray: [-15.791, -17.534, -17.825, -17.859]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [111767.0, 112304.0, 112840.0, 113377.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.25579093045
actual_array: [ 1.93 -15.84]
varray: [0.466, 0.478, 0.358, 0.361]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.41575
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [111767.0, 112304.0, 112840.0, 113377.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -33.8153332847
actual_array: [ 18.1 -119. ]
varray: [-15.791, -17.534, -17.825, -17.859]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -17.25225
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [111767.0, 112304.0, 112840.0, 113377.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.71045
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 895.99155
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 47 - Sample 47 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 47
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.69825
varray: [2.507, 2.782, 2.689, 2.815]
inst: mean_all
operation: mean_all(residuals)
tarray: [113964.0, 114501.0, 115037.0, 115574.0]
source_col: residuals
SAMP_residuals_STD | 0.138210889585
varray: [2.507, 2.782, 2.689, 2.815]
inst: std_all
operation: std_all(residuals)
tarray: [113964.0, 114501.0, 115037.0, 115574.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.88375
varray: [-1.059, -0.91, -0.838, -0.728]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [113964.0, 114501.0, 115037.0, 115574.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.138747672653
varray: [-1.059, -0.91, -0.838, -0.728]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [113964.0, 114501.0, 115037.0, 115574.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1301.47525
varray: [1301.747, 1303.79, 1299.291, 1301.073]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [113964.0, 114501.0, 115037.0, 115574.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.85873332765
varray: [1301.747, 1303.79, 1299.291, 1301.073]
inst: std_all
operation: std_all(baseline_shift)
tarray: [113964.0, 114501.0, 115037.0, 115574.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -2.79025
varray: [-3.104, -3.079, -2.816, -2.162]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [113964.0, 114501.0, 115037.0, 115574.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.438625409661
varray: [-3.104, -3.079, -2.816, -2.162]
inst: std_all
operation: std_all(slope_shift)
tarray: [113964.0, 114501.0, 115037.0, 115574.0]
source_col: slope_shift
SAMP_d18O_STD | 0.099513818136
varray: [2.54, 2.695, 2.702, 2.777]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [113964.0, 114501.0, 115037.0, 115574.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.03965603286
varray: [-8.366, -6.372, -6.445, -6.105]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [113964.0, 114501.0, 115037.0, 115574.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 107.84725
varray: [100.938, 110.486, 108.893, 111.072]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [113964.0, 114501.0, 115037.0, 115574.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 4.69728087124
varray: [100.938, 110.486, 108.893, 111.072]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [113964.0, 114501.0, 115037.0, 115574.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.69825
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.88375
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1301.47525
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.79025
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1971
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1971
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.8074671456
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.51255
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [2.54, 2.695, 2.702, 2.777]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [113964.0, 114501.0, 115037.0, 115574.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-8.366, -6.372, -6.445, -6.105]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [113964.0, 114501.0, 115037.0, 115574.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.6785
varray: [2.54, 2.695, 2.702, 2.777]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [113964.0, 114501.0, 115037.0, 115574.0]
source_col: d(18_16)Mean
dD_mean | -6.822
varray: [-8.366, -6.372, -6.445, -6.105]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [113964.0, 114501.0, 115037.0, 115574.0]
source_col: d(D_H)Mean
d18O_stdev | 0.099513818136
varray: [2.54, 2.695, 2.702, 2.777]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [113964.0, 114501.0, 115037.0, 115574.0]
source_col: d(18_16)Mean
dD_stdev | 1.03965603286
varray: [-8.366, -6.372, -6.445, -6.105]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [113964.0, 114501.0, 115037.0, 115574.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.6785
varray: [2.54, 2.695, 2.702, 2.777]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [113964.0, 114501.0, 115037.0, 115574.0]
source_col: d(18_16)Mean
measured_dD_mean | -6.822
varray: [-8.366, -6.372, -6.445, -6.105]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [113964.0, 114501.0, 115037.0, 115574.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.099513818136
varray: [2.54, 2.695, 2.702, 2.777]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [113964.0, 114501.0, 115037.0, 115574.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.03965603286
varray: [-8.366, -6.372, -6.445, -6.105]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [113964.0, 114501.0, 115037.0, 115574.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 0.00395793993346
actual_array: [ 1.93 -15.84]
varray: [2.54, 2.695, 2.702, 2.777]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 2.6785
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [113964.0, 114501.0, 115037.0, 115574.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -24.0618777683
actual_array: [ 18.1 -119. ]
varray: [-8.366, -6.372, -6.445, -6.105]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -6.822
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [113964.0, 114501.0, 115037.0, 115574.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.7827
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 5.9943
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 48 - Sample 48 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 48
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.83225
varray: [2.666, 2.882, 2.919, 2.862]
inst: mean_all
operation: mean_all(residuals)
tarray: [116161.0, 116698.0, 117234.0, 117770.0]
source_col: residuals
SAMP_residuals_STD | 0.113320710081
varray: [2.666, 2.882, 2.919, 2.862]
inst: std_all
operation: std_all(residuals)
tarray: [116161.0, 116698.0, 117234.0, 117770.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4225
varray: [0.306, 0.372, 0.496, 0.516]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [116161.0, 116698.0, 117234.0, 117770.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.100444014257
varray: [0.306, 0.372, 0.496, 0.516]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [116161.0, 116698.0, 117234.0, 117770.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1293.6305
varray: [1289.604, 1294.886, 1295.765, 1294.267]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [116161.0, 116698.0, 117234.0, 117770.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.75379768078
varray: [1289.604, 1294.886, 1295.765, 1294.267]
inst: std_all
operation: std_all(baseline_shift)
tarray: [116161.0, 116698.0, 117234.0, 117770.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.38875
varray: [1.36, 1.267, 1.703, 1.225]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [116161.0, 116698.0, 117234.0, 117770.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.216961402097
varray: [1.36, 1.267, 1.703, 1.225]
inst: std_all
operation: std_all(slope_shift)
tarray: [116161.0, 116698.0, 117234.0, 117770.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0259743463184
varray: [0.581, 0.573, 0.587, 0.631]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [116161.0, 116698.0, 117234.0, 117770.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.888538641066
varray: [-0.75, 0.853, 0.77, 1.284]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [116161.0, 116698.0, 117234.0, 117770.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.42625
varray: [104.964, 110.539, 111.096, 111.106]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [116161.0, 116698.0, 117234.0, 117770.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.98660971393
varray: [104.964, 110.539, 111.096, 111.106]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [116161.0, 116698.0, 117234.0, 117770.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.83225
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1293.6305
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.38875
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0631
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0631
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.17935655448
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.93355
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [0.581, 0.573, 0.587, 0.631]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [116161.0, 116698.0, 117234.0, 117770.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-0.75, 0.853, 0.77, 1.284]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [116161.0, 116698.0, 117234.0, 117770.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.593
varray: [0.581, 0.573, 0.587, 0.631]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [116161.0, 116698.0, 117234.0, 117770.0]
source_col: d(18_16)Mean
dD_mean | 0.53925
varray: [-0.75, 0.853, 0.77, 1.284]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [116161.0, 116698.0, 117234.0, 117770.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0259743463184
varray: [0.581, 0.573, 0.587, 0.631]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [116161.0, 116698.0, 117234.0, 117770.0]
source_col: d(18_16)Mean
dD_stdev | 0.888538641066
varray: [-0.75, 0.853, 0.77, 1.284]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [116161.0, 116698.0, 117234.0, 117770.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.593
varray: [0.581, 0.573, 0.587, 0.631]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [116161.0, 116698.0, 117234.0, 117770.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.53925
varray: [-0.75, 0.853, 0.77, 1.284]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [116161.0, 116698.0, 117234.0, 117770.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0259743463184
varray: [0.581, 0.573, 0.587, 0.631]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [116161.0, 116698.0, 117234.0, 117770.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.888538641066
varray: [-0.75, 0.853, 0.77, 1.284]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [116161.0, 116698.0, 117234.0, 117770.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.07877602059
actual_array: [ 1.93 -15.84]
varray: [0.581, 0.573, 0.587, 0.631]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.593
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [116161.0, 116698.0, 117234.0, 117770.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -17.1782820329
actual_array: [ 18.1 -119. ]
varray: [-0.75, 0.853, 0.77, 1.284]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 0.53925
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [116161.0, 116698.0, 117234.0, 117770.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.6487
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.85045
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 49 - Sample 49 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 49
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.84375
varray: [2.815, 2.863, 2.866, 2.831]
inst: mean_all
operation: mean_all(residuals)
tarray: [118357.0, 118894.0, 119431.0, 119967.0]
source_col: residuals
SAMP_residuals_STD | 0.0248646335183
varray: [2.815, 2.863, 2.866, 2.831]
inst: std_all
operation: std_all(residuals)
tarray: [118357.0, 118894.0, 119431.0, 119967.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.45025
varray: [0.491, 0.443, 0.568, 0.299]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [118357.0, 118894.0, 119431.0, 119967.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.113217710629
varray: [0.491, 0.443, 0.568, 0.299]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [118357.0, 118894.0, 119431.0, 119967.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1293.031
varray: [1291.151, 1293.374, 1293.156, 1294.443]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [118357.0, 118894.0, 119431.0, 119967.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.37373190009
varray: [1291.151, 1293.374, 1293.156, 1294.443]
inst: std_all
operation: std_all(baseline_shift)
tarray: [118357.0, 118894.0, 119431.0, 119967.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.80275
varray: [2.064, 1.899, 1.718, 1.53]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [118357.0, 118894.0, 119431.0, 119967.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.230282978094
varray: [2.064, 1.899, 1.718, 1.53]
inst: std_all
operation: std_all(slope_shift)
tarray: [118357.0, 118894.0, 119431.0, 119967.0]
source_col: slope_shift
SAMP_d18O_STD | 0.220211072988
varray: [-7.047, -7.414, -7.478, -7.536]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [118357.0, 118894.0, 119431.0, 119967.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.99440526928
varray: [-52.158, -57.07, -58.167, -58.724]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [118357.0, 118894.0, 119431.0, 119967.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.3165
varray: [108.055, 109.293, 110.493, 109.425]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [118357.0, 118894.0, 119431.0, 119967.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.997974448571
varray: [108.055, 109.293, 110.493, 109.425]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [118357.0, 118894.0, 119431.0, 119967.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.84375
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.45025
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1293.031
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.80275
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0516
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0516
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.78216795897
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.0433
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [-7.047, -7.414, -7.478, -7.536]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [118357.0, 118894.0, 119431.0, 119967.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-52.158, -57.07, -58.167, -58.724]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [118357.0, 118894.0, 119431.0, 119967.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.36875
varray: [-7.047, -7.414, -7.478, -7.536]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [118357.0, 118894.0, 119431.0, 119967.0]
source_col: d(18_16)Mean
dD_mean | -56.52975
varray: [-52.158, -57.07, -58.167, -58.724]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [118357.0, 118894.0, 119431.0, 119967.0]
source_col: d(D_H)Mean
d18O_stdev | 0.220211072988
varray: [-7.047, -7.414, -7.478, -7.536]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [118357.0, 118894.0, 119431.0, 119967.0]
source_col: d(18_16)Mean
dD_stdev | 2.99440526928
varray: [-52.158, -57.07, -58.167, -58.724]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [118357.0, 118894.0, 119431.0, 119967.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.36875
varray: [-7.047, -7.414, -7.478, -7.536]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [118357.0, 118894.0, 119431.0, 119967.0]
source_col: d(18_16)Mean
measured_dD_mean | -56.52975
varray: [-52.158, -57.07, -58.167, -58.724]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [118357.0, 118894.0, 119431.0, 119967.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.220211072988
varray: [-7.047, -7.414, -7.478, -7.536]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [118357.0, 118894.0, 119431.0, 119967.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.99440526928
varray: [-52.158, -57.07, -58.167, -58.724]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [118357.0, 118894.0, 119431.0, 119967.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.0299661957
actual_array: [ 1.93 -15.84]
varray: [-7.047, -7.414, -7.478, -7.536]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.36875
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [118357.0, 118894.0, 119431.0, 119967.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -70.5442083398
actual_array: [ 18.1 -119. ]
varray: [-52.158, -57.07, -58.167, -58.724]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -56.52975
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [118357.0, 118894.0, 119431.0, 119967.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.6372
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -2.44995
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 50 - Sample 50 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 50
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.80525
varray: [2.796, 2.792, 2.818, 2.815]
inst: mean_all
operation: mean_all(residuals)
tarray: [120552.0, 121089.0, 121625.0, 122162.0]
source_col: residuals
SAMP_residuals_STD | 0.0131497781984
varray: [2.796, 2.792, 2.818, 2.815]
inst: std_all
operation: std_all(residuals)
tarray: [120552.0, 121089.0, 121625.0, 122162.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.39675
varray: [0.373, 0.242, 0.334, 0.638]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [120552.0, 121089.0, 121625.0, 122162.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.169951708827
varray: [0.373, 0.242, 0.334, 0.638]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [120552.0, 121089.0, 121625.0, 122162.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1290.48275
varray: [1291.316, 1291.182, 1290.065, 1289.368]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [120552.0, 121089.0, 121625.0, 122162.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.931027881788
varray: [1291.316, 1291.182, 1290.065, 1289.368]
inst: std_all
operation: std_all(baseline_shift)
tarray: [120552.0, 121089.0, 121625.0, 122162.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.4275
varray: [1.428, 1.241, 1.335, 1.706]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [120552.0, 121089.0, 121625.0, 122162.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.200749429223
varray: [1.428, 1.241, 1.335, 1.706]
inst: std_all
operation: std_all(slope_shift)
tarray: [120552.0, 121089.0, 121625.0, 122162.0]
source_col: slope_shift
SAMP_d18O_STD | 0.232429738487
varray: [0.771, 1.125, 1.225, 1.294]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [120552.0, 121089.0, 121625.0, 122162.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.96236328798
varray: [-4.059, 0.738, 1.914, 2.43]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [120552.0, 121089.0, 121625.0, 122162.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 107.66125
varray: [107.514, 107.933, 107.544, 107.654]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [120552.0, 121089.0, 121625.0, 122162.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.190902025483
varray: [107.514, 107.933, 107.544, 107.654]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [120552.0, 121089.0, 121625.0, 122162.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.80525
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.39675
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1290.48275
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.4275
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0901
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0901
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.11188630045
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.69855
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [0.771, 1.125, 1.225, 1.294]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [120552.0, 121089.0, 121625.0, 122162.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-4.059, 0.738, 1.914, 2.43]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [120552.0, 121089.0, 121625.0, 122162.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.10375
varray: [0.771, 1.125, 1.225, 1.294]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [120552.0, 121089.0, 121625.0, 122162.0]
source_col: d(18_16)Mean
dD_mean | 0.25575
varray: [-4.059, 0.738, 1.914, 2.43]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [120552.0, 121089.0, 121625.0, 122162.0]
source_col: d(D_H)Mean
d18O_stdev | 0.232429738487
varray: [0.771, 1.125, 1.225, 1.294]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [120552.0, 121089.0, 121625.0, 122162.0]
source_col: d(18_16)Mean
dD_stdev | 2.96236328798
varray: [-4.059, 0.738, 1.914, 2.43]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [120552.0, 121089.0, 121625.0, 122162.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.10375
varray: [0.771, 1.125, 1.225, 1.294]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [120552.0, 121089.0, 121625.0, 122162.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.25575
varray: [-4.059, 0.738, 1.914, 2.43]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [120552.0, 121089.0, 121625.0, 122162.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.232429738487
varray: [0.771, 1.125, 1.225, 1.294]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [120552.0, 121089.0, 121625.0, 122162.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.96236328798
varray: [-4.059, 0.738, 1.914, 2.43]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [120552.0, 121089.0, 121625.0, 122162.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.56870343831
actual_array: [ 1.93 -15.84]
varray: [0.771, 1.125, 1.225, 1.294]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 1.10375
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [120552.0, 121089.0, 121625.0, 122162.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -17.4433863821
actual_array: [ 18.1 -119. ]
varray: [-4.059, 0.738, 1.914, 2.43]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 0.25575
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [120552.0, 121089.0, 121625.0, 122162.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.6757
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -4.9982
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 51 - Sample 51 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Sample Variables
Sample 51
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.82075
varray: [2.817, 2.777, 2.842, 2.847]
inst: mean_all
operation: mean_all(residuals)
tarray: [122748.0, 123285.0, 123821.0, 124358.0]
source_col: residuals
SAMP_residuals_STD | 0.0319830684373
varray: [2.817, 2.777, 2.842, 2.847]
inst: std_all
operation: std_all(residuals)
tarray: [122748.0, 123285.0, 123821.0, 124358.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.33675
varray: [0.446, 0.219, 0.275, 0.407]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [122748.0, 123285.0, 123821.0, 124358.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.107313792217
varray: [0.446, 0.219, 0.275, 0.407]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [122748.0, 123285.0, 123821.0, 124358.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1293.6605
varray: [1293.216, 1293.955, 1293.737, 1293.734]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [122748.0, 123285.0, 123821.0, 124358.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.313881612502
varray: [1293.216, 1293.955, 1293.737, 1293.734]
inst: std_all
operation: std_all(baseline_shift)
tarray: [122748.0, 123285.0, 123821.0, 124358.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.8425
varray: [1.803, 1.779, 1.531, 2.257]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [122748.0, 123285.0, 123821.0, 124358.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.302453853229
varray: [1.803, 1.779, 1.531, 2.257]
inst: std_all
operation: std_all(slope_shift)
tarray: [122748.0, 123285.0, 123821.0, 124358.0]
source_col: slope_shift
SAMP_d18O_STD | 0.22313430186
varray: [-7.03, -7.398, -7.508, -7.491]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [122748.0, 123285.0, 123821.0, 124358.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.22373747897
varray: [-51.715, -56.879, -58.282, -58.736]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [122748.0, 123285.0, 123821.0, 124358.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.21125
varray: [109.65, 108.499, 109.353, 109.343]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [122748.0, 123285.0, 123821.0, 124358.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.495732706876
varray: [109.65, 108.499, 109.353, 109.343]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [122748.0, 123285.0, 123821.0, 124358.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.82075
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.33675
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1293.6605
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.8425
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0746
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0746
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.57654514998
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.14855
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1865
varray: [-7.03, -7.398, -7.508, -7.491]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [122748.0, 123285.0, 123821.0, 124358.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.3479
varray: [-51.715, -56.879, -58.282, -58.736]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [122748.0, 123285.0, 123821.0, 124358.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.35675
varray: [-7.03, -7.398, -7.508, -7.491]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [122748.0, 123285.0, 123821.0, 124358.0]
source_col: d(18_16)Mean
dD_mean | -56.403
varray: [-51.715, -56.879, -58.282, -58.736]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [122748.0, 123285.0, 123821.0, 124358.0]
source_col: d(D_H)Mean
d18O_stdev | 0.22313430186
varray: [-7.03, -7.398, -7.508, -7.491]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [122748.0, 123285.0, 123821.0, 124358.0]
source_col: d(18_16)Mean
dD_stdev | 3.22373747897
varray: [-51.715, -56.879, -58.282, -58.736]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [122748.0, 123285.0, 123821.0, 124358.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.35675
varray: [-7.03, -7.398, -7.508, -7.491]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [122748.0, 123285.0, 123821.0, 124358.0]
source_col: d(18_16)Mean
measured_dD_mean | -56.403
varray: [-51.715, -56.879, -58.282, -58.736]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [122748.0, 123285.0, 123821.0, 124358.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.22313430186
varray: [-7.03, -7.398, -7.508, -7.491]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [122748.0, 123285.0, 123821.0, 124358.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.22373747897
varray: [-51.715, -56.879, -58.282, -58.736]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [122748.0, 123285.0, 123821.0, 124358.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.0179821115
actual_array: [ 1.93 -15.84]
varray: [-7.03, -7.398, -7.508, -7.491]
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.35675
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: [122748.0, 123285.0, 123821.0, 124358.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -70.4256828503
actual_array: [ 18.1 -119. ]
varray: [-51.715, -56.879, -58.282, -58.736]
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -56.403
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: [122748.0, 123285.0, 123821.0, 124358.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.6602
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.82045
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 1 - Sample 1 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 1 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.995
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.765
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.6553855881
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.995
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.1802046061
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -65.765
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 1 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.635
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.473
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2958630631
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.635
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.1667024297
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -61.473
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 1 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.553
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.565
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2139718213
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.553
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3176204221
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -60.565
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 1 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.494
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.185
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1550500742
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.494
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.9622777318
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -60.185
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 1 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.464
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.953
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1250898638
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.464
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.7453316682
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -59.953
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 1 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.417
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.925
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0781522008
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.417
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.7191485226
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -59.925
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 1 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.498
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.02
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1590447689
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.498
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.8079841952
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -60.02
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 1 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.429
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.898
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.090136285
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.429
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.6939004893
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -59.898
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 1 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.447
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.796
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1081124112
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.447
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.5985190303
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -59.796
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 1 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.463
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.74
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1240911901
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.463
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.5461527391
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -59.74
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 2 - Sample 2 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 2 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.427
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6071
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 38.2655
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.939
operation: col(d(18_16)Mean)
inst: col
dD_actual | 27.899
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.26278611411
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 3.939
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 8.40615789553
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 27.899
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 2 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.427
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6071
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 38.2655
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.421
operation: col(d(18_16)Mean)
inst: col
dD_actual | 36.041
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.74414682807
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 4.421
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 16.0198425928
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 36.041
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 2 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.427
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6071
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 38.2655
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.625
operation: col(d(18_16)Mean)
inst: col
dD_actual | 38.497
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.94787625888
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 4.625
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 18.3164785074
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 38.497
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 2 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.427
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6071
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 38.2655
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.658
operation: col(d(18_16)Mean)
inst: col
dD_actual | 39.122
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.98083249033
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 4.658
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 18.9009237218
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 39.122
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 2 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.427
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6071
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 38.2655
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.724
operation: col(d(18_16)Mean)
inst: col
dD_actual | 39.579
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.04674495324
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 4.724
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.3282700626
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 39.579
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 2 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.427
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6071
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 38.2655
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.653
operation: col(d(18_16)Mean)
inst: col
dD_actual | 39.6
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.97583912193
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 4.653
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.3479074218
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 39.6
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 2 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.427
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6071
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 38.2655
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.781
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.205
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.10366935303
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 4.781
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.9136503894
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 40.205
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 2 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.427
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6071
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 38.2655
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.725
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.48
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.04774362692
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 4.725
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.1708062837
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 40.48
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 2 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.427
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6071
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 38.2655
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.792
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.805
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.11465476351
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 4.792
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.4747177952
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 40.805
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 2 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.427
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6071
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 38.2655
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.753
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.427
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.07570648997
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 4.753
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.1212453296
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 40.427
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 3 - Sample 3 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 3 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.197
operation: col(d(18_16)Mean)
inst: col
dD_actual | -92.886
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.8518123932
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -12.197
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -104.541386463
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -92.886
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 3 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.002
operation: col(d(18_16)Mean)
inst: col
dD_actual | -105.698
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.655744706
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -13.002
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -116.522045802
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -105.698
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 3 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.187
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.406
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.8404993368
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -13.187
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -119.054330027
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -108.406
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 3 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.345
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.117
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9982897783
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -13.345
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.654307246
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -110.117
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 3 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.214
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.983
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.8674635262
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -13.214
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.529002192
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -109.983
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 3 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.363
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.818
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0162659046
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -13.363
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.309820999
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -110.818
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 3 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.338
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.128
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9912990626
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -13.338
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.599705825
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -111.128
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 3 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.4
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.395
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0532168308
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -13.4
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.849380821
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -111.395
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 3 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.416
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.414
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0691956097
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -13.416
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.867147955
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -111.414
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 3 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.403
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.634
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0562128518
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -13.403
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.072872671
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -111.634
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 4 - Sample 4 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 4 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.005
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.361
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.6653723249
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -8.005
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.7375315626
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -66.361
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 4 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.739
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.962
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3997251259
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.739
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.6239723654
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -61.962
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 4 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.69
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.034
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3507901155
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.69
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.756188111
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -61.034
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 4 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.661
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.792
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3218285788
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.661
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.529890924
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -60.792
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 4 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.628
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.472
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2888723474
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.628
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.2306549742
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -60.472
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 4 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.602
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.476
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2629068317
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.602
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.2343954236
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -60.476
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 4 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.632
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.23
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2928670421
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.632
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.0043577872
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -60.23
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 4 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.616
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.175
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2768882632
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.616
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.9529266083
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -60.175
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 4 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.609
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.101
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2698975474
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.609
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.883728295
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -60.101
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 4 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.605
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.237
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2659028527
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.605
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.0109035736
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -60.237
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 5 - Sample 5 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 5 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.375
operation: col(d(18_16)Mean)
inst: col
dD_actual | -113.704
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.0269236692
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -14.375
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -124.008555221
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -113.704
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 5 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.684
operation: col(d(18_16)Mean)
inst: col
dD_actual | -118.569
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.3355138364
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -14.684
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -128.55787677
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -118.569
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 5 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.757
operation: col(d(18_16)Mean)
inst: col
dD_actual | -119.794
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.4084170151
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -14.757
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -129.70338939
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -119.794
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 5 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.742
operation: col(d(18_16)Mean)
inst: col
dD_actual | -120.13
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.3934369099
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -14.742
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -130.017587137
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -120.13
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 6 - Sample 6 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 6 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.788
operation: col(d(18_16)Mean)
inst: col
dD_actual | -120.079
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.4393758992
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -14.788
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -129.969896408
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -120.079
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 6 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.843
operation: col(d(18_16)Mean)
inst: col
dD_actual | -120.316
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.4943029516
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -14.843
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -130.191518033
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -120.316
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 6 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.844
operation: col(d(18_16)Mean)
inst: col
dD_actual | -120.509
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.4953016253
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -14.844
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -130.371994715
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -120.509
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 6 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.812
operation: col(d(18_16)Mean)
inst: col
dD_actual | -120.725
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.4633440675
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -14.812
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -130.573978981
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -120.725
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 7 - Sample 7 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 7 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.742
operation: col(d(18_16)Mean)
inst: col
dD_actual | -120.455
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.3934369099
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -14.742
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -130.321498649
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -120.455
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 7 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.426
operation: col(d(18_16)Mean)
inst: col
dD_actual | -119.345
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.0778560269
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -14.426
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -129.283523948
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -119.345
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 7 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.684
operation: col(d(18_16)Mean)
inst: col
dD_actual | -117.046
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.336840156
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -13.684
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -127.133700671
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -117.046
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 7 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.61
operation: col(d(18_16)Mean)
inst: col
dD_actual | -120.266
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.2616119841
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -14.61
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -130.144762416
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -120.266
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 8 - Sample 8 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 8 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.76
operation: col(d(18_16)Mean)
inst: col
dD_actual | -120.746
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.4114130361
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -14.76
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -130.593616341
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -120.746
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 8 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.746
operation: col(d(18_16)Mean)
inst: col
dD_actual | -120.691
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.3974316046
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -14.746
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -130.542185162
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -120.691
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 8 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.845
operation: col(d(18_16)Mean)
inst: col
dD_actual | -121.06
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.496300299
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -14.845
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -130.887241616
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -121.06
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 8 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.895
operation: col(d(18_16)Mean)
inst: col
dD_actual | -121.381
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.546233983
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -14.895
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -131.187412679
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -121.381
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 9 - Sample 9 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 9 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.757
operation: col(d(18_16)Mean)
inst: col
dD_actual | -113.459
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.4097433347
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -13.757
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -123.779452697
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -113.459
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 9 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.65
operation: col(d(18_16)Mean)
inst: col
dD_actual | -112.405
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.3028852509
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -13.65
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.793844287
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -112.405
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 9 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.627
operation: col(d(18_16)Mean)
inst: col
dD_actual | -112.295
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.2799157562
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -13.627
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.690981929
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -112.295
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 9 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.614
operation: col(d(18_16)Mean)
inst: col
dD_actual | -112.082
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.2669329984
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -13.614
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.491803
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -112.082
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 10 - Sample 10 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 10 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.732
operation: col(d(18_16)Mean)
inst: col
dD_actual | -119.584
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.3834501731
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -14.732
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -129.507015798
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -119.584
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 10 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.73
operation: col(d(18_16)Mean)
inst: col
dD_actual | -120.273
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.3814528257
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -14.73
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -130.151308202
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -120.273
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 10 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.779
operation: col(d(18_16)Mean)
inst: col
dD_actual | -120.357
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.4303878361
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -14.779
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -130.229857639
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -120.357
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 10 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.758
operation: col(d(18_16)Mean)
inst: col
dD_actual | -120.685
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.4094156888
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -14.758
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -130.536574488
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -120.685
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 11 - Sample 11 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 11 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.658
operation: col(d(18_16)Mean)
inst: col
dD_actual | -120.02
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.3095483207
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -14.658
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -129.91472478
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -120.02
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 11 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.725
operation: col(d(18_16)Mean)
inst: col
dD_actual | -120.195
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.3764594573
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -14.725
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -130.07836944
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -120.195
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 11 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.778
operation: col(d(18_16)Mean)
inst: col
dD_actual | -120.31
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.4293891624
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -14.778
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -130.185907359
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -120.31
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 11 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.671
operation: col(d(18_16)Mean)
inst: col
dD_actual | -119.867
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.3225310786
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -14.671
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -129.771652591
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -119.867
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 12 - Sample 12 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 12 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.951
operation: col(d(18_16)Mean)
inst: col
dD_actual | -17.514
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.27477019827
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 3.951
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -34.0600989405
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -17.514
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 12 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.605
operation: col(d(18_16)Mean)
inst: col
dD_actual | -8.337
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.92790278527
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 4.605
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.4785729681
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -8.337
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 12 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.844
operation: col(d(18_16)Mean)
inst: col
dD_actual | -5.756
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.16658579489
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 4.844
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -23.0650480106
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -5.756
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 12 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.979
operation: col(d(18_16)Mean)
inst: col
dD_actual | -4.701
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.30140674175
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 4.979
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -22.0785044887
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -4.701
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 13 - Sample 13 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 13 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.193
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.529
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.85444929638
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.193
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.6083946624
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -55.529
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 13 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.569
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.379
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2299506002
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.569
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.2085771833
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -59.379
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 13 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.543
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.459
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2039850845
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.543
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.2833861707
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -59.459
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 13 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.597
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.051
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2579134633
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.597
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.8369726778
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -60.051
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 14 - Sample 14 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 14 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.562
operation: col(d(18_16)Mean)
inst: col
dD_actual | -13.722
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.112387543836
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 2.562
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -30.5141529355
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -13.722
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 14 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.971
operation: col(d(18_16)Mean)
inst: col
dD_actual | -9.62
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.296069991458
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 2.971
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -26.6783221043
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -9.62
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 14 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.053
operation: col(d(18_16)Mean)
inst: col
dD_actual | -8.46
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.377961233252
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 3.053
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.5935917863
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -8.46
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 14 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.116
operation: col(d(18_16)Mean)
inst: col
dD_actual | -7.868
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.440877675119
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 3.116
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.0400052792
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -7.868
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 15 - Sample 15 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 15 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.107
operation: col(d(18_16)Mean)
inst: col
dD_actual | -2.966
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.566784068429
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 2.107
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -20.4560845734
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -2.966
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 15 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.147
operation: col(d(18_16)Mean)
inst: col
dD_actual | -2.109
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.526837121212
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 2.147
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -19.6546932954
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -2.109
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 15 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.147
operation: col(d(18_16)Mean)
inst: col
dD_actual | -1.978
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.526837121212
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 2.147
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -19.5321935785
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -1.978
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 15 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.111
operation: col(d(18_16)Mean)
inst: col
dD_actual | -1.688
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.562789373707
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 2.111
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -19.261010999
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -1.688
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 16 - Sample 16 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 16 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.122
operation: col(d(18_16)Mean)
inst: col
dD_actual | 0.706
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.55047764365
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 1.122
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.0223520497
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 0.706
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 16 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.156
operation: col(d(18_16)Mean)
inst: col
dD_actual | 1.296
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.51652273851
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 1.156
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.4706357673
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 1.296
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 16 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.056
operation: col(d(18_16)Mean)
inst: col
dD_actual | 0.911
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.61639010656
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 1.056
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.8306540193
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 0.911
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 16 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.122
operation: col(d(18_16)Mean)
inst: col
dD_actual | 1.286
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.55047764365
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 1.122
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.4799868907
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 1.286
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 17 - Sample 17 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 17 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.791
operation: col(d(18_16)Mean)
inst: col
dD_actual | -4.835
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.4609403943
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -0.791
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -22.2038095426
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -4.835
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 17 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.918
operation: col(d(18_16)Mean)
inst: col
dD_actual | -5.616
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.58777195171
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -0.918
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -22.9341322826
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -5.616
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 17 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.792
operation: col(d(18_16)Mean)
inst: col
dD_actual | -5.324
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.46193906798
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -0.792
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -22.6610794784
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -5.324
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 17 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.826
operation: col(d(18_16)Mean)
inst: col
dD_actual | -5.197
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.49589397311
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -0.826
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -22.5423202108
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -5.197
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 18 - Sample 18 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 18 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.414
operation: col(d(18_16)Mean)
inst: col
dD_actual | -13.56
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.0831140972
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -1.414
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -30.362664736
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -13.56
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 18 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.398
operation: col(d(18_16)Mean)
inst: col
dD_actual | -13.956
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.06713531832
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -1.398
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -30.7329692238
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -13.956
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 18 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.445
operation: col(d(18_16)Mean)
inst: col
dD_actual | -14.284
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.1140729813
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -1.445
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -31.0396860724
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -14.284
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 18 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.454
operation: col(d(18_16)Mean)
inst: col
dD_actual | -14.411
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.12306104442
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -1.454
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -31.1584453399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -14.411
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 19 - Sample 19 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 19 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.043
operation: col(d(18_16)Mean)
inst: col
dD_actual | 0.601
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.62804654483
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.043
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.1205388457
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 0.601
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 19 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.264
operation: col(d(18_16)Mean)
inst: col
dD_actual | 1.905
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.40733966145
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.264
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -15.9011523503
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 1.905
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 19 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.193
operation: col(d(18_16)Mean)
inst: col
dD_actual | 2.15
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.47824549276
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.193
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -15.6720498263
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 2.15
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 19 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.167
operation: col(d(18_16)Mean)
inst: col
dD_actual | 2.324
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.50421100845
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.167
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -15.5093402786
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 2.324
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 20 - Sample 20 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 20 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.396
operation: col(d(18_16)Mean)
inst: col
dD_actual | -5.383
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.06646429053
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -0.396
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -22.7162511067
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -5.383
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 20 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.43
operation: col(d(18_16)Mean)
inst: col
dD_actual | -5.521
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.10041919567
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -0.43
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -22.84529661
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -5.521
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 20 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.425
operation: col(d(18_16)Mean)
inst: col
dD_actual | -5.898
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.09542582726
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -0.425
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -23.1978339633
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -5.898
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 20 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.409
operation: col(d(18_16)Mean)
inst: col
dD_actual | -6.172
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.07944704838
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -0.409
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -23.4540547453
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -6.172
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 21 - Sample 21 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 21 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.926
operation: col(d(18_16)Mean)
inst: col
dD_actual | -1.17
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.74621768501
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.926
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.7766228053
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -1.17
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 21 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.933
operation: col(d(18_16)Mean)
inst: col
dD_actual | -0.831
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.73922696925
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.933
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.459619721
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -0.831
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 21 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.923
operation: col(d(18_16)Mean)
inst: col
dD_actual | -0.947
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.74921370605
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.923
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.5680927528
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -0.947
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 21 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.944
operation: col(d(18_16)Mean)
inst: col
dD_actual | -0.805
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.72824155876
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.944
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.4353068001
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -0.805
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 22 - Sample 22 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 22 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.047
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.913
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.70864293903
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.047
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.2270284299
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -51.913
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 22 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.333
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.301
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.99426361164
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.333
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.2654137344
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -57.301
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 22 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.422
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.885
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0831455692
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.422
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.8115193427
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -57.885
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 22 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.389
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.618
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0501893377
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.389
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.4969566902
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -58.618
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 23 - Sample 23 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 23 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.502
operation: col(d(18_16)Mean)
inst: col
dD_actual | -15.701
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.17232370066
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -0.502
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -32.3647402625
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -15.701
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 23 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.179
operation: col(d(18_16)Mean)
inst: col
dD_actual | -12.088
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.84975210188
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -0.179
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.986179367
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -12.088
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 23 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.075
operation: col(d(18_16)Mean)
inst: col
dD_actual | -11.057
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.74589003912
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -0.075
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.0220785413
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -11.057
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 23 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.078
operation: col(d(18_16)Mean)
inst: col
dD_actual | -10.671
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.74888606016
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -0.078
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.6611251768
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -10.671
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 24 - Sample 24 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 24 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.761
operation: col(d(18_16)Mean)
inst: col
dD_actual | -10.435
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.91099884228
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.761
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.4404386639
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -10.435
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 24 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.735
operation: col(d(18_16)Mean)
inst: col
dD_actual | -10.418
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.93696435797
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.735
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.424541754
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -10.418
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 24 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.77
operation: col(d(18_16)Mean)
inst: col
dD_actual | -10.697
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.90201077916
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.77
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.6854380977
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -10.697
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 24 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.785
operation: col(d(18_16)Mean)
inst: col
dD_actual | -10.893
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.88703067395
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.785
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.868720117
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -10.893
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 25 - Sample 25 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 25 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.619
operation: col(d(18_16)Mean)
inst: col
dD_actual | -8.177
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.0528105049
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.619
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.3289549932
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -8.177
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 25 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.565
operation: col(d(18_16)Mean)
inst: col
dD_actual | -8.442
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.10673888364
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.565
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.5767597641
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -8.442
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 25 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.586
operation: col(d(18_16)Mean)
inst: col
dD_actual | -7.97
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.08576673635
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.586
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.1353867382
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -7.97
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 25 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.647
operation: col(d(18_16)Mean)
inst: col
dD_actual | -8.159
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.02484764185
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.647
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.312122971
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -8.159
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 26 - Sample 26 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 26 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.774
operation: col(d(18_16)Mean)
inst: col
dD_actual | -2.4
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.89934240401
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 1.774
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -19.9268109873
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -2.4
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 26 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.683
operation: col(d(18_16)Mean)
inst: col
dD_actual | -2.303
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.990221708929
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 1.683
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -19.83610509
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -2.303
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 26 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.691
operation: col(d(18_16)Mean)
inst: col
dD_actual | -2.151
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.982232319486
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 1.691
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -19.6939680138
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -2.151
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 26 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.645
operation: col(d(18_16)Mean)
inst: col
dD_actual | -2.578
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.02817130879
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 1.645
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -20.0932609843
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -2.578
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 27 - Sample 27 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 27 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.458
operation: col(d(18_16)Mean)
inst: col
dD_actual | -12.03
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.21359696745
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.458
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.9319428511
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -12.03
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 27 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.357
operation: col(d(18_16)Mean)
inst: col
dD_actual | -13.412
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.31446300917
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.357
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -30.2242681092
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -13.412
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 27 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.411
operation: col(d(18_16)Mean)
inst: col
dD_actual | -13.602
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.26053463043
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.411
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -30.4019394544
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -13.602
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 27 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.332
operation: col(d(18_16)Mean)
inst: col
dD_actual | -13.677
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.33942985118
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.332
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -30.4720728801
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -13.677
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 28 - Sample 28 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 28 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.67098951308
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.6825413639
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 28 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.889
operation: col(d(18_16)Mean)
inst: col
dD_actual | -3.167
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.78316861119
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.889
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -20.6440421544
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -3.167
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 28 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.889
operation: col(d(18_16)Mean)
inst: col
dD_actual | -3.167
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.78316861119
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.889
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -20.6440421544
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -3.167
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 28 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.889
operation: col(d(18_16)Mean)
inst: col
dD_actual | -3.167
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.78316861119
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.889
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -20.6440421544
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -3.167
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 29 - Sample 29 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 29 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.926
operation: col(d(18_16)Mean)
inst: col
dD_actual | -5.473
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.74621768501
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.926
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -22.8004112175
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -5.473
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 29 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.897
operation: col(d(18_16)Mean)
inst: col
dD_actual | -5.966
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.77517922174
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.897
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -23.2614216027
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -5.966
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 29 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.885
operation: col(d(18_16)Mean)
inst: col
dD_actual | -5.805
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.78716330591
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.885
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -23.1108685154
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -5.805
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 29 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.924
operation: col(d(18_16)Mean)
inst: col
dD_actual | -6.072
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.74821503237
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.924
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -23.360543511
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -6.072
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 30 - Sample 30 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 30 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.647
operation: col(d(18_16)Mean)
inst: col
dD_actual | -6.037
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.0275002809999
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 2.647
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -23.327814579
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -6.037
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 30 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.732
operation: col(d(18_16)Mean)
inst: col
dD_actual | -6.208
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.0573869818362
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 2.732
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -23.4877187897
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -6.208
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 30 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.74
operation: col(d(18_16)Mean)
inst: col
dD_actual | -6.135
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.0653763712796
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 2.74
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -23.4194555886
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -6.135
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 30 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.803
operation: col(d(18_16)Mean)
inst: col
dD_actual | -6.186
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.128292813146
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 2.803
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -23.4671463181
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -6.186
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 31 - Sample 31 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 31 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.96
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.752
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.62175832884
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -6.96
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.0115876857
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -52.752
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 31 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.51
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.563
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1710288531
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.51
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.5104131682
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -57.563
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 31 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.454
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.131
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.115103127
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.454
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.0415569791
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -58.131
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 31 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.512
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.151
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1730262004
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.512
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.995371569
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -59.151
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 32 - Sample 32 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 32 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.837
operation: col(d(18_16)Mean)
inst: col
dD_actual | -13.665
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.83509964257
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.837
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -30.460851532
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -13.665
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 32 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.199
operation: col(d(18_16)Mean)
inst: col
dD_actual | -9.517
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.47357977025
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 1.199
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -26.5820055329
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -9.517
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 32 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.259
operation: col(d(18_16)Mean)
inst: col
dD_actual | -8.597
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.41365934943
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 1.259
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.7217021773
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -8.597
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 32 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.352
operation: col(d(18_16)Mean)
inst: col
dD_actual | -8.175
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.32078269715
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 1.352
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.3270847685
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -8.175
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 33 - Sample 33 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 33 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.201
operation: col(d(18_16)Mean)
inst: col
dD_actual | -22.053
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.87039660327
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -1.201
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -38.3045738657
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -22.053
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 33 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.335
operation: col(d(18_16)Mean)
inst: col
dD_actual | -23.571
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.00421887645
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -1.335
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -39.7240744025
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -23.571
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 33 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.306
operation: col(d(18_16)Mean)
inst: col
dD_actual | -23.493
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.97525733972
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -1.306
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -39.6511356397
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -23.493
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 33 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.329
operation: col(d(18_16)Mean)
inst: col
dD_actual | -23.638
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.99822683437
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -1.329
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -39.7867269295
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -23.638
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 34 - Sample 34 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 34 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.224
operation: col(d(18_16)Mean)
inst: col
dD_actual | -13.244
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.89336609792
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -1.224
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -30.0671692356
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -13.244
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 34 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.379
operation: col(d(18_16)Mean)
inst: col
dD_actual | -13.724
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.04816051839
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -1.379
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -30.5160231602
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -13.724
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 34 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.453
operation: col(d(18_16)Mean)
inst: col
dD_actual | -14.251
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.12206237074
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -1.453
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -31.008827365
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -14.251
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 34 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.354
operation: col(d(18_16)Mean)
inst: col
dD_actual | -13.587
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.02319367638
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -1.354
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -30.3879127692
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -13.587
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 35 - Sample 35 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 35 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.116
operation: col(d(18_16)Mean)
inst: col
dD_actual | 0.786
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.55646968573
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 1.116
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9475430622
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 0.786
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 35 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.442
operation: col(d(18_16)Mean)
inst: col
dD_actual | 3.445
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.23090206591
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 1.442
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -14.461079342
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 3.445
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 35 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.337
operation: col(d(18_16)Mean)
inst: col
dD_actual | 2.844
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.33576280236
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 1.337
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -15.0230818602
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 2.844
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 35 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.4
operation: col(d(18_16)Mean)
inst: col
dD_actual | 3.238
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.27284636049
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 1.4
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -14.654647597
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 3.238
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 36 - Sample 36 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 36 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.225
operation: col(d(18_16)Mean)
inst: col
dD_actual | -8.736
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.8943647716
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -1.225
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.851682793
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -8.736
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 36 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.305
operation: col(d(18_16)Mean)
inst: col
dD_actual | -9.957
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.97425866604
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -1.305
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -26.9934549639
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -9.957
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 36 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.382
operation: col(d(18_16)Mean)
inst: col
dD_actual | -10.134
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.05115653943
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -1.382
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.1589698486
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -10.134
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 36 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.405
operation: col(d(18_16)Mean)
inst: col
dD_actual | -10.339
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.07412603408
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -1.405
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.3506678789
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -10.339
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 37 - Sample 37 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 37 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.321
operation: col(d(18_16)Mean)
inst: col
dD_actual | -10.021
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.9915637645
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -0.321
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.0533021538
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -10.021
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 37 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.262
operation: col(d(18_16)Mean)
inst: col
dD_actual | -9.715
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.93264201735
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -0.262
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -26.7671577768
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -9.715
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 37 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.264
operation: col(d(18_16)Mean)
inst: col
dD_actual | -9.941
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.93463936472
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -0.264
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -26.9784931664
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -9.941
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 37 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.273
operation: col(d(18_16)Mean)
inst: col
dD_actual | -10.102
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.94362742784
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -0.273
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.1290462536
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -10.102
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 38 - Sample 38 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 38 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.426
operation: col(d(18_16)Mean)
inst: col
dD_actual | -2.621
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.24555452522
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.426
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -20.1334708151
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -2.621
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 38 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.46
operation: col(d(18_16)Mean)
inst: col
dD_actual | -2.221
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.21159962009
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.46
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -19.7594258779
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -2.221
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 38 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.425
operation: col(d(18_16)Mean)
inst: col
dD_actual | -1.777
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.2465531989
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.425
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -19.3442359975
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -1.777
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 38 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.44
operation: col(d(18_16)Mean)
inst: col
dD_actual | -1.79
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.2315730937
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.44
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -19.356392458
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -1.79
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 39 - Sample 39 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 39 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.53
operation: col(d(18_16)Mean)
inst: col
dD_actual | -1.939
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.14169246246
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.53
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -19.4957241971
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -1.939
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 39 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.533
operation: col(d(18_16)Mean)
inst: col
dD_actual | -1.665
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.13869644142
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.533
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -19.2395034151
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -1.665
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 39 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.645
operation: col(d(18_16)Mean)
inst: col
dD_actual | -1.316
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.02684498921
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.645
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.9131492074
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -1.316
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 39 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.593
operation: col(d(18_16)Mean)
inst: col
dD_actual | -1.66
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.07877602059
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.593
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -19.2348278534
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -1.66
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 40 - Sample 40 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 40 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.095
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.246
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.75657927569
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.095
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.5384208401
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -52.246
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 40 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.389
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.759
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0501893377
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.389
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.7585828444
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -56.759
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 40 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.467
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.265
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1280858848
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.467
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.1668620331
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -58.265
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 40 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.497
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.513
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1580460952
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.497
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.3987698942
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -58.513
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 41 - Sample 41 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 41 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.02
operation: col(d(18_16)Mean)
inst: col
dD_actual | -5.214
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.69096298669
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -0.02
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -22.5582171207
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -5.214
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 41 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.387
operation: col(d(18_16)Mean)
inst: col
dD_actual | 0.345
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.28450279876
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.387
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3599276055
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 0.345
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 41 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.496
operation: col(d(18_16)Mean)
inst: col
dD_actual | 1.58
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.17564736759
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.496
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.2050638618
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 1.58
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 41 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.438
operation: col(d(18_16)Mean)
inst: col
dD_actual | 1.596
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.23357044106
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.438
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.1901020643
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 1.596
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 42 - Sample 42 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 42 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.227
operation: col(d(18_16)Mean)
inst: col
dD_actual | 3.019
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.551730453646
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 3.227
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -14.8594372001
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 3.019
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 42 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.216
operation: col(d(18_16)Mean)
inst: col
dD_actual | 2.803
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.540745043162
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 3.216
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -15.0614214663
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 2.803
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 42 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.25
operation: col(d(18_16)Mean)
inst: col
dD_actual | 3.082
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.574699948296
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 3.25
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -14.8005251225
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 3.082
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 42 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.232
operation: col(d(18_16)Mean)
inst: col
dD_actual | 3.06
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.556723822048
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 3.232
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -14.8210975941
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 3.06
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 43 - Sample 43 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 43 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.908
operation: col(d(18_16)Mean)
inst: col
dD_actual | -12.754
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.57778521491
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -0.908
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -29.6089641874
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -12.754
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 43 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.931
operation: col(d(18_16)Mean)
inst: col
dD_actual | -13.562
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.60075470956
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -0.931
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -30.3645349607
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -13.562
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 43 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.0
operation: col(d(18_16)Mean)
inst: col
dD_actual | -14.201
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.66966319351
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -1.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -30.9620717479
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -14.201
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 43 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.055
operation: col(d(18_16)Mean)
inst: col
dD_actual | -14.566
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.72459024593
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -1.055
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -31.3033877531
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -14.566
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 44 - Sample 44 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 44 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.981
operation: col(d(18_16)Mean)
inst: col
dD_actual | 6.986
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.692616952163
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 1.981
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -11.1498465352
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 6.986
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 44 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.1
operation: col(d(18_16)Mean)
inst: col
dD_actual | 8.691
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.573774784192
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 2.1
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -9.55547999023
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 8.691
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 44 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.095
operation: col(d(18_16)Mean)
inst: col
dD_actual | 8.95
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.578768152594
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 2.095
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -9.31328589338
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 8.95
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 44 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.145
operation: col(d(18_16)Mean)
inst: col
dD_actual | 9.161
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.528834468573
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 2.145
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -9.11597718899
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 9.161
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 45 - Sample 45 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 45 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.892
operation: col(d(18_16)Mean)
inst: col
dD_actual | -2.363
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.78017259014
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.892
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -19.8922118306
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -2.363
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 45 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.795
operation: col(d(18_16)Mean)
inst: col
dD_actual | -3.728
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.87704393715
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.795
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -21.1686401789
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -3.728
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 45 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.806
operation: col(d(18_16)Mean)
inst: col
dD_actual | -3.685
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.86605852666
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.806
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -21.1284303481
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -3.685
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 45 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.753
operation: col(d(18_16)Mean)
inst: col
dD_actual | -4.124
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.91898823172
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.753
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -21.5389446667
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -4.124
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 46 - Sample 46 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 46 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.466
operation: col(d(18_16)Mean)
inst: col
dD_actual | -15.791
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.20560757801
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.466
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -32.4489003734
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -15.791
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 46 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.478
operation: col(d(18_16)Mean)
inst: col
dD_actual | -17.534
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.19362349384
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.478
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -34.0788011873
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -17.534
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 46 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.358
operation: col(d(18_16)Mean)
inst: col
dD_actual | -17.825
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.31346433549
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.358
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -34.3509188792
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -17.825
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 46 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.361
operation: col(d(18_16)Mean)
inst: col
dD_actual | -17.859
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.31046831445
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.361
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -34.3827126988
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -17.859
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 47 - Sample 47 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 47 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.54
operation: col(d(18_16)Mean)
inst: col
dD_actual | -8.366
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.134358364805
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 2.54
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.505691226
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -8.366
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 47 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.695
operation: col(d(18_16)Mean)
inst: col
dD_actual | -6.372
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.0204360556605
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 2.695
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -23.641077214
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -6.372
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 47 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.702
operation: col(d(18_16)Mean)
inst: col
dD_actual | -6.445
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.0274267714234
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 2.702
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -23.709340415
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -6.445
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 47 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.777
operation: col(d(18_16)Mean)
inst: col
dD_actual | -6.105
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.102327297455
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 2.777
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -23.3914022184
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -6.105
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 48 - Sample 48 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 48 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.581
operation: col(d(18_16)Mean)
inst: col
dD_actual | -0.75
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.09076010476
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.581
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.3838756212
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -0.75
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 48 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.573
operation: col(d(18_16)Mean)
inst: col
dD_actual | 0.853
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.0987494942
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.573
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.8848905352
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 0.853
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 48 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.587
operation: col(d(18_16)Mean)
inst: col
dD_actual | 0.77
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.08476806267
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.587
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9625048597
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 0.77
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 48 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.631
operation: col(d(18_16)Mean)
inst: col
dD_actual | 1.284
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.04082642074
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.631
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.4818571154
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 1.284
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 49 - Sample 49 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 49 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.047
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.158
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.70864293903
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.047
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.4561309539
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -52.158
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 49 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.414
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.07
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0751561798
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.414
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.0494027831
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -57.07
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 49 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.478
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.167
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1390712953
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.478
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.0752210235
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -58.167
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 49 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.536
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.724
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1969943688
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.536
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.5960785985
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -58.724
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 50 - Sample 50 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 50 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.771
operation: col(d(18_16)Mean)
inst: col
dD_actual | -4.059
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.90101210548
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 0.771
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -21.4781623644
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -4.059
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 50 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.125
operation: col(d(18_16)Mean)
inst: col
dD_actual | 0.738
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.54748162261
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 1.125
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.9924284547
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 0.738
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 50 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.225
operation: col(d(18_16)Mean)
inst: col
dD_actual | 1.914
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.44761425456
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 1.225
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -15.8927363392
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 1.914
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 50 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.294
operation: col(d(18_16)Mean)
inst: col
dD_actual | 2.43
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.37870577061
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: 1.294
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -15.4102183702
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: 2.43
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean

## 51 - Sample 51 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:25:05
Detail Variables
Detail 51 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.03
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.715
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.69166548647
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.03
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.0418761859
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -51.715
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 51 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.398
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.879
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0591774009
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.398
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.8707963256
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -56.879
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 51 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.508
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.282
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1690315057
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.508
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.1827589429
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -58.282
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean
Detail 51 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6071
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.2655
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.403
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1865
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.3479
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.491
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.736
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1520540531
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.6745368036
measure_array: [ 4.6071 -13.1865]
inst: fit_type1
actual_mean: -7.491
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00132808104
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.6072999467
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.9095369074
measure_array: [ 38.2655 -108.3479]
inst: fit_type1
actual_mean: -58.736
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06939022611
tarray: []
source_col: d(D_H)Mean