## Summary
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
221005_22-231_22-233_22-245_HIDS2052
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -10.2863 | -74.6601 |  | 38.5296 | 7.2919 | 0.18715 | 1.8669 |  | 7.4897
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.25325 | 3.9487
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.38261 | 5.7905
4 |  | EV | -10.3668 | -74.9209 |  |  |  | 0.096755 | 1.8627
5 |  | A00000195900 | -2.1054 | -7.2544 |  |  |  | 0.21947 | 3.7669
6 |  | A00000195853 | -2.334 | -6.6953 |  |  |  | 0.033655 | 0.13132
7 |  | A00000195600 | -2.2665 | -6.5426 |  |  |  | 0.036473 | 0.212
8 |  | A00000195608 | -2.1231 | -5.6113 |  |  |  | 0.010214 | 0.070249
9 |  | A00000195783 | -2.6639 | -8.7853 |  |  |  | 0.014592 | 0.12598
10 |  | A00000195899 | -2.4123 | -8.0722 |  |  |  | 0.017635 | 0.014526
11 |  | A00000196379 | -2.7913 | -9.0581 |  |  |  | 0.041041 | 0.16201
12 |  | A00000195459 | -2.7285 | -8.9895 |  |  |  | 0.026147 | 0.29495
13 |  | EV | -10.0291 | -69.898 |  |  |  | 0.18443 | 3.4181
14 |  | A00000195884 | -2.2964 | -8.085 |  |  |  | 0.21674 | 3.6888
15 |  | A00000195866 | -2.3712 | -7.7694 |  |  |  | 0.039803 | 0.22974
16 |  | A00000195850 | -2.6369 | -8.4138 |  |  |  | 0.018037 | 0.042899
17 | * | A00000195836 | -2.6047 | -9.902 |  |  |  | 0.055518 | 4.4505
18 |  | A00000194866 | -0.99563 | -1.6282 |  |  |  | 0.026166 | 0.46955
19 |  | A00000194883 | -0.99364 | 5.3672 |  |  |  | 0.035223 | 0.53973
20 |  | A00000194857 | 4.7766 | 28.3511 |  | 2.2228 | 2.6133 | 0.18072 | 1.3473 |  | 2.4766
21 |  | A00000194855 | 1.4189 | 14.4648 |  |  |  | 0.087632 | 0.75135
22 |  | EV | -9.8563 | -68.2378 |  |  |  | 0.31667 | 4.7788
23 |  | A00000194867 | -1.8129 | -6.9175 |  | 5.5118 | 6.112 | 0.24509 | 3.7926 |  | 5.3641
24 |  | A00000196463 | -8.9373 | -62.9825 |  | 3.4983 |  | 0.2211 | 3.24
25 |  | A00000194885 | -1.8142 | -7.1684 |  | 5.9603 | 5.5408 | 0.1892 | 3.1525 |  | 5.5876
26 |  | A00000196461 | -1.324 | -0.18994 |  | 6.74 | 6.5226 | 0.021422 | 0.22829 |  | 5.7394
27 |  | A00000320061 | -5.0131 | -30.5215 |  | 6.728 | 5.9067 | 0.11986 | 1.9564 |  | 5.7604
28 |  | A00000320079 | -5.4798 | -34.8966 |  | 7.9765 | 7.7096 | 0.046428 | 0.29866 |  | 7.1696
29 |  | A00000320088 | -5.6302 | -35.3674 |  | 7.9307 | 7.3526 | 0.022211 | 0.032847 |  | 7.4308
30 |  | A00000320123 | -5.3968 | -34.6536 |  | 7.2097 | 6.7635 | 0.061606 | 0.28691 |  | 6.9533
31 |  | EV | -10.0288 | -70.6893 |  | 7.2337 | 6.7992 | 0.15763 | 2.1591 |  | 7.1651
32 |  | A00000347319 | -5.4718 | -36.1194 |  | 6.7725 | 6.1566 | 0.090142 | 1.5784 |  | 6.6624
33 |  | A00000320113 | -5.7651 | -35.9211 |  | 6.8895 | 6.353 | 0.041161 | 0.4396 |  | 5.9488
34 |  | A00000320119 | -5.8964 | -37.1362 |  | 4.959 | 4.1306 | 0.066035 | 0.33973 |  | 4.7821
35 |  | A00000320114 | -5.5491 | -35.878 |  | 6.5457 | 6.2905 | 0.029056 | 0.19938 |  | 5.9821
36 |  | A00000320120 | -5.8827 | -37.0133 |  | 6.6543 | 5.5497 | 0.047563 | 0.22563 |  | 5.6598
37 |  | A00000320050 | -4.8161 | -31.9551 |  | 5.9015 | 6.2637 | 0.046907 | 0.31241 |  | 6.4166
38 |  | A00000320064 | -4.0489 | -25.4324 |  | 6.2663 | 6.6386 | 0.033372 | 0.38583 |  | 6.6781
39 |  | A00000320071 | -4.5359 | -28.4285 |  | 6.2852 | 5.6925 | 0.067282 | 0.6022 |  | 6.2578
40 |  | EV | -10.0428 | -70.6397 |  | 4.0142 | 3.9342 | 0.16162 | 2.2999 |  | 3.9958
41 |  | A00000320106 | -5.3527 | -35.3124 |  | 7.2158 | 6.5047 | 0.12078 | 1.9909 |  | 6.5346
42 |  | A00000320107 | -5.0298 | -33.2395 |  | 6.3873 | 7.4597 | 0.030757 | 0.19537 |  | 6.5623
43 |  | A00000320108 | -4.8882 | -32.491 |  | 6.7835 | 5.7193 | 0.11827 | 0.63352 |  | 6.2226
44 |  | A00000347310 | -3.9484 | -23.1104 |  | 6.2428 | 5.5408 | 0.033866 | 0.55391 |  | 5.9028
45 |  | A00000347323 | -3.8836 | -24.0096 |  | 5.6608 | 5.6657 | 0.033758 | 0.27063 |  | 5.2461
46 |  | A00000347301 | -5.2789 | -33.4598 |  | 4.9108 | 3.9432 | 0.059118 | 0.59704 |  | 4.5276
47 |  | A00000347307 | -5.3487 | -34.3314 |  | 6.073 | 6.3262 | 0.062516 | 0.45733 |  | 6.4176
48 |  | A00000347306 | -4.7224 | -30.0809 |  | 5.947 | 5.7104 | 0.05621 | 0.3274 |  | 5.2088
49 |  | EV | -10.0179 | -70.5428 |  | 5.204 | 4.6036 | 0.15868 | 2.2094 |  | 4.3351
50 |  | A00000347300 | -5.5491 | -36.8389 |  |  |  | 0.091569 | 1.5373
51 |  | A00000319827 | -5.8753 | -39.2429 |  |  |  | 0.038621 | 0.27872
52 |  | A00000319824 | -6.8469 | -50.6503 |  | 2.9908 | 3.8182 | 0.090877 | 0.95834 |  | 3.8006
53 |  | A00000347227 | -4.1546 | -24.1592 |  | 5.5148 | 4.0592 | 0.062942 | 1.3685 |  | 4.5421
54 |  | A00000347481 | -2.1096 | -2.2233 |  | 2.4308 | 2.4437 | 0.041737 | 0.97587 |  | 2.9561
55 |  | A00000347479 | -0.90363 | 4.7847 |  | 4.0868 | 3.6843 | 0.035584 | 0.34633 |  | 3.5273
56 |  | A00000319825 | -5.6279 | -32.0162 |  | 3.6115 | 3.1488 | 0.17529 | 2.1898 |  | 3.5356
57 |  | A00000319822 | -2.633 | -8.598 |  | 3.531 | 3.5237 | 0.066284 | 1.3579 |  | 3.9804
58 |  | EV | -9.9904 | -69.614 |  | 3.763 | 2.5598 | 0.23392 | 3.7283 |  | 2.9886
59 |  | A00000347483 | -3.3284 | -19.0673 |  | 4.4793 | 3.5237 | 0.16284 | 2.4179 |  | 3.6294
60 |  | TALL4-C-4 | -1.702 | -19.9366 |  | 807.9165 |  | 0.10246 | 0.22275 | -28.5245 | -5.8357
61 |  | EBC 02 | -8.9654 | -55.0626 |  | 8.066 |  | 0.14971 | 1.8187
62 |  | EBC 05 | -10.7993 | -73.4124 |  | 5.8102 |  | 0.047455 | 0.91572
63 |  | EBC 01 | -6.9578 | -39.9625 |  | 5.9028 | 5.7461 | 0.33983 | 2.4986 |  | 5.4573
64 |  | EBC 04 | -1.2701 | 6.129 |  | 9.2825 | 7.9238 | 0.1246 | 0.38006 |  | 7.4896
65 |  | EV | -9.9281 | -69.1113 |  | 3.5003 | 2.4973 | 0.25181 | 4.3281 |  | 3.2311

## Detail
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
1 |  | EV | -8.089 | -65.076 | -10.6776 | -78.2815 | 0.416 | 1365.224 | 2022/10/05 14:06:14
1 |  | EV | -7.945 | -63.796 | -10.534 | -77.0828 | 0.46 | 1456.668 | 2022/10/05 14:15:11
1 |  | EV | -7.774 | -62.106 | -10.3634 | -75.5001 | 0.597 | 1365.829 | 2022/10/05 14:24:08
1 |  | EV | -7.704 | -61.064 | -10.2936 | -74.5243 | 0.391 | 1317.71 | 2022/10/05 14:33:05
1 |  | EV | -7.622 | -60.397 | -10.2118 | -73.8996 | 0.389 | 1306.263 | 2022/10/05 14:42:02
1 |  | EV | -7.559 | -60.071 | -10.149 | -73.5943 | 0.456 | 1301.046 | 2022/10/05 14:50:59
1 |  | EV | -7.539 | -59.778 | -10.1291 | -73.3199 | 0.64 | 1299.419 | 2022/10/05 14:59:56
1 |  | EV | -7.616 | -59.828 | -10.2059 | -73.3668 | 0.409 | 1298.017 | 2022/10/05 15:08:53
1 |  | EV | -7.553 | -59.984 | -10.143 | -73.5129 | 0.518 | 1297.468 | 2022/10/05 15:17:49
1 |  | EV | -7.566 | -59.99 | -10.156 | -73.5185 | 0.419 | 1296.599 | 2022/10/05 15:26:46
Sample 1 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -10.2863 | -74.6601 |  | 38.5296 | 7.2919 | 0.18715 | 1.8669 |  | 7.4897
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
2 |  | PZ | 3.875 | 27.309 | 1.2543 | 8.2364 | 0.336 | 1291.907 | 2022/10/05 15:36:35
2 |  | PZ | 4.425 | 35.67 | 1.8028 | 16.0664 | 0.281 | 1291.6 | 2022/10/05 15:45:31
2 |  | PZ | 4.604 | 37.859 | 1.9814 | 18.1164 | 0.132 | 1292.005 | 2022/10/05 15:54:29
2 |  | PZ | 4.634 | 38.931 | 2.0113 | 19.1203 | 0.21 | 1294.707 | 2022/10/05 16:03:26
2 |  | PZ | 4.566 | 39.076 | 1.9435 | 19.2561 | 0.145 | 1291.157 | 2022/10/05 16:12:24
2 |  | PZ | 4.657 | 39.468 | 2.0342 | 19.6232 | 0.498 | 1291.603 | 2022/10/05 16:21:21
2 |  | PZ | 4.647 | 39.646 | 2.0242 | 19.7899 | 0.475 | 1291.201 | 2022/10/05 16:30:19
2 |  | PZ | 4.683 | 40.008 | 2.0601 | 20.1289 | 0.696 | 1291.453 | 2022/10/05 16:39:16
2 |  | PZ | 4.744 | 40.12 | 2.121 | 20.2338 | 0.525 | 1291.769 | 2022/10/05 16:48:13
2 |  | PZ | 4.69 | 40.328 | 2.0671 | 20.4286 | 0.539 | 1291.332 | 2022/10/05 16:57:10
Sample 2 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.25325 | 3.9487
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
3 |  | UT2 | -12.217 | -92.943 | -14.7945 | -104.3787 | 0.458 | 1293.919 | 2022/10/05 17:06:59
3 |  | UT2 | -13.128 | -105.802 | -15.7031 | -116.4211 | 0.419 | 1291.052 | 2022/10/05 17:15:56
3 |  | UT2 | -13.347 | -108.88 | -15.9215 | -119.3036 | 0.333 | 1289.881 | 2022/10/05 17:24:53
3 |  | UT2 | -13.31 | -109.777 | -15.8846 | -120.1436 | 0.109 | 1291.838 | 2022/10/05 17:33:51
3 |  | UT2 | -13.406 | -110.581 | -15.9803 | -120.8966 | 0.598 | 1291.057 | 2022/10/05 17:42:49
3 |  | UT2 | -13.456 | -110.82 | -16.0302 | -121.1204 | 0.558 | 1292.033 | 2022/10/05 17:51:46
3 |  | UT2 | -13.424 | -111.487 | -15.9983 | -121.745 | 0.391 | 1292.275 | 2022/10/05 18:00:42
3 |  | UT2 | -13.422 | -111.33 | -15.9963 | -121.598 | 0.466 | 1292.443 | 2022/10/05 18:09:40
3 |  | UT2 | -13.464 | -111.891 | -16.0382 | -122.1234 | 0.208 | 1292.308 | 2022/10/05 18:18:37
3 |  | UT2 | -13.479 | -112.047 | -16.0531 | -122.2695 | 0.112 | 1292.355 | 2022/10/05 18:27:34
Sample 3 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.38261 | 5.7905
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
4 |  | EV | -8.031 | -66.537 | -10.6197 | -79.6497 | 0.678 | 1293.104 | 2022/10/05 18:37:24
4 |  | EV | -7.811 | -62.33 | -10.4003 | -75.7099 | 0.035 | 1290.864 | 2022/10/05 18:46:22
4 |  | EV | -7.75 | -61.365 | -10.3395 | -74.8062 | 0.447 | 1291.705 | 2022/10/05 18:55:18
4 |  | EV | -7.768 | -60.864 | -10.3574 | -74.337 | 0.592 | 1291.385 | 2022/10/05 19:04:16
4 |  | EV | -7.769 | -60.764 | -10.3584 | -74.2433 | 0.39 | 1292.669 | 2022/10/05 19:13:12
4 |  | EV | -7.691 | -60.768 | -10.2807 | -74.2471 | 0.307 | 1291.47 | 2022/10/05 19:22:09
4 |  | EV | -7.695 | -60.403 | -10.2846 | -73.9053 | 0.09 | 1292.106 | 2022/10/05 19:31:06
4 |  | EV | -7.732 | -60.553 | -10.3215 | -74.0457 | 0.3 | 1291.855 | 2022/10/05 19:40:03
4 |  | EV | -7.788 | -60.877 | -10.3774 | -74.3492 | 0.085 | 1292.031 | 2022/10/05 19:49:01
4 |  | EV | -7.739 | -60.414 | -10.3285 | -73.9156 | 0.105 | 1291.936 | 2022/10/05 19:57:59
Sample 4 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
4 |  | EV | -10.3668 | -74.9209 |  |  |  | 0.096755 | 1.8627
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
5 |  | A00000195900 | 0.189 | 5.349 | -2.4218 | -12.329 | 0.113 | 1293.942 | 2022/10/05 20:07:49
5 |  | A00000195900 | 0.535 | 11.125 | -2.0767 | -6.9198 | 0.16 | 1293.802 | 2022/10/05 20:16:46
5 |  | A00000195900 | 0.624 | 12.932 | -1.988 | -5.2276 | 0.316 | 1295.021 | 2022/10/05 20:25:43
5 |  | A00000195900 | 0.677 | 13.665 | -1.9351 | -4.5411 | 0.11 | 1293.179 | 2022/10/05 20:34:40
Sample 5 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
5 |  | A00000195900 | -2.1054 | -7.2544 |  |  |  | 0.21947 | 3.7669
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
6 |  | A00000195853 | 0.245 | 11.243 | -2.3659 | -6.8093 | 0.259 | 1293.944 | 2022/10/05 20:44:30
6 |  | A00000195853 | 0.282 | 11.35 | -2.329 | -6.7091 | 0.313 | 1292.932 | 2022/10/05 20:53:28
6 |  | A00000195853 | 0.259 | 11.316 | -2.352 | -6.7409 | 0.281 | 1293.093 | 2022/10/05 21:02:25
6 |  | A00000195853 | 0.322 | 11.55 | -2.2892 | -6.5218 | 0.268 | 1293.137 | 2022/10/05 21:11:22
Sample 6 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
6 |  | A00000195853 | -2.334 | -6.6953 |  |  |  | 0.033655 | 0.13132
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
7 |  | A00000195600 | 0.316 | 11.441 | -2.2951 | -6.6239 | 0.373 | 1294.234 | 2022/10/05 21:21:13
7 |  | A00000195600 | 0.311 | 11.281 | -2.3001 | -6.7737 | 0.346 | 1291.731 | 2022/10/05 21:30:10
7 |  | A00000195600 | 0.382 | 11.621 | -2.2293 | -6.4553 | 0.09 | 1292.999 | 2022/10/05 21:39:08
7 |  | A00000195600 | 0.37 | 11.768 | -2.2413 | -6.3176 | 0.374 | 1293.915 | 2022/10/05 21:48:05
Sample 7 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
7 |  | A00000195600 | -2.2665 | -6.5426 |  |  |  | 0.036473 | 0.212
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
8 |  | A00000195608 | 0.489 | 12.429 | -2.1226 | -5.6986 | 0.229 | 1293.606 | 2022/10/05 21:57:53
8 |  | A00000195608 | 0.497 | 12.508 | -2.1146 | -5.6246 | 0.295 | 1293.664 | 2022/10/05 22:06:50
8 |  | A00000195608 | 0.474 | 12.568 | -2.1376 | -5.5684 | 0.341 | 1294.425 | 2022/10/05 22:15:47
8 |  | A00000195608 | 0.494 | 12.584 | -2.1176 | -5.5535 | -0.012 | 1294.045 | 2022/10/05 22:24:44
Sample 8 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
8 |  | A00000195608 | -2.1231 | -5.6113 |  |  |  | 0.010214 | 0.070249
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
9 |  | A00000195783 | -0.071 | 9.294 | -2.6811 | -8.6345 | 0.074 | 1292.49 | 2022/10/05 22:34:32
9 |  | A00000195783 | -0.059 | 9.137 | -2.6691 | -8.7816 | 0.312 | 1293.337 | 2022/10/05 22:43:28
9 |  | A00000195783 | -0.048 | 9.114 | -2.6582 | -8.8031 | 0.277 | 1293.865 | 2022/10/05 22:52:26
9 |  | A00000195783 | -0.037 | 8.987 | -2.6472 | -8.922 | 0.296 | 1293.712 | 2022/10/05 23:01:23
Sample 9 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
9 |  | A00000195783 | -2.6639 | -8.7853 |  |  |  | 0.014592 | 0.12598
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
10 |  | A00000195899 | 0.194 | 9.902 | -2.4168 | -8.0651 | 0.363 | 1292.621 | 2022/10/05 23:11:12
10 |  | A00000195899 | 0.219 | 9.91 | -2.3919 | -8.0576 | 0.095 | 1290.366 | 2022/10/05 23:20:09
10 |  | A00000195899 | 0.177 | 9.877 | -2.4338 | -8.0885 | 0.245 | 1289.449 | 2022/10/05 23:29:06
10 |  | A00000195899 | 0.204 | 9.889 | -2.4068 | -8.0773 | 0.337 | 1291.066 | 2022/10/05 23:38:04
Sample 10 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
10 |  | A00000195899 | -2.4123 | -8.0722 |  |  |  | 0.017635 | 0.014526
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
11 |  | A00000196379 | -0.152 | 8.947 | -2.7619 | -8.9595 | 0.468 | 1293.225 | 2022/10/05 23:47:53
11 |  | A00000196379 | -0.205 | 8.711 | -2.8147 | -9.1805 | -0.039 | 1291.625 | 2022/10/05 23:56:51
11 |  | A00000196379 | -0.227 | 8.696 | -2.8367 | -9.1945 | 0.557 | 1292.178 | 2022/10/06 00:05:48
11 |  | A00000196379 | -0.142 | 9.013 | -2.7519 | -8.8977 | 0.133 | 1292.511 | 2022/10/06 00:14:45
Sample 11 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
11 |  | A00000196379 | -2.7913 | -9.0581 |  |  |  | 0.041041 | 0.16201
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
12 |  | A00000195459 | -0.15 | 8.747 | -2.7599 | -9.1468 | 0.435 | 1292.604 | 2022/10/06 00:24:34
12 |  | A00000195459 | -0.13 | 8.927 | -2.7399 | -8.9782 | -0.136 | 1293.168 | 2022/10/06 00:33:32
12 |  | A00000195459 | -0.096 | 9.325 | -2.706 | -8.6055 | 0.194 | 1292.37 | 2022/10/06 00:42:29
12 |  | A00000195459 | -0.098 | 8.661 | -2.708 | -9.2273 | 0.366 | 1292.81 | 2022/10/06 00:51:26
Sample 12 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
12 |  | A00000195459 | -2.7285 | -8.9895 |  |  |  | 0.026147 | 0.29495
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
13 |  | EV | -7.179 | -51.202 | -9.77 | -65.2886 | 0.552 | 1291.498 | 2022/10/06 01:01:16
13 |  | EV | -7.438 | -56.523 | -10.0283 | -70.2717 | 0.231 | 1292.261 | 2022/10/06 01:10:13
13 |  | EV | -7.551 | -57.919 | -10.141 | -71.579 | 0.34 | 1292.869 | 2022/10/06 01:19:11
13 |  | EV | -7.587 | -58.852 | -10.1769 | -72.4528 | 0.17 | 1292.807 | 2022/10/06 01:28:08
Sample 13 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
13 |  | EV | -10.0291 | -69.898 |  |  |  | 0.18443 | 3.4181
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
14 |  | A00000195884 | -0.008 | 4.49 | -2.6183 | -13.1334 | 0.19 | 1292.796 | 2022/10/06 01:37:58
14 |  | A00000195884 | 0.386 | 10.547 | -2.2253 | -7.4611 | 0.456 | 1292.406 | 2022/10/06 01:46:56
14 |  | A00000195884 | 0.446 | 11.963 | -2.1655 | -6.135 | 0.511 | 1292.773 | 2022/10/06 01:55:53
14 |  | A00000195884 | 0.435 | 12.523 | -2.1765 | -5.6106 | 0.622 | 1292.441 | 2022/10/06 02:04:51
Sample 14 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
14 |  | A00000195884 | -2.2964 | -8.085 |  |  |  | 0.21674 | 3.6888
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
15 |  | A00000195866 | 0.19 | 9.879 | -2.4208 | -8.0867 | 0.059 | 1292.397 | 2022/10/06 02:14:38
15 |  | A00000195866 | 0.262 | 10.275 | -2.349 | -7.7158 | 0.228 | 1293.917 | 2022/10/06 02:23:35
15 |  | A00000195866 | 0.227 | 10.34 | -2.3839 | -7.655 | 0.268 | 1293.588 | 2022/10/06 02:32:32
15 |  | A00000195866 | 0.28 | 10.377 | -2.331 | -7.6203 | 0.235 | 1294.02 | 2022/10/06 02:41:29
Sample 15 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
15 |  | A00000195866 | -2.3712 | -7.7694 |  |  |  | 0.039803 | 0.22974
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
16 |  | A00000195850 | -0.008 | 9.481 | -2.6183 | -8.4594 | 0.259 | 1291.9 | 2022/10/06 02:51:16
16 |  | A00000195850 | -0.028 | 9.546 | -2.6382 | -8.3985 | 0.395 | 1290.757 | 2022/10/06 03:00:14
16 |  | A00000195850 | -0.044 | 9.562 | -2.6542 | -8.3835 | 0.243 | 1290.441 | 2022/10/06 03:09:11
Sample 16 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
16 |  | A00000195850 | -2.6369 | -8.4138 |  |  |  | 0.018037 | 0.042899
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
17 |  | A00000195836 |  |  | -2.6103 | -17.3383 |  |  | 2022/10/06 03:18:33
17 |  | A00000195836 | -0.049 | 9.78 | -2.6592 | -8.1794 | 0.476 | 1290.514 | 2022/10/06 03:29:25
17 |  | A00000195836 | 0.094 | 10.452 | -2.5165 | -7.5501 | -0.05 | 1294.095 | 2022/10/06 03:38:22
17 |  | A00000195836 | 0.015 | 9.88 | -2.5953 | -8.0857 | 0.369 | 1293.358 | 2022/10/06 03:47:19
17 |  | A00000195836 | -0.032 | 9.591 | -2.6422 | -8.3564 | 0.223 | 1292.899 | 2022/10/06 03:56:16
Sample 17 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
17 | * | A00000195836 | -2.6047 | -9.902 |  |  |  | 0.055518 | 4.4505
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
18 |  | A00000194866 | 1.586 | 16.181 | -1.0285 | -2.1849 | 0.243 | 1291.716 | 2022/10/06 04:06:04
18 |  | A00000194866 | 1.621 | 16.766 | -0.99364 | -1.6371 | 0.204 | 1291.716 | 2022/10/06 04:15:01
18 |  | A00000194866 | 1.619 | 16.827 | -0.99563 | -1.5799 | 0.189 | 1291.789 | 2022/10/06 04:23:58
18 |  | A00000194866 | 1.65 | 17.328 | -0.96471 | -1.1107 | 0.249 | 1294.361 | 2022/10/06 04:32:55
Sample 18 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
18 |  | A00000194866 | -0.99563 | -1.6282 |  |  |  | 0.026166 | 0.46955
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
19 |  | A00000194883 | 1.645 | 23.472 | -0.9697 | 4.6431 | 0.438 | 1291.86 | 2022/10/06 04:42:44
19 |  | A00000194883 | 1.644 | 24.434 | -0.9707 | 5.544 | 0.272 | 1290.353 | 2022/10/06 04:51:41
19 |  | A00000194883 | 1.625 | 24.724 | -0.98965 | 5.8156 | 0.134 | 1292.794 | 2022/10/06 05:00:38
19 |  | A00000194883 | 1.57 | 24.351 | -1.0445 | 5.4662 | 0.173 | 1293.104 | 2022/10/06 05:09:35
Sample 19 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
19 |  | A00000194883 | -0.99364 | 5.3672 |  |  |  | 0.035223 | 0.53973
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
20 |  | A00000194857 | 7.155 | 46.787 | 4.5255 | 26.4774 | 0.327 | 1293.742 | 2022/10/06 05:19:24
20 |  | A00000194857 | 7.397 | 49.196 | 4.7669 | 28.7334 | 0.191 | 1294.175 | 2022/10/06 05:28:21
20 |  | A00000194857 | 7.525 | 49.521 | 4.8945 | 29.0378 | 0.441 | 1294.317 | 2022/10/06 05:37:18
20 |  | A00000194857 | 7.55 | 49.647 | 4.9195 | 29.1558 | 0.11 | 1294.236 | 2022/10/06 05:46:15
Sample 20 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
20 |  | A00000194857 | 4.7766 | 28.3511 |  | 2.2228 | 2.6133 | 0.18072 | 1.3473 |  | 2.4766
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
21 |  | A00000194855 | 4.17 | 35.066 | 1.5485 | 15.5008 | 0.152 | 1290.999 | 2022/10/06 05:56:05
21 |  | A00000194855 | 4.015 | 33.611 | 1.3939 | 14.1382 | 0.282 | 1291.274 | 2022/10/06 06:05:03
21 |  | A00000194855 | 3.987 | 33.406 | 1.366 | 13.9462 | 0.258 | 1291.507 | 2022/10/06 06:14:00
21 |  | A00000194855 | 3.988 | 33.756 | 1.367 | 14.274 | 0.478 | 1291.33 | 2022/10/06 06:22:57
Sample 21 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
21 |  | A00000194855 | 1.4189 | 14.4648 |  |  |  | 0.087632 | 0.75135
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
22 |  | EV | -6.802 | -47.37 | -9.394 | -61.6999 | 0.385 | 1292.892 | 2022/10/06 06:32:42
22 |  | EV | -7.331 | -55.247 | -9.9216 | -69.0767 | 0.143 | 1292.497 | 2022/10/06 06:41:38
22 |  | EV | -7.429 | -56.918 | -10.0194 | -70.6416 | 0.184 | 1293.376 | 2022/10/06 06:50:35
22 |  | EV | -7.5 | -57.87 | -10.0902 | -71.5331 | 0.144 | 1292.73 | 2022/10/06 06:59:31
Sample 22 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
22 |  | EV | -9.8563 | -68.2378 |  |  |  | 0.31667 | 4.7788
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
23 |  | A00000194867 | 0.437 | 5.506 | -2.1745 | -12.182 | 0.517 | 1294.691 | 2022/10/06 07:09:19
23 |  | A00000194867 | 0.863 | 12.182 | -1.7496 | -5.9299 | 0.329 | 1298.909 | 2022/10/06 07:18:16
23 |  | A00000194867 | 0.943 | 13.342 | -1.6698 | -4.8436 | 0.381 | 1299.7 | 2022/10/06 07:27:13
23 |  | A00000194867 | 0.955 | 13.48 | -1.6578 | -4.7144 | 0.214 | 1296.326 | 2022/10/06 07:36:09
Sample 23 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
23 |  | A00000194867 | -1.8129 | -6.9175 |  | 5.5118 | 6.112 | 0.24509 | 3.7926 |  | 5.3641
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
24 |  | A00000196463 | -6.019 | -43.966 | -8.6131 | -58.5121 | 0.17 | 1297.903 | 2022/10/06 07:45:56
24 |  | A00000196463 | -6.403 | -49.623 | -8.9961 | -63.8099 | 0.29 | 1293.715 | 2022/10/06 07:54:53
24 |  | A00000196463 | -6.444 | -50.26 | -9.037 | -64.4064 | 0.417 | 1297.632 | 2022/10/06 08:03:50
24 |  | A00000196463 | -6.51 | -51.109 | -9.1028 | -65.2015 | 0.166 | 1292.322 | 2022/10/06 08:12:47
Sample 24 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
24 |  | A00000196463 | -8.9373 | -62.9825 |  | 3.4983 |  | 0.2211 | 3.24
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
25 |  | A00000194885 | 0.524 | 6.335 | -2.0877 | -11.4056 | 0.359 | 1299.306 | 2022/10/06 08:22:34
25 |  | A00000194885 | 0.821 | 11.141 | -1.7915 | -6.9048 | 0.369 | 1295.033 | 2022/10/06 08:31:31
25 |  | A00000194885 | 0.928 | 12.619 | -1.6848 | -5.5207 | 0.112 | 1297.654 | 2022/10/06 08:40:29
25 |  | A00000194885 | 0.92 | 13.343 | -1.6928 | -4.8427 | 0.254 | 1299.427 | 2022/10/06 08:49:25
Sample 25 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
25 |  | A00000194885 | -1.8142 | -7.1684 |  | 5.9603 | 5.5408 | 0.1892 | 3.1525 |  | 5.5876
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
26 |  | A00000196461 | 1.258 | 18.069 | -1.3557 | -0.4168 | 0.479 | 1300.688 | 2022/10/06 14:48:05
26 |  | A00000196461 | 1.301 | 18.164 | -1.3128 | -0.32783 | 0.219 | 1298.562 | 2022/10/06 14:57:01
26 |  | A00000196461 | 1.304 | 18.498 | -1.3098 | -0.015046 | 0.247 | 1296.754 | 2022/10/06 15:05:57
26 |  | A00000196461 | 1.296 | 18.514 | -1.3178 | -6.2296e-05 | 0.255 | 1298.535 | 2022/10/06 15:14:54
Sample 26 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
26 |  | A00000196461 | -1.324 | -0.18994 |  | 6.74 | 6.5226 | 0.021422 | 0.22829 |  | 5.7394
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
27 |  | A00000320061 | -2.231 | -11.241 | -4.8353 | -27.8654 | 0.481 | 1300.227 | 2022/10/06 15:24:42
27 |  | A00000320061 | -2.448 | -14.322 | -5.0517 | -30.7507 | 0.52 | 1299.964 | 2022/10/06 15:33:38
27 |  | A00000320061 | -2.486 | -15.273 | -5.0896 | -31.6413 | -0.039 | 1297.115 | 2022/10/06 15:42:34
27 |  | A00000320061 | -2.472 | -15.473 | -5.0757 | -31.8286 | 0.21 | 1297.185 | 2022/10/06 15:51:30
Sample 27 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
27 |  | A00000320061 | -5.0131 | -30.5215 |  | 6.728 | 5.9067 | 0.11986 | 1.9564 |  | 5.7604
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
28 |  | A00000320079 | -2.816 | -18.346 | -5.4187 | -34.5192 | 0.36 | 1301.057 | 2022/10/06 16:01:19
28 |  | A00000320079 | -2.889 | -18.969 | -5.4915 | -35.1026 | 0.41 | 1299.701 | 2022/10/06 16:10:15
28 |  | A00000320079 | -2.876 | -18.699 | -5.4786 | -34.8498 | 0.462 | 1299.553 | 2022/10/06 16:19:11
28 |  | A00000320079 | -2.928 | -18.982 | -5.5304 | -35.1148 | 0.473 | 1299.174 | 2022/10/06 16:28:07
Sample 28 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
28 |  | A00000320079 | -5.4798 | -34.8966 |  | 7.9765 | 7.7096 | 0.046428 | 0.29866 |  | 7.1696
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
29 |  | A00000320088 | -3.012 | -19.248 | -5.6142 | -35.3639 | 0.378 | 1301.297 | 2022/10/06 16:37:52
29 |  | A00000320088 | -3.044 | -19.266 | -5.6461 | -35.3808 | 0.381 | 1300.33 | 2022/10/06 16:46:48
29 |  | A00000320088 | -3.05 | -19.285 | -5.6521 | -35.3985 | 0.498 | 1299.653 | 2022/10/06 16:55:44
29 |  | A00000320088 | -3.006 | -19.208 | -5.6082 | -35.3264 | 0.441 | 1298.022 | 2022/10/06 17:04:40
Sample 29 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
29 |  | A00000320088 | -5.6302 | -35.3674 |  | 7.9307 | 7.3526 | 0.022211 | 0.032847 |  | 7.4308
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
30 |  | A00000320123 | -2.727 | -18.218 | -5.33 | -34.3993 | 0.54 | 1301.779 | 2022/10/06 17:14:25
30 |  | A00000320123 | -2.781 | -18.506 | -5.3838 | -34.669 | 0.243 | 1298.659 | 2022/10/06 17:23:21
30 |  | A00000320123 | -2.792 | -18.352 | -5.3948 | -34.5248 | 0.362 | 1298.244 | 2022/10/06 17:32:17
30 |  | A00000320123 | -2.876 | -18.882 | -5.4786 | -35.0211 | 0.542 | 1297.736 | 2022/10/06 17:41:14
Sample 30 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
30 |  | A00000320123 | -5.3968 | -34.6536 |  | 7.2097 | 6.7635 | 0.061606 | 0.28691 |  | 6.9533
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
31 |  | EV | -7.21 | -53.831 | -9.8009 | -67.7506 | 0.564 | 1301.183 | 2022/10/06 17:50:59
31 |  | EV | -7.467 | -57.29 | -10.0573 | -70.99 | 0.602 | 1298.78 | 2022/10/06 17:59:55
31 |  | EV | -7.511 | -58.194 | -10.1011 | -71.8365 | 0.487 | 1298.293 | 2022/10/06 18:08:51
31 |  | EV | -7.566 | -58.561 | -10.156 | -72.1802 | 0.261 | 1298.258 | 2022/10/06 18:17:47
Sample 31 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
31 |  | EV | -10.0288 | -70.6893 |  | 7.2337 | 6.7992 | 0.15763 | 2.1591 |  | 7.1651
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
32 |  | A00000347319 | -3.001 | -22.316 | -5.6032 | -38.2371 | 0.495 | 1300.986 | 2022/10/06 18:27:33
32 |  | A00000347319 | -2.854 | -19.929 | -5.4566 | -36.0017 | 0.089 | 1298.557 | 2022/10/06 18:36:29
32 |  | A00000347319 | -2.811 | -18.807 | -5.4137 | -34.9509 | 0.409 | 1297.362 | 2022/10/06 18:45:25
32 |  | A00000347319 | -2.811 | -19.167 | -5.4137 | -35.288 | 0.398 | 1297.764 | 2022/10/06 18:54:21
Sample 32 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
32 |  | A00000347319 | -5.4718 | -36.1194 |  | 6.7725 | 6.1566 | 0.090142 | 1.5784 |  | 6.6624
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
33 |  | A00000320113 | -3.122 | -19.661 | -5.7239 | -35.7507 | 0.221 | 1300.606 | 2022/10/06 19:04:08
33 |  | A00000320113 | -3.134 | -19.322 | -5.7359 | -35.4332 | 0.345 | 1301.326 | 2022/10/06 19:13:04
33 |  | A00000320113 | -3.203 | -20.313 | -5.8047 | -36.3613 | 0.402 | 1296.243 | 2022/10/06 19:22:00
33 |  | A00000320113 | -3.194 | -20.076 | -5.7957 | -36.1393 | 0.116 | 1296.962 | 2022/10/06 19:30:56
Sample 33 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
33 |  | A00000320113 | -5.7651 | -35.9211 |  | 6.8895 | 6.353 | 0.041161 | 0.4396 |  | 5.9488
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
34 |  | A00000320119 | -3.27 | -21.026 | -5.8715 | -37.029 | 0.185 | 1300.029 | 2022/10/06 19:40:42
34 |  | A00000320119 | -3.364 | -21.532 | -5.9653 | -37.5028 | 0.092 | 1292.544 | 2022/10/06 19:49:39
34 |  | A00000320119 | -3.215 | -20.736 | -5.8167 | -36.7574 | 0.511 | 1299.278 | 2022/10/06 19:58:35
34 |  | A00000320119 | -3.331 | -21.268 | -5.9324 | -37.2556 | 0.441 | 1295.564 | 2022/10/06 20:07:31
Sample 34 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
34 |  | A00000320119 | -5.8964 | -37.1362 |  | 4.959 | 4.1306 | 0.066035 | 0.33973 |  | 4.7821
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
35 |  | A00000320114 | -2.909 | -19.506 | -5.5115 | -35.6055 | 0.128 | 1300.674 | 2022/10/06 20:17:18
35 |  | A00000320114 | -2.94 | -19.958 | -5.5424 | -36.0288 | 0.274 | 1298.696 | 2022/10/06 20:26:14
35 |  | A00000320114 | -2.975 | -19.852 | -5.5773 | -35.9295 | 0.474 | 1297.482 | 2022/10/06 20:35:10
35 |  | A00000320114 | -2.963 | -19.872 | -5.5653 | -35.9483 | 0.391 | 1296.91 | 2022/10/06 20:44:07
Sample 35 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
35 |  | A00000320114 | -5.5491 | -35.878 |  | 6.5457 | 6.2905 | 0.029056 | 0.19938 |  | 5.9821
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
36 |  | A00000320120 | -3.212 | -20.785 | -5.8137 | -36.8033 | 0.231 | 1300.608 | 2022/10/06 20:53:49
36 |  | A00000320120 | -3.305 | -20.845 | -5.9064 | -36.8595 | 0.134 | 1298.19 | 2022/10/06 21:02:46
36 |  | A00000320120 | -3.29 | -21.203 | -5.8915 | -37.1947 | 0.163 | 1297.859 | 2022/10/06 21:11:42
36 |  | A00000320120 | -3.318 | -21.204 | -5.9194 | -37.1957 | 0.498 | 1297.539 | 2022/10/06 21:20:38
Sample 36 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
36 |  | A00000320120 | -5.8827 | -37.0133 |  | 6.6543 | 5.5497 | 0.047563 | 0.22563 |  | 5.6598
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
37 |  | A00000320050 | -2.278 | -15.919 | -4.8822 | -32.2463 | 0.468 | 1300.09 | 2022/10/06 21:30:22
37 |  | A00000320050 | -2.212 | -15.693 | -4.8164 | -32.0347 | 0.304 | 1297.071 | 2022/10/06 21:39:18
37 |  | A00000320050 | -2.178 | -15.175 | -4.7824 | -31.5496 | -0.07 | 1297.591 | 2022/10/06 21:48:14
37 |  | A00000320050 | -2.179 | -15.645 | -4.7834 | -31.9897 | 0.997 | 1296.433 | 2022/10/06 21:57:10
Sample 37 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
37 |  | A00000320050 | -4.8161 | -31.9551 |  | 5.9015 | 6.2637 | 0.046907 | 0.31241 |  | 6.4166
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
38 |  | A00000320064 | -1.47 | -9.086 | -4.0763 | -25.8473 | 0.602 | 1300.515 | 2022/10/06 22:06:54
38 |  | A00000320064 | -1.441 | -8.774 | -4.0474 | -25.5551 | 0.362 | 1297.457 | 2022/10/06 22:15:50
38 |  | A00000320064 | -1.396 | -8.171 | -4.0025 | -24.9904 | 0.348 | 1296.802 | 2022/10/06 22:24:46
38 |  | A00000320064 | -1.463 | -8.541 | -4.0694 | -25.3369 | 0.4 | 1297.87 | 2022/10/06 22:33:42
Sample 38 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
38 |  | A00000320064 | -4.0489 | -25.4324 |  | 6.2663 | 6.6386 | 0.033372 | 0.38583 |  | 6.6781
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
39 |  | A00000320071 | -1.851 | -11.004 | -4.4563 | -27.6435 | 0.391 | 1300.2 | 2022/10/06 22:43:26
39 |  | A00000320071 | -1.921 | -11.917 | -4.5261 | -28.4985 | 0.409 | 1296.495 | 2022/10/06 22:52:23
39 |  | A00000320071 | -2.015 | -12.434 | -4.6199 | -28.9826 | 0.185 | 1297.902 | 2022/10/06 23:01:18
39 |  | A00000320071 | -1.936 | -12.014 | -4.5411 | -28.5893 | 0.261 | 1298.123 | 2022/10/06 23:10:14
Sample 39 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
39 |  | A00000320071 | -4.5359 | -28.4285 |  | 6.2852 | 5.6925 | 0.067282 | 0.6022 |  | 6.2578
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
40 |  | EV | -7.223 | -53.526 | -9.8139 | -67.465 | 0.307 | 1297.846 | 2022/10/06 23:19:59
40 |  | EV | -7.522 | -57.585 | -10.1121 | -71.2662 | 0.424 | 1293.916 | 2022/10/06 23:28:55
40 |  | EV | -7.469 | -57.94 | -10.0592 | -71.5987 | 0.406 | 1296.329 | 2022/10/06 23:37:51
40 |  | EV | -7.596 | -58.613 | -10.1859 | -72.2289 | 0.452 | 1295.545 | 2022/10/06 23:46:47
Sample 40 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
40 |  | EV | -10.0428 | -70.6397 |  | 4.0142 | 3.9342 | 0.16162 | 2.2999 |  | 3.9958
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
41 |  | A00000320106 | -2.92 | -22.067 | -5.5225 | -38.0039 | 0.258 | 1300.63 | 2022/10/06 23:56:33
41 |  | A00000320106 | -2.728 | -19 | -5.331 | -35.1316 | 0.395 | 1296.725 | 2022/10/07 00:05:29
41 |  | A00000320106 | -2.716 | -17.844 | -5.319 | -34.0491 | 0.616 | 1299.498 | 2022/10/07 00:14:25
41 |  | A00000320106 | -2.635 | -17.861 | -5.2382 | -34.065 | 0.173 | 1299.589 | 2022/10/07 00:23:21
Sample 41 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
41 |  | A00000320106 | -5.3527 | -35.3124 |  | 7.2158 | 6.5047 | 0.12078 | 1.9909 |  | 6.5346
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
42 |  | A00000320107 | -2.396 | -16.701 | -4.9999 | -32.9787 | 0.487 | 1300.719 | 2022/10/07 00:33:06
42 |  | A00000320107 | -2.418 | -17.129 | -5.0218 | -33.3795 | 0.371 | 1297.846 | 2022/10/07 00:42:02
42 |  | A00000320107 | -2.421 | -16.988 | -5.0248 | -33.2474 | 0.729 | 1297.325 | 2022/10/07 00:50:58
42 |  | A00000320107 | -2.469 | -17.1 | -5.0727 | -33.3523 | 0.317 | 1297.238 | 2022/10/07 00:59:54
Sample 42 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
42 |  | A00000320107 | -5.0298 | -33.2395 |  | 6.3873 | 7.4597 | 0.030757 | 0.19537 |  | 6.5623
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
43 |  | A00000320108 | -2.39 | -16.514 | -4.9939 | -32.8035 | 0.345 | 1298.109 | 2022/10/07 01:09:38
43 |  | A00000320108 | -2.116 | -15.251 | -4.7206 | -31.6207 | 0.275 | 1301.067 | 2022/10/07 01:18:34
43 |  | A00000320108 | -2.298 | -16.317 | -4.9021 | -32.619 | 0.406 | 1297.318 | 2022/10/07 01:27:30
43 |  | A00000320108 | -2.332 | -16.639 | -4.936 | -32.9206 | 0.228 | 1298.219 | 2022/10/07 01:36:25
Sample 43 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
43 |  | A00000320108 | -4.8882 | -32.491 |  | 6.7835 | 5.7193 | 0.11827 | 0.63352 |  | 6.2226
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
44 |  | A00000347310 | -1.368 | -6.95 | -3.9746 | -23.8469 | 0.401 | 1300.107 | 2022/10/07 01:46:10
44 |  | A00000347310 | -1.292 | -5.717 | -3.8988 | -22.6922 | 0.299 | 1298.009 | 2022/10/07 01:55:06
44 |  | A00000347310 | -1.353 | -5.844 | -3.9597 | -22.8112 | 0.194 | 1297.225 | 2022/10/07 02:04:02
44 |  | A00000347310 | -1.354 | -6.143 | -3.9607 | -23.0912 | 0.561 | 1297.209 | 2022/10/07 02:12:57
Sample 44 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
44 |  | A00000347310 | -3.9484 | -23.1104 |  | 6.2428 | 5.5408 | 0.033866 | 0.55391 |  | 5.9028
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
45 |  | A00000347323 | -1.25 | -6.857 | -3.8569 | -23.7598 | 0.279 | 1300.253 | 2022/10/07 02:22:43
45 |  | A00000347323 | -1.326 | -7.499 | -3.9327 | -24.361 | 0.326 | 1294.432 | 2022/10/07 02:31:38
45 |  | A00000347323 | -1.262 | -7.039 | -3.8689 | -23.9303 | 0.599 | 1297.561 | 2022/10/07 02:40:34
45 |  | A00000347323 | -1.269 | -7.1 | -3.8759 | -23.9874 | 0.362 | 1297.976 | 2022/10/07 02:49:30
Sample 45 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
45 |  | A00000347323 | -3.8836 | -24.0096 |  | 5.6608 | 5.6657 | 0.033758 | 0.27063 |  | 5.2461
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
46 |  | A00000347301 | -2.593 | -16.354 | -5.1963 | -32.6537 | 0.163 | 1298.446 | 2022/10/07 02:59:16
46 |  | A00000347301 | -2.702 | -17.466 | -5.305 | -33.6951 | 0.362 | 1293.672 | 2022/10/07 03:08:11
46 |  | A00000347301 | -2.678 | -17.32 | -5.2811 | -33.5583 | 0.334 | 1298.739 | 2022/10/07 03:17:07
46 |  | A00000347301 | -2.73 | -17.719 | -5.333 | -33.932 | 0.485 | 1296.365 | 2022/10/07 03:26:03
Sample 46 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
46 |  | A00000347301 | -5.2789 | -33.4598 |  | 4.9108 | 3.9432 | 0.059118 | 0.59704 |  | 4.5276
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
47 |  | A00000347307 | -2.7 | -17.524 | -5.303 | -33.7494 | 0.6 | 1301.139 | 2022/10/07 03:35:49
47 |  | A00000347307 | -2.838 | -18.625 | -5.4407 | -34.7805 | 0.621 | 1295.001 | 2022/10/07 03:44:46
47 |  | A00000347307 | -2.718 | -18.192 | -5.321 | -34.375 | 0.402 | 1298.009 | 2022/10/07 03:53:41
47 |  | A00000347307 | -2.727 | -18.241 | -5.33 | -34.4209 | 0.38 | 1297.722 | 2022/10/07 04:02:37
Sample 47 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
47 |  | A00000347307 | -5.3487 | -34.3314 |  | 6.073 | 6.3262 | 0.062516 | 0.45733 |  | 6.4176
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
48 |  | A00000347306 | -2.187 | -14.029 | -4.7914 | -30.4763 | 0.445 | 1299.028 | 2022/10/07 04:12:24
48 |  | A00000347306 | -2.131 | -13.561 | -4.7356 | -30.0381 | 0.188 | 1294.794 | 2022/10/07 04:21:20
48 |  | A00000347306 | -2.053 | -13.231 | -4.6578 | -29.729 | 0.143 | 1299.415 | 2022/10/07 04:30:16
48 |  | A00000347306 | -2.1 | -13.606 | -4.7047 | -30.0802 | 0.571 | 1298.13 | 2022/10/07 04:39:12
Sample 48 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
48 |  | A00000347306 | -4.7224 | -30.0809 |  | 5.947 | 5.7104 | 0.05621 | 0.3274 |  | 5.2088
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
49 |  | EV | -7.192 | -53.577 | -9.783 | -67.5128 | 0.292 | 1298.209 | 2022/10/07 04:48:59
49 |  | EV | -7.476 | -57.22 | -10.0662 | -70.9244 | 0.247 | 1294.997 | 2022/10/07 04:57:55
49 |  | EV | -7.532 | -58.151 | -10.1221 | -71.7963 | 0.241 | 1297.648 | 2022/10/07 05:06:51
49 |  | EV | -7.51 | -58.302 | -10.1001 | -71.9377 | 0.284 | 1297.541 | 2022/10/07 05:15:47
Sample 49 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
49 |  | EV | -10.0179 | -70.5428 |  | 5.204 | 4.6036 | 0.15868 | 2.2094 |  | 4.3351
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
50 |  | A00000347300 | -3.058 | -22.99 | -5.6601 | -38.8683 | 0.475 | 1300.724 | 2022/10/07 05:25:31
50 |  | A00000347300 | -2.984 | -20.835 | -5.5863 | -36.8501 | 0.248 | 1289.995 | 2022/10/07 05:34:27
50 |  | A00000347300 | -2.859 | -19.833 | -5.4616 | -35.9117 | 0.327 | 1292.703 | 2022/10/07 05:43:23
50 |  | A00000347300 | -2.886 | -19.634 | -5.4885 | -35.7254 | 0.033 | 1292.304 | 2022/10/07 05:52:19
Sample 50 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
50 |  | A00000347300 | -5.5491 | -36.8389 |  |  |  | 0.091569 | 1.5373
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
51 |  | A00000319827 | -3.23 | -23.023 | -5.8316 | -38.8992 | 0.631 | 1292.684 | 2022/10/07 06:02:04
51 |  | A00000319827 | -3.311 | -23.466 | -5.9124 | -39.314 | 0.222 | 1293.102 | 2022/10/07 06:11:00
51 |  | A00000319827 | -3.253 | -23.377 | -5.8546 | -39.2307 | 0.389 | 1293.646 | 2022/10/07 06:19:56
51 |  | A00000319827 | -3.301 | -23.694 | -5.9024 | -39.5275 | 0.287 | 1290.847 | 2022/10/07 06:28:51
Sample 51 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
51 |  | A00000319827 | -5.8753 | -39.2429 |  |  |  | 0.038621 | 0.27872
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
52 |  | A00000319824 | -4.134 | -34.299 | -6.7332 | -49.4591 | 0.671 | 1298.581 | 2022/10/07 06:38:37
52 |  | A00000319824 | -4.284 | -36.064 | -6.8828 | -51.112 | 0.457 | 1292.253 | 2022/10/07 06:47:33
52 |  | A00000319824 | -4.226 | -35.417 | -6.8249 | -50.506 | 0.216 | 1295.506 | 2022/10/07 06:56:29
52 |  | A00000319824 | -4.348 | -36.504 | -6.9466 | -51.524 | 0.319 | 1293.202 | 2022/10/07 07:05:25
Sample 52 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
52 |  | A00000319824 | -6.8469 | -50.6503 |  | 2.9908 | 3.8182 | 0.090877 | 0.95834 |  | 3.8006
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
53 |  | A00000347227 | -1.593 | -9.214 | -4.199 | -25.9671 | 0.168 | 1300.183 | 2022/10/07 07:15:12
53 |  | A00000347227 | -1.612 | -7.29 | -4.218 | -24.1653 | 0.3 | 1294.881 | 2022/10/07 07:24:07
53 |  | A00000347227 | -1.499 | -6.402 | -4.1053 | -23.3337 | 0.267 | 1297.651 | 2022/10/07 07:33:03
53 |  | A00000347227 | -1.49 | -6.228 | -4.0963 | -23.1708 | 0.184 | 1296.923 | 2022/10/07 07:41:59
Sample 53 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
53 |  | A00000347227 | -4.1546 | -24.1592 |  | 5.5148 | 4.0592 | 0.062942 | 1.3685 |  | 4.5421
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
54 |  | A00000347481 | 0.476 | 14.768 | -2.1356 | -3.5082 | 0.55 | 1298.255 | 2022/10/07 07:51:46
54 |  | A00000347481 | 0.464 | 16.122 | -2.1475 | -2.2402 | 0.128 | 1292.528 | 2022/10/07 08:00:42
54 |  | A00000347481 | 0.557 | 16.899 | -2.0548 | -1.5125 | 0.419 | 1294.766 | 2022/10/07 08:09:38
54 |  | A00000347481 | 0.511 | 16.771 | -2.1007 | -1.6324 | 0.309 | 1291.753 | 2022/10/07 08:18:34
Sample 54 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
54 |  | A00000347481 | -2.1096 | -2.2233 |  | 2.4308 | 2.4437 | 0.041737 | 0.97587 |  | 2.9561
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
55 |  | A00000347479 | 1.708 | 23.244 | -0.90687 | 4.4295 | 0.341 | 1298.412 | 2022/10/07 08:28:21
55 |  | A00000347479 | 1.712 | 23.419 | -0.90288 | 4.5934 | 0.131 | 1293.421 | 2022/10/07 08:37:17
55 |  | A00000347479 | 1.669 | 23.87 | -0.94576 | 5.0158 | 0.114 | 1295.507 | 2022/10/07 08:46:13
55 |  | A00000347479 | 1.756 | 23.96 | -0.859 | 5.1001 | 0.352 | 1296.586 | 2022/10/07 08:55:09
Sample 55 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
55 |  | A00000347479 | -0.90363 | 4.7847 |  | 4.0868 | 3.6843 | 0.035584 | 0.34633 |  | 3.5273
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
56 |  | A00000319825 | -2.767 | -12.421 | -5.3699 | -28.9705 | 0.153 | 1297.239 | 2022/10/07 09:04:56
56 |  | A00000319825 | -3.108 | -16.323 | -5.7099 | -32.6247 | 0.422 | 1292.055 | 2022/10/07 09:13:52
56 |  | A00000319825 | -3.152 | -16.979 | -5.7538 | -33.239 | 0.375 | 1295.649 | 2022/10/07 09:22:49
56 |  | A00000319825 | -3.076 | -16.97 | -5.678 | -33.2306 | 0.395 | 1297.082 | 2022/10/07 09:31:44
Sample 56 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
56 |  | A00000319825 | -5.6279 | -32.0162 |  | 3.6115 | 3.1488 | 0.17529 | 2.1898 |  | 3.5356
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
57 |  | A00000319822 | -0.101 | 7.477 | -2.711 | -10.3361 | 0.348 | 1297.685 | 2022/10/07 09:41:30
57 |  | A00000319822 | -0.045 | 9.222 | -2.6552 | -8.702 | 0.35 | 1292.476 | 2022/10/07 09:50:26
57 |  | A00000319822 | 0 | 10.052 | -2.6103 | -7.9247 | 0.275 | 1296.2 | 2022/10/07 09:59:22
57 |  | A00000319822 | 0.055 | 10.581 | -2.5554 | -7.4293 | 0.398 | 1295.342 | 2022/10/07 10:08:18
Sample 57 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
57 |  | A00000319822 | -2.633 | -8.598 |  | 3.531 | 3.5237 | 0.066284 | 1.3579 |  | 3.9804
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
58 |  | EV | -7.052 | -50.308 | -9.6434 | -64.4514 | 0.408 | 1297.843 | 2022/10/07 10:18:04
58 |  | EV | -7.476 | -56.79 | -10.0662 | -70.5217 | 0.106 | 1293.242 | 2022/10/07 10:27:00
58 |  | EV | -7.548 | -57.936 | -10.138 | -71.5949 | 0.37 | 1295.208 | 2022/10/07 10:35:56
58 |  | EV | -7.524 | -58.249 | -10.1141 | -71.888 | 0.267 | 1296.338 | 2022/10/07 10:44:51
Sample 58 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
58 |  | EV | -9.9904 | -69.614 |  | 3.763 | 2.5598 | 0.23392 | 3.7283 |  | 2.9886
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
59 |  | A00000347483 | -0.902 | -5.226 | -3.5099 | -22.2324 | -0.107 | 1300.231 | 2022/10/07 10:54:38
59 |  | A00000347483 | -0.775 | -1.791 | -3.3832 | -19.0155 | 0.258 | 1292.902 | 2022/10/07 11:03:34
59 |  | A00000347483 | -0.689 | -0.726 | -3.2974 | -18.0182 | 0.198 | 1294.953 | 2022/10/07 11:12:30
59 |  | A00000347483 | -0.514 | 0.358 | -3.1229 | -17.003 | 0.078 | 1297.41 | 2022/10/07 11:21:26
Sample 59 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
59 |  | A00000347483 | -3.3284 | -19.0673 |  | 4.4793 | 3.5237 | 0.16284 | 2.4179 |  | 3.6294
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
60 |  | TALL4-C-4 | 0.791 | -2.923 | -1.8214 | -20.0757 | -0.866 | 2078.645 | 2022/10/07 11:31:13
60 |  | TALL4-C-4 | 0.875 | -2.977 | -1.7376 | -20.1262 | -0.82 | 2087.595 | 2022/10/07 11:40:08
60 |  | TALL4-C-4 | 0.945 | -2.709 | -1.6678 | -19.8752 | -0.686 | 2115.845 | 2022/10/07 11:49:04
60 |  | TALL4-C-4 | 1.032 | -2.489 | -1.5811 | -19.6692 | -0.856 | 2117.16 | 2022/10/07 11:58:00
Sample 60 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
60 |  | TALL4-C-4 | -1.702 | -19.9366 |  | 807.9165 |  | 0.10246 | 0.22275 | -28.5245 | -5.8357
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
61 |  | EBC 02 | -6.158 | -37.606 | -8.7518 | -52.556 | 0.262 | 1303.878 | 2022/10/07 12:07:48
61 |  | EBC 02 | -6.399 | -40.677 | -8.9921 | -55.432 | 0.474 | 1298.745 | 2022/10/07 12:16:44
61 |  | EBC 02 | -6.505 | -41.426 | -9.0978 | -56.1334 | 0.26 | 1297.097 | 2022/10/07 12:25:39
61 |  | EBC 02 | -6.427 | -41.421 | -9.02 | -56.1288 | 0.363 | 1300.123 | 2022/10/07 12:34:35
Sample 61 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
61 |  | EBC 02 | -8.9654 | -55.0626 |  | 8.066 |  | 0.14971 | 1.8187
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
62 |  | EBC 05 | -8.18 | -58.646 | -10.7683 | -72.2598 | 0.308 | 1293.456 | 2022/10/07 12:44:23
62 |  | EBC 05 | -8.228 | -60.001 | -10.8162 | -73.5288 | 0.132 | 1298.044 | 2022/10/07 12:53:19
62 |  | EBC 05 | -8.166 | -59.998 | -10.7544 | -73.526 | 0.4 | 1298.807 | 2022/10/07 13:02:15
62 |  | EBC 05 | -8.27 | -60.862 | -10.8581 | -74.3351 | 0.054 | 1300.513 | 2022/10/07 13:11:10
Sample 62 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
62 |  | EBC 05 | -10.7993 | -73.4124 |  | 5.8102 |  | 0.047455 | 0.91572
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
63 |  | EBC 01 | -4.752 | -27.864 | -7.3495 | -43.4327 | 0.409 | 1294.448 | 2022/10/07 13:20:58
63 |  | EBC 01 | -3.922 | -22.421 | -6.5218 | -38.3354 | 0.391 | 1301.469 | 2022/10/07 13:29:54
63 |  | EBC 01 | -4.385 | -23.321 | -6.9835 | -39.1782 | 0.333 | 1297.648 | 2022/10/07 13:38:50
63 |  | EBC 01 | -4.378 | -23.028 | -6.9765 | -38.9038 | 0.582 | 1297.625 | 2022/10/07 13:47:46
Sample 63 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
63 |  | EBC 01 | -6.9578 | -39.9625 |  | 5.9028 | 5.7461 | 0.33983 | 2.4986 |  | 5.4573
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
64 |  | EBC 04 | 1.3 | 25.258 | -1.3138 | 6.3156 | 0.119 | 1298.982 | 2022/10/07 13:57:32
64 |  | EBC 04 | 1.527 | 25.23 | -1.0874 | 6.2894 | 0.368 | 1309.888 | 2022/10/07 14:06:27
64 |  | EBC 04 | 1.248 | 24.489 | -1.3656 | 5.5955 | 0.275 | 1296.857 | 2022/10/07 14:15:23
64 |  | EBC 04 | 1.3 | 25.258 | -1.3138 | 6.3156 | 0.119 | 1298.982 | 2022/10/07 14:24:19
Sample 64 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
64 |  | EBC 04 | -1.2701 | 6.129 |  | 9.2825 | 7.9238 | 0.1246 | 0.38006 |  | 7.4896
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
65 |  | EV | -6.972 | -48.944 | -9.5636 | -63.174 | 0.202 | 1297.92 | 2022/10/07 14:34:06
65 |  | EV | -7.373 | -56.126 | -9.9635 | -69.8999 | 0.545 | 1294.542 | 2022/10/07 14:43:01
65 |  | EV | -7.481 | -57.735 | -10.0712 | -71.4067 | 0.466 | 1293.968 | 2022/10/07 14:51:57
65 |  | EV | -7.524 | -58.331 | -10.1141 | -71.9648 | 0.318 | 1295.15 | 2022/10/07 15:00:52
Sample 65 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
65 |  | EV | -9.9281 | -69.1113 |  | 3.5003 | 2.4973 | 0.25181 | 4.3281 |  | 3.2311

## Source
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Line | Analysis | Time Code | Port | Inj Nr | d(18_16)Mean | d(D_H)Mean | H2O_Mean | Ignore | Good | Identifier 1 | Identifier 2 | Gas Configuration | Timestamp Mean | d(18_16)_SD | d(D_H)_SD | H2O_SD | d(18_16)_Sl | d(D_H)_Sl | H2O_Sl | baseline_shift | slope_shift | residuals | baseline_curvature | interval | ch4_ppm | h16od_adjust | h16od_shift | n2_flag | DAS Temp | Tray | Sample | Job | Method | Error Code
1 | P-1343 | 2022/10/05 14:06:14 | 2-01 | 1 | -8.089 | -65.076 | 20250 | -1 | 1 | EV | standard | H2O | 1664979081 | 0.327 | 0.729 | 98.687 | 0 | 0.001 | 2.326 | 1365.224 | 0.338 | 3.246 | 124.19 | 0.868 | 0.416 | 0 | 0 | 1 | 43.062 | 1 | 1 | 1 | GJB_Rinse | 0
2 | P-1343 | 2022/10/05 14:15:11 | 2-01 | 2 | -7.945 | -63.796 | 20804 | -1 | 1 | EV | standard | H2O | 1664979617 | 0.3 | 0.732 | 93.019 | 0 | 0.002 | 2.195 | 1456.668 | -3.555 | 3.371 | 127.809 | 0.876 | 0.46 | 0 | 0 | 1 | 42.562 | 1 | 1 | 1 | GJB_Rinse | 0
3 | P-1343 | 2022/10/05 14:24:08 | 2-01 | 3 | -7.774 | -62.106 | 19230 | -1 | 1 | EV | standard | H2O | 1664980154 | 0.343 | 0.713 | 86.817 | 0.001 | 0.002 | 2.053 | 1365.829 | -0.733 | 3.086 | 119.681 | 0.87 | 0.597 | 0 | 0 | 1 | 42.25 | 1 | 1 | 1 | GJB_Rinse | 0
4 | P-1343 | 2022/10/05 14:33:05 | 2-01 | 4 | -7.704 | -61.064 | 18235 | 0 | 1 | EV | standard | H2O | 1664980692 | 0.33 | 0.76 | 85.647 | 0 | 0.002 | 2.033 | 1317.71 | 1.433 | 2.917 | 111.725 | 0.867 | 0.391 | 0 | 0 | 1 | 42.125 | 1 | 1 | 1 | GJB_Rinse | 0
5 | P-1343 | 2022/10/05 14:42:02 | 2-01 | 5 | -7.622 | -60.397 | 18212 | 0 | 1 | EV | standard | H2O | 1664981229 | 0.343 | 0.784 | 81.403 | 0.001 | 0 | 1.941 | 1306.263 | 2.074 | 2.87 | 111.387 | 0.862 | 0.389 | 0 | 0 | 1 | 42.125 | 1 | 1 | 1 | GJB_Rinse | 0
6 | P-1343 | 2022/10/05 14:50:59 | 2-01 | 6 | -7.559 | -60.071 | 18093 | 0 | 1 | EV | standard | H2O | 1664981765 | 0.299 | 0.698 | 83.02 | 0.001 | 0.003 | 1.977 | 1301.046 | 1.315 | 2.925 | 112.543 | 0.864 | 0.456 | 0 | 0 | 1 | 42.125 | 1 | 1 | 1 | GJB_Rinse | 0
7 | P-1343 | 2022/10/05 14:59:56 | 2-01 | 7 | -7.539 | -59.778 | 18179 | 0 | 1 | EV | standard | H2O | 1664982302 | 0.319 | 0.747 | 80.226 | 0.001 | 0.001 | 1.903 | 1299.419 | 1.902 | 2.943 | 113.046 | 0.867 | 0.64 | 0 | 0 | 1 | 42.125 | 1 | 1 | 1 | GJB_Rinse | 0
8 | P-1343 | 2022/10/05 15:08:53 | 2-01 | 8 | -7.616 | -59.828 | 18187 | 0 | 1 | EV | standard | H2O | 1664982839 | 0.309 | 0.735 | 82.677 | 0.001 | 0.004 | 1.961 | 1298.017 | 1.689 | 2.907 | 112.084 | 0.863 | 0.409 | 0 | 0 | 1 | 42.188 | 1 | 1 | 1 | GJB_Rinse | 0
9 | P-1343 | 2022/10/05 15:17:49 | 2-01 | 9 | -7.553 | -59.984 | 18166 | 0 | 1 | EV | standard | H2O | 1664983376 | 0.284 | 0.787 | 81.78 | 0.001 | 0.004 | 1.949 | 1297.468 | 1.915 | 2.908 | 112.127 | 0.864 | 0.518 | 0 | 0 | 1 | 42.25 | 1 | 1 | 1 | GJB_Rinse | 0
10 | P-1343 | 2022/10/05 15:26:46 | 2-01 | 10 | -7.566 | -59.99 | 18101 | 0 | 1 | EV | standard | H2O | 1664983913 | 0.294 | 0.768 | 77.084 | 0.001 | 0.003 | 1.832 | 1296.599 | 1.715 | 2.88 | 111.857 | 0.866 | 0.419 | 0 | 0 | 1 | 42.312 | 1 | 1 | 1 | GJB_Rinse | 0
11 | P-1344 | 2022/10/05 15:36:35 | 2-02 | 1 | 3.875 | 27.309 | 17508 | -1 | 1 | PZ | standard | H2O | 1664984501 | 0.298 | 0.708 | 79.066 | 0 | -0.004 | 1.88 | 1291.907 | 1.273 | 2.802 | 107.978 | 0.865 | 0.336 | 0 | 0 | 1 | 42.375 | 1 | 2 | 1 | GJB_Rinse | 0
12 | P-1344 | 2022/10/05 15:45:31 | 2-02 | 2 | 4.425 | 35.67 | 17443 | -1 | 1 | PZ | standard | H2O | 1664985038 | 0.299 | 0.767 | 74.072 | -0.001 | -0.006 | 1.775 | 1291.6 | 1.891 | 2.803 | 106.59 | 0.864 | 0.281 | 0 | 0 | 1 | 42.312 | 1 | 2 | 1 | GJB_Rinse | 0
13 | P-1344 | 2022/10/05 15:54:29 | 2-02 | 3 | 4.604 | 37.859 | 17476 | -1 | 1 | PZ | standard | H2O | 1664985575 | 0.311 | 0.693 | 75.9 | 0.001 | -0.004 | 1.817 | 1292.005 | 1.49 | 2.772 | 107.235 | 0.864 | 0.132 | 0 | 0 | 1 | 42.312 | 1 | 2 | 1 | GJB_Rinse | 0
14 | P-1344 | 2022/10/05 16:03:26 | 2-02 | 4 | 4.634 | 38.931 | 17839 | 0 | 1 | PZ | standard | H2O | 1664986113 | 0.299 | 0.724 | 78.249 | 0.001 | -0.004 | 1.867 | 1294.707 | 1.248 | 2.829 | 109.297 | 0.867 | 0.21 | 0 | 0 | 1 | 42.312 | 1 | 2 | 1 | GJB_Rinse | 0
15 | P-1344 | 2022/10/05 16:12:24 | 2-02 | 5 | 4.566 | 39.076 | 17386 | 0 | 1 | PZ | standard | H2O | 1664986650 | 0.265 | 0.757 | 73.843 | 0 | -0.004 | 1.759 | 1291.157 | 1.174 | 2.813 | 107.073 | 0.868 | 0.145 | 0 | 0 | 1 | 42.312 | 1 | 2 | 1 | GJB_Rinse | 0
16 | P-1344 | 2022/10/05 16:21:21 | 2-02 | 6 | 4.657 | 39.468 | 17486 | 0 | 1 | PZ | standard | H2O | 1664987187 | 0.293 | 0.735 | 78.97 | 0.001 | -0.002 | 1.891 | 1291.603 | 1.434 | 2.824 | 108.432 | 0.864 | 0.498 | 0 | 0 | 1 | 42.312 | 1 | 2 | 1 | GJB_Rinse | 0
17 | P-1344 | 2022/10/05 16:30:19 | 2-02 | 7 | 4.647 | 39.646 | 17426 | 0 | 1 | PZ | standard | H2O | 1664987725 | 0.323 | 0.748 | 77.162 | 0 | -0.001 | 1.843 | 1291.201 | 0.992 | 2.844 | 108.953 | 0.866 | 0.475 | 0 | 0 | 1 | 42.25 | 1 | 2 | 1 | GJB_Rinse | 0
18 | P-1344 | 2022/10/05 16:39:16 | 2-02 | 8 | 4.683 | 40.008 | 17493 | 0 | 1 | PZ | standard | H2O | 1664988263 | 0.311 | 0.758 | 77.147 | 0.001 | -0.003 | 1.843 | 1291.453 | 1.028 | 2.851 | 109.662 | 0.866 | 0.696 | 0 | 0 | 1 | 42.25 | 1 | 2 | 1 | GJB_Rinse | 0
19 | P-1344 | 2022/10/05 16:48:13 | 2-02 | 9 | 4.744 | 40.12 | 17537 | 0 | 1 | PZ | standard | H2O | 1664988800 | 0.336 | 0.769 | 76.062 | 0.001 | -0.003 | 1.812 | 1291.769 | 1.258 | 2.864 | 109.929 | 0.863 | 0.525 | 0 | 0 | 1 | 42.188 | 1 | 2 | 1 | GJB_Rinse | 0
20 | P-1344 | 2022/10/05 16:57:10 | 2-02 | 10 | 4.69 | 40.328 | 17470 | 0 | 1 | PZ | standard | H2O | 1664989336 | 0.293 | 0.736 | 76.084 | 0.001 | -0.002 | 1.822 | 1291.332 | 1.095 | 2.879 | 109.435 | 0.863 | 0.539 | 0 | 0 | 1 | 42.188 | 1 | 2 | 1 | GJB_Rinse | 0
21 | P-1345 | 2022/10/05 17:06:59 | 2-03 | 1 | -12.217 | -92.943 | 17881 | -1 | 1 | UT2 | standard | H2O | 1664989925 | 0.324 | 0.922 | 74.947 | 0.003 | 0.011 | 1.783 | 1293.919 | 2.628 | 2.811 | 109.522 | 0.863 | 0.458 | 0 | 0 | 1 | 42.25 | 1 | 3 | 1 | GJB_Rinse | 0
22 | P-1345 | 2022/10/05 17:15:56 | 2-03 | 2 | -13.128 | -105.802 | 17456 | -1 | 1 | UT2 | standard | H2O | 1664990463 | 0.311 | 0.913 | 73.031 | 0.002 | 0.013 | 1.743 | 1291.052 | 2.44 | 2.733 | 106.723 | 0.862 | 0.419 | 0 | 0 | 1 | 42.25 | 1 | 3 | 1 | GJB_Rinse | 0
23 | P-1345 | 2022/10/05 17:24:53 | 2-03 | 3 | -13.347 | -108.88 | 17296 | -1 | 1 | UT2 | standard | H2O | 1664991000 | 0.269 | 0.867 | 76.939 | 0.001 | 0.011 | 1.838 | 1289.881 | 2.299 | 2.719 | 105.858 | 0.861 | 0.333 | 0 | 0 | 1 | 42.25 | 1 | 3 | 1 | GJB_Rinse | 0
24 | P-1345 | 2022/10/05 17:33:51 | 2-03 | 4 | -13.31 | -109.777 | 17520 | 0 | 1 | UT2 | standard | H2O | 1664991538 | 0.303 | 0.758 | 73.548 | 0.003 | 0.008 | 1.754 | 1291.838 | 2.335 | 2.73 | 105.837 | 0.863 | 0.109 | 0 | 0 | 1 | 42.188 | 1 | 3 | 1 | GJB_Rinse | 0
25 | P-1345 | 2022/10/05 17:42:49 | 2-03 | 5 | -13.406 | -110.581 | 17485 | 0 | 1 | UT2 | standard | H2O | 1664992075 | 0.292 | 0.815 | 75.874 | 0.001 | 0.009 | 1.812 | 1291.057 | 2.092 | 2.81 | 108.517 | 0.862 | 0.598 | 0 | 0 | 1 | 42.125 | 1 | 3 | 1 | GJB_Rinse | 0
26 | P-1345 | 2022/10/05 17:51:46 | 2-03 | 6 | -13.456 | -110.82 | 17592 | 0 | 1 | UT2 | standard | H2O | 1664992612 | 0.308 | 0.785 | 76.437 | 0.001 | 0.006 | 1.825 | 1292.033 | 1.81 | 2.819 | 109.136 | 0.866 | 0.558 | 0 | 0 | 1 | 42.188 | 1 | 3 | 1 | GJB_Rinse | 0
27 | P-1345 | 2022/10/05 18:00:42 | 2-03 | 7 | -13.424 | -111.487 | 17637 | 0 | 1 | UT2 | standard | H2O | 1664993149 | 0.309 | 0.825 | 76.107 | 0.001 | 0.004 | 1.813 | 1292.275 | 2.16 | 2.801 | 108.698 | 0.869 | 0.391 | 0 | 0 | 1 | 42.188 | 1 | 3 | 1 | GJB_Rinse | 0
28 | P-1345 | 2022/10/05 18:09:40 | 2-03 | 8 | -13.422 | -111.33 | 17666 | 0 | 1 | UT2 | standard | H2O | 1664993686 | 0.309 | 0.72 | 76.593 | 0.002 | 0.007 | 1.817 | 1292.443 | 2.018 | 2.808 | 109.188 | 0.861 | 0.466 | 0 | 0 | 1 | 42.25 | 1 | 3 | 1 | GJB_Rinse | 0
29 | P-1345 | 2022/10/05 18:18:37 | 2-03 | 9 | -13.464 | -111.891 | 17626 | 0 | 1 | UT2 | standard | H2O | 1664994223 | 0.295 | 0.765 | 76.46 | 0.001 | 0.006 | 1.822 | 1292.308 | 1.966 | 2.753 | 108.154 | 0.864 | 0.208 | 0 | 0 | 1 | 42.375 | 1 | 3 | 1 | GJB_Rinse | 0
30 | P-1345 | 2022/10/05 18:27:34 | 2-03 | 10 | -13.479 | -112.047 | 17620 | 0 | 1 | UT2 | standard | H2O | 1664994761 | 0.309 | 0.783 | 78.074 | 0.001 | 0.006 | 1.865 | 1292.355 | 2.484 | 2.756 | 106.886 | 0.866 | 0.112 | 0 | 0 | 1 | 42.375 | 1 | 3 | 1 | GJB_Rinse | 0
31 | P-1346 | 2022/10/05 18:37:24 | 2-04 | 1 | -8.031 | -66.537 | 17779 | -1 | 1 | EV | standard | H2O | 1664995351 | 0.319 | 0.8 | 77.509 | 0.001 | 0 | 1.845 | 1293.104 | 2.128 | 2.886 | 110.74 | 0.863 | 0.678 | 0 | 0 | 1 | 42.562 | 1 | 4 | 1 | GJB_Rinse | 0
32 | P-1346 | 2022/10/05 18:46:22 | 2-04 | 2 | -7.811 | -62.33 | 17416 | -1 | 1 | EV | standard | H2O | 1664995888 | 0.327 | 0.811 | 78.004 | 0.001 | 0.003 | 1.867 | 1290.864 | 1.994 | 2.71 | 105.678 | 0.869 | 0.035 | 0 | 0 | 1 | 42.312 | 1 | 4 | 1 | GJB_Rinse | 0
33 | P-1346 | 2022/10/05 18:55:18 | 2-04 | 3 | -7.75 | -61.365 | 17554 | -1 | 1 | EV | standard | H2O | 1664996425 | 0.312 | 0.72 | 73.171 | 0.002 | 0 | 1.748 | 1291.705 | 1.966 | 2.818 | 107.905 | 0.865 | 0.447 | 0 | 0 | 1 | 42.25 | 1 | 4 | 1 | GJB_Rinse | 0
34 | P-1346 | 2022/10/05 19:04:16 | 2-04 | 4 | -7.768 | -60.864 | 17529 | 0 | 1 | EV | standard | H2O | 1664996962 | 0.305 | 0.686 | 76.561 | 0.002 | 0.001 | 1.82 | 1291.385 | 1.742 | 2.815 | 108.896 | 0.865 | 0.592 | 0 | 0 | 1 | 42.25 | 1 | 4 | 1 | GJB_Rinse | 0
35 | P-1346 | 2022/10/05 19:13:12 | 2-04 | 5 | -7.769 | -60.764 | 17696 | 0 | 1 | EV | standard | H2O | 1664997499 | 0.322 | 0.703 | 75.882 | 0.001 | 0.004 | 1.821 | 1292.669 | 1.619 | 2.828 | 109.69 | 0.862 | 0.39 | 0 | 0 | 1 | 42.25 | 1 | 4 | 1 | GJB_Rinse | 0
36 | P-1346 | 2022/10/05 19:22:09 | 2-04 | 6 | -7.691 | -60.768 | 17540 | 0 | 1 | EV | standard | H2O | 1664998036 | 0.321 | 0.731 | 77.101 | 0.002 | 0.003 | 1.841 | 1291.47 | 2.305 | 2.764 | 107.469 | 0.861 | 0.307 | 0 | 0 | 1 | 42.125 | 1 | 4 | 1 | GJB_Rinse | 0
37 | P-1346 | 2022/10/05 19:31:06 | 2-04 | 7 | -7.695 | -60.403 | 17578 | 0 | 1 | EV | standard | H2O | 1664998573 | 0.285 | 0.751 | 76.618 | -0.001 | 0.003 | 1.837 | 1292.106 | 1.652 | 2.741 | 107.541 | 0.864 | 0.09 | 0 | 0 | 1 | 42.125 | 1 | 4 | 1 | GJB_Rinse | 0
38 | P-1346 | 2022/10/05 19:40:03 | 2-04 | 8 | -7.732 | -60.553 | 17563 | 0 | 1 | EV | standard | H2O | 1664999110 | 0.305 | 0.81 | 76.325 | 0.001 | 0.001 | 1.808 | 1291.855 | 1.514 | 2.769 | 108.451 | 0.864 | 0.3 | 0 | 0 | 1 | 42.125 | 1 | 4 | 1 | GJB_Rinse | 0
39 | P-1346 | 2022/10/05 19:49:01 | 2-04 | 9 | -7.788 | -60.877 | 17565 | 0 | 1 | EV | standard | H2O | 1664999648 | 0.324 | 0.749 | 74.241 | 0 | 0.001 | 1.771 | 1292.031 | 1.416 | 2.811 | 107.942 | 0.867 | 0.085 | 0 | 0 | 1 | 42.188 | 1 | 4 | 1 | GJB_Rinse | 0
40 | P-1346 | 2022/10/05 19:57:59 | 2-04 | 10 | -7.739 | -60.414 | 17576 | 0 | 1 | EV | standard | H2O | 1665000185 | 0.295 | 0.72 | 73.723 | 0.001 | 0.003 | 1.767 | 1291.936 | 1.574 | 2.815 | 107.714 | 0.868 | 0.105 | 0 | 0 | 1 | 42.25 | 1 | 4 | 1 | GJB_Rinse | 0
41 | P-1347 | 2022/10/05 20:07:49 | 2-05 | 1 | 0.189 | 5.349 | 17788 | -1 | 1 | A00000195900 | 22-245 | H2O | 1665000775 | 0.331 | 0.768 | 78.936 | 0 | -0.004 | 1.876 | 1293.942 | 1.161 | 2.833 | 110.518 | 0.865 | 0.113 | 0 | 0 | 1 | 42.312 | 1 | 5 | 1 | GJB_Rinse | 0
42 | P-1347 | 2022/10/05 20:16:46 | 2-05 | 2 | 0.535 | 11.125 | 17739 | -1 | 1 | A00000195900 | 22-245 | H2O | 1665001313 | 0.302 | 0.78 | 75.403 | 0.001 | -0.006 | 1.791 | 1293.802 | 0.82 | 2.855 | 110.354 | 0.865 | 0.16 | 0 | 0 | 1 | 42.312 | 1 | 5 | 1 | GJB_Rinse | 0
43 | P-1347 | 2022/10/05 20:25:43 | 2-05 | 3 | 0.624 | 12.932 | 17950 | -1 | 1 | A00000195900 | 22-245 | H2O | 1665001850 | 0.297 | 0.672 | 78.372 | 0 | -0.003 | 1.867 | 1295.021 | 1.116 | 2.903 | 112.051 | 0.864 | 0.316 | 0 | 0 | 1 | 42.25 | 1 | 5 | 1 | GJB_Rinse | 0
44 | P-1347 | 2022/10/05 20:34:40 | 2-05 | 4 | 0.677 | 13.665 | 17660 | 0 | 1 | A00000195900 | 22-245 | H2O | 1665002386 | 0.303 | 0.779 | 75.74 | 0.001 | -0.005 | 1.795 | 1293.179 | 1.346 | 2.853 | 108.82 | 0.867 | 0.11 | 0 | 0 | 1 | 42.375 | 1 | 5 | 1 | GJB_Rinse | 0
45 | P-1348 | 2022/10/05 20:44:30 | 2-06 | 1 | 0.245 | 11.243 | 17825 | -1 | 1 | A00000195853 | 22-245 | H2O | 1665002977 | 0.306 | 0.709 | 75.394 | 0 | -0.004 | 1.803 | 1293.944 | 1.627 | 2.86 | 109.367 | 0.865 | 0.259 | 0 | 0 | 1 | 42.312 | 1 | 6 | 1 | GJB_Rinse | 0
46 | P-1348 | 2022/10/05 20:53:28 | 2-06 | 2 | 0.282 | 11.35 | 17702 | -1 | 1 | A00000195853 | 22-245 | H2O | 1665003514 | 0.357 | 0.691 | 75.102 | 0.002 | -0.002 | 1.786 | 1292.932 | 1.861 | 2.792 | 108.62 | 0.865 | 0.313 | 0 | 0 | 1 | 42.25 | 1 | 6 | 1 | GJB_Rinse | 0
47 | P-1348 | 2022/10/05 21:02:25 | 2-06 | 3 | 0.259 | 11.316 | 17698 | -1 | 1 | A00000195853 | 22-245 | H2O | 1665004052 | 0.305 | 0.724 | 77.762 | 0.001 | 0.002 | 1.84 | 1293.093 | 1.631 | 2.816 | 108.498 | 0.864 | 0.281 | 0 | 0 | 1 | 42.25 | 1 | 6 | 1 | GJB_Rinse | 0
48 | P-1348 | 2022/10/05 21:11:22 | 2-06 | 4 | 0.322 | 11.55 | 17703 | 0 | 1 | A00000195853 | 22-245 | H2O | 1665004589 | 0.321 | 0.738 | 74.974 | 0 | -0.001 | 1.776 | 1293.137 | 1.154 | 2.832 | 109.419 | 0.863 | 0.268 | 0 | 0 | 1 | 42.25 | 1 | 6 | 1 | GJB_Rinse | 0
49 | P-1349 | 2022/10/05 21:21:13 | 2-07 | 1 | 0.316 | 11.441 | 17889 | -1 | 1 | A00000195600 | 22-245 | H2O | 1665005180 | 0.327 | 0.745 | 75.761 | 0.001 | 0 | 1.808 | 1294.234 | 1.12 | 2.882 | 111.257 | 0.868 | 0.373 | 0 | 0 | 1 | 42.312 | 1 | 7 | 1 | GJB_Rinse | 0
50 | P-1349 | 2022/10/05 21:30:10 | 2-07 | 2 | 0.311 | 11.281 | 17531 | -1 | 1 | A00000195600 | 22-245 | H2O | 1665005717 | 0.308 | 0.805 | 73.632 | 0.001 | 0.003 | 1.765 | 1291.731 | 1.408 | 2.771 | 108.009 | 0.868 | 0.346 | 0 | 0 | 1 | 42.312 | 1 | 7 | 1 | GJB_Rinse | 0
51 | P-1349 | 2022/10/05 21:39:08 | 2-07 | 3 | 0.382 | 11.621 | 17702 | -1 | 1 | A00000195600 | 22-245 | H2O | 1665006254 | 0.317 | 0.701 | 75.296 | 0.001 | -0.002 | 1.792 | 1292.999 | 1.559 | 2.811 | 108.172 | 0.875 | 0.09 | 0 | 0 | 1 | 42.312 | 1 | 7 | 1 | GJB_Rinse | 0
52 | P-1349 | 2022/10/05 21:48:05 | 2-07 | 4 | 0.37 | 11.768 | 17876 | 0 | 1 | A00000195600 | 22-245 | H2O | 1665006792 | 0.33 | 0.774 | 74.284 | 0 | -0.001 | 1.765 | 1293.915 | 1.615 | 2.847 | 110.279 | 0.875 | 0.374 | 0 | 0 | 1 | 42.25 | 1 | 7 | 1 | GJB_Rinse | 0
53 | P-1350 | 2022/10/05 21:57:53 | 2-08 | 1 | 0.489 | 12.429 | 17720 | -1 | 1 | A00000195608 | 22-245 | H2O | 1665007380 | 0.309 | 0.757 | 76.64 | 0.001 | 0 | 1.822 | 1293.606 | 0.871 | 2.832 | 110.03 | 0.871 | 0.229 | 0 | 0 | 1 | 42.25 | 1 | 8 | 1 | GJB_Rinse | 0
54 | P-1350 | 2022/10/05 22:06:50 | 2-08 | 2 | 0.497 | 12.508 | 17759 | -1 | 1 | A00000195608 | 22-245 | H2O | 1665007917 | 0.33 | 0.744 | 76.998 | 0 | -0.001 | 1.829 | 1293.664 | 1.029 | 2.887 | 110.654 | 0.877 | 0.295 | 0 | 0 | 1 | 42.25 | 1 | 8 | 1 | GJB_Rinse | 0
55 | P-1350 | 2022/10/05 22:15:47 | 2-08 | 3 | 0.474 | 12.568 | 17919 | -1 | 1 | A00000195608 | 22-245 | H2O | 1665008453 | 0.314 | 0.691 | 73.932 | 0.001 | 0.001 | 1.763 | 1294.425 | 1.692 | 2.914 | 111.268 | 0.872 | 0.341 | 0 | 0 | 1 | 42.375 | 1 | 8 | 1 | GJB_Rinse | 0
56 | P-1350 | 2022/10/05 22:24:44 | 2-08 | 4 | 0.494 | 12.584 | 17783 | 0 | 1 | A00000195608 | 22-245 | H2O | 1665008991 | 0.332 | 0.771 | 74.682 | 0.002 | 0.001 | 1.781 | 1294.045 | 1.373 | 2.783 | 108.778 | 0.873 | -0.012 | 0 | 0 | 1 | 42.375 | 1 | 8 | 1 | GJB_Rinse | 0
57 | P-1351 | 2022/10/05 22:34:32 | 2-09 | 1 | -0.071 | 9.294 | 17596 | -1 | 1 | A00000195783 | 22-245 | H2O | 1665009578 | 0.333 | 0.776 | 76.364 | 0.001 | 0.001 | 1.806 | 1292.49 | 1.112 | 2.771 | 108.355 | 0.874 | 0.074 | 0 | 0 | 1 | 42.438 | 1 | 9 | 1 | GJB_Rinse | 0
58 | P-1351 | 2022/10/05 22:43:28 | 2-09 | 2 | -0.059 | 9.137 | 17692 | -1 | 1 | A00000195783 | 22-245 | H2O | 1665010115 | 0.324 | 0.748 | 73.646 | 0.001 | 0.001 | 1.752 | 1293.337 | 0.235 | 2.838 | 111.149 | 0.875 | 0.312 | 0 | 0 | 1 | 42.375 | 1 | 9 | 1 | GJB_Rinse | 0
59 | P-1351 | 2022/10/05 22:52:26 | 2-09 | 3 | -0.048 | 9.114 | 17811 | -1 | 1 | A00000195783 | 22-245 | H2O | 1665010653 | 0.31 | 0.707 | 72.168 | 0.001 | -0.002 | 1.718 | 1293.865 | 0.944 | 2.847 | 111.194 | 0.868 | 0.277 | 0 | 0 | 1 | 42.438 | 1 | 9 | 1 | GJB_Rinse | 0
60 | P-1351 | 2022/10/05 23:01:23 | 2-09 | 4 | -0.037 | 8.987 | 17786 | 0 | 1 | A00000195783 | 22-245 | H2O | 1665011190 | 0.277 | 0.724 | 74.26 | 0 | 0.003 | 1.772 | 1293.712 | 1.23 | 2.862 | 110.438 | 0.872 | 0.296 | 0 | 0 | 1 | 42.5 | 1 | 9 | 1 | GJB_Rinse | 0
61 | P-1352 | 2022/10/05 23:11:12 | 2-10 | 1 | 0.194 | 9.902 | 17661 | -1 | 1 | A00000195899 | 22-245 | H2O | 1665011778 | 0.31 | 0.781 | 75.167 | 0 | 0.001 | 1.79 | 1292.621 | 0.977 | 2.811 | 109.897 | 0.869 | 0.363 | 0 | 0 | 1 | 42.438 | 1 | 10 | 1 | GJB_Rinse | 0
62 | P-1352 | 2022/10/05 23:20:09 | 2-10 | 2 | 0.219 | 9.91 | 17346 | -1 | 1 | A00000195899 | 22-245 | H2O | 1665012316 | 0.327 | 0.773 | 75.732 | 0 | 0.003 | 1.822 | 1290.366 | 1.217 | 2.747 | 107.024 | 0.87 | 0.095 | 0 | 0 | 1 | 42.438 | 1 | 10 | 1 | GJB_Rinse | 0
63 | P-1352 | 2022/10/05 23:29:06 | 2-10 | 3 | 0.177 | 9.877 | 17231 | -1 | 1 | A00000195899 | 22-245 | H2O | 1665012853 | 0.337 | 0.766 | 73.906 | 0.002 | 0.002 | 1.771 | 1289.449 | 1.369 | 2.773 | 106.399 | 0.879 | 0.245 | 0 | 0 | 1 | 42.5 | 1 | 10 | 1 | GJB_Rinse | 0
64 | P-1352 | 2022/10/05 23:38:04 | 2-10 | 4 | 0.204 | 9.889 | 17493 | 0 | 1 | A00000195899 | 22-245 | H2O | 1665013390 | 0.298 | 0.773 | 74.129 | 0.002 | -0.001 | 1.769 | 1291.066 | 1.767 | 2.797 | 108.053 | 0.877 | 0.337 | 0 | 0 | 1 | 42.438 | 1 | 10 | 1 | GJB_Rinse | 0
65 | P-1353 | 2022/10/05 23:47:53 | 2-11 | 1 | -0.152 | 8.947 | 17774 | -1 | 1 | A00000196379 | 22-245 | H2O | 1665013980 | 0.305 | 0.706 | 76.599 | 0.001 | 0.001 | 1.808 | 1293.225 | 1.315 | 2.88 | 110.75 | 0.876 | 0.468 | 0 | 0 | 1 | 42.312 | 1 | 11 | 1 | GJB_Rinse | 0
66 | P-1353 | 2022/10/05 23:56:51 | 2-11 | 2 | -0.205 | 8.711 | 17510 | -1 | 1 | A00000196379 | 22-245 | H2O | 1665014517 | 0.309 | 0.699 | 74 | 0 | 0.001 | 1.772 | 1291.625 | 1.829 | 2.758 | 106.48 | 0.874 | -0.039 | 0 | 0 | 1 | 42.312 | 1 | 11 | 1 | GJB_Rinse | 0
67 | P-1353 | 2022/10/06 00:05:48 | 2-11 | 3 | -0.227 | 8.696 | 17654 | -1 | 1 | A00000196379 | 22-245 | H2O | 1665015055 | 0.326 | 0.728 | 77.172 | 0.001 | 0.001 | 1.838 | 1292.178 | 1.728 | 2.846 | 109.842 | 0.874 | 0.557 | 0 | 0 | 1 | 42.25 | 1 | 11 | 1 | GJB_Rinse | 0
68 | P-1353 | 2022/10/06 00:14:45 | 2-11 | 4 | -0.142 | 9.013 | 17640 | 0 | 1 | A00000196379 | 22-245 | H2O | 1665015592 | 0.279 | 0.785 | 74.929 | 0.001 | 0 | 1.775 | 1292.511 | 1.706 | 2.794 | 107.635 | 0.868 | 0.133 | 0 | 0 | 1 | 42.312 | 1 | 11 | 1 | GJB_Rinse | 0
69 | P-1354 | 2022/10/06 00:24:34 | 2-12 | 1 | -0.15 | 8.747 | 17672 | -1 | 1 | A00000195459 | 22-245 | H2O | 1665016181 | 0.289 | 0.721 | 72.86 | 0.001 | 0 | 1.74 | 1292.604 | 1.292 | 2.883 | 109.767 | 0.867 | 0.435 | 0 | 0 | 1 | 42.375 | 1 | 12 | 1 | GJB_Rinse | 0
70 | P-1354 | 2022/10/06 00:33:32 | 2-12 | 2 | -0.13 | 8.927 | 17685 | -1 | 1 | A00000195459 | 22-245 | H2O | 1665016718 | 0.302 | 0.646 | 76.173 | 0 | 0.002 | 1.815 | 1293.168 | 1.054 | 2.76 | 107.662 | 0.868 | -0.136 | 0 | 0 | 1 | 42.375 | 1 | 12 | 1 | GJB_Rinse | 0
71 | P-1354 | 2022/10/06 00:42:29 | 2-12 | 3 | -0.096 | 9.325 | 17621 | -1 | 1 | A00000195459 | 22-245 | H2O | 1665017256 | 0.334 | 0.783 | 74.834 | 0.002 | -0.002 | 1.795 | 1292.37 | 1.608 | 2.824 | 107.526 | 0.868 | 0.194 | 0 | 0 | 1 | 42.375 | 1 | 12 | 1 | GJB_Rinse | 0
72 | P-1354 | 2022/10/06 00:51:26 | 2-12 | 4 | -0.098 | 8.661 | 17687 | 0 | 1 | A00000195459 | 22-245 | H2O | 1665017793 | 0.339 | 0.702 | 76.237 | 0.001 | 0.001 | 1.813 | 1292.81 | 1.283 | 2.826 | 109.082 | 0.869 | 0.366 | 0 | 0 | 1 | 42.375 | 1 | 12 | 1 | GJB_Rinse | 0
73 | P-1355 | 2022/10/06 01:01:16 | 2-13 | 1 | -7.179 | -51.202 | 17544 | -1 | 1 | EV | standard | H2O | 1665018382 | 0.277 | 0.854 | 75.214 | 0.001 | 0.008 | 1.828 | 1291.498 | 1.796 | 2.782 | 108.772 | 0.873 | 0.552 | 0 | 0 | 1 | 42.5 | 1 | 13 | 1 | GJB_Rinse | 0
74 | P-1355 | 2022/10/06 01:10:13 | 2-13 | 2 | -7.438 | -56.523 | 17616 | -1 | 1 | EV | standard | H2O | 1665018920 | 0.291 | 0.813 | 75.27 | 0.001 | 0.009 | 1.791 | 1292.261 | 1.771 | 2.771 | 107.982 | 0.87 | 0.231 | 0 | 0 | 1 | 42.5 | 1 | 13 | 1 | GJB_Rinse | 0
75 | P-1355 | 2022/10/06 01:19:11 | 2-13 | 3 | -7.551 | -57.919 | 17710 | -1 | 1 | EV | standard | H2O | 1665019457 | 0.318 | 0.815 | 75.207 | 0.001 | 0.006 | 1.801 | 1292.869 | 1.71 | 2.818 | 109.192 | 0.875 | 0.34 | 0 | 0 | 1 | 42.562 | 1 | 13 | 1 | GJB_Rinse | 0
76 | P-1355 | 2022/10/06 01:28:08 | 2-13 | 4 | -7.587 | -58.852 | 17691 | 0 | 1 | EV | standard | H2O | 1665019995 | 0.322 | 0.825 | 75.406 | 0 | 0.008 | 1.802 | 1292.807 | 1.767 | 2.814 | 108.175 | 0.871 | 0.17 | 0 | 0 | 1 | 42.438 | 1 | 13 | 1 | GJB_Rinse | 0
77 | P-1356 | 2022/10/06 01:37:58 | 2-14 | 1 | -0.008 | 4.49 | 17680 | -1 | 1 | A00000195884 | 22-245 | H2O | 1665020584 | 0.31 | 0.677 | 73.208 | 0 | 0 | 1.751 | 1292.796 | 1.516 | 2.802 | 107.724 | 0.866 | 0.19 | 0 | 0 | 1 | 42.375 | 1 | 14 | 1 | GJB_Rinse | 0
78 | P-1356 | 2022/10/06 01:46:56 | 2-14 | 2 | 0.386 | 10.547 | 17692 | -1 | 1 | A00000195884 | 22-245 | H2O | 1665021123 | 0.323 | 0.751 | 77.805 | 0.001 | 0.001 | 1.848 | 1292.406 | 1.546 | 2.89 | 110.618 | 0.87 | 0.456 | 0 | 0 | 1 | 42.375 | 1 | 14 | 1 | GJB_Rinse | 0
79 | P-1356 | 2022/10/06 01:55:53 | 2-14 | 3 | 0.446 | 11.963 | 17732 | -1 | 1 | A00000195884 | 22-245 | H2O | 1665021660 | 0.323 | 0.749 | 74.401 | 0.001 | -0.001 | 1.769 | 1292.773 | 1.692 | 2.901 | 109.775 | 0.87 | 0.511 | 0 | 0 | 1 | 42.438 | 1 | 14 | 1 | GJB_Rinse | 0
80 | P-1356 | 2022/10/06 02:04:51 | 2-14 | 4 | 0.435 | 12.523 | 17702 | 0 | 1 | A00000195884 | 22-245 | H2O | 1665022197 | 0.325 | 0.745 | 76.608 | -0.001 | 0 | 1.823 | 1292.441 | 1.396 | 2.866 | 110.998 | 0.874 | 0.622 | 0 | 0 | 1 | 42.438 | 1 | 14 | 1 | GJB_Rinse | 0
81 | P-1357 | 2022/10/06 02:14:38 | 2-15 | 1 | 0.19 | 9.879 | 17610 | -1 | 1 | A00000195866 | 22-245 | H2O | 1665022785 | 0.29 | 0.713 | 75.573 | 0 | 0.001 | 1.792 | 1292.397 | 1.345 | 2.779 | 107.377 | 0.867 | 0.059 | 0 | 0 | 1 | 42.5 | 1 | 15 | 1 | GJB_Rinse | 0
82 | P-1357 | 2022/10/06 02:23:35 | 2-15 | 2 | 0.262 | 10.275 | 17823 | -1 | 1 | A00000195866 | 22-245 | H2O | 1665023322 | 0.303 | 0.749 | 74.92 | 0.001 | 0.001 | 1.788 | 1293.917 | 1.043 | 2.832 | 110.163 | 0.872 | 0.228 | 0 | 0 | 1 | 42.438 | 1 | 15 | 1 | GJB_Rinse | 0
83 | P-1357 | 2022/10/06 02:32:32 | 2-15 | 3 | 0.227 | 10.34 | 17792 | -1 | 1 | A00000195866 | 22-245 | H2O | 1665023859 | 0.327 | 0.777 | 72.986 | 0.001 | -0.002 | 1.737 | 1293.588 | 1.348 | 2.84 | 109.36 | 0.864 | 0.268 | 0 | 0 | 1 | 42.375 | 1 | 15 | 1 | GJB_Rinse | 0
84 | P-1357 | 2022/10/06 02:41:29 | 2-15 | 4 | 0.28 | 10.377 | 17880 | 0 | 1 | A00000195866 | 22-245 | H2O | 1665024395 | 0.279 | 0.728 | 74.295 | 0.001 | -0.001 | 1.777 | 1294.02 | 1.872 | 2.809 | 109.398 | 0.871 | 0.235 | 0 | 0 | 1 | 42.312 | 1 | 15 | 1 | GJB_Rinse | 0
85 | P-1358 | 2022/10/06 02:51:16 | 2-16 | 1 | -0.008 | 9.481 | 17549 | -1 | 1 | A00000195850 | 22-245 | H2O | 1665024982 | 0.315 | 0.821 | 75.782 | 0 | 0 | 1.804 | 1291.9 | 1.155 | 2.806 | 108.081 | 0.869 | 0.259 | 0 | 0 | 1 | 42.25 | 1 | 16 | 1 | GJB_Rinse | 0
86 | P-1358 | 2022/10/06 03:00:14 | 2-16 | 2 | -0.028 | 9.546 | 17447 | -1 | 1 | A00000195850 | 22-245 | H2O | 1665025520 | 0.304 | 0.849 | 74.576 | 0 | -0.004 | 1.789 | 1290.757 | 1.669 | 2.807 | 107.683 | 0.868 | 0.395 | 0 | 0 | 1 | 42.25 | 1 | 16 | 1 | GJB_Rinse | 0
87 | P-1358 | 2022/10/06 03:09:11 | 2-16 | 3 | -0.044 | 9.562 | 17372 | -1 | 1 | A00000195850 | 22-245 | H2O | 1665026057 | 0.302 | 0.74 | 73.763 | 0 | -0.001 | 1.755 | 1290.441 | 1.573 | 2.776 | 106.551 | 0.864 | 0.243 | 0 | 0 | 1 | 42.375 | 1 | 16 | 1 | GJB_Rinse | 0
88 | P-1359 | 2022/10/06 03:18:33 | 2-17 | 1 | [5 empty] | A00000195836 | 22-245 | H2O | [17 empty] | 1 | 17 | 1 | GJB_Rinse | 0
89 | P-1359 | 2022/10/06 03:29:25 | 2-17 | 2 | -0.049 | 9.78 | 17385 | -1 | 1 | A00000195836 | 22-245 | H2O | 1665027272 | 0.287 | 0.799 | 71.938 | 0 | 0.002 | 1.709 | 1290.514 | 1.212 | 2.797 | 108.038 | 0.875 | 0.476 | 0 | 0 | 1 | 42.438 | 1 | 17 | 1 | GJB_Rinse | 0
90 | P-1359 | 2022/10/06 03:38:22 | 2-17 | 3 | 0.094 | 10.452 | 17868 | -1 | 1 | A00000195836 | 22-245 | H2O | 1665027809 | 0.314 | 0.75 | 74.638 | 0.001 | 0.002 | 1.772 | 1294.095 | 1.723 | 2.807 | 108.457 | 0.87 | -0.05 | 0 | 0 | 1 | 42.375 | 1 | 17 | 1 | GJB_Rinse | 0
91 | P-1359 | 2022/10/06 03:47:19 | 2-17 | 4 | 0.015 | 9.88 | 17783 | 0 | 1 | A00000195836 | 22-245 | H2O | 1665028346 | 0.32 | 0.754 | 75.662 | 0.001 | 0.002 | 1.794 | 1293.358 | 1.293 | 2.838 | 110.558 | 0.873 | 0.369 | 0 | 0 | 1 | 42.375 | 1 | 17 | 1 | GJB_Rinse | 0
92 | P-1359 | 2022/10/06 03:56:16 | 2-17 | 5 | -0.032 | 9.591 | 17708 | 0 | 1 | A00000195836 | 22-245 | H2O | 1665028883 | 0.314 | 0.728 | 76.23 | 0 | 0 | 1.822 | 1292.899 | 1.259 | 2.89 | 109.735 | 0.875 | 0.223 | 0 | 0 | 1 | 42.312 | 1 | 17 | 1 | GJB_Rinse | 0
93 | P-1360 | 2022/10/06 04:06:04 | 2-18 | 1 | 1.586 | 16.181 | 17567 | -1 | 1 | A00000194866 | 22-245 | H2O | 1665029471 | 0.293 | 0.728 | 76.01 | 0 | -0.001 | 1.807 | 1291.716 | 1.777 | 2.802 | 107.813 | 0.876 | 0.243 | 0 | 0 | 1 | 42.375 | 1 | 18 | 1 | GJB_Rinse | 0
94 | P-1360 | 2022/10/06 04:15:01 | 2-18 | 2 | 1.621 | 16.766 | 17550 | -1 | 1 | A00000194866 | 22-245 | H2O | 1665030008 | 0.294 | 0.76 | 75.08 | 0.001 | -0.002 | 1.789 | 1291.716 | 1.636 | 2.779 | 107.121 | 0.874 | 0.204 | 0 | 0 | 1 | 42.438 | 1 | 18 | 1 | GJB_Rinse | 0
95 | P-1360 | 2022/10/06 04:23:58 | 2-18 | 3 | 1.619 | 16.827 | 17570 | -1 | 1 | A00000194866 | 22-245 | H2O | 1665030545 | 0.319 | 0.718 | 74.154 | 0 | -0.001 | 1.766 | 1291.789 | 1.347 | 2.83 | 108.001 | 0.868 | 0.189 | 0 | 0 | 1 | 42.375 | 1 | 18 | 1 | GJB_Rinse | 0
96 | P-1360 | 2022/10/06 04:32:55 | 2-18 | 4 | 1.65 | 17.328 | 17917 | 0 | 1 | A00000194866 | 22-245 | H2O | 1665031081 | 0.315 | 0.738 | 76.836 | 0 | 0.004 | 1.83 | 1294.361 | 1.309 | 2.888 | 110.639 | 0.869 | 0.249 | 0 | 0 | 1 | 42.375 | 1 | 18 | 1 | GJB_Rinse | 0
97 | P-1361 | 2022/10/06 04:42:44 | 2-19 | 1 | 1.645 | 23.472 | 17585 | -1 | 1 | A00000194883 | 22-245 | H2O | 1665031670 | 0.295 | 0.677 | 76.637 | 0.001 | 0.003 | 1.832 | 1291.86 | 1.389 | 2.837 | 109.119 | 0.871 | 0.438 | 0 | 0 | 1 | 42.5 | 1 | 19 | 1 | GJB_Rinse | 0
98 | P-1361 | 2022/10/06 04:51:41 | 2-19 | 2 | 1.644 | 24.434 | 17373 | -1 | 1 | A00000194883 | 22-245 | H2O | 1665032208 | 0.33 | 0.768 | 73.802 | 0.002 | 0.001 | 1.765 | 1290.353 | 1.604 | 2.762 | 106.922 | 0.875 | 0.272 | 0 | 0 | 1 | 42.438 | 1 | 19 | 1 | GJB_Rinse | 0
99 | P-1361 | 2022/10/06 05:00:38 | 2-19 | 3 | 1.625 | 24.724 | 17697 | -1 | 1 | A00000194883 | 22-245 | H2O | 1665032745 | 0.323 | 0.789 | 73.101 | 0.002 | 0 | 1.735 | 1292.794 | 1.255 | 2.825 | 108.886 | 0.871 | 0.134 | 0 | 0 | 1 | 42.312 | 1 | 19 | 1 | GJB_Rinse | 0
100 | P-1361 | 2022/10/06 05:09:35 | 2-19 | 4 | 1.57 | 24.351 | 17759 | 0 | 1 | A00000194883 | 22-245 | H2O | 1665033281 | 0.32 | 0.699 | 76.616 | 0.001 | 0.001 | 1.822 | 1293.104 | 1.405 | 2.844 | 109.629 | 0.875 | 0.173 | 0 | 0 | 1 | 42.25 | 1 | 19 | 1 | GJB_Rinse | 0
101 | P-1362 | 2022/10/06 05:19:24 | 2-20 | 1 | 7.155 | 46.787 | 17852 | -1 | 1 | A00000194857 | 22-245 | H2O | 1665033871 | 0.322 | 0.699 | 75.925 | 0.001 | -0.001 | 1.806 | 1293.742 | 1.594 | 2.872 | 109.789 | 0.865 | 0.327 | 0 | 0 | 1 | 42.188 | 1 | 20 | 1 | GJB_Rinse | 0
102 | P-1362 | 2022/10/06 05:28:21 | 2-20 | 2 | 7.397 | 49.196 | 17867 | -1 | 1 | A00000194857 | 22-245 | H2O | 1665034407 | 0.293 | 0.786 | 74.779 | 0.001 | -0.001 | 1.79 | 1294.175 | 0.738 | 2.882 | 111.156 | 0.874 | 0.191 | 0 | 0 | 1 | 42.25 | 1 | 20 | 1 | GJB_Rinse | 0
103 | P-1362 | 2022/10/06 05:37:18 | 2-20 | 3 | 7.525 | 49.521 | 17914 | -1 | 1 | A00000194857 | 22-245 | H2O | 1665034945 | 0.308 | 0.759 | 75.193 | 0 | 0.001 | 1.784 | 1294.317 | 1.07 | 2.883 | 111.828 | 0.876 | 0.441 | 0 | 0 | 1 | 42.25 | 1 | 20 | 1 | GJB_Rinse | 0
104 | P-1362 | 2022/10/06 05:46:15 | 2-20 | 4 | 7.55 | 49.647 | 17892 | 0 | 1 | A00000194857 | 22-245 | H2O | 1665035482 | 0.298 | 0.747 | 76.046 | 0 | 0 | 1.802 | 1294.236 | 1.331 | 2.86 | 109.754 | 0.873 | 0.11 | 0 | 0 | 1 | 42.25 | 1 | 20 | 1 | GJB_Rinse | 0
105 | P-1363 | 2022/10/06 05:56:05 | 2-21 | 1 | 4.17 | 35.066 | 17428 | -1 | 1 | A00000194855 | 22-245 | H2O | 1665036072 | 0.308 | 0.794 | 75.287 | 0 | 0.001 | 1.792 | 1290.999 | 1.468 | 2.787 | 107.158 | 0.87 | 0.152 | 0 | 0 | 1 | 42.25 | 1 | 21 | 1 | GJB_Rinse | 0
106 | P-1363 | 2022/10/06 06:05:03 | 2-21 | 2 | 4.015 | 33.611 | 17473 | -1 | 1 | A00000194855 | 22-245 | H2O | 1665036609 | 0.297 | 0.742 | 77.71 | 0.001 | 0.002 | 1.847 | 1291.274 | 1.564 | 2.794 | 107.052 | 0.866 | 0.282 | 0 | 0 | 1 | 42.312 | 1 | 21 | 1 | GJB_Rinse | 0
107 | P-1363 | 2022/10/06 06:14:00 | 2-21 | 3 | 3.987 | 33.406 | 17491 | -1 | 1 | A00000194855 | 22-245 | H2O | 1665037147 | 0.307 | 0.692 | 75.593 | 0.001 | 0.002 | 1.808 | 1291.507 | 0.973 | 2.81 | 108.818 | 0.87 | 0.258 | 0 | 0 | 1 | 42.375 | 1 | 21 | 1 | GJB_Rinse | 0
108 | P-1363 | 2022/10/06 06:22:57 | 2-21 | 4 | 3.988 | 33.756 | 17518 | 0 | 1 | A00000194855 | 22-245 | H2O | 1665037683 | 0.312 | 0.746 | 75.101 | 0.001 | 0.002 | 1.792 | 1291.33 | 1.416 | 2.845 | 109.329 | 0.867 | 0.478 | 0 | 0 | 1 | 42.438 | 1 | 21 | 1 | GJB_Rinse | 0
109 | P-1364 | 2022/10/06 06:32:42 | 2-22 | 1 | -6.802 | -47.37 | 17738 | -1 | 1 | EV | standard | H2O | 1665038269 | 0.308 | 0.843 | 76.354 | 0.002 | 0.008 | 1.813 | 1292.892 | 1.819 | 2.804 | 109.671 | 0.871 | 0.385 | 0 | 0 | 1 | 42.25 | 1 | 22 | 1 | GJB_Rinse | 0
110 | P-1364 | 2022/10/06 06:41:38 | 2-22 | 2 | -7.331 | -55.247 | 17643 | -1 | 1 | EV | standard | H2O | 1665038804 | 0.323 | 0.789 | 77.146 | 0.002 | 0.01 | 1.836 | 1292.497 | 1.675 | 2.778 | 107.615 | 0.869 | 0.143 | 0 | 0 | 1 | 42.188 | 1 | 22 | 1 | GJB_Rinse | 0
111 | P-1364 | 2022/10/06 06:50:35 | 2-22 | 3 | -7.429 | -56.918 | 17786 | -1 | 1 | EV | standard | H2O | 1665039341 | 0.307 | 0.793 | 77.06 | 0.001 | 0.007 | 1.844 | 1293.376 | 1.784 | 2.82 | 108.925 | 0.875 | 0.184 | 0 | 0 | 1 | 42.312 | 1 | 22 | 1 | GJB_Rinse | 0
112 | P-1364 | 2022/10/06 06:59:31 | 2-22 | 4 | -7.5 | -57.87 | 17668 | 0 | 1 | EV | standard | H2O | 1665039878 | 0.316 | 0.738 | 78.176 | 0.003 | 0.006 | 1.856 | 1292.73 | 1.501 | 2.785 | 108.248 | 0.866 | 0.144 | 0 | 0 | 1 | 42.375 | 1 | 22 | 1 | GJB_Rinse | 0
113 | P-1365 | 2022/10/06 07:09:19 | 2-23 | 1 | 0.437 | 5.506 | 18018 | -1 | 1 | A00000194867 | 22-245 | H2O | 1665040465 | 0.316 | 0.717 | 79.317 | 0 | 0.001 | 1.879 | 1294.691 | 1.673 | 2.926 | 112.114 | 0.873 | 0.517 | 0 | 0 | 1 | 42.375 | 1 | 23 | 1 | GJB_Rinse | 0
114 | P-1365 | 2022/10/06 07:18:16 | 2-23 | 2 | 0.863 | 12.182 | 18574 | -1 | 1 | A00000194867 | 22-245 | H2O | 1665041002 | 0.322 | 0.714 | 76.276 | 0.001 | 0.001 | 1.812 | 1298.909 | 1.658 | 3.007 | 114.218 | 0.866 | 0.329 | 0 | 0 | 1 | 42.25 | 1 | 23 | 1 | GJB_Rinse | 0
115 | P-1365 | 2022/10/06 07:27:13 | 2-23 | 3 | 0.943 | 13.342 | 18695 | -1 | 1 | A00000194867 | 22-245 | H2O | 1665041539 | 0.318 | 0.664 | 78.856 | 0 | 0.002 | 1.875 | 1299.7 | 1.256 | 3.064 | 116.448 | 0.865 | 0.381 | 0 | 0 | 1 | 42.188 | 1 | 23 | 1 | GJB_Rinse | 0
116 | P-1365 | 2022/10/06 07:36:09 | 2-23 | 4 | 0.955 | 13.48 | 18198 | 0 | 1 | A00000194867 | 22-245 | H2O | 1665042076 | 0.35 | 0.751 | 77.361 | 0.001 | 0.002 | 1.843 | 1296.326 | 1.535 | 2.892 | 111.297 | 0.868 | 0.214 | 0 | 0 | 1 | 42.125 | 1 | 23 | 1 | GJB_Rinse | 0
117 | P-1366 | 2022/10/06 07:45:56 | 2-24 | 1 | -6.019 | -43.966 | 18410 | -1 | 1 | A00000196463 | 22-245 | H2O | 1665042663 | 0.316 | 0.81 | 76.378 | 0.001 | 0.007 | 1.818 | 1297.903 | 1.811 | 2.92 | 112.225 | 0.868 | 0.17 | 0 | 0 | 1 | 42.25 | 1 | 24 | 1 | GJB_Rinse | 0
118 | P-1366 | 2022/10/06 07:54:53 | 2-24 | 2 | -6.403 | -49.623 | 17828 | -1 | 1 | A00000196463 | 22-245 | H2O | 1665043200 | 0.275 | 0.789 | 75.072 | 0.001 | 0.007 | 1.797 | 1293.715 | 1.745 | 2.801 | 109.321 | 0.869 | 0.29 | 0 | 0 | 1 | 42.312 | 1 | 24 | 1 | GJB_Rinse | 0
119 | P-1366 | 2022/10/06 08:03:50 | 2-24 | 3 | -6.444 | -50.26 | 18393 | -1 | 1 | A00000196463 | 22-245 | H2O | 1665043737 | 0.304 | 0.776 | 78.383 | 0.001 | 0.007 | 1.859 | 1297.632 | 1.563 | 2.924 | 113.595 | 0.867 | 0.417 | 0 | 0 | 1 | 42.312 | 1 | 24 | 1 | GJB_Rinse | 0
120 | P-1366 | 2022/10/06 08:12:47 | 2-24 | 4 | -6.51 | -51.109 | 17639 | 0 | 1 | A00000196463 | 22-245 | H2O | 1665044273 | 0.313 | 0.764 | 76.717 | 0 | 0.003 | 1.832 | 1292.322 | 1.963 | 2.78 | 107.172 | 0.866 | 0.166 | 0 | 0 | 1 | 42.312 | 1 | 24 | 1 | GJB_Rinse | 0
121 | P-1367 | 2022/10/06 08:22:34 | 2-25 | 1 | 0.524 | 6.335 | 18592 | -1 | 1 | A00000194885 | 22-245 | H2O | 1665044861 | 0.297 | 0.756 | 74.657 | 0 | -0.002 | 1.77 | 1299.306 | 1.286 | 2.98 | 115.409 | 0.867 | 0.359 | 0 | 0 | 1 | 42.5 | 1 | 25 | 1 | GJB_Rinse | 0
122 | P-1367 | 2022/10/06 08:31:31 | 2-25 | 2 | 0.821 | 11.141 | 18006 | -1 | 1 | A00000194885 | 22-245 | H2O | 1665045398 | 0.329 | 0.729 | 76.446 | 0.001 | 0 | 1.814 | 1295.033 | 1.55 | 2.882 | 111.241 | 0.872 | 0.369 | 0 | 0 | 1 | 42.625 | 1 | 25 | 1 | GJB_Rinse | 0
123 | P-1367 | 2022/10/06 08:40:29 | 2-25 | 3 | 0.928 | 12.619 | 18328 | -1 | 1 | A00000194885 | 22-245 | H2O | 1665045936 | 0.337 | 0.722 | 76.684 | 0 | -0.003 | 1.821 | 1297.654 | 1.223 | 2.943 | 112.611 | 0.867 | 0.112 | 0 | 0 | 1 | 42.75 | 1 | 25 | 1 | GJB_Rinse | 0
124 | P-1367 | 2022/10/06 08:49:25 | 2-25 | 4 | 0.92 | 13.343 | 18577 | 0 | 1 | A00000194885 | 22-245 | H2O | 1665046472 | 0.307 | 0.808 | 79.905 | 0.001 | 0.001 | 1.905 | 1299.427 | 0.717 | 3.02 | 115.71 | 0.863 | 0.254 | 0 | 0 | 1 | 42.688 | 1 | 25 | 1 | GJB_Rinse | 0
125 | P-1370 | 2022/10/06 14:48:05 | 2-26 | 1 | 1.258 | 18.069 | 18820 | -1 | 1 | A00000196461 | 22-245 | H2O | 1665067992 | 0.304 | 0.763 | 83.719 | 0.001 | 0.007 | 1.981 | 1300.688 | 1.811 | 3.065 | 116.252 | 0.87 | 0.479 | 0 | 0 | 1 | 42.312 | 1 | 26 | 1 | GJB_Rinse | 0
126 | P-1370 | 2022/10/06 14:57:01 | 2-26 | 2 | 1.301 | 18.164 | 18455 | -1 | 1 | A00000196461 | 22-245 | H2O | 1665068528 | 0.331 | 0.768 | 77.499 | 0.001 | 0.001 | 1.83 | 1298.562 | 1.306 | 2.957 | 113.594 | 0.865 | 0.219 | 0 | 0 | 1 | 42.375 | 1 | 26 | 1 | GJB_Rinse | 0
127 | P-1370 | 2022/10/06 15:05:57 | 2-26 | 3 | 1.304 | 18.498 | 18245 | -1 | 1 | A00000196461 | 22-245 | H2O | 1665069064 | 0.346 | 0.765 | 76.12 | 0 | -0.001 | 1.807 | 1296.754 | 1.762 | 2.92 | 112.264 | 0.872 | 0.247 | 0 | 0 | 1 | 42.438 | 1 | 26 | 1 | GJB_Rinse | 0
128 | P-1370 | 2022/10/06 15:14:54 | 2-26 | 4 | 1.296 | 18.514 | 18494 | 0 | 1 | A00000196461 | 22-245 | H2O | 1665069600 | 0.31 | 0.732 | 76.211 | 0.002 | 0.001 | 1.815 | 1298.535 | 1.574 | 2.993 | 113.468 | 0.864 | 0.255 | 0 | 0 | 1 | 42.75 | 1 | 26 | 1 | GJB_Rinse | 0
129 | P-1371 | 2022/10/06 15:24:42 | 2-27 | 1 | -2.231 | -11.241 | 18732 | -1 | 1 | A00000320061 | 22-245 | H2O | 1665070188 | 0.326 | 0.698 | 79.106 | 0.001 | 0.005 | 1.876 | 1300.227 | 1.649 | 3.017 | 115.755 | 0.867 | 0.481 | 0 | 0 | 1 | 42.625 | 1 | 27 | 1 | GJB_Rinse | 0
130 | P-1371 | 2022/10/06 15:33:38 | 2-27 | 2 | -2.448 | -14.322 | 18685 | -1 | 1 | A00000320061 | 22-245 | H2O | 1665070725 | 0.335 | 0.695 | 76.124 | 0 | 0.004 | 1.803 | 1299.964 | 1.257 | 3.006 | 115.949 | 0.868 | 0.52 | 0 | 0 | 1 | 42.438 | 1 | 27 | 1 | GJB_Rinse | 0
131 | P-1371 | 2022/10/06 15:42:34 | 2-27 | 3 | -2.486 | -15.273 | 18277 | -1 | 1 | A00000320061 | 22-245 | H2O | 1665071261 | 0.322 | 0.755 | 79.172 | 0.001 | 0.006 | 1.872 | 1297.115 | 1.774 | 2.899 | 111.4 | 0.874 | -0.039 | 0 | 0 | 1 | 42.375 | 1 | 27 | 1 | GJB_Rinse | 0
132 | P-1371 | 2022/10/06 15:51:30 | 2-27 | 4 | -2.472 | -15.473 | 18290 | 0 | 1 | A00000320061 | 22-245 | H2O | 1665071796 | 0.291 | 0.724 | 75.854 | 0.001 | -0.001 | 1.804 | 1297.185 | 1.303 | 2.944 | 112.558 | 0.87 | 0.21 | 0 | 0 | 1 | 42.375 | 1 | 27 | 1 | GJB_Rinse | 0
133 | P-1372 | 2022/10/06 16:01:19 | 2-28 | 1 | -2.816 | -18.346 | 18849 | -1 | 1 | A00000320079 | 22-245 | H2O | 1665072385 | 0.311 | 0.675 | 77.558 | 0.001 | 0.002 | 1.834 | 1301.057 | 1.526 | 3.055 | 116.33 | 0.865 | 0.36 | 0 | 0 | 1 | 42.312 | 1 | 28 | 1 | GJB_Rinse | 0
134 | P-1372 | 2022/10/06 16:10:15 | 2-28 | 2 | -2.889 | -18.969 | 18679 | -1 | 1 | A00000320079 | 22-245 | H2O | 1665072921 | 0.323 | 0.689 | 79.078 | 0 | 0.004 | 1.874 | 1299.701 | 1.961 | 2.997 | 114.81 | 0.869 | 0.41 | 0 | 0 | 1 | 42.25 | 1 | 28 | 1 | GJB_Rinse | 0
135 | P-1372 | 2022/10/06 16:19:11 | 2-28 | 3 | -2.876 | -18.699 | 18648 | -1 | 1 | A00000320079 | 22-245 | H2O | 1665073458 | 0.312 | 0.77 | 75.992 | 0 | 0.005 | 1.798 | 1299.553 | 1.599 | 3.002 | 115.106 | 0.864 | 0.462 | 0 | 0 | 1 | 42.188 | 1 | 28 | 1 | GJB_Rinse | 0
136 | P-1372 | 2022/10/06 16:28:07 | 2-28 | 4 | -2.928 | -18.982 | 18599 | 0 | 1 | A00000320079 | 22-245 | H2O | 1665073994 | 0.304 | 0.772 | 78.432 | 0.001 | 0.003 | 1.862 | 1299.174 | 1.497 | 3.014 | 115.053 | 0.866 | 0.473 | 0 | 0 | 1 | 42.25 | 1 | 28 | 1 | GJB_Rinse | 0
137 | P-1373 | 2022/10/06 16:37:52 | 2-29 | 1 | -3.012 | -19.248 | 18861 | -1 | 1 | A00000320088 | 22-245 | H2O | 1665074579 | 0.329 | 0.665 | 77.104 | 0 | 0.003 | 1.83 | 1301.297 | 1.444 | 3.031 | 115.666 | 0.87 | 0.378 | 0 | 0 | 1 | 42.25 | 1 | 29 | 1 | GJB_Rinse | 0
138 | P-1373 | 2022/10/06 16:46:48 | 2-29 | 2 | -3.044 | -19.266 | 18757 | -1 | 1 | A00000320088 | 22-245 | H2O | 1665075114 | 0.347 | 0.748 | 78.209 | 0.002 | 0.004 | 1.845 | 1300.33 | 1.415 | 3.003 | 115.977 | 0.867 | 0.381 | 0 | 0 | 1 | 42.25 | 1 | 29 | 1 | GJB_Rinse | 0
139 | P-1373 | 2022/10/06 16:55:44 | 2-29 | 3 | -3.05 | -19.285 | 18694 | -1 | 1 | A00000320088 | 22-245 | H2O | 1665075651 | 0.314 | 0.721 | 78.164 | 0.001 | 0.004 | 1.848 | 1299.653 | 1.856 | 3.019 | 115.978 | 0.865 | 0.498 | 0 | 0 | 1 | 42.312 | 1 | 29 | 1 | GJB_Rinse | 0
140 | P-1373 | 2022/10/06 17:04:40 | 2-29 | 4 | -3.006 | -19.208 | 18455 | 0 | 1 | A00000320088 | 22-245 | H2O | 1665076187 | 0.31 | 0.699 | 79.373 | 0.001 | 0.004 | 1.877 | 1298.022 | 1.524 | 2.975 | 114.723 | 0.869 | 0.441 | 0 | 0 | 1 | 42.312 | 1 | 29 | 1 | GJB_Rinse | 0
141 | P-1374 | 2022/10/06 17:14:25 | 2-30 | 1 | -2.727 | -18.218 | 18964 | -1 | 1 | A00000320123 | 22-245 | H2O | 1665076772 | 0.317 | 0.788 | 80.644 | 0.001 | 0.006 | 1.902 | 1301.779 | 1.652 | 3.05 | 117.301 | 0.872 | 0.54 | 0 | 0 | 1 | 42.312 | 1 | 30 | 1 | GJB_Rinse | 0
142 | P-1374 | 2022/10/06 17:23:21 | 2-30 | 2 | -2.781 | -18.506 | 18505 | -1 | 1 | A00000320123 | 22-245 | H2O | 1665077307 | 0.307 | 0.714 | 75.363 | 0 | 0.002 | 1.795 | 1298.659 | 1.471 | 2.956 | 114.149 | 0.874 | 0.243 | 0 | 0 | 1 | 42.375 | 1 | 30 | 1 | GJB_Rinse | 0
143 | P-1374 | 2022/10/06 17:32:17 | 2-30 | 3 | -2.792 | -18.352 | 18447 | -1 | 1 | A00000320123 | 22-245 | H2O | 1665077844 | 0.316 | 0.68 | 77.295 | 0.001 | 0 | 1.843 | 1298.244 | 1.029 | 2.976 | 114.663 | 0.868 | 0.362 | 0 | 0 | 1 | 42.375 | 1 | 30 | 1 | GJB_Rinse | 0
144 | P-1374 | 2022/10/06 17:41:14 | 2-30 | 4 | -2.876 | -18.882 | 18461 | 0 | 1 | A00000320123 | 22-245 | H2O | 1665078380 | 0.301 | 0.745 | 76.566 | 0.001 | 0.001 | 1.813 | 1297.736 | 2.129 | 2.98 | 114.321 | 0.869 | 0.542 | 0 | 0 | 1 | 42.375 | 1 | 30 | 1 | GJB_Rinse | 0
145 | P-1375 | 2022/10/06 17:50:59 | 2-31 | 1 | -7.21 | -53.831 | 18905 | -1 | 1 | EV | standard | H2O | 1665078965 | 0.328 | 0.778 | 78.812 | 0.001 | 0.005 | 1.879 | 1301.183 | 1.753 | 3.035 | 117.464 | 0.869 | 0.564 | 0 | 0 | 1 | 42.312 | 1 | 31 | 1 | GJB_Rinse | 0
146 | P-1375 | 2022/10/06 17:59:55 | 2-31 | 2 | -7.467 | -57.29 | 18548 | -1 | 1 | EV | standard | H2O | 1665079501 | 0.306 | 0.737 | 78.466 | 0.002 | 0.006 | 1.853 | 1298.78 | 1.324 | 3.033 | 115.433 | 0.866 | 0.602 | 0 | 0 | 1 | 42.25 | 1 | 31 | 1 | GJB_Rinse | 0
147 | P-1375 | 2022/10/06 18:08:51 | 2-31 | 3 | -7.511 | -58.194 | 18504 | -1 | 1 | EV | standard | H2O | 1665080038 | 0.268 | 0.739 | 77.679 | 0.001 | 0.005 | 1.838 | 1298.293 | 1.743 | 2.951 | 114.685 | 0.864 | 0.487 | 0 | 0 | 1 | 42.25 | 1 | 31 | 1 | GJB_Rinse | 0
148 | P-1375 | 2022/10/06 18:17:47 | 2-31 | 4 | -7.566 | -58.561 | 18477 | 0 | 1 | EV | standard | H2O | 1665080574 | 0.302 | 0.7 | 77.74 | 0 | 0.003 | 1.842 | 1298.258 | 1.67 | 2.947 | 113.699 | 0.868 | 0.261 | 0 | 0 | 1 | 42.312 | 1 | 31 | 1 | GJB_Rinse | 0
149 | P-1376 | 2022/10/06 18:27:33 | 2-32 | 1 | -3.001 | -22.316 | 18872 | -1 | 1 | A00000347319 | 22-245 | H2O | 1665081159 | 0.303 | 0.742 | 77.522 | 0.001 | 0.001 | 1.831 | 1300.986 | 1.712 | 3.044 | 117.1 | 0.87 | 0.495 | 0 | 0 | 1 | 42.375 | 1 | 32 | 1 | GJB_Rinse | 0
150 | P-1376 | 2022/10/06 18:36:29 | 2-32 | 2 | -2.854 | -19.929 | 18503 | -1 | 1 | A00000347319 | 22-245 | H2O | 1665081695 | 0.302 | 0.769 | 77.051 | 0.001 | 0.002 | 1.824 | 1298.557 | 1.713 | 2.932 | 112.781 | 0.869 | 0.089 | 0 | 0 | 1 | 42.625 | 1 | 32 | 1 | GJB_Rinse | 0
151 | P-1376 | 2022/10/06 18:45:25 | 2-32 | 3 | -2.811 | -18.807 | 18381 | -1 | 1 | A00000347319 | 22-245 | H2O | 1665082232 | 0.311 | 0.677 | 77.951 | 0.001 | 0.001 | 1.852 | 1297.362 | 1.311 | 2.946 | 114.409 | 0.871 | 0.409 | 0 | 0 | 1 | 42.438 | 1 | 32 | 1 | GJB_Rinse | 0
152 | P-1376 | 2022/10/06 18:54:21 | 2-32 | 4 | -2.811 | -19.167 | 18441 | 0 | 1 | A00000347319 | 22-245 | H2O | 1665082768 | 0.331 | 0.707 | 74.87 | 0.001 | 0.001 | 1.771 | 1297.764 | 1.459 | 2.972 | 114.98 | 0.868 | 0.398 | 0 | 0 | 1 | 42.312 | 1 | 32 | 1 | GJB_Rinse | 0
153 | P-1377 | 2022/10/06 19:04:08 | 2-33 | 1 | -3.122 | -19.661 | 18806 | -1 | 1 | A00000320113 | 22-245 | H2O | 1665083354 | 0.296 | 0.734 | 77.474 | 0.001 | 0.003 | 1.845 | 1300.606 | 1.701 | 2.987 | 115.019 | 0.868 | 0.221 | 0 | 0 | 1 | 42.312 | 1 | 33 | 1 | GJB_Rinse | 0
154 | P-1377 | 2022/10/06 19:13:04 | 2-33 | 2 | -3.134 | -19.322 | 18905 | -1 | 1 | A00000320113 | 22-245 | H2O | 1665083890 | 0.313 | 0.701 | 80.136 | 0 | 0 | 1.894 | 1301.326 | 1.497 | 3.061 | 117.179 | 0.865 | 0.345 | 0 | 0 | 1 | 42 | 1 | 33 | 1 | GJB_Rinse | 0
155 | P-1377 | 2022/10/06 19:22:00 | 2-33 | 3 | -3.203 | -20.313 | 18219 | -1 | 1 | A00000320113 | 22-245 | H2O | 1665084426 | 0.301 | 0.811 | 75.982 | 0.001 | 0.004 | 1.809 | 1296.243 | 1.905 | 2.926 | 112.112 | 0.869 | 0.402 | 0 | 0 | 1 | 42.188 | 1 | 33 | 1 | GJB_Rinse | 0
156 | P-1377 | 2022/10/06 19:30:56 | 2-33 | 4 | -3.194 | -20.076 | 18281 | 0 | 1 | A00000320113 | 22-245 | H2O | 1665084962 | 0.295 | 0.746 | 72.474 | 0 | -0.001 | 1.737 | 1296.962 | 1.425 | 2.942 | 112.106 | 0.869 | 0.116 | 0 | 0 | 1 | 42.312 | 1 | 33 | 1 | GJB_Rinse | 0
157 | P-1378 | 2022/10/06 19:40:42 | 2-34 | 1 | -3.27 | -21.026 | 18731 | -1 | 1 | A00000320119 | 22-245 | H2O | 1665085549 | 0.294 | 0.745 | 78.555 | 0.001 | 0.005 | 1.864 | 1300.029 | 1.623 | 2.957 | 114.649 | 0.872 | 0.185 | 0 | 0 | 1 | 42.312 | 1 | 34 | 1 | GJB_Rinse | 0
158 | P-1378 | 2022/10/06 19:49:39 | 2-34 | 2 | -3.364 | -21.532 | 17684 | -1 | 1 | A00000320119 | 22-245 | H2O | 1665086085 | 0.328 | 0.728 | 76.12 | 0.001 | -0.001 | 1.805 | 1292.544 | 1.951 | 2.802 | 107.658 | 0.866 | 0.092 | 0 | 0 | 1 | 42.312 | 1 | 34 | 1 | GJB_Rinse | 0
159 | P-1378 | 2022/10/06 19:58:35 | 2-34 | 3 | -3.215 | -20.736 | 18613 | -1 | 1 | A00000320119 | 22-245 | H2O | 1665086621 | 0.307 | 0.8 | 77.293 | -0.001 | 0.001 | 1.836 | 1299.278 | 0.794 | 2.991 | 116.724 | 0.87 | 0.511 | 0 | 0 | 1 | 42.375 | 1 | 34 | 1 | GJB_Rinse | 0
160 | P-1378 | 2022/10/06 20:07:31 | 2-34 | 4 | -3.331 | -21.268 | 18118 | 0 | 1 | A00000320119 | 22-245 | H2O | 1665087157 | 0.322 | 0.711 | 75.858 | 0.001 | 0.002 | 1.807 | 1295.564 | 1.495 | 2.917 | 112.718 | 0.873 | 0.441 | 0 | 0 | 1 | 42.375 | 1 | 34 | 1 | GJB_Rinse | 0
161 | P-1379 | 2022/10/06 20:17:18 | 2-35 | 1 | -2.909 | -19.506 | 18817 | -1 | 1 | A00000320114 | 22-245 | H2O | 1665087744 | 0.305 | 0.666 | 77.705 | 0.001 | 0 | 1.847 | 1300.674 | 2.037 | 2.994 | 114.743 | 0.87 | 0.128 | 0 | 0 | 1 | 42.375 | 1 | 35 | 1 | GJB_Rinse | 0
162 | P-1379 | 2022/10/06 20:26:14 | 2-35 | 2 | -2.94 | -19.958 | 18529 | -1 | 1 | A00000320114 | 22-245 | H2O | 1665088280 | 0.322 | 0.897 | 80.116 | 0.001 | 0.002 | 1.904 | 1298.696 | 1.346 | 3.002 | 114.674 | 0.875 | 0.274 | 0 | 0 | 1 | 42.438 | 1 | 35 | 1 | GJB_Rinse | 0
163 | P-1379 | 2022/10/06 20:35:10 | 2-35 | 3 | -2.975 | -19.852 | 18418 | -1 | 1 | A00000320114 | 22-245 | H2O | 1665088817 | 0.291 | 0.704 | 78.996 | 0 | 0.001 | 1.886 | 1297.482 | 2.071 | 2.975 | 113.557 | 0.872 | 0.474 | 0 | 0 | 1 | 42.375 | 1 | 35 | 1 | GJB_Rinse | 0
164 | P-1379 | 2022/10/06 20:44:07 | 2-35 | 4 | -2.963 | -19.872 | 18297 | 0 | 1 | A00000320114 | 22-245 | H2O | 1665089353 | 0.293 | 0.769 | 76.214 | 0.001 | 0.005 | 1.807 | 1296.91 | 1.334 | 2.938 | 113.575 | 0.872 | 0.391 | 0 | 0 | 1 | 42.312 | 1 | 35 | 1 | GJB_Rinse | 0
165 | P-1380 | 2022/10/06 20:53:49 | 2-36 | 1 | -3.212 | -20.785 | 18803 | -1 | 1 | A00000320120 | 22-245 | H2O | 1665089937 | 0.32 | 0.751 | 77.686 | 0.001 | 0.004 | 1.847 | 1300.608 | 1.698 | 3 | 115.071 | 0.87 | 0.231 | 0 | 0 | 1 | 42.312 | 1 | 36 | 1 | GJB_Rinse | 0
166 | P-1380 | 2022/10/06 21:02:46 | 2-36 | 2 | -3.305 | -20.845 | 18441 | -1 | 1 | A00000320120 | 22-245 | H2O | 1665090472 | 0.305 | 0.713 | 81.629 | 0.001 | 0.006 | 1.93 | 1298.19 | 1.444 | 2.939 | 112.54 | 0.87 | 0.134 | 0 | 0 | 1 | 42.375 | 1 | 36 | 1 | GJB_Rinse | 0
167 | P-1380 | 2022/10/06 21:11:42 | 2-36 | 3 | -3.29 | -21.203 | 18405 | -1 | 1 | A00000320120 | 22-245 | H2O | 1665091008 | 0.309 | 0.736 | 76.814 | 0.001 | 0.002 | 1.814 | 1297.859 | 1.354 | 2.932 | 113.025 | 0.87 | 0.163 | 0 | 0 | 1 | 42.375 | 1 | 36 | 1 | GJB_Rinse | 0
168 | P-1380 | 2022/10/06 21:20:38 | 2-36 | 4 | -3.318 | -21.204 | 18415 | 0 | 1 | A00000320120 | 22-245 | H2O | 1665091545 | 0.306 | 0.725 | 76.162 | 0.001 | 0.001 | 1.798 | 1297.539 | 1.409 | 2.955 | 114.624 | 0.87 | 0.498 | 0 | 0 | 1 | 42.438 | 1 | 36 | 1 | GJB_Rinse | 0
169 | P-1381 | 2022/10/06 21:30:22 | 2-37 | 1 | -2.278 | -15.919 | 18774 | -1 | 1 | A00000320050 | 22-245 | H2O | 1665092128 | 0.345 | 0.778 | 79.026 | 0 | 0.002 | 1.876 | 1300.09 | 1.886 | 3.008 | 115.891 | 0.881 | 0.468 | 0 | 0 | 1 | 42.5 | 1 | 37 | 1 | GJB_Rinse | 0
170 | P-1381 | 2022/10/06 21:39:18 | 2-37 | 2 | -2.212 | -15.693 | 18307 | -1 | 1 | A00000320050 | 22-245 | H2O | 1665092664 | 0.31 | 0.751 | 76.604 | 0.001 | 0.001 | 1.823 | 1297.071 | 1.147 | 2.948 | 113.678 | 0.875 | 0.304 | 0 | 0 | 1 | 42.5 | 1 | 37 | 1 | GJB_Rinse | 0
171 | P-1381 | 2022/10/06 21:48:14 | 2-37 | 3 | -2.178 | -15.175 | 18318 | -1 | 1 | A00000320050 | 22-245 | H2O | 1665093201 | 0.291 | 0.748 | 78.128 | 0 | 0.003 | 1.857 | 1297.591 | 0.846 | 2.903 | 112.215 | 0.876 | -0.07 | 0 | 0 | 1 | 42.5 | 1 | 37 | 1 | GJB_Rinse | 0
172 | P-1381 | 2022/10/06 21:57:10 | 2-37 | 4 | -2.179 | -15.645 | 18311 | 0 | 1 | A00000320050 | 22-245 | H2O | 1665093736 | 0.29 | 0.738 | 74.769 | 0.001 | 0.003 | 1.78 | 1296.433 | 1.616 | 3.047 | 116.503 | 0.875 | 0.997 | 0 | 0 | 1 | 42.5 | 1 | 37 | 1 | GJB_Rinse | 0
173 | P-1382 | 2022/10/06 22:06:54 | 2-38 | 1 | -1.47 | -9.086 | 18838 | -1 | 1 | A00000320064 | 22-245 | H2O | 1665094321 | 0.318 | 0.736 | 75.838 | 0 | -0.001 | 1.806 | 1300.515 | 1.657 | 3.076 | 117.565 | 0.874 | 0.602 | 0 | 0 | 1 | 42.5 | 1 | 38 | 1 | GJB_Rinse | 0
174 | P-1382 | 2022/10/06 22:15:50 | 2-38 | 2 | -1.441 | -8.774 | 18395 | -1 | 1 | A00000320064 | 22-245 | H2O | 1665094856 | 0.295 | 0.762 | 77.132 | 0.002 | -0.001 | 1.825 | 1297.457 | 1.406 | 2.957 | 114.42 | 0.875 | 0.362 | 0 | 0 | 1 | 42.562 | 1 | 38 | 1 | GJB_Rinse | 0
175 | P-1382 | 2022/10/06 22:24:46 | 2-38 | 3 | -1.396 | -8.171 | 18311 | -1 | 1 | A00000320064 | 22-245 | H2O | 1665095392 | 0.326 | 0.65 | 74.656 | 0.002 | 0 | 1.768 | 1296.802 | 1.617 | 2.957 | 113.6 | 0.873 | 0.348 | 0 | 0 | 1 | 42.5 | 1 | 38 | 1 | GJB_Rinse | 0
176 | P-1382 | 2022/10/06 22:33:42 | 2-38 | 4 | -1.463 | -8.541 | 18441 | 0 | 1 | A00000320064 | 22-245 | H2O | 1665095929 | 0.305 | 0.776 | 78.16 | 0.001 | 0 | 1.855 | 1297.87 | 1.625 | 2.958 | 113.748 | 0.867 | 0.4 | 0 | 0 | 1 | 42.5 | 1 | 38 | 1 | GJB_Rinse | 0
177 | P-1383 | 2022/10/06 22:43:26 | 2-39 | 1 | -1.851 | -11.004 | 18748 | -1 | 1 | A00000320071 | 22-245 | H2O | 1665096513 | 0.32 | 0.783 | 77.874 | 0 | 0.003 | 1.843 | 1300.2 | 1.137 | 3.066 | 117.132 | 0.878 | 0.391 | 0 | 0 | 1 | 42.438 | 1 | 39 | 1 | GJB_Rinse | 0
178 | P-1383 | 2022/10/06 22:52:23 | 2-39 | 2 | -1.921 | -11.917 | 18281 | -1 | 1 | A00000320071 | 22-245 | H2O | 1665097049 | 0.318 | 0.793 | 76.999 | 0 | 0 | 1.835 | 1296.495 | 1.965 | 2.9 | 113.072 | 0.879 | 0.409 | 0 | 0 | 1 | 42.438 | 1 | 39 | 1 | GJB_Rinse | 0
179 | P-1383 | 2022/10/06 23:01:18 | 2-39 | 3 | -2.015 | -12.434 | 18449 | -1 | 1 | A00000320071 | 22-245 | H2O | 1665097584 | 0.285 | 0.727 | 76.804 | 0.001 | 0.003 | 1.821 | 1297.902 | 1.687 | 2.946 | 113.366 | 0.878 | 0.185 | 0 | 0 | 1 | 42.438 | 1 | 39 | 1 | GJB_Rinse | 0
180 | P-1383 | 2022/10/06 23:10:14 | 2-39 | 4 | -1.936 | -12.014 | 18457 | 0 | 1 | A00000320071 | 22-245 | H2O | 1665098121 | 0.294 | 0.776 | 79.35 | 0.002 | 0.003 | 1.886 | 1298.123 | 1.308 | 2.93 | 114.082 | 0.87 | 0.261 | 0 | 0 | 1 | 42.438 | 1 | 39 | 1 | GJB_Rinse | 0
181 | P-1384 | 2022/10/06 23:19:59 | 2-40 | 1 | -7.223 | -53.526 | 18452 | -1 | 1 | EV | standard | H2O | 1665098706 | 0.302 | 0.808 | 76.588 | 0.001 | 0.005 | 1.822 | 1297.846 | 1.967 | 2.946 | 113.43 | 0.869 | 0.307 | 0 | 0 | 1 | 42.438 | 1 | 40 | 1 | GJB_Rinse | 0
182 | P-1384 | 2022/10/06 23:28:55 | 2-40 | 2 | -7.522 | -57.585 | 17888 | -1 | 1 | EV | standard | H2O | 1665099242 | 0.339 | 0.752 | 76.839 | 0.001 | 0.004 | 1.834 | 1293.916 | 1.561 | 2.866 | 111.025 | 0.872 | 0.424 | 0 | 0 | 1 | 42.5 | 1 | 40 | 1 | GJB_Rinse | 0
183 | P-1384 | 2022/10/06 23:37:51 | 2-40 | 3 | -7.469 | -57.94 | 18267 | -1 | 1 | EV | standard | H2O | 1665099777 | 0.315 | 0.793 | 77.405 | 0.001 | 0.007 | 1.842 | 1296.329 | 2.556 | 2.931 | 111.956 | 0.875 | 0.406 | 0 | 0 | 1 | 42.438 | 1 | 40 | 1 | GJB_Rinse | 0
184 | P-1384 | 2022/10/06 23:46:47 | 2-40 | 4 | -7.596 | -58.613 | 18113 | 0 | 1 | EV | standard | H2O | 1665100314 | 0.285 | 0.712 | 76.854 | 0.001 | 0.002 | 1.836 | 1295.545 | 1.751 | 2.902 | 112.193 | 0.876 | 0.452 | 0 | 0 | 1 | 42.438 | 1 | 40 | 1 | GJB_Rinse | 0
185 | P-1385 | 2022/10/06 23:56:33 | 2-41 | 1 | -2.92 | -22.067 | 18798 | -1 | 1 | A00000320106 | 22-245 | H2O | 1665100899 | 0.328 | 0.722 | 75.217 | 0 | 0.002 | 1.789 | 1300.63 | 1.667 | 2.959 | 115.544 | 0.874 | 0.258 | 0 | 0 | 1 | 42.438 | 1 | 41 | 1 | GJB_Rinse | 0
186 | P-1385 | 2022/10/07 00:05:29 | 2-41 | 2 | -2.728 | -19 | 18257 | -1 | 1 | A00000320106 | 22-245 | H2O | 1665101435 | 0.298 | 0.704 | 77.774 | 0 | 0.002 | 1.844 | 1296.725 | 1.413 | 2.963 | 113.327 | 0.873 | 0.395 | 0 | 0 | 1 | 42.438 | 1 | 41 | 1 | GJB_Rinse | 0
187 | P-1385 | 2022/10/07 00:14:25 | 2-41 | 3 | -2.716 | -17.844 | 18695 | -1 | 1 | A00000320106 | 22-245 | H2O | 1665101971 | 0.331 | 0.81 | 81.104 | 0.002 | 0.006 | 1.92 | 1299.498 | 1.908 | 3.054 | 115.599 | 0.873 | 0.616 | 0 | 0 | 1 | 42.438 | 1 | 41 | 1 | GJB_Rinse | 0
188 | P-1385 | 2022/10/07 00:23:21 | 2-41 | 4 | -2.635 | -17.861 | 18654 | 0 | 1 | A00000320106 | 22-245 | H2O | 1665102507 | 0.306 | 0.735 | 76.746 | 0 | 0 | 1.829 | 1299.589 | 1.762 | 2.957 | 114.289 | 0.879 | 0.173 | 0 | 0 | 1 | 42.5 | 1 | 41 | 1 | GJB_Rinse | 0
189 | P-1386 | 2022/10/07 00:33:06 | 2-42 | 1 | -2.396 | -16.701 | 18829 | -1 | 1 | A00000320107 | 22-245 | H2O | 1665103092 | 0.354 | 0.757 | 76.888 | 0.002 | 0.003 | 1.82 | 1300.719 | 1.42 | 3.083 | 117.107 | 0.873 | 0.487 | 0 | 0 | 1 | 42.562 | 1 | 42 | 1 | GJB_Rinse | 0
190 | P-1386 | 2022/10/07 00:42:02 | 2-42 | 2 | -2.418 | -17.129 | 18423 | -1 | 1 | A00000320107 | 22-245 | H2O | 1665103629 | 0.316 | 0.73 | 75.808 | 0 | -0.001 | 1.814 | 1297.846 | 1.619 | 2.977 | 113.408 | 0.87 | 0.371 | 0 | 0 | 1 | 42.5 | 1 | 42 | 1 | GJB_Rinse | 0
191 | P-1386 | 2022/10/07 00:50:58 | 2-42 | 3 | -2.421 | -16.988 | 18405 | -1 | 1 | A00000320107 | 22-245 | H2O | 1665104165 | 0.35 | 0.833 | 78.257 | 0.001 | 0.005 | 1.871 | 1297.325 | 1.538 | 3.014 | 115.273 | 0.872 | 0.729 | 0 | 0 | 1 | 42.5 | 1 | 42 | 1 | GJB_Rinse | 0
192 | P-1386 | 2022/10/07 00:59:54 | 2-42 | 4 | -2.469 | -17.1 | 18381 | 0 | 1 | A00000320107 | 22-245 | H2O | 1665104700 | 0.315 | 0.818 | 77.738 | 0 | 0.002 | 1.85 | 1297.238 | 1.917 | 2.966 | 113.082 | 0.87 | 0.317 | 0 | 0 | 1 | 42.438 | 1 | 42 | 1 | GJB_Rinse | 0
193 | P-1387 | 2022/10/07 01:09:38 | 2-43 | 1 | -2.39 | -16.514 | 18494 | -1 | 1 | A00000320108 | 22-245 | H2O | 1665105285 | 0.288 | 0.725 | 78.504 | 0.001 | 0 | 1.864 | 1298.109 | 1.762 | 2.939 | 114.166 | 0.876 | 0.345 | 0 | 0 | 1 | 42.25 | 1 | 43 | 1 | GJB_Rinse | 0
194 | P-1387 | 2022/10/07 01:18:34 | 2-43 | 2 | -2.116 | -15.251 | 18852 | -1 | 1 | A00000320108 | 22-245 | H2O | 1665105820 | 0.307 | 0.688 | 76.74 | 0.001 | 0 | 1.828 | 1301.067 | 1.432 | 3.021 | 116.377 | 0.874 | 0.275 | 0 | 0 | 1 | 42.312 | 1 | 43 | 1 | GJB_Rinse | 0
195 | P-1387 | 2022/10/07 01:27:30 | 2-43 | 3 | -2.298 | -16.317 | 18384 | -1 | 1 | A00000320108 | 22-245 | H2O | 1665106356 | 0.306 | 0.76 | 76.586 | 0.001 | 0.001 | 1.817 | 1297.318 | 1.947 | 2.943 | 113.5 | 0.867 | 0.406 | 0 | 0 | 1 | 42.312 | 1 | 43 | 1 | GJB_Rinse | 0
196 | P-1387 | 2022/10/07 01:36:25 | 2-43 | 4 | -2.332 | -16.639 | 18441 | 0 | 1 | A00000320108 | 22-245 | H2O | 1665106892 | 0.313 | 0.709 | 77.047 | 0.001 | 0.001 | 1.838 | 1298.219 | 1.304 | 2.942 | 113.468 | 0.877 | 0.228 | 0 | 0 | 1 | 42.438 | 1 | 43 | 1 | GJB_Rinse | 0
197 | P-1388 | 2022/10/07 01:46:10 | 2-44 | 1 | -1.368 | -6.95 | 18718 | -1 | 1 | A00000347310 | 22-245 | H2O | 1665107476 | 0.297 | 0.735 | 76.223 | 0.001 | 0.002 | 1.812 | 1300.107 | 0.989 | 2.975 | 116.173 | 0.876 | 0.401 | 0 | 0 | 1 | 42.562 | 1 | 44 | 1 | GJB_Rinse | 0
198 | P-1388 | 2022/10/07 01:55:06 | 2-44 | 2 | -1.292 | -5.717 | 18479 | -1 | 1 | A00000347310 | 22-245 | H2O | 1665108012 | 0.338 | 0.776 | 77.608 | 0.001 | 0 | 1.834 | 1298.009 | 2.258 | 2.959 | 112.952 | 0.876 | 0.299 | 0 | 0 | 1 | 42.562 | 1 | 44 | 1 | GJB_Rinse | 0
199 | P-1388 | 2022/10/07 02:04:02 | 2-44 | 3 | -1.353 | -5.844 | 18348 | -1 | 1 | A00000347310 | 22-245 | H2O | 1665108548 | 0.318 | 0.673 | 75.561 | 0.001 | 0 | 1.795 | 1297.225 | 1.86 | 2.908 | 112.325 | 0.866 | 0.194 | 0 | 0 | 1 | 42.5 | 1 | 44 | 1 | GJB_Rinse | 0
200 | P-1388 | 2022/10/07 02:12:57 | 2-44 | 4 | -1.354 | -6.143 | 18361 | 0 | 1 | A00000347310 | 22-245 | H2O | 1665109084 | 0.34 | 0.74 | 78.553 | 0.001 | 0.002 | 1.87 | 1297.209 | 1.223 | 2.983 | 114.782 | 0.869 | 0.561 | 0 | 0 | 1 | 42.375 | 1 | 44 | 1 | GJB_Rinse | 0
201 | P-1389 | 2022/10/07 02:22:43 | 2-45 | 1 | -1.25 | -6.857 | 18797 | -1 | 1 | A00000347323 | 22-245 | H2O | 1665109669 | 0.311 | 0.691 | 76.234 | 0 | 0.001 | 1.805 | 1300.253 | 2.146 | 3.023 | 115.509 | 0.874 | 0.279 | 0 | 0 | 1 | 42.375 | 1 | 45 | 1 | GJB_Rinse | 0
202 | P-1389 | 2022/10/07 02:31:38 | 2-45 | 2 | -1.326 | -7.499 | 17971 | -1 | 1 | A00000347323 | 22-245 | H2O | 1665110205 | 0.362 | 0.782 | 78.201 | 0.001 | 0 | 1.862 | 1294.432 | 1.939 | 2.876 | 110.234 | 0.869 | 0.326 | 0 | 0 | 1 | 42.312 | 1 | 45 | 1 | GJB_Rinse | 0
203 | P-1389 | 2022/10/07 02:40:34 | 2-45 | 3 | -1.262 | -7.039 | 18435 | -1 | 1 | A00000347323 | 22-245 | H2O | 1665110741 | 0.299 | 0.687 | 76.361 | 0 | 0.002 | 1.823 | 1297.561 | 2.176 | 2.972 | 113.622 | 0.872 | 0.599 | 0 | 0 | 1 | 42.312 | 1 | 45 | 1 | GJB_Rinse | 0
204 | P-1389 | 2022/10/07 02:49:30 | 2-45 | 4 | -1.269 | -7.1 | 18431 | 0 | 1 | A00000347323 | 22-245 | H2O | 1665111276 | 0.284 | 0.748 | 79.12 | 0.001 | 0.002 | 1.88 | 1297.976 | 1.327 | 2.968 | 114.24 | 0.87 | 0.362 | 0 | 0 | 1 | 42.375 | 1 | 45 | 1 | GJB_Rinse | 0
205 | P-1390 | 2022/10/07 02:59:16 | 2-46 | 1 | -2.593 | -16.354 | 18503 | -1 | 1 | A00000347301 | 22-245 | H2O | 1665111862 | 0.327 | 0.729 | 77.031 | 0 | 0.002 | 1.827 | 1298.446 | 1.986 | 2.94 | 112.885 | 0.871 | 0.163 | 0 | 0 | 1 | 42.438 | 1 | 46 | 1 | GJB_Rinse | 0
206 | P-1390 | 2022/10/07 03:08:11 | 2-46 | 2 | -2.702 | -17.466 | 17859 | -1 | 1 | A00000347301 | 22-245 | H2O | 1665112398 | 0.315 | 0.704 | 77.753 | 0 | 0.003 | 1.846 | 1293.672 | 1.599 | 2.842 | 110.961 | 0.872 | 0.362 | 0 | 0 | 1 | 42.438 | 1 | 46 | 1 | GJB_Rinse | 0
207 | P-1390 | 2022/10/07 03:17:07 | 2-46 | 3 | -2.678 | -17.32 | 18548 | -1 | 1 | A00000347301 | 22-245 | H2O | 1665112934 | 0.323 | 0.674 | 77.051 | 0.001 | 0.002 | 1.83 | 1298.739 | 1.914 | 2.939 | 113.36 | 0.871 | 0.334 | 0 | 0 | 1 | 42.438 | 1 | 46 | 1 | GJB_Rinse | 0
208 | P-1390 | 2022/10/07 03:26:03 | 2-46 | 4 | -2.73 | -17.719 | 18233 | 0 | 1 | A00000347301 | 22-245 | H2O | 1665113470 | 0.275 | 0.787 | 75.651 | 0 | 0.001 | 1.789 | 1296.365 | 1.242 | 2.925 | 113.525 | 0.875 | 0.485 | 0 | 0 | 1 | 42.438 | 1 | 46 | 1 | GJB_Rinse | 0
209 | P-1391 | 2022/10/07 03:35:49 | 2-47 | 1 | -2.7 | -17.524 | 18863 | -1 | 1 | A00000347307 | 22-245 | H2O | 1665114056 | 0.288 | 0.712 | 78.061 | 0.001 | 0.001 | 1.853 | 1301.139 | 0.95 | 3.066 | 118.207 | 0.876 | 0.6 | 0 | 0 | 1 | 42.438 | 1 | 47 | 1 | GJB_Rinse | 0
210 | P-1391 | 2022/10/07 03:44:46 | 2-47 | 2 | -2.838 | -18.625 | 18026 | -1 | 1 | A00000347307 | 22-245 | H2O | 1665114592 | 0.316 | 0.754 | 75.895 | 0.001 | 0 | 1.797 | 1295.001 | 0.888 | 2.918 | 113.184 | 0.873 | 0.621 | 0 | 0 | 1 | 42.438 | 1 | 47 | 1 | GJB_Rinse | 0
211 | P-1391 | 2022/10/07 03:53:41 | 2-47 | 3 | -2.718 | -18.192 | 18436 | -1 | 1 | A00000347307 | 22-245 | H2O | 1665115128 | 0.32 | 0.707 | 74.651 | 0 | 0 | 1.781 | 1298.009 | 1.628 | 2.954 | 113.184 | 0.871 | 0.402 | 0 | 0 | 1 | 42.438 | 1 | 47 | 1 | GJB_Rinse | 0
212 | P-1391 | 2022/10/07 04:02:37 | 2-47 | 4 | -2.727 | -18.241 | 18419 | 0 | 1 | A00000347307 | 22-245 | H2O | 1665115664 | 0.31 | 0.755 | 74.797 | 0 | 0.002 | 1.787 | 1297.722 | 1.691 | 2.975 | 113.716 | 0.871 | 0.38 | 0 | 0 | 1 | 42.812 | 1 | 47 | 1 | GJB_Rinse | 0
213 | P-1392 | 2022/10/07 04:12:24 | 2-48 | 1 | -2.187 | -14.029 | 18611 | -1 | 1 | A00000347306 | 22-245 | H2O | 1665116251 | 0.312 | 0.655 | 77.673 | 0 | 0.002 | 1.847 | 1299.028 | 1.867 | 2.988 | 114.754 | 0.869 | 0.445 | 0 | 0 | 1 | 42.5 | 1 | 48 | 1 | GJB_Rinse | 0
214 | P-1392 | 2022/10/07 04:21:20 | 2-48 | 2 | -2.131 | -13.561 | 17967 | -1 | 1 | A00000347306 | 22-245 | H2O | 1665116786 | 0.295 | 0.7 | 77.168 | 0 | 0.003 | 1.838 | 1294.794 | 1.636 | 2.869 | 109.884 | 0.869 | 0.188 | 0 | 0 | 1 | 42.5 | 1 | 48 | 1 | GJB_Rinse | 0
215 | P-1392 | 2022/10/07 04:30:16 | 2-48 | 3 | -2.053 | -13.231 | 18624 | -1 | 1 | A00000347306 | 22-245 | H2O | 1665117323 | 0.315 | 0.715 | 76.695 | 0.001 | -0.001 | 1.824 | 1299.415 | 1.812 | 2.965 | 113.885 | 0.869 | 0.143 | 0 | 0 | 1 | 42.438 | 1 | 48 | 1 | GJB_Rinse | 0
216 | P-1392 | 2022/10/07 04:39:12 | 2-48 | 4 | -2.1 | -13.606 | 18467 | 0 | 1 | A00000347306 | 22-245 | H2O | 1665117859 | 0.322 | 0.752 | 76.839 | 0.001 | 0 | 1.82 | 1298.13 | 1.514 | 3.022 | 114.933 | 0.872 | 0.571 | 0 | 0 | 1 | 42.375 | 1 | 48 | 1 | GJB_Rinse | 0
217 | P-1393 | 2022/10/07 04:48:59 | 2-49 | 1 | -7.192 | -53.577 | 18444 | -1 | 1 | EV | standard | H2O | 1665118446 | 0.318 | 0.755 | 78.715 | 0.001 | 0.003 | 1.86 | 1298.209 | 1.71 | 2.943 | 112.899 | 0.875 | 0.292 | 0 | 0 | 1 | 42.375 | 1 | 49 | 1 | GJB_Rinse | 0
218 | P-1393 | 2022/10/07 04:57:55 | 2-49 | 2 | -7.476 | -57.22 | 18016 | -1 | 1 | EV | standard | H2O | 1665118982 | 0.336 | 0.731 | 77.91 | 0.001 | 0.004 | 1.85 | 1294.997 | 1.679 | 2.879 | 111.004 | 0.871 | 0.247 | 0 | 0 | 1 | 42.438 | 1 | 49 | 1 | GJB_Rinse | 0
219 | P-1393 | 2022/10/07 05:06:51 | 2-49 | 3 | -7.532 | -58.151 | 18371 | -1 | 1 | EV | standard | H2O | 1665119517 | 0.302 | 0.742 | 77.955 | 0 | 0.006 | 1.854 | 1297.648 | 1.501 | 2.94 | 113.056 | 0.87 | 0.241 | 0 | 0 | 1 | 42.375 | 1 | 49 | 1 | GJB_Rinse | 0
220 | P-1393 | 2022/10/07 05:15:47 | 2-49 | 4 | -7.51 | -58.302 | 18385 | 0 | 1 | EV | standard | H2O | 1665120053 | 0.296 | 0.74 | 77.26 | 0.002 | 0.005 | 1.834 | 1297.541 | 1.945 | 2.958 | 113.002 | 0.872 | 0.284 | 0 | 0 | 1 | 42.312 | 1 | 49 | 1 | GJB_Rinse | 0
221 | P-1394 | 2022/10/07 05:25:31 | 2-50 | 1 | -3.058 | -22.99 | 18821 | -1 | 1 | A00000347300 | 22-245 | H2O | 1665120638 | 0.292 | 0.745 | 79.337 | 0.001 | 0.002 | 1.879 | 1300.724 | 1.554 | 3.061 | 116.755 | 0.868 | 0.475 | 0 | 0 | 1 | 42.312 | 1 | 50 | 1 | GJB_Rinse | 0
222 | P-1394 | 2022/10/07 05:34:27 | 2-50 | 2 | -2.984 | -20.835 | 17319 | -1 | 1 | A00000347300 | 22-245 | H2O | 1665121173 | 0.322 | 0.791 | 75.852 | 0 | 0.001 | 1.815 | 1289.995 | 1.408 | 2.759 | 106.932 | 0.87 | 0.248 | 0 | 0 | 1 | 42.375 | 1 | 50 | 1 | GJB_Rinse | 0
223 | P-1394 | 2022/10/07 05:43:23 | 2-50 | 3 | -2.859 | -19.833 | 17694 | -1 | 1 | A00000347300 | 22-245 | H2O | 1665121709 | 0.329 | 0.707 | 77.363 | 0.001 | 0.003 | 1.837 | 1292.703 | 1.59 | 2.85 | 109.432 | 0.866 | 0.327 | 0 | 0 | 1 | 42.375 | 1 | 50 | 1 | GJB_Rinse | 0
224 | P-1394 | 2022/10/07 05:52:19 | 2-50 | 4 | -2.886 | -19.634 | 17613 | 0 | 1 | A00000347300 | 22-245 | H2O | 1665122245 | 0.303 | 0.757 | 75.453 | -0.001 | -0.001 | 1.784 | 1292.304 | 1.455 | 2.789 | 108.196 | 0.87 | 0.033 | 0 | 0 | 1 | 42.438 | 1 | 50 | 1 | GJB_Rinse | 0
225 | P-1395 | 2022/10/07 06:02:04 | 2-51 | 1 | -3.23 | -23.023 | 17733 | -1 | 1 | A00000319827 | 22-245 | H2O | 1665122830 | 0.265 | 0.743 | 77.213 | 0 | 0.002 | 1.829 | 1292.684 | 1.793 | 2.867 | 110.109 | 0.863 | 0.631 | 0 | 0 | 1 | 42.312 | 1 | 51 | 1 | GJB_Rinse | 0
226 | P-1395 | 2022/10/07 06:11:00 | 2-51 | 2 | -3.311 | -23.466 | 17756 | -1 | 1 | A00000319827 | 22-245 | H2O | 1665123366 | 0.326 | 0.711 | 75.862 | 0.002 | 0 | 1.794 | 1293.102 | 1.624 | 2.852 | 109.857 | 0.868 | 0.222 | 0 | 0 | 1 | 42.312 | 1 | 51 | 1 | GJB_Rinse | 0
227 | P-1395 | 2022/10/07 06:19:56 | 2-51 | 3 | -3.253 | -23.377 | 17829 | -1 | 1 | A00000319827 | 22-245 | H2O | 1665123902 | 0.344 | 0.765 | 74.725 | 0.001 | 0.004 | 1.779 | 1293.646 | 1.229 | 2.868 | 110.781 | 0.874 | 0.389 | 0 | 0 | 1 | 42.25 | 1 | 51 | 1 | GJB_Rinse | 0
228 | P-1395 | 2022/10/07 06:28:51 | 2-51 | 4 | -3.301 | -23.694 | 17436 | 0 | 1 | A00000319827 | 22-245 | H2O | 1665124438 | 0.301 | 0.749 | 74.352 | 0 | 0.004 | 1.773 | 1290.847 | 1.479 | 2.803 | 107.389 | 0.872 | 0.287 | 0 | 0 | 1 | 42.312 | 1 | 51 | 1 | GJB_Rinse | 0
229 | P-1396 | 2022/10/07 06:38:37 | 2-52 | 1 | -4.134 | -34.299 | 18552 | -1 | 1 | A00000319824 | 22-245 | H2O | 1665125024 | 0.315 | 0.862 | 77.49 | 0.001 | 0.003 | 1.841 | 1298.581 | 1.757 | 3.064 | 116.398 | 0.865 | 0.671 | 0 | 0 | 1 | 42.375 | 1 | 52 | 1 | GJB_Rinse | 0
230 | P-1396 | 2022/10/07 06:47:33 | 2-52 | 2 | -4.284 | -36.064 | 17630 | -1 | 1 | A00000319824 | 22-245 | H2O | 1665125560 | 0.301 | 0.74 | 75.623 | 0 | 0.003 | 1.796 | 1292.253 | 1.254 | 2.863 | 110.274 | 0.866 | 0.457 | 0 | 0 | 1 | 42.312 | 1 | 52 | 1 | GJB_Rinse | 0
231 | P-1396 | 2022/10/07 06:56:29 | 2-52 | 3 | -4.226 | -35.417 | 18062 | -1 | 1 | A00000319824 | 22-245 | H2O | 1665126096 | 0.271 | 0.777 | 78.184 | 0.001 | 0.002 | 1.846 | 1295.506 | 1.493 | 2.878 | 111.723 | 0.866 | 0.216 | 0 | 0 | 1 | 42.375 | 1 | 52 | 1 | GJB_Rinse | 0
232 | P-1396 | 2022/10/07 07:05:25 | 2-52 | 4 | -4.348 | -36.504 | 17746 | 0 | 1 | A00000319824 | 22-245 | H2O | 1665126633 | 0.327 | 0.66 | 77.259 | 0.002 | 0.003 | 1.831 | 1293.202 | 1.496 | 2.827 | 109.428 | 0.867 | 0.319 | 0 | 0 | 1 | 42.312 | 1 | 52 | 1 | GJB_Rinse | 0
233 | P-1397 | 2022/10/07 07:15:12 | 2-53 | 1 | -1.593 | -9.214 | 18700 | -1 | 1 | A00000347227 | 22-245 | H2O | 1665127219 | 0.362 | 0.706 | 78.903 | 0.001 | 0 | 1.862 | 1300.183 | 1.328 | 2.985 | 115.028 | 0.871 | 0.168 | 0 | 0 | 1 | 42.25 | 1 | 53 | 1 | GJB_Rinse | 0
234 | P-1397 | 2022/10/07 07:24:07 | 2-53 | 2 | -1.612 | -7.29 | 17991 | -1 | 1 | A00000347227 | 22-245 | H2O | 1665127753 | 0.326 | 0.752 | 76.902 | 0.001 | 0.001 | 1.835 | 1294.881 | 1.499 | 2.842 | 110.822 | 0.866 | 0.3 | 0 | 0 | 1 | 42.312 | 1 | 53 | 1 | GJB_Rinse | 0
235 | P-1397 | 2022/10/07 07:33:03 | 2-53 | 3 | -1.499 | -6.402 | 18379 | -1 | 1 | A00000347227 | 22-245 | H2O | 1665128290 | 0.292 | 0.658 | 75.552 | 0.001 | -0.001 | 1.798 | 1297.651 | 1.662 | 2.958 | 113.389 | 0.864 | 0.267 | 0 | 0 | 1 | 42.125 | 1 | 53 | 1 | GJB_Rinse | 0
236 | P-1397 | 2022/10/07 07:41:59 | 2-53 | 4 | -1.49 | -6.228 | 18250 | 0 | 1 | A00000347227 | 22-245 | H2O | 1665128825 | 0.306 | 0.702 | 77.206 | 0.001 | 0.002 | 1.823 | 1296.923 | 1.535 | 2.874 | 111.55 | 0.87 | 0.184 | 0 | 0 | 1 | 42.188 | 1 | 53 | 1 | GJB_Rinse | 0
237 | P-1398 | 2022/10/07 07:51:46 | 2-54 | 1 | 0.476 | 14.768 | 18509 | -1 | 1 | A00000347481 | 22-245 | H2O | 1665129413 | 0.323 | 0.798 | 77.178 | 0 | -0.001 | 1.836 | 1298.255 | 1.76 | 2.997 | 114.856 | 0.87 | 0.55 | 0 | 0 | 1 | 42.438 | 1 | 54 | 1 | GJB_Rinse | 0
238 | P-1398 | 2022/10/07 08:00:42 | 2-54 | 2 | 0.464 | 16.122 | 17642 | -1 | 1 | A00000347481 | 22-245 | H2O | 1665129948 | 0.281 | 0.687 | 74.742 | -0.001 | -0.001 | 1.789 | 1292.528 | 1.086 | 2.795 | 109.213 | 0.875 | 0.128 | 0 | 0 | 1 | 42.438 | 1 | 54 | 1 | GJB_Rinse | 0
239 | P-1398 | 2022/10/07 08:09:38 | 2-54 | 3 | 0.557 | 16.899 | 17987 | -1 | 1 | A00000347481 | 22-245 | H2O | 1665130484 | 0.317 | 0.719 | 75.619 | 0.002 | 0.003 | 1.797 | 1294.766 | 1.262 | 2.879 | 112.027 | 0.865 | 0.419 | 0 | 0 | 1 | 42.375 | 1 | 54 | 1 | GJB_Rinse | 0
240 | P-1398 | 2022/10/07 08:18:34 | 2-54 | 4 | 0.511 | 16.771 | 17558 | 0 | 1 | A00000347481 | 22-245 | H2O | 1665131021 | 0.314 | 0.816 | 83.952 | 0.001 | 0.007 | 1.987 | 1291.753 | 1.492 | 2.807 | 108.349 | 0.869 | 0.309 | 0 | 0 | 1 | 42.5 | 1 | 54 | 1 | GJB_Rinse | 0
241 | P-1399 | 2022/10/07 08:28:21 | 2-55 | 1 | 1.708 | 23.244 | 18500 | -1 | 1 | A00000347479 | 22-245 | H2O | 1665131608 | 0.317 | 0.679 | 78.296 | 0 | 0.001 | 1.859 | 1298.412 | 1.458 | 2.95 | 114.134 | 0.866 | 0.341 | 0 | 0 | 1 | 42.75 | 1 | 55 | 1 | GJB_Rinse | 0
242 | P-1399 | 2022/10/07 08:37:17 | 2-55 | 2 | 1.712 | 23.419 | 17763 | -1 | 1 | A00000347479 | 22-245 | H2O | 1665132143 | 0.292 | 0.701 | 76.852 | 0.001 | 0 | 1.825 | 1293.421 | 0.861 | 2.839 | 109.6 | 0.876 | 0.131 | 0 | 0 | 1 | 42.812 | 1 | 55 | 1 | GJB_Rinse | 0
243 | P-1399 | 2022/10/07 08:46:13 | 2-55 | 3 | 1.669 | 23.87 | 18081 | -1 | 1 | A00000347479 | 22-245 | H2O | 1665132680 | 0.337 | 0.698 | 78.536 | 0.001 | 0 | 1.866 | 1295.507 | 1.768 | 2.847 | 110.051 | 0.87 | 0.114 | 0 | 0 | 1 | 42.812 | 1 | 55 | 1 | GJB_Rinse | 0
244 | P-1399 | 2022/10/07 08:55:09 | 2-55 | 4 | 1.756 | 23.96 | 18261 | 0 | 1 | A00000347479 | 22-245 | H2O | 1665133215 | 0.314 | 0.777 | 73.344 | 0.001 | -0.001 | 1.743 | 1296.586 | 1.538 | 2.981 | 112.945 | 0.865 | 0.352 | 0 | 0 | 1 | 42.812 | 1 | 55 | 1 | GJB_Rinse | 0
245 | P-1400 | 2022/10/07 09:04:56 | 2-56 | 1 | -2.767 | -12.421 | 18318 | -1 | 1 | A00000319825 | 22-245 | H2O | 1665133803 | 0.313 | 0.763 | 77.101 | 0.001 | 0.001 | 1.832 | 1297.239 | 1.775 | 2.899 | 111.83 | 0.87 | 0.153 | 0 | 0 | 1 | 42.688 | 1 | 56 | 1 | GJB_Rinse | 0
246 | P-1400 | 2022/10/07 09:13:52 | 2-56 | 2 | -3.108 | -16.323 | 17644 | -1 | 1 | A00000319825 | 22-245 | H2O | 1665134339 | 0.328 | 0.784 | 76.704 | 0.001 | 0.003 | 1.832 | 1292.055 | 1.714 | 2.833 | 109.028 | 0.866 | 0.422 | 0 | 0 | 1 | 42.5 | 1 | 56 | 1 | GJB_Rinse | 0
247 | P-1400 | 2022/10/07 09:22:49 | 2-56 | 3 | -3.152 | -16.979 | 18122 | -1 | 1 | A00000319825 | 22-245 | H2O | 1665134875 | 0.295 | 0.761 | 77.659 | 0 | 0.002 | 1.841 | 1295.649 | 1.335 | 2.875 | 112.574 | 0.863 | 0.375 | 0 | 0 | 1 | 42.438 | 1 | 56 | 1 | GJB_Rinse | 0
248 | P-1400 | 2022/10/07 09:31:44 | 2-56 | 4 | -3.076 | -16.97 | 18306 | 0 | 1 | A00000319825 | 22-245 | H2O | 1665135411 | 0.323 | 0.699 | 74.749 | 0 | 0.003 | 1.776 | 1297.082 | 1.447 | 2.95 | 113.331 | 0.865 | 0.395 | 0 | 0 | 1 | 42.188 | 1 | 56 | 1 | GJB_Rinse | 0
249 | P-1401 | 2022/10/07 09:41:30 | 2-57 | 1 | -0.101 | 7.477 | 18396 | -1 | 1 | A00000319822 | 22-245 | H2O | 1665135997 | 0.285 | 0.729 | 76.446 | 0.001 | 0.002 | 1.817 | 1297.685 | 0.926 | 2.948 | 114.552 | 0.866 | 0.348 | 0 | 0 | 1 | 42.125 | 1 | 57 | 1 | GJB_Rinse | 0
250 | P-1401 | 2022/10/07 09:50:26 | 2-57 | 2 | -0.045 | 9.222 | 17699 | -1 | 1 | A00000319822 | 22-245 | H2O | 1665136533 | 0.313 | 0.884 | 74.338 | 0.001 | 0.001 | 1.773 | 1292.476 | 1.952 | 2.855 | 109.066 | 0.862 | 0.35 | 0 | 0 | 1 | 42.188 | 1 | 57 | 1 | GJB_Rinse | 0
251 | P-1401 | 2022/10/07 09:59:22 | 2-57 | 3 | 0 | 10.052 | 18179 | -1 | 1 | A00000319822 | 22-245 | H2O | 1665137069 | 0.336 | 0.796 | 77.907 | 0.001 | 0.002 | 1.849 | 1296.2 | 1.305 | 2.902 | 112.346 | 0.877 | 0.275 | 0 | 0 | 1 | 42.25 | 1 | 57 | 1 | GJB_Rinse | 0
252 | P-1401 | 2022/10/07 10:08:18 | 2-57 | 4 | 0.055 | 10.581 | 18081 | 0 | 1 | A00000319822 | 22-245 | H2O | 1665137604 | 0.331 | 0.775 | 75.747 | 0.002 | 0.003 | 1.809 | 1295.342 | 1.235 | 2.894 | 112.578 | 0.871 | 0.398 | 0 | 0 | 1 | 42.188 | 1 | 57 | 1 | GJB_Rinse | 0
253 | P-1402 | 2022/10/07 10:18:04 | 2-58 | 1 | -7.052 | -50.308 | 18434 | -1 | 1 | EV | standard | H2O | 1665138191 | 0.294 | 0.746 | 76.974 | 0.002 | 0.006 | 1.824 | 1297.843 | 2.122 | 2.92 | 112.848 | 0.867 | 0.408 | 0 | 0 | 1 | 42.125 | 1 | 58 | 1 | GJB_Rinse | 0
254 | P-1402 | 2022/10/07 10:27:00 | 2-58 | 2 | -7.476 | -56.79 | 17792 | -1 | 1 | EV | standard | H2O | 1665138726 | 0.313 | 0.76 | 75.073 | 0.001 | 0.006 | 1.778 | 1293.242 | 1.972 | 2.788 | 109.096 | 0.863 | 0.106 | 0 | 0 | 1 | 42.188 | 1 | 58 | 1 | GJB_Rinse | 0
255 | P-1402 | 2022/10/07 10:35:56 | 2-58 | 3 | -7.548 | -57.936 | 18069 | -1 | 1 | EV | standard | H2O | 1665139261 | 0.336 | 0.735 | 73.84 | 0 | 0.003 | 1.789 | 1295.208 | 2.173 | 2.87 | 110.258 | 0.864 | 0.37 | 0 | 0 | 1 | 42.25 | 1 | 58 | 1 | GJB_Rinse | 0
256 | P-1402 | 2022/10/07 10:44:51 | 2-58 | 4 | -7.524 | -58.249 | 18221 | 0 | 1 | EV | standard | H2O | 1665139798 | 0.314 | 0.784 | 75.174 | 0.001 | 0.006 | 1.788 | 1296.338 | 1.71 | 2.913 | 112.373 | 0.864 | 0.267 | 0 | 0 | 1 | 42.125 | 1 | 58 | 1 | GJB_Rinse | 0
257 | P-1403 | 2022/10/07 10:54:38 | 2-59 | 1 | -0.902 | -5.226 | 18643 | -1 | 1 | A00000347483 | 22-245 | H2O | 1665140385 | 0.296 | 0.691 | 76.999 | 0.001 | -0.001 | 1.82 | 1300.231 | 0.963 | 2.955 | 114.412 | 0.868 | -0.107 | 0 | 0 | 1 | 42.062 | 1 | 59 | 1 | GJB_Rinse | 0
258 | P-1403 | 2022/10/07 11:03:34 | 2-59 | 2 | -0.775 | -1.791 | 17678 | -1 | 1 | A00000347483 | 22-245 | H2O | 1665140921 | 0.313 | 0.722 | 76.435 | 0.001 | -0.001 | 1.806 | 1292.902 | 1.572 | 2.861 | 108.866 | 0.864 | 0.258 | 0 | 0 | 1 | 42.062 | 1 | 59 | 1 | GJB_Rinse | 0
259 | P-1403 | 2022/10/07 11:12:30 | 2-59 | 3 | -0.689 | -0.726 | 17943 | -1 | 1 | A00000347483 | 22-245 | H2O | 1665141456 | 0.27 | 0.799 | 77.118 | 0 | 0 | 1.835 | 1294.953 | 1.217 | 2.867 | 111.115 | 0.864 | 0.198 | 0 | 0 | 1 | 42.188 | 1 | 59 | 1 | GJB_Rinse | 0
260 | P-1403 | 2022/10/07 11:21:26 | 2-59 | 4 | -0.514 | 0.358 | 18254 | 0 | 1 | A00000347483 | 22-245 | H2O | 1665141992 | 0.326 | 0.666 | 78.852 | 0.001 | 0.002 | 1.876 | 1297.41 | 0.988 | 2.916 | 112.745 | 0.864 | 0.078 | 0 | 0 | 1 | 42.25 | 1 | 59 | 1 | GJB_Rinse | 0
261 | P-1404 | 2022/10/07 11:31:13 | 2-60 | 1 | 0.791 | -2.923 | 17297 | -1 | 1 | TALL4-C-4 | 22-233 | H2O | 1665142579 | 0.359 | 1.186 | 74.35 | 0 | 0.002 | 1.78 | 2078.645 | -27.628 | 2.747 | 100.566 | 0.869 | -0.866 | 0 | 0 | 1 | 42.312 | 1 | 60 | 1 | GJB_Rinse | 0
262 | P-1404 | 2022/10/07 11:40:08 | 2-60 | 2 | 0.875 | -2.977 | 17283 | -1 | 1 | TALL4-C-4 | 22-233 | H2O | 1665143115 | 0.368 | 1.138 | 76.068 | 0 | 0 | 1.814 | 2087.595 | -28.461 | 2.758 | 100.423 | 0.879 | -0.82 | 0 | 0 | 1 | 42.375 | 1 | 60 | 1 | GJB_Rinse | 0
263 | P-1404 | 2022/10/07 11:49:04 | 2-60 | 3 | 0.945 | -2.709 | 17765 | -1 | 1 | TALL4-C-4 | 22-233 | H2O | 1665143651 | 0.399 | 1.165 | 72.153 | -0.002 | -0.001 | 1.707 | 2115.845 | -29.259 | 2.873 | 104.907 | 0.873 | -0.686 | 0 | 0 | 1 | 42.438 | 1 | 60 | 1 | GJB_Rinse | 0
264 | P-1404 | 2022/10/07 11:58:00 | 2-60 | 4 | 1.032 | -2.489 | 17750 | 0 | 1 | TALL4-C-4 | 22-233 | H2O | 1665144186 | 0.382 | 1.16 | 74.575 | 0.001 | 0.003 | 1.774 | 2117.16 | -28.75 | 2.81 | 103.382 | 0.874 | -0.856 | 0 | 0 | 1 | 42.375 | 1 | 60 | 1 | GJB_Rinse | 0
265 | P-1405 | 2022/10/07 12:07:48 | 2-61 | 1 | -6.158 | -37.606 | 17405 | -1 | 1 | EBC 02 | 22-231 | H2O | 1665144774 | 0.28 | 0.834 | 73.975 | 0.001 | 0.005 | 1.752 | 1303.878 | 1.542 | 2.742 | 107.536 | 0.862 | 0.262 | 0 | 0 | 1 | 42.25 | 1 | 61 | 1 | GJB_Rinse | 0
266 | P-1405 | 2022/10/07 12:16:44 | 2-61 | 2 | -6.399 | -40.677 | 17996 | -1 | 1 | EBC 02 | 22-231 | H2O | 1665145310 | 0.334 | 0.763 | 75.659 | 0.002 | 0.004 | 1.795 | 1298.745 | 1.523 | 2.893 | 112.127 | 0.866 | 0.474 | 0 | 0 | 1 | 42.125 | 1 | 61 | 1 | GJB_Rinse | 0
267 | P-1405 | 2022/10/07 12:25:39 | 2-61 | 3 | -6.505 | -41.426 | 17947 | -1 | 1 | EBC 02 | 22-231 | H2O | 1665145846 | 0.297 | 0.663 | 76.27 | 0 | 0.001 | 1.82 | 1297.097 | 1.284 | 2.839 | 110.876 | 0.862 | 0.26 | 0 | 0 | 1 | 42.188 | 1 | 61 | 1 | GJB_Rinse | 0
268 | P-1405 | 2022/10/07 12:34:35 | 2-61 | 4 | -6.427 | -41.421 | 18512 | 0 | 1 | EBC 02 | 22-231 | H2O | 1665146382 | 0.316 | 0.775 | 76.764 | 0.001 | 0.006 | 1.823 | 1300.123 | 1.988 | 2.941 | 114.318 | 0.861 | 0.363 | 0 | 0 | 1 | 42.062 | 1 | 61 | 1 | GJB_Rinse | 0
269 | P-1406 | 2022/10/07 12:44:23 | 2-62 | 1 | -8.18 | -58.646 | 17676 | -1 | 1 | EBC 05 | 22-231 | H2O | 1665146970 | 0.299 | 0.799 | 75.822 | 0.001 | 0.004 | 1.813 | 1293.456 | 1.537 | 2.785 | 108.859 | 0.861 | 0.308 | 0 | 0 | 1 | 41.875 | 1 | 62 | 1 | GJB_Rinse | 0
270 | P-1406 | 2022/10/07 12:53:19 | 2-62 | 2 | -8.228 | -60.001 | 18320 | -1 | 1 | EBC 05 | 22-231 | H2O | 1665147505 | 0.321 | 0.718 | 73.395 | 0.002 | 0.002 | 1.751 | 1298.044 | 2.408 | 2.867 | 111.308 | 0.866 | 0.132 | 0 | 0 | 1 | 41.938 | 1 | 62 | 1 | GJB_Rinse | 0
271 | P-1406 | 2022/10/07 13:02:15 | 2-62 | 3 | -8.166 | -59.998 | 18456 | -1 | 1 | EBC 05 | 22-231 | H2O | 1665148042 | 0.293 | 0.691 | 75.86 | 0.001 | 0.002 | 1.797 | 1298.807 | 2.434 | 2.914 | 112.698 | 0.868 | 0.4 | 0 | 0 | 1 | 42 | 1 | 62 | 1 | GJB_Rinse | 0
272 | P-1406 | 2022/10/07 13:11:10 | 2-62 | 4 | -8.27 | -60.862 | 18623 | 0 | 1 | EBC 05 | 22-231 | H2O | 1665148577 | 0.312 | 0.719 | 77.113 | 0 | 0.003 | 1.826 | 1300.513 | 1.544 | 2.902 | 113.233 | 0.868 | 0.054 | 0 | 0 | 1 | 42.062 | 1 | 62 | 1 | GJB_Rinse | 0
273 | P-1407 | 2022/10/07 13:20:58 | 2-63 | 1 | -4.752 | -27.864 | 17890 | -1 | 1 | EBC 01 | 22-231 | H2O | 1665149164 | 0.312 | 0.599 | 76.885 | 0.001 | 0.003 | 1.826 | 1294.448 | 1.504 | 2.879 | 110.469 | 0.87 | 0.409 | 0 | 0 | 1 | 42.125 | 1 | 63 | 1 | GJB_Rinse | 0
274 | P-1407 | 2022/10/07 13:29:54 | 2-63 | 2 | -3.922 | -22.421 | 18869 | -1 | 1 | EBC 01 | 22-231 | H2O | 1665149701 | 0.323 | 0.715 | 78.533 | 0.002 | 0.001 | 1.865 | 1301.469 | 1.76 | 3.06 | 116.809 | 0.871 | 0.391 | 0 | 0 | 1 | 42.25 | 1 | 63 | 1 | GJB_Rinse | 0
275 | P-1407 | 2022/10/07 13:38:50 | 2-63 | 3 | -4.385 | -23.321 | 18317 | -1 | 1 | EBC 01 | 22-231 | H2O | 1665150236 | 0.299 | 0.748 | 78.067 | 0.001 | 0 | 1.857 | 1297.648 | 1.5 | 2.936 | 112.52 | 0.875 | 0.333 | 0 | 0 | 1 | 42.188 | 1 | 63 | 1 | GJB_Rinse | 0
276 | P-1407 | 2022/10/07 13:47:46 | 2-63 | 4 | -4.378 | -23.028 | 18342 | 0 | 1 | EBC 01 | 22-231 | H2O | 1665150772 | 0.3 | 0.702 | 76 | 0.001 | 0.002 | 1.802 | 1297.625 | 1.174 | 2.973 | 114.652 | 0.871 | 0.582 | 0 | 0 | 1 | 42.25 | 1 | 63 | 1 | GJB_Rinse | 0
277 | P-1408 | 2022/10/07 13:57:32 | 2-64 | 1 | 1.3 | 25.258 | 18528 | 0 | 1 | EBC 04 | 22-231 | H2O | 1665152966 | 0.292 | 0.721 | 79.015 | 0 | 0 | 1.876 | 1298.982 | 0.99 | 2.987 | 114.277 | 0.871 | 0.119 | 0 | 0 | 1 | 42.188 | 1 | 64 | 1 | GJB_Rinse | 0
278 | P-1408 | 2022/10/07 14:06:27 | 2-64 | 2 | 1.527 | 25.23 | 20014 | -1 | 1 | EBC 04 | 22-231 | H2O | 1665151894 | 0.328 | 0.778 | 79.905 | 0.001 | -0.001 | 1.882 | 1309.888 | 1.745 | 3.202 | 121.983 | 0.873 | 0.368 | 0 | 0 | 1 | 42.25 | 1 | 64 | 1 | GJB_Rinse | 0
279 | P-1408 | 2022/10/07 14:15:23 | 2-64 | 3 | 1.248 | 24.489 | 18233 | -1 | 1 | EBC 04 | 22-231 | H2O | 1665152430 | 0.288 | 0.701 | 77.536 | 0 | -0.001 | 1.841 | 1296.857 | 1.251 | 2.916 | 112.042 | 0.867 | 0.275 | 0 | 0 | 1 | 42.25 | 1 | 64 | 1 | GJB_Rinse | 0
280 | P-1408 | 2022/10/07 14:24:19 | 2-64 | 4 | 1.3 | 25.258 | 18528 | 0 | 1 | EBC 04 | 22-231 | H2O | 1665152966 | 0.292 | 0.721 | 79.015 | 0 | 0 | 1.876 | 1298.982 | 0.99 | 2.987 | 114.277 | 0.871 | 0.119 | 0 | 0 | 1 | 42.188 | 1 | 64 | 1 | GJB_Rinse | 0
281 | P-1409 | 2022/10/07 14:34:06 | 2-65 | 1 | -6.972 | -48.944 | 18411 | -1 | 1 | EV | standard | H2O | 1665153552 | 0.31 | 0.761 | 77.138 | 0.001 | 0.008 | 1.83 | 1297.92 | 1.928 | 2.886 | 112.343 | 0.866 | 0.202 | 0 | 0 | 1 | 42.25 | 1 | 65 | 1 | GJB_Rinse | 0
282 | P-1409 | 2022/10/07 14:43:01 | 2-65 | 2 | -7.373 | -56.126 | 17963 | -1 | 1 | EV | standard | H2O | 1665154088 | 0.31 | 0.705 | 77.243 | 0.001 | 0.005 | 1.832 | 1294.542 | 1.856 | 2.857 | 111.346 | 0.867 | 0.545 | 0 | 0 | 1 | 42.188 | 1 | 65 | 1 | GJB_Rinse | 0
283 | P-1409 | 2022/10/07 14:51:57 | 2-65 | 3 | -7.481 | -57.735 | 17879 | -1 | 1 | EV | standard | H2O | 1665154623 | 0.3 | 0.714 | 75.655 | 0.001 | 0.003 | 1.799 | 1293.968 | 1.84 | 2.874 | 110.45 | 0.865 | 0.466 | 0 | 0 | 1 | 42.188 | 1 | 65 | 1 | GJB_Rinse | 0
284 | P-1409 | 2022/10/07 15:00:52 | 2-65 | 4 | -7.524 | -58.331 | 18033 | 0 | 1 | EV | standard | H2O | 1665155159 | 0.306 | 0.733 | 76.986 | 0.001 | 0.004 | 1.824 | 1295.15 | 1.566 | 2.867 | 111.406 | 0.863 | 0.318 | 0 | 0 | 1 | 42.25 | 1 | 65 | 1 | GJB_Rinse | 0

## Instructions
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Variable | Instruction | Comment
StandardsFile | standards file.csv
OutputBase | ChemCorrect
CorrectionType | M
timecode | Time Code
identifier | Identifier 1
sample | Sample
injection | Inj Nr | Inj Nr
line | Line
job | Job
Variable | Instruction | Comment
baseline_shift_MEAN | mean_standards(baseline_shift) | The mean of the organic baseline of the standards.
baseline_shift_STD | std_standards(baseline_shift) | The standard deviation of the organic baseline of the standards.
ch4_ppm_MEAN | mean_standards(ch4_ppm) | The mean methane concentration of the standards.
ch4_ppm_STD | std_standards(ch4_ppm) | The standard deviation of the methane concentrations of the standards.
residuals_MEAN | mean_standards(residuals) | The mean difference between the measured and modeled spectra of the standards.
residuals_STD | std_standards(residuals) | The standard deviation of the difference between measured and modeled spectra of the standards.
slope_shift_MEAN | mean_standards(slope_shift) | The mean organic slope of the standards.
slope_shift_STD | std_standards(slope_shift) | The standard deviation of the organic slope of the standards.
baseline_curvature_MEAN | mean_standards(baseline_curvature) | The mean of the baseline curvature of the standards.
baseline_curvature_STD | std_standards(baseline_curvature) | The standard deviation of the baseline curvature of the standards.
residuals_MEAN_ABS | abs(residuals_MEAN) | The absolute value of the mean difference between the measured and modeled spectra of the standards.
baseline_shift_MEAN_ABS | abs(baseline_shift_MEAN) | The absolute value of the mean of the organic baseline of the standards.
ch4_ppm_MEAN_ABS | abs(ch4_ppm_MEAN) | The absolute value of the mean methane concentration of the standards.
slope_shift_MEAN_ABS | abs(slope_shift_MEAN) | The absolute value of the mean organic slope of the standards.
residuals_15SIGMA | evaluation(residuals_STD*1.5) | One and half times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_15SIGMA | evaluation(residuals_MEAN+residuals_15SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_2SIGMA | evaluation(residuals_STD*2) | Twice the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_2SIGMA | evaluation(residuals_MEAN+residuals_2SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_3SIGMA | evaluation(residuals_STD*3) | Three times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_3SIGMA | evaluation(residuals_MEAN+residuals_3SIGMA) | The mean difference between measured and modeled spectra of the standards plus three standard deviation.
residuals_10PCT | evaluation(residuals_MEAN*0.1) | 30% of the difference between measured and modeled spectra of the standards.
ch4_ppm_3SIGMA | evaluation(ch4_ppm_STD*3) | Three times the standard deviation of the methanol amplitudes of the standards.
ch4_ppm_PLUS_3SIGMA | evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA) | The mean methanol amplitude of the standards plus three standard deviation.
HALF | float(0.5) | A fixed value of 0.5 or half.
baseline_shift_HALFSIGMA | evaluation(baseline_shift_STD*HALF) | Half the standard deviation of the organic baseline of the standards.
baseline_shift_PLUS_HALFSIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA) | The absolute value of the mean organic baseline of the standards plus one-half standard deviation.
baseline_shift_PLUS_1SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD) | The absolute value of the mean organic baseline of the standards plus one standard deviation.
baseline_shift_3SIGMA | evaluation(baseline_shift_STD*3) | Three times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_3SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA) | The absolute value of the mean organic baseline of the standards plus three standard deviation.
baseline_shift_18SIGMA | evaluation(baseline_shift_STD*18) | Eighteen times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_18SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA) | The absolute value of the mean organic baseline of the standards plus eighteen standard deviation.
slope_shift_5SIGMA | evaluation(slope_shift_STD*5) | Five times the standard deviation of the organic slope of the standards.
slope_shift_PLUS_5SIGMA | evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA) | The absolute value of the mean organic slope of the standards plus five standard deviation.
baseline_curvature_3SIGMA | evaluation(baseline_curvature_STD*3) | Three times the standard deviation of the baseline curvature of the standards.
baseline_curvature_PLUS_3SIGMA | evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
TRUE | evaluation(1>0)
FALSE | evaluation(1<0)
PROC_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the uncalibrated d18O value in order to calibrate it.
PROC_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the uncalibrated dD value in order to calibrate it.
Variable | Instruction | Comment
IS_STANDARD | is_standard_sample() | Standard recognition.
IS_NOT_STANDARD | bool_xor(IS_STANDARD,TRUE)
SAMP_residuals_MEAN | mean_all(residuals) | The mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_residuals_STD | std_all(residuals) | One standard deviation of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN | mean_all(ch4_ppm) | The mean methane concentration of a sample.
SAMP_ch4_ppm_STD | std_all(ch4_ppm) | The standard deviation of the methane concentrations of a sample.
SAMP_baseline_shift_MEAN | mean_all(baseline_shift) | The mean of the organic baseline of a sample.
SAMP_baseline_shift_STD | std_all(baseline_shift) | The standard deviation of the organic baseline of a sample.
SAMP_slope_shift_MEAN | mean_all(slope_shift) | The mean organic slope of a sample.
SAMP_slope_shift_STD | std_all(slope_shift) | The standard deviation of the organic slope of a sample.
SAMP_d18O_STD | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
SAMP_dD_STD | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_baseline_curvature_MEAN | mean_all(baseline_curvature) | The mean of the baseline curvature of the ssamples
SAMP_baseline_curvature_STD | std_all(baseline_curvature) | The standard deviation of the baseline curvature of the samples
SAMP_residuals_MEAN_ABS | abs(SAMP_residuals_MEAN) | Absolute value of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN_ABS | abs(SAMP_ch4_ppm_MEAN) | Absolute value of the mean methane concentration of a sample.
SAMP_baseline_shift_MEAN_ABS | abs(SAMP_baseline_shift_MEAN) | Absolute value of the mean organic baseline of a sample.
SAMP_slope_shift_MEAN_ABS | abs(SAMP_slope_shift_MEAN) | Absolute value of the mean organic slope of a sample.
SAMP_residuals_DIFFERENTIAL | evaluation(SAMP_residuals_MEAN-residuals_MEAN)
SAMP_residuals_DIFFERENTIAL_ABS | abs(SAMP_residuals_DIFFERENTIAL) | Absolute value of the difference between the mean of the sample residual and the mean of all standards residuals
SAMP_residuals_DIFFERENTIAL_PCT | evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
SAMP_baseline_curvature_DIFFERENTIAL | evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
SAMP_STD_DELTA_18_16O_CURR_MEAN | current_standards_mean(d(18_16)Mean) | The d18O mean of the last run standard.
SAMP_STD_DELTA_DH_CURR_MEAN | current_standards_mean(d(D_H)Mean) | The dD mean of the last run standard.
known_dD | known_standard(d(D_H)Mean) | The dD value of the known Standard.
known_d18O | known_standard(d(18_16)Mean) | The d18O value of the known Standard.
d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
measured_d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
measured_dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
measured_d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
measured_dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the mean uncalibrated d18O value in order to calibrate it.
SAMP_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the mean uncalibrated dD value in order to calibrate it.
SAMP_TEST_VALUE | evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
residuals_GT_3SIGMA | evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA) | Threshold of red trigger. If the mean difference between the measured spectra of the sample and the modeled spectra of the standard is greater than that of the mean difference of the standard plus three standard deviations, this will trigger.
residuals_GE_2SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GE_15SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GT_10PCT | evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
ch4_ppm_GE_3SIGMA | evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA) | True if the mean methane concentration of the sample is greater than or equal to that of the standard plus three standard deviations, otherwise false.
baseline_shift_GT_1SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus one standard deviation, otherwise false.
baseline_shift_GT_3SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus three standard deviation, otherwise false. Sets yellow trigger for "other", trigger threshold is approximately sample to sample dD precision specification
baseline_shift_GT_18SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus eighteen standard deviation, otherwise false. Sets red trigger for "other", trigger threshold is approximately 24 hour dD drift specification
slope_shift_GT_5SIGMA | evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA) | True if the mean organic slope of the sample is greater than that of the standard plus five standard deviations, otherwise false.
baseline_curvature_GT_3SIGMA | evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
RED | truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA)) | Red trigger. If the sample is not a standard and the threshold is triggered (measured-modeled > standard+3stdev's), then RED.
NOT_RED | bool_xor(RED,TRUE) | If the sample is determined to be "red" than it can not be "not red."
YELLOW | truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA))) | Truth evaluation used to determine whether or not a given sample is yellow.
NOT_YELLOW | bool_xor(YELLOW,TRUE) | If sample is determined to be "yellow" than it can not be "not yellow."
GREEN | truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW)) | This is the pre-requisite for a sample to be colored green. It must not be a standard, must not be yellow, and must not be red.
NOT_GREEN | bool_xor(GREEN,TRUE) | Pre-requisite for being "not-green."
CYAN | truth_evaluation((IS_STANDARD)) | This is the pre-requisite for a sample to be a standard and therefore cyan colored.
NOT_CYAN | bool_xor(CYAN,TRUE) | Pre-requisite for being not a standard.
METHANE | truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA)) | This is the threshold for amounts of methane being included in the main ChemCorrect spreadsheet. Methane will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured methane concentration must be greater than or equal to that of the standard plus three standard deviations.
OTHER1 | truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA)) | This is the threshold for amounts of "OTHER1" being included in the main ChemCorrect spreadsheet. This column (OTHER1) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic baseline must be greater than that of the standard plus one standard deviation.
OTHER2 | truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA)) | This is the threshold for amounts of "OTHER2" being included in the main ChemCorrect spreadsheet. This column (OTHER2) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic slope must be greater than that of the standard plus three standard deviations.
OTHER3 | truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | truth_evaluation(RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN)) | Organic baseline differential showing the difference between the measured baseline and the standard baseline.
Variable | Instruction | Comment
DTL_TEST | math_evaluation((1 + 2)) | Test evaluation of 1+2 being equal to 3.
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample | FALSE | FALSE | FALSE | FALSE | Testing Sample
B | identifier | TRUE | Name | RED | YELLOW | GREEN | CYAN | Testing Name
C | SAMP_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d<sup>18</sup>O Mean | RED | YELLOW | GREEN | CYAN | Testing18 16 Mean
D | SAMP_DELTA_DH_CALIBRATE | TRUE | Calibrated d<sup>2</sup>H Mean | RED | YELLOW | GREEN | CYAN | This is a test to see that I can use tool tips
E | SAMP_ch4_ppm_MEAN | METHANE | CH<sub>4</sub> | RED | YELLOW | FALSE | FALSE
F | baseline_shift_DIFF | OTHER1 | C<sub>2+</sub> alcohols | RED | YELLOW | FALSE | FALSE
G | SAMP_residuals_DIFFERENTIAL_PCT | OTHER3 | relative %deviation | RED | YELLOW | FALSE | FALSE
H | SAMP_d18O_STD | TRUE | Uncalibrated d<sup>18</sup>O precision | RED | YELLOW | GREEN | CYAN
I | SAMP_dD_STD | TRUE | Uncalibrated d<sup>2</sup>H precision | RED | YELLOW | GREEN | CYAN
J | SAMP_slope_shift_MEAN | OTHER2 | slope | RED | YELLOW | FALSE | FALSE
K | SAMP_baseline_curvature_DIFFERENTIAL | OTHER4 | curvature | RED | YELLOW | FALSE | FALSE
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample
B | identifier | TRUE | Name
C | col(d(18_16)Mean) | TRUE | Measured d(18_16) Mean
D | col(d(D_H)Mean) | TRUE | Measured d(D_H) Mean
E | DTL_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d(18_16) Mean
F | DTL_DELTA_DH_CALIBRATE | TRUE | Calibrated d(D_H) Mean
G | col(ch4_ppm) | METHANE | CH<sub>4</sub>
H | col(baseline_shift) | OTHER1 | Other
I | col(Time Code) | TRUE | Time Code
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | d18O_mean | TRUE | d(18_16) Mean | RED | YELLOW | GREEN | CYAN
B | dD_mean | TRUE | d(D_H) Mean | RED | YELLOW | GREEN | CYAN
C | XY_X(d18O_mean) XY_Y(dD_mean) | TRUE | d18O x dD | gmwl | best_fit

## Process Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Process Variables
Variable | Value | Operations
baseline_shift_MEAN | 1291.89475
varray: [1291.907, 1291.6, 1292.005, 1294.707, 1291.157, 1291.603, 1291.201, 1291.453, 1291.769, 1291.332, 1293.919, 1291.052, 1289.881, 1291.838, 1291.057, 1292.033, 1292.275, 1292.443, 1292.308, 1292.355]
inst: mean_standards
operation: mean_standards(baseline_shift)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0, 10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: baseline_shift
baseline_shift_STD | 1.0275467911
varray: [1291.907, 1291.6, 1292.005, 1294.707, 1291.157, 1291.603, 1291.201, 1291.453, 1291.769, 1291.332, 1293.919, 1291.052, 1289.881, 1291.838, 1291.057, 1292.033, 1292.275, 1292.443, 1292.308, 1292.355]
inst: std_standards
operation: std_standards(baseline_shift)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0, 10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: baseline_shift
ch4_ppm_MEAN | 0.37445
varray: [0.336, 0.281, 0.132, 0.21, 0.145, 0.498, 0.475, 0.696, 0.525, 0.539, 0.458, 0.419, 0.333, 0.109, 0.598, 0.558, 0.391, 0.466, 0.208, 0.112]
inst: mean_standards
operation: mean_standards(ch4_ppm)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0, 10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: ch4_ppm
ch4_ppm_STD | 0.177296799924
varray: [0.336, 0.281, 0.132, 0.21, 0.145, 0.498, 0.475, 0.696, 0.525, 0.539, 0.458, 0.419, 0.333, 0.109, 0.598, 0.558, 0.391, 0.466, 0.208, 0.112]
inst: std_standards
operation: std_standards(ch4_ppm)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0, 10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: ch4_ppm
residuals_MEAN | 2.80105
varray: [2.802, 2.803, 2.772, 2.829, 2.813, 2.824, 2.844, 2.851, 2.864, 2.879, 2.811, 2.733, 2.719, 2.73, 2.81, 2.819, 2.801, 2.808, 2.753, 2.756]
inst: mean_standards
operation: mean_standards(residuals)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0, 10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: residuals
residuals_STD | 0.0447077939749
varray: [2.802, 2.803, 2.772, 2.829, 2.813, 2.824, 2.844, 2.851, 2.864, 2.879, 2.811, 2.733, 2.719, 2.73, 2.81, 2.819, 2.801, 2.808, 2.753, 2.756]
inst: std_standards
operation: std_standards(residuals)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0, 10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: residuals
slope_shift_MEAN | 1.75575
varray: [1.273, 1.891, 1.49, 1.248, 1.174, 1.434, 0.992, 1.028, 1.258, 1.095, 2.628, 2.44, 2.299, 2.335, 2.092, 1.81, 2.16, 2.018, 1.966, 2.484]
inst: mean_standards
operation: mean_standards(slope_shift)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0, 10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: slope_shift
slope_shift_STD | 0.5429661107
varray: [1.273, 1.891, 1.49, 1.248, 1.174, 1.434, 0.992, 1.028, 1.258, 1.095, 2.628, 2.44, 2.299, 2.335, 2.092, 1.81, 2.16, 2.018, 1.966, 2.484]
inst: std_standards
operation: std_standards(slope_shift)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0, 10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: slope_shift
baseline_curvature_MEAN | 108.15515
varray: [107.978, 106.59, 107.235, 109.297, 107.073, 108.432, 108.953, 109.662, 109.929, 109.435, 109.522, 106.723, 105.858, 105.837, 108.517, 109.136, 108.698, 109.188, 108.154, 106.886]
inst: mean_standards
operation: mean_standards(baseline_curvature)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0, 10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: baseline_curvature
baseline_curvature_STD | 1.3019469338
varray: [107.978, 106.59, 107.235, 109.297, 107.073, 108.432, 108.953, 109.662, 109.929, 109.435, 109.522, 106.723, 105.858, 105.837, 108.517, 109.136, 108.698, 109.188, 108.154, 106.886]
inst: std_standards
operation: std_standards(baseline_curvature)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0, 10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: baseline_curvature
residuals_MEAN_ABS | 2.80105
operation: abs(residuals_MEAN)
inst: abs
source_col: residuals_MEAN
baseline_shift_MEAN_ABS | 1291.89475
operation: abs(baseline_shift_MEAN)
inst: abs
source_col: baseline_shift_MEAN
ch4_ppm_MEAN_ABS | 0.37445
operation: abs(ch4_ppm_MEAN)
inst: abs
source_col: ch4_ppm_MEAN
slope_shift_MEAN_ABS | 1.75575
operation: abs(slope_shift_MEAN)
inst: abs
source_col: slope_shift_MEAN
residuals_15SIGMA | 0.0670616909624
operation: evaluation(residuals_STD*1.5)
inst: evaluation
factors: residuals_STD*1.5
residuals_PLUS_15SIGMA | 2.86811169096
operation: evaluation(residuals_MEAN+residuals_15SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_15SIGMA
residuals_2SIGMA | 0.0894155879499
operation: evaluation(residuals_STD*2)
inst: evaluation
factors: residuals_STD*2
residuals_PLUS_2SIGMA | 2.89046558795
operation: evaluation(residuals_MEAN+residuals_2SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_2SIGMA
residuals_3SIGMA | 0.134123381925
operation: evaluation(residuals_STD*3)
inst: evaluation
factors: residuals_STD*3
residuals_PLUS_3SIGMA | 2.93517338192
operation: evaluation(residuals_MEAN+residuals_3SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_3SIGMA
residuals_10PCT | 0.280105
operation: evaluation(residuals_MEAN*0.1)
inst: evaluation
factors: residuals_MEAN*0.1
ch4_ppm_3SIGMA | 0.531890399771
operation: evaluation(ch4_ppm_STD*3)
inst: evaluation
factors: ch4_ppm_STD*3
ch4_ppm_PLUS_3SIGMA | 0.906340399771
operation: evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA)
inst: evaluation
factors: ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA
HALF | 0.5
operation: float(0.5)
inst: float
baseline_shift_HALFSIGMA | 0.513773395549
operation: evaluation(baseline_shift_STD*HALF)
inst: evaluation
factors: baseline_shift_STD*HALF
baseline_shift_PLUS_HALFSIGMA | 1292.4085234
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA
baseline_shift_PLUS_1SIGMA | 1292.92229679
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_STD
baseline_shift_3SIGMA | 3.08264037329
operation: evaluation(baseline_shift_STD*3)
inst: evaluation
factors: baseline_shift_STD*3
baseline_shift_PLUS_3SIGMA | 1294.97739037
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_3SIGMA
baseline_shift_18SIGMA | 18.4958422398
operation: evaluation(baseline_shift_STD*18)
inst: evaluation
factors: baseline_shift_STD*18
baseline_shift_PLUS_18SIGMA | 1310.39059224
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_18SIGMA
slope_shift_5SIGMA | 2.7148305535
operation: evaluation(slope_shift_STD*5)
inst: evaluation
factors: slope_shift_STD*5
slope_shift_PLUS_5SIGMA | 4.4705805535
operation: evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA)
inst: evaluation
factors: slope_shift_MEAN_ABS+slope_shift_5SIGMA
baseline_curvature_3SIGMA | 3.90584080139
operation: evaluation(baseline_curvature_STD*3)
inst: evaluation
factors: baseline_curvature_STD*3
baseline_curvature_PLUS_3SIGMA | 112.060990801
operation: evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
inst: evaluation
factors: baseline_curvature_3SIGMA+baseline_curvature_MEAN
TRUE | True
operation: evaluation(1>0)
inst: evaluation
factors: 1>0
FALSE | False
operation: evaluation(1<0)
inst: evaluation
factors: 1<0
PROC_DELTA_18_16O_CALIBRATE | -5.12635343253
actual_array: [ 1.93 -15.84]
varray: [-8.089, -7.945, -7.774, -7.704, -7.622, -7.559, -7.539, -7.616, -7.553, -7.566, 3.875, 4.425, 4.604, 4.634, 4.566, 4.657, 4.647, 4.683, 4.744, 4.69, -12.217, -13.128, -13.347, -13.31, -13.406, -13.456, -13.424, -13.422, -13.464, -13.479, -8.031, -7.811, -7.75, -7.768, -7.769, -7.691, -7.695, -7.732, -7.788, -7.739, 0.189, 0.535, 0.624, 0.677, 0.245, 0.282, 0.259, 0.322, 0.316, 0.311, 0.382, 0.37, 0.489, 0.497, 0.474, 0.494, -0.071, -0.059, -0.048, -0.037, 0.194, 0.219, 0.177, 0.204, -0.152, -0.205, -0.227, -0.142, -0.15, -0.13, -0.096, -0.098, -7.179, -7.438, -7.551, -7.587, -0.008, 0.386, 0.446, 0.435, 0.19, 0.262, 0.227, 0.28, -0.008, -0.028, -0.044, 0.0, -0.049, 0.094, 0.015, -0.032, 1.586, 1.621, 1.619, 1.65, 1.645, 1.644, 1.625, 1.57, 7.155, 7.397, 7.525, 7.55, 4.17, 4.015, 3.987, 3.988, -6.802, -7.331, -7.429, -7.5, 0.437, 0.863, 0.943, 0.955, -6.019, -6.403, -6.444, -6.51, 0.524, 0.821, 0.928, 0.92, 1.258, 1.301, 1.304, 1.296, -2.231, -2.448, -2.486, -2.472, -2.816, -2.889, -2.876, -2.928, -3.012, -3.044, -3.05, -3.006, -2.727, -2.781, -2.792, -2.876, -7.21, -7.467, -7.511, -7.566, -3.001, -2.854, -2.811, -2.811, -3.122, -3.134, -3.203, -3.194, -3.27, -3.364, -3.215, -3.331, -2.909, -2.94, -2.975, -2.963, -3.212, -3.305, -3.29, -3.318, -2.278, -2.212, -2.178, -2.179, -1.47, -1.441, -1.396, -1.463, -1.851, -1.921, -2.015, -1.936, -7.223, -7.522, -7.469, -7.596, -2.92, -2.728, -2.716, -2.635, -2.396, -2.418, -2.421, -2.469, -2.39, -2.116, -2.298, -2.332, -1.368, -1.292, -1.353, -1.354, -1.25, -1.326, -1.262, -1.269, -2.593, -2.702, -2.678, -2.73, -2.7, -2.838, -2.718, -2.727, -2.187, -2.131, -2.053, -2.1, -7.192, -7.476, -7.532, -7.51, -3.058, -2.984, -2.859, -2.886, -3.23, -3.311, -3.253, -3.301, -4.134, -4.284, -4.226, -4.348, -1.593, -1.612, -1.499, -1.49, 0.476, 0.464, 0.557, 0.511, 1.708, 1.712, 1.669, 1.756, -2.767, -3.108, -3.152, -3.076, -0.101, -0.045, 0.0, 0.055, -7.052, -7.476, -7.548, -7.524, -0.902, -0.775, -0.689, -0.514, 0.791, 0.875, 0.945, 1.032, -6.158, -6.399, -6.505, -6.427, -8.18, -8.228, -8.166, -8.27, -4.752, -3.922, -4.385, -4.378, 1.3, 1.527, 1.248, 1.3, -6.972, -7.373, -7.481, -7.524]
fit_rsq: 1.0
fit_intercept: 2.61730844119
except: True
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.52283450704
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [1.0, 538.0, 1075.0, 1612.0, 2149.0, 2686.0, 3223.0, 3760.0, 4296.0, 4833.0, 5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0, 10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0, 16271.0, 16809.0, 17345.0, 17883.0, 18419.0, 18956.0, 19493.0, 20030.0, 20568.0, 21106.0, 21696.0, 22233.0, 22770.0, 23307.0, 23897.0, 24435.0, 24972.0, 25509.0, 26100.0, 26637.0, 27175.0, 27712.0, 28300.0, 28837.0, 29374.0, 29911.0, 30499.0, 31035.0, 31573.0, 32110.0, 32699.0, 33236.0, 33773.0, 34311.0, 34900.0, 35438.0, 35975.0, 36512.0, 37101.0, 37639.0, 38176.0, 38713.0, 39303.0, 39840.0, 40378.0, 40915.0, 41505.0, 42043.0, 42580.0, 43118.0, 43705.0, 44242.0, 44779.0, 45316.0, 45903.0, 46441.0, 46978.0, 47540.0, 48192.0, 48729.0, 49266.0, 49803.0, 50391.0, 50928.0, 51465.0, 52002.0, 52591.0, 53128.0, 53665.0, 54202.0, 54791.0, 55328.0, 55865.0, 56402.0, 56992.0, 57530.0, 58067.0, 58604.0, 59189.0, 59725.0, 60262.0, 60798.0, 61386.0, 61923.0, 62460.0, 62996.0, 63583.0, 64120.0, 64657.0, 65194.0, 65781.0, 66318.0, 66856.0, 67392.0, 88912.0, 89448.0, 89984.0, 90521.0, 91109.0, 91645.0, 92181.0, 92717.0, 93306.0, 93842.0, 94378.0, 94914.0, 95499.0, 96035.0, 96571.0, 97107.0, 97692.0, 98228.0, 98764.0, 99301.0, 99886.0, 100422.0, 100958.0, 101494.0, 102080.0, 102616.0, 103152.0, 103688.0, 104275.0, 104811.0, 105347.0, 105883.0, 106469.0, 107006.0, 107542.0, 108078.0, 108665.0, 109201.0, 109737.0, 110274.0, 110856.0, 111393.0, 111929.0, 112465.0, 113049.0, 113585.0, 114121.0, 114657.0, 115241.0, 115777.0, 116313.0, 116849.0, 117433.0, 117970.0, 118505.0, 119041.0, 119626.0, 120162.0, 120698.0, 121234.0, 121820.0, 122356.0, 122892.0, 123428.0, 124013.0, 124549.0, 125085.0, 125621.0, 126205.0, 126741.0, 127277.0, 127812.0, 128397.0, 128933.0, 129469.0, 130004.0, 130590.0, 131125.0, 131661.0, 132197.0, 132783.0, 133318.0, 133854.0, 134390.0, 134976.0, 135513.0, 136048.0, 136584.0, 137171.0, 137707.0, 138243.0, 138779.0, 139366.0, 139902.0, 140438.0, 140974.0, 141558.0, 142094.0, 142630.0, 143166.0, 143751.0, 144287.0, 144823.0, 145358.0, 145944.0, 146480.0, 147016.0, 147552.0, 148139.0, 148674.0, 149210.0, 149746.0, 150333.0, 150869.0, 151405.0, 151941.0, 152528.0, 153064.0, 153600.0, 154136.0, 154723.0, 155259.0, 155796.0, 156331.0, 156917.0, 157453.0, 157989.0, 158525.0, 159111.0, 159647.0, 160183.0, 160718.0, 161305.0, 161841.0, 162377.0, 162913.0, 163500.0, 164035.0, 164571.0, 165107.0, 165695.0, 166231.0, 166766.0, 167302.0, 167890.0, 168426.0, 168962.0, 169497.0, 170085.0, 170621.0, 171157.0, 171693.0, 172279.0, 172814.0, 173350.0, 173886.0, 174473.0, 175008.0, 175544.0, 176079.0]
source_col: d(18_16)Mean
PROC_DELTA_DH_CALIBRATE | -32.1240002709
actual_array: [ 18.1 -119. ]
varray: [-65.076, -63.796, -62.106, -61.064, -60.397, -60.071, -59.778, -59.828, -59.984, -59.99, 27.309, 35.67, 37.859, 38.931, 39.076, 39.468, 39.646, 40.008, 40.12, 40.328, -92.943, -105.802, -108.88, -109.777, -110.581, -110.82, -111.487, -111.33, -111.891, -112.047, -66.537, -62.33, -61.365, -60.864, -60.764, -60.768, -60.403, -60.553, -60.877, -60.414, 5.349, 11.125, 12.932, 13.665, 11.243, 11.35, 11.316, 11.55, 11.441, 11.281, 11.621, 11.768, 12.429, 12.508, 12.568, 12.584, 9.294, 9.137, 9.114, 8.987, 9.902, 9.91, 9.877, 9.889, 8.947, 8.711, 8.696, 9.013, 8.747, 8.927, 9.325, 8.661, -51.202, -56.523, -57.919, -58.852, 4.49, 10.547, 11.963, 12.523, 9.879, 10.275, 10.34, 10.377, 9.481, 9.546, 9.562, 0.0, 9.78, 10.452, 9.88, 9.591, 16.181, 16.766, 16.827, 17.328, 23.472, 24.434, 24.724, 24.351, 46.787, 49.196, 49.521, 49.647, 35.066, 33.611, 33.406, 33.756, -47.37, -55.247, -56.918, -57.87, 5.506, 12.182, 13.342, 13.48, -43.966, -49.623, -50.26, -51.109, 6.335, 11.141, 12.619, 13.343, 18.069, 18.164, 18.498, 18.514, -11.241, -14.322, -15.273, -15.473, -18.346, -18.969, -18.699, -18.982, -19.248, -19.266, -19.285, -19.208, -18.218, -18.506, -18.352, -18.882, -53.831, -57.29, -58.194, -58.561, -22.316, -19.929, -18.807, -19.167, -19.661, -19.322, -20.313, -20.076, -21.026, -21.532, -20.736, -21.268, -19.506, -19.958, -19.852, -19.872, -20.785, -20.845, -21.203, -21.204, -15.919, -15.693, -15.175, -15.645, -9.086, -8.774, -8.171, -8.541, -11.004, -11.917, -12.434, -12.014, -53.526, -57.585, -57.94, -58.613, -22.067, -19.0, -17.844, -17.861, -16.701, -17.129, -16.988, -17.1, -16.514, -15.251, -16.317, -16.639, -6.95, -5.717, -5.844, -6.143, -6.857, -7.499, -7.039, -7.1, -16.354, -17.466, -17.32, -17.719, -17.524, -18.625, -18.192, -18.241, -14.029, -13.561, -13.231, -13.606, -53.577, -57.22, -58.151, -58.302, -22.99, -20.835, -19.833, -19.634, -23.023, -23.466, -23.377, -23.694, -34.299, -36.064, -35.417, -36.504, -9.214, -7.29, -6.402, -6.228, 14.768, 16.122, 16.899, 16.771, 23.244, 23.419, 23.87, 23.96, -12.421, -16.323, -16.979, -16.97, 7.477, 9.222, 10.052, 10.581, -50.308, -56.79, -57.936, -58.249, -5.226, -1.791, -0.726, 0.358, -2.923, -2.977, -2.709, -2.489, -37.606, -40.677, -41.426, -41.421, -58.646, -60.001, -59.998, -60.862, -27.864, -22.421, -23.321, -23.028, 25.258, 25.23, 24.489, 25.258, -48.944, -56.126, -57.735, -58.331]
fit_rsq: 1.0
fit_intercept: 18.5140665208
except: True
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -15.7883908451
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [1.0, 538.0, 1075.0, 1612.0, 2149.0, 2686.0, 3223.0, 3760.0, 4296.0, 4833.0, 5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0, 10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0, 16271.0, 16809.0, 17345.0, 17883.0, 18419.0, 18956.0, 19493.0, 20030.0, 20568.0, 21106.0, 21696.0, 22233.0, 22770.0, 23307.0, 23897.0, 24435.0, 24972.0, 25509.0, 26100.0, 26637.0, 27175.0, 27712.0, 28300.0, 28837.0, 29374.0, 29911.0, 30499.0, 31035.0, 31573.0, 32110.0, 32699.0, 33236.0, 33773.0, 34311.0, 34900.0, 35438.0, 35975.0, 36512.0, 37101.0, 37639.0, 38176.0, 38713.0, 39303.0, 39840.0, 40378.0, 40915.0, 41505.0, 42043.0, 42580.0, 43118.0, 43705.0, 44242.0, 44779.0, 45316.0, 45903.0, 46441.0, 46978.0, 47540.0, 48192.0, 48729.0, 49266.0, 49803.0, 50391.0, 50928.0, 51465.0, 52002.0, 52591.0, 53128.0, 53665.0, 54202.0, 54791.0, 55328.0, 55865.0, 56402.0, 56992.0, 57530.0, 58067.0, 58604.0, 59189.0, 59725.0, 60262.0, 60798.0, 61386.0, 61923.0, 62460.0, 62996.0, 63583.0, 64120.0, 64657.0, 65194.0, 65781.0, 66318.0, 66856.0, 67392.0, 88912.0, 89448.0, 89984.0, 90521.0, 91109.0, 91645.0, 92181.0, 92717.0, 93306.0, 93842.0, 94378.0, 94914.0, 95499.0, 96035.0, 96571.0, 97107.0, 97692.0, 98228.0, 98764.0, 99301.0, 99886.0, 100422.0, 100958.0, 101494.0, 102080.0, 102616.0, 103152.0, 103688.0, 104275.0, 104811.0, 105347.0, 105883.0, 106469.0, 107006.0, 107542.0, 108078.0, 108665.0, 109201.0, 109737.0, 110274.0, 110856.0, 111393.0, 111929.0, 112465.0, 113049.0, 113585.0, 114121.0, 114657.0, 115241.0, 115777.0, 116313.0, 116849.0, 117433.0, 117970.0, 118505.0, 119041.0, 119626.0, 120162.0, 120698.0, 121234.0, 121820.0, 122356.0, 122892.0, 123428.0, 124013.0, 124549.0, 125085.0, 125621.0, 126205.0, 126741.0, 127277.0, 127812.0, 128397.0, 128933.0, 129469.0, 130004.0, 130590.0, 131125.0, 131661.0, 132197.0, 132783.0, 133318.0, 133854.0, 134390.0, 134976.0, 135513.0, 136048.0, 136584.0, 137171.0, 137707.0, 138243.0, 138779.0, 139366.0, 139902.0, 140438.0, 140974.0, 141558.0, 142094.0, 142630.0, 143166.0, 143751.0, 144287.0, 144823.0, 145358.0, 145944.0, 146480.0, 147016.0, 147552.0, 148139.0, 148674.0, 149210.0, 149746.0, 150333.0, 150869.0, 151405.0, 151941.0, 152528.0, 153064.0, 153600.0, 154136.0, 154723.0, 155259.0, 155796.0, 156331.0, 156917.0, 157453.0, 157989.0, 158525.0, 159111.0, 159647.0, 160183.0, 160718.0, 161305.0, 161841.0, 162377.0, 162913.0, 163500.0, 164035.0, 164571.0, 165107.0, 165695.0, 166231.0, 166766.0, 167302.0, 167890.0, 168426.0, 168962.0, 169497.0, 170085.0, 170621.0, 171157.0, 171693.0, 172279.0, 172814.0, 173350.0, 173886.0, 174473.0, 175008.0, 175544.0, 176079.0]
source_col: d(D_H)Mean

## 1 - Sample 1 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 1
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.0053
varray: [3.246, 3.371, 3.086, 2.917, 2.87, 2.925, 2.943, 2.907, 2.908, 2.88]
inst: mean_all
operation: mean_all(residuals)
tarray: [1.0, 538.0, 1075.0, 1612.0, 2149.0, 2686.0, 3223.0, 3760.0, 4296.0, 4833.0]
source_col: residuals
SAMP_residuals_STD | 0.173027968195
varray: [3.246, 3.371, 3.086, 2.917, 2.87, 2.925, 2.943, 2.907, 2.908, 2.88]
inst: std_all
operation: std_all(residuals)
tarray: [1.0, 538.0, 1075.0, 1612.0, 2149.0, 2686.0, 3223.0, 3760.0, 4296.0, 4833.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4695
varray: [0.416, 0.46, 0.597, 0.391, 0.389, 0.456, 0.64, 0.409, 0.518, 0.419]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [1.0, 538.0, 1075.0, 1612.0, 2149.0, 2686.0, 3223.0, 3760.0, 4296.0, 4833.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0879813742915
varray: [0.416, 0.46, 0.597, 0.391, 0.389, 0.456, 0.64, 0.409, 0.518, 0.419]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [1.0, 538.0, 1075.0, 1612.0, 2149.0, 2686.0, 3223.0, 3760.0, 4296.0, 4833.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1330.4243
varray: [1365.224, 1456.668, 1365.829, 1317.71, 1306.263, 1301.046, 1299.419, 1298.017, 1297.468, 1296.599]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [1.0, 538.0, 1075.0, 1612.0, 2149.0, 2686.0, 3223.0, 3760.0, 4296.0, 4833.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 51.9105217295
varray: [1365.224, 1456.668, 1365.829, 1317.71, 1306.263, 1301.046, 1299.419, 1298.017, 1297.468, 1296.599]
inst: std_all
operation: std_all(baseline_shift)
tarray: [1.0, 538.0, 1075.0, 1612.0, 2149.0, 2686.0, 3223.0, 3760.0, 4296.0, 4833.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.8093
varray: [0.338, -3.555, -0.733, 1.433, 2.074, 1.315, 1.902, 1.689, 1.915, 1.715]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [1.0, 538.0, 1075.0, 1612.0, 2149.0, 2686.0, 3223.0, 3760.0, 4296.0, 4833.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.76077186673
varray: [0.338, -3.555, -0.733, 1.433, 2.074, 1.315, 1.902, 1.689, 1.915, 1.715]
inst: std_all
operation: std_all(slope_shift)
tarray: [1.0, 538.0, 1075.0, 1612.0, 2149.0, 2686.0, 3223.0, 3760.0, 4296.0, 4833.0]
source_col: slope_shift
SAMP_d18O_STD | 0.187152967739
varray: [-8.089, -7.945, -7.774, -7.704, -7.622, -7.559, -7.539, -7.616, -7.553, -7.566]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1075.0, 1612.0, 2149.0, 2686.0, 3223.0, 3760.0, 4296.0, 4833.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.86689843799
varray: [-65.076, -63.796, -62.106, -61.064, -60.397, -60.071, -59.778, -59.828, -59.984, -59.99]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1075.0, 1612.0, 2149.0, 2686.0, 3223.0, 3760.0, 4296.0, 4833.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 115.6449
varray: [124.19, 127.809, 119.681, 111.725, 111.387, 112.543, 113.046, 112.084, 112.127, 111.857]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [1.0, 538.0, 1075.0, 1612.0, 2149.0, 2686.0, 3223.0, 3760.0, 4296.0, 4833.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 6.02355907066
varray: [124.19, 127.809, 119.681, 111.725, 111.387, 112.543, 113.046, 112.084, 112.127, 111.857]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [1.0, 538.0, 1075.0, 1612.0, 2149.0, 2686.0, 3223.0, 3760.0, 4296.0, 4833.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.0053
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4695
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1330.4243
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.8093
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.20425
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.20425
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 7.2919083915
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 7.48975
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 0
varray: [-8.089, -7.945, -7.774, -7.704, -7.622, -7.559, -7.539, -7.616, -7.553, -7.566]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [1.0, 538.0, 1075.0, 1612.0, 2149.0, 2686.0, 3223.0, 3760.0, 4296.0, 4833.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 0
varray: [-65.076, -63.796, -62.106, -61.064, -60.397, -60.071, -59.778, -59.828, -59.984, -59.99]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [1.0, 538.0, 1075.0, 1612.0, 2149.0, 2686.0, 3223.0, 3760.0, 4296.0, 4833.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.6967
varray: [-8.089, -7.945, -7.774, -7.704, -7.622, -7.559, -7.539, -7.616, -7.553, -7.566]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1075.0, 1612.0, 2149.0, 2686.0, 3223.0, 3760.0, 4296.0, 4833.0]
source_col: d(18_16)Mean
dD_mean | -61.209
varray: [-65.076, -63.796, -62.106, -61.064, -60.397, -60.071, -59.778, -59.828, -59.984, -59.99]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1075.0, 1612.0, 2149.0, 2686.0, 3223.0, 3760.0, 4296.0, 4833.0]
source_col: d(D_H)Mean
d18O_stdev | 0.187152967739
varray: [-8.089, -7.945, -7.774, -7.704, -7.622, -7.559, -7.539, -7.616, -7.553, -7.566]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1075.0, 1612.0, 2149.0, 2686.0, 3223.0, 3760.0, 4296.0, 4833.0]
source_col: d(18_16)Mean
dD_stdev | 1.86689843799
varray: [-65.076, -63.796, -62.106, -61.064, -60.397, -60.071, -59.778, -59.828, -59.984, -59.99]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1075.0, 1612.0, 2149.0, 2686.0, 3223.0, 3760.0, 4296.0, 4833.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.6967
varray: [-8.089, -7.945, -7.774, -7.704, -7.622, -7.559, -7.539, -7.616, -7.553, -7.566]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1075.0, 1612.0, 2149.0, 2686.0, 3223.0, 3760.0, 4296.0, 4833.0]
source_col: d(18_16)Mean
measured_dD_mean | -61.209
varray: [-65.076, -63.796, -62.106, -61.064, -60.397, -60.071, -59.778, -59.828, -59.984, -59.99]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1075.0, 1612.0, 2149.0, 2686.0, 3223.0, 3760.0, 4296.0, 4833.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.187152967739
varray: [-8.089, -7.945, -7.774, -7.704, -7.622, -7.559, -7.539, -7.616, -7.553, -7.566]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1075.0, 1612.0, 2149.0, 2686.0, 3223.0, 3760.0, 4296.0, 4833.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.86689843799
varray: [-65.076, -63.796, -62.106, -61.064, -60.397, -60.071, -59.778, -59.828, -59.984, -59.99]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1075.0, 1612.0, 2149.0, 2686.0, 3223.0, 3760.0, 4296.0, 4833.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.286338942
actual_array: [ 1.93 -15.84]
varray: [-8.089, -7.945, -7.774, -7.704, -7.622, -7.559, -7.539, -7.616, -7.553, -7.566]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.6967
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [1.0, 538.0, 1075.0, 1612.0, 2149.0, 2686.0, 3223.0, 3760.0, 4296.0, 4833.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -74.6600683209
actual_array: [ 18.1 -119. ]
varray: [-65.076, -63.796, -62.106, -61.064, -60.397, -60.071, -59.778, -59.828, -59.984, -59.99]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -61.209
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [1.0, 538.0, 1075.0, 1612.0, 2149.0, 2686.0, 3223.0, 3760.0, 4296.0, 4833.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.88945
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 38.52955
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 2 - Sample 2 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 2
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8281
varray: [2.802, 2.803, 2.772, 2.829, 2.813, 2.824, 2.844, 2.851, 2.864, 2.879]
inst: mean_all
operation: mean_all(residuals)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0]
source_col: residuals
SAMP_residuals_STD | 0.0322505813901
varray: [2.802, 2.803, 2.772, 2.829, 2.813, 2.824, 2.844, 2.851, 2.864, 2.879]
inst: std_all
operation: std_all(residuals)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3837
varray: [0.336, 0.281, 0.132, 0.21, 0.145, 0.498, 0.475, 0.696, 0.525, 0.539]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.190420616297
varray: [0.336, 0.281, 0.132, 0.21, 0.145, 0.498, 0.475, 0.696, 0.525, 0.539]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1291.8734
varray: [1291.907, 1291.6, 1292.005, 1294.707, 1291.157, 1291.603, 1291.201, 1291.453, 1291.769, 1291.332]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.03503625272
varray: [1291.907, 1291.6, 1292.005, 1294.707, 1291.157, 1291.603, 1291.201, 1291.453, 1291.769, 1291.332]
inst: std_all
operation: std_all(baseline_shift)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.2883
varray: [1.273, 1.891, 1.49, 1.248, 1.174, 1.434, 0.992, 1.028, 1.258, 1.095]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.265395742242
varray: [1.273, 1.891, 1.49, 1.248, 1.174, 1.434, 0.992, 1.028, 1.258, 1.095]
inst: std_all
operation: std_all(slope_shift)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0]
source_col: slope_shift
SAMP_d18O_STD | 0.253245269781
varray: [3.875, 4.425, 4.604, 4.634, 4.566, 4.657, 4.647, 4.683, 4.744, 4.69]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.94869830169
varray: [27.309, 35.67, 37.859, 38.931, 39.076, 39.468, 39.646, 40.008, 40.12, 40.328]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.4584
varray: [107.978, 106.59, 107.235, 109.297, 107.073, 108.432, 108.953, 109.662, 109.929, 109.435]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.18562615431
varray: [107.978, 106.59, 107.235, 109.297, 107.073, 108.432, 108.953, 109.662, 109.929, 109.435]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8281
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3837
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1291.8734
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.2883
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.02705
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.02705
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.965709287589
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.30325
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 4.5525
varray: [3.875, 4.425, 4.604, 4.634, 4.566, 4.657, 4.647, 4.683, 4.744, 4.69]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 37.8415
varray: [27.309, 35.67, 37.859, 38.931, 39.076, 39.468, 39.646, 40.008, 40.12, 40.328]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0]
source_col: d(D_H)Mean
known_dD | 18.1
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 1.93
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 4.5525
varray: [3.875, 4.425, 4.604, 4.634, 4.566, 4.657, 4.647, 4.683, 4.744, 4.69]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0]
source_col: d(18_16)Mean
dD_mean | 37.8415
varray: [27.309, 35.67, 37.859, 38.931, 39.076, 39.468, 39.646, 40.008, 40.12, 40.328]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0]
source_col: d(D_H)Mean
d18O_stdev | 0.253245269781
varray: [3.875, 4.425, 4.604, 4.634, 4.566, 4.657, 4.647, 4.683, 4.744, 4.69]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0]
source_col: d(18_16)Mean
dD_stdev | 3.94869830169
varray: [27.309, 35.67, 37.859, 38.931, 39.076, 39.468, 39.646, 40.008, 40.12, 40.328]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0]
source_col: d(D_H)Mean
measured_d18O_mean | 4.5525
varray: [3.875, 4.425, 4.604, 4.634, 4.566, 4.657, 4.647, 4.683, 4.744, 4.69]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0]
source_col: d(18_16)Mean
measured_dD_mean | 37.8415
varray: [27.309, 35.67, 37.859, 38.931, 39.076, 39.468, 39.646, 40.008, 40.12, 40.328]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.253245269781
varray: [3.875, 4.425, 4.604, 4.634, 4.566, 4.657, 4.647, 4.683, 4.744, 4.69]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.94869830169
varray: [27.309, 35.67, 37.859, 38.931, 39.076, 39.468, 39.646, 40.008, 40.12, 40.328]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.93
actual_array: [ 1.93 -15.84]
varray: [3.875, 4.425, 4.604, 4.634, 4.566, 4.657, 4.647, 4.683, 4.744, 4.69]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 4.5525
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 18.1
actual_array: [ 18.1 -119. ]
varray: [27.309, 35.67, 37.859, 38.931, 39.076, 39.468, 39.646, 40.008, 40.12, 40.328]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 37.8415
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [5422.0, 5958.0, 6496.0, 7033.0, 7571.0, 8108.0, 8646.0, 9183.0, 9720.0, 10257.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.06665
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.02135
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 3 - Sample 3 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 3
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.774
varray: [2.811, 2.733, 2.719, 2.73, 2.81, 2.819, 2.801, 2.808, 2.753, 2.756]
inst: mean_all
operation: mean_all(residuals)
tarray: [10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: residuals
SAMP_residuals_STD | 0.0394151691047
varray: [2.811, 2.733, 2.719, 2.73, 2.81, 2.819, 2.801, 2.808, 2.753, 2.756]
inst: std_all
operation: std_all(residuals)
tarray: [10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3652
varray: [0.458, 0.419, 0.333, 0.109, 0.598, 0.558, 0.391, 0.466, 0.208, 0.112]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.172947519336
varray: [0.458, 0.419, 0.333, 0.109, 0.598, 0.558, 0.391, 0.466, 0.208, 0.112]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1291.9161
varray: [1293.919, 1291.052, 1289.881, 1291.838, 1291.057, 1292.033, 1292.275, 1292.443, 1292.308, 1292.355]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.07550395939
varray: [1293.919, 1291.052, 1289.881, 1291.838, 1291.057, 1292.033, 1292.275, 1292.443, 1292.308, 1292.355]
inst: std_all
operation: std_all(baseline_shift)
tarray: [10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.2232
varray: [2.628, 2.44, 2.299, 2.335, 2.092, 1.81, 2.16, 2.018, 1.966, 2.484]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.257623239117
varray: [2.628, 2.44, 2.299, 2.335, 2.092, 1.81, 2.16, 2.018, 1.966, 2.484]
inst: std_all
operation: std_all(slope_shift)
tarray: [10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: slope_shift
SAMP_d18O_STD | 0.382609475982
varray: [-12.217, -13.128, -13.347, -13.31, -13.406, -13.456, -13.424, -13.422, -13.464, -13.479]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 5.79048612045
varray: [-92.943, -105.802, -108.88, -109.777, -110.581, -110.82, -111.487, -111.33, -111.891, -112.047]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 107.8519
varray: [109.522, 106.723, 105.858, 105.837, 108.517, 109.136, 108.698, 109.188, 108.154, 106.886]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.4029989665
varray: [109.522, 106.723, 105.858, 105.837, 108.517, 109.136, 108.698, 109.188, 108.154, 106.886]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.774
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3652
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1291.9161
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.2232
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.02705
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.02705
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.965709287589
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.30325
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-12.217, -13.128, -13.347, -13.31, -13.406, -13.456, -13.424, -13.422, -13.464, -13.479]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-92.943, -105.802, -108.88, -109.777, -110.581, -110.82, -111.487, -111.33, -111.891, -112.047]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: d(D_H)Mean
known_dD | -119.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | -15.84
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -13.2653
varray: [-12.217, -13.128, -13.347, -13.31, -13.406, -13.456, -13.424, -13.422, -13.464, -13.479]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: d(18_16)Mean
dD_mean | -108.5558
varray: [-92.943, -105.802, -108.88, -109.777, -110.581, -110.82, -111.487, -111.33, -111.891, -112.047]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: d(D_H)Mean
d18O_stdev | 0.382609475982
varray: [-12.217, -13.128, -13.347, -13.31, -13.406, -13.456, -13.424, -13.422, -13.464, -13.479]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: d(18_16)Mean
dD_stdev | 5.79048612045
varray: [-92.943, -105.802, -108.88, -109.777, -110.581, -110.82, -111.487, -111.33, -111.891, -112.047]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: d(D_H)Mean
measured_d18O_mean | -13.2653
varray: [-12.217, -13.128, -13.347, -13.31, -13.406, -13.456, -13.424, -13.422, -13.464, -13.479]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: d(18_16)Mean
measured_dD_mean | -108.5558
varray: [-92.943, -105.802, -108.88, -109.777, -110.581, -110.82, -111.487, -111.33, -111.891, -112.047]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.382609475982
varray: [-12.217, -13.128, -13.347, -13.31, -13.406, -13.456, -13.424, -13.422, -13.464, -13.479]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: d(18_16)Mean
measured_dD_stdev | 5.79048612045
varray: [-92.943, -105.802, -108.88, -109.777, -110.581, -110.82, -111.487, -111.33, -111.891, -112.047]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -15.84
actual_array: [ 1.93 -15.84]
varray: [-12.217, -13.128, -13.347, -13.31, -13.406, -13.456, -13.424, -13.422, -13.464, -13.479]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -13.2653
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -119.0
actual_array: [ 18.1 -119. ]
varray: [-92.943, -105.802, -108.88, -109.777, -110.581, -110.82, -111.487, -111.33, -111.891, -112.047]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -108.5558
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [10846.0, 11383.0, 11920.0, 12458.0, 12996.0, 13533.0, 14069.0, 14607.0, 15144.0, 15681.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.12075
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.0213499999998
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 4 - Sample 4 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 4
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.7957
varray: [2.886, 2.71, 2.818, 2.815, 2.828, 2.764, 2.741, 2.769, 2.811, 2.815]
inst: mean_all
operation: mean_all(residuals)
tarray: [16271.0, 16809.0, 17345.0, 17883.0, 18419.0, 18956.0, 19493.0, 20030.0, 20568.0, 21106.0]
source_col: residuals
SAMP_residuals_STD | 0.050252805339
varray: [2.886, 2.71, 2.818, 2.815, 2.828, 2.764, 2.741, 2.769, 2.811, 2.815]
inst: std_all
operation: std_all(residuals)
tarray: [16271.0, 16809.0, 17345.0, 17883.0, 18419.0, 18956.0, 19493.0, 20030.0, 20568.0, 21106.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3029
varray: [0.678, 0.035, 0.447, 0.592, 0.39, 0.307, 0.09, 0.3, 0.085, 0.105]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [16271.0, 16809.0, 17345.0, 17883.0, 18419.0, 18956.0, 19493.0, 20030.0, 20568.0, 21106.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.225190511937
varray: [0.678, 0.035, 0.447, 0.592, 0.39, 0.307, 0.09, 0.3, 0.085, 0.105]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [16271.0, 16809.0, 17345.0, 17883.0, 18419.0, 18956.0, 19493.0, 20030.0, 20568.0, 21106.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1291.9125
varray: [1293.104, 1290.864, 1291.705, 1291.385, 1292.669, 1291.47, 1292.106, 1291.855, 1292.031, 1291.936]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [16271.0, 16809.0, 17345.0, 17883.0, 18419.0, 18956.0, 19493.0, 20030.0, 20568.0, 21106.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.638033306341
varray: [1293.104, 1290.864, 1291.705, 1291.385, 1292.669, 1291.47, 1292.106, 1291.855, 1292.031, 1291.936]
inst: std_all
operation: std_all(baseline_shift)
tarray: [16271.0, 16809.0, 17345.0, 17883.0, 18419.0, 18956.0, 19493.0, 20030.0, 20568.0, 21106.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.791
varray: [2.128, 1.994, 1.966, 1.742, 1.619, 2.305, 1.652, 1.514, 1.416, 1.574]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [16271.0, 16809.0, 17345.0, 17883.0, 18419.0, 18956.0, 19493.0, 20030.0, 20568.0, 21106.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.291613899989
varray: [2.128, 1.994, 1.966, 1.742, 1.619, 2.305, 1.652, 1.514, 1.416, 1.574]
inst: std_all
operation: std_all(slope_shift)
tarray: [16271.0, 16809.0, 17345.0, 17883.0, 18419.0, 18956.0, 19493.0, 20030.0, 20568.0, 21106.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0967553616085
varray: [-8.031, -7.811, -7.75, -7.768, -7.769, -7.691, -7.695, -7.732, -7.788, -7.739]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16271.0, 16809.0, 17345.0, 17883.0, 18419.0, 18956.0, 19493.0, 20030.0, 20568.0, 21106.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.86269489659
varray: [-66.537, -62.33, -61.365, -60.864, -60.764, -60.768, -60.403, -60.553, -60.877, -60.414]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16271.0, 16809.0, 17345.0, 17883.0, 18419.0, 18956.0, 19493.0, 20030.0, 20568.0, 21106.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.2026
varray: [110.74, 105.678, 107.905, 108.896, 109.69, 107.469, 107.541, 108.451, 107.942, 107.714]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [16271.0, 16809.0, 17345.0, 17883.0, 18419.0, 18956.0, 19493.0, 20030.0, 20568.0, 21106.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.37185844751
varray: [110.74, 105.678, 107.905, 108.896, 109.69, 107.469, 107.541, 108.451, 107.942, 107.714]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [16271.0, 16809.0, 17345.0, 17883.0, 18419.0, 18956.0, 19493.0, 20030.0, 20568.0, 21106.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.7957
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3029
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1291.9125
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.791
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.00535
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.00535
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.190999803645
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.04745
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-8.031, -7.811, -7.75, -7.768, -7.769, -7.691, -7.695, -7.732, -7.788, -7.739]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [16271.0, 16809.0, 17345.0, 17883.0, 18419.0, 18956.0, 19493.0, 20030.0, 20568.0, 21106.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-66.537, -62.33, -61.365, -60.864, -60.764, -60.768, -60.403, -60.553, -60.877, -60.414]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [16271.0, 16809.0, 17345.0, 17883.0, 18419.0, 18956.0, 19493.0, 20030.0, 20568.0, 21106.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.7774
varray: [-8.031, -7.811, -7.75, -7.768, -7.769, -7.691, -7.695, -7.732, -7.788, -7.739]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16271.0, 16809.0, 17345.0, 17883.0, 18419.0, 18956.0, 19493.0, 20030.0, 20568.0, 21106.0]
source_col: d(18_16)Mean
dD_mean | -61.4875
varray: [-66.537, -62.33, -61.365, -60.864, -60.764, -60.768, -60.403, -60.553, -60.877, -60.414]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16271.0, 16809.0, 17345.0, 17883.0, 18419.0, 18956.0, 19493.0, 20030.0, 20568.0, 21106.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0967553616085
varray: [-8.031, -7.811, -7.75, -7.768, -7.769, -7.691, -7.695, -7.732, -7.788, -7.739]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16271.0, 16809.0, 17345.0, 17883.0, 18419.0, 18956.0, 19493.0, 20030.0, 20568.0, 21106.0]
source_col: d(18_16)Mean
dD_stdev | 1.86269489659
varray: [-66.537, -62.33, -61.365, -60.864, -60.764, -60.768, -60.403, -60.553, -60.877, -60.414]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16271.0, 16809.0, 17345.0, 17883.0, 18419.0, 18956.0, 19493.0, 20030.0, 20568.0, 21106.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.7774
varray: [-8.031, -7.811, -7.75, -7.768, -7.769, -7.691, -7.695, -7.732, -7.788, -7.739]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16271.0, 16809.0, 17345.0, 17883.0, 18419.0, 18956.0, 19493.0, 20030.0, 20568.0, 21106.0]
source_col: d(18_16)Mean
measured_dD_mean | -61.4875
varray: [-66.537, -62.33, -61.365, -60.864, -60.764, -60.768, -60.403, -60.553, -60.877, -60.414]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16271.0, 16809.0, 17345.0, 17883.0, 18419.0, 18956.0, 19493.0, 20030.0, 20568.0, 21106.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0967553616085
varray: [-8.031, -7.811, -7.75, -7.768, -7.769, -7.691, -7.695, -7.732, -7.788, -7.739]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16271.0, 16809.0, 17345.0, 17883.0, 18419.0, 18956.0, 19493.0, 20030.0, 20568.0, 21106.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.86269489659
varray: [-66.537, -62.33, -61.365, -60.864, -60.764, -60.768, -60.403, -60.553, -60.877, -60.414]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16271.0, 16809.0, 17345.0, 17883.0, 18419.0, 18956.0, 19493.0, 20030.0, 20568.0, 21106.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.3668224472
actual_array: [ 1.93 -15.84]
varray: [-8.031, -7.811, -7.75, -7.768, -7.769, -7.691, -7.695, -7.732, -7.788, -7.739]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.7774
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [16271.0, 16809.0, 17345.0, 17883.0, 18419.0, 18956.0, 19493.0, 20030.0, 20568.0, 21106.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -74.9208815327
actual_array: [ 18.1 -119. ]
varray: [-66.537, -62.33, -61.365, -60.864, -60.764, -60.768, -60.403, -60.553, -60.877, -60.414]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -61.4875
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [16271.0, 16809.0, 17345.0, 17883.0, 18419.0, 18956.0, 19493.0, 20030.0, 20568.0, 21106.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.09905
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.01775
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 5 - Sample 5 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 5
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.861
varray: [2.833, 2.855, 2.903, 2.853]
inst: mean_all
operation: mean_all(residuals)
tarray: [21696.0, 22233.0, 22770.0, 23307.0]
source_col: residuals
SAMP_residuals_STD | 0.0297097066069
varray: [2.833, 2.855, 2.903, 2.853]
inst: std_all
operation: std_all(residuals)
tarray: [21696.0, 22233.0, 22770.0, 23307.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.17475
varray: [0.113, 0.16, 0.316, 0.11]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [21696.0, 22233.0, 22770.0, 23307.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0969101817836
varray: [0.113, 0.16, 0.316, 0.11]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [21696.0, 22233.0, 22770.0, 23307.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1293.986
varray: [1293.942, 1293.802, 1295.021, 1293.179]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [21696.0, 22233.0, 22770.0, 23307.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.765564279905
varray: [1293.942, 1293.802, 1295.021, 1293.179]
inst: std_all
operation: std_all(baseline_shift)
tarray: [21696.0, 22233.0, 22770.0, 23307.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.11075
varray: [1.161, 0.82, 1.116, 1.346]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [21696.0, 22233.0, 22770.0, 23307.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.217891983943
varray: [1.161, 0.82, 1.116, 1.346]
inst: std_all
operation: std_all(slope_shift)
tarray: [21696.0, 22233.0, 22770.0, 23307.0]
source_col: slope_shift
SAMP_d18O_STD | 0.219465069354
varray: [0.189, 0.535, 0.624, 0.677]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21696.0, 22233.0, 22770.0, 23307.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.76689592591
varray: [5.349, 11.125, 12.932, 13.665]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21696.0, 22233.0, 22770.0, 23307.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.43575
varray: [110.518, 110.354, 112.051, 108.82]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [21696.0, 22233.0, 22770.0, 23307.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.3207483674
varray: [110.518, 110.354, 112.051, 108.82]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [21696.0, 22233.0, 22770.0, 23307.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.861
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.17475
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1293.986
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.11075
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.05995
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.05995
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.14026882776
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 2.2806
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [0.189, 0.535, 0.624, 0.677]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [21696.0, 22233.0, 22770.0, 23307.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [5.349, 11.125, 12.932, 13.665]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [21696.0, 22233.0, 22770.0, 23307.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.50625
varray: [0.189, 0.535, 0.624, 0.677]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21696.0, 22233.0, 22770.0, 23307.0]
source_col: d(18_16)Mean
dD_mean | 10.76775
varray: [5.349, 11.125, 12.932, 13.665]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21696.0, 22233.0, 22770.0, 23307.0]
source_col: d(D_H)Mean
d18O_stdev | 0.219465069354
varray: [0.189, 0.535, 0.624, 0.677]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21696.0, 22233.0, 22770.0, 23307.0]
source_col: d(18_16)Mean
dD_stdev | 3.76689592591
varray: [5.349, 11.125, 12.932, 13.665]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21696.0, 22233.0, 22770.0, 23307.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.50625
varray: [0.189, 0.535, 0.624, 0.677]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21696.0, 22233.0, 22770.0, 23307.0]
source_col: d(18_16)Mean
measured_dD_mean | 10.76775
varray: [5.349, 11.125, 12.932, 13.665]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21696.0, 22233.0, 22770.0, 23307.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.219465069354
varray: [0.189, 0.535, 0.624, 0.677]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21696.0, 22233.0, 22770.0, 23307.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.76689592591
varray: [5.349, 11.125, 12.932, 13.665]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21696.0, 22233.0, 22770.0, 23307.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.10539508245
actual_array: [ 1.93 -15.84]
varray: [0.189, 0.535, 0.624, 0.677]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.50625
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [21696.0, 22233.0, 22770.0, 23307.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -7.25436872811
actual_array: [ 18.1 -119. ]
varray: [5.349, 11.125, 12.932, 13.665]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 10.76775
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [21696.0, 22233.0, 22770.0, 23307.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.03375
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 2.09125
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 6 - Sample 6 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 6
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.825
varray: [2.86, 2.792, 2.816, 2.832]
inst: mean_all
operation: mean_all(residuals)
tarray: [23897.0, 24435.0, 24972.0, 25509.0]
source_col: residuals
SAMP_residuals_STD | 0.0285423661715
varray: [2.86, 2.792, 2.816, 2.832]
inst: std_all
operation: std_all(residuals)
tarray: [23897.0, 24435.0, 24972.0, 25509.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.28025
varray: [0.259, 0.313, 0.281, 0.268]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [23897.0, 24435.0, 24972.0, 25509.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0236273147014
varray: [0.259, 0.313, 0.281, 0.268]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [23897.0, 24435.0, 24972.0, 25509.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1293.2765
varray: [1293.944, 1292.932, 1293.093, 1293.137]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [23897.0, 24435.0, 24972.0, 25509.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.453640459689
varray: [1293.944, 1292.932, 1293.093, 1293.137]
inst: std_all
operation: std_all(baseline_shift)
tarray: [23897.0, 24435.0, 24972.0, 25509.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.56825
varray: [1.627, 1.861, 1.631, 1.154]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [23897.0, 24435.0, 24972.0, 25509.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.297038016647
varray: [1.627, 1.861, 1.631, 1.154]
inst: std_all
operation: std_all(slope_shift)
tarray: [23897.0, 24435.0, 24972.0, 25509.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0336551135292
varray: [0.245, 0.282, 0.259, 0.322]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23897.0, 24435.0, 24972.0, 25509.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.13131990202
varray: [11.243, 11.35, 11.316, 11.55]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23897.0, 24435.0, 24972.0, 25509.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.976
varray: [109.367, 108.62, 108.498, 109.419]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [23897.0, 24435.0, 24972.0, 25509.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.484544459604
varray: [109.367, 108.62, 108.498, 109.419]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [23897.0, 24435.0, 24972.0, 25509.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.825
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.28025
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1293.2765
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.56825
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.02395
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.02395
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.855036504168
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.82085
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [0.245, 0.282, 0.259, 0.322]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [23897.0, 24435.0, 24972.0, 25509.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [11.243, 11.35, 11.316, 11.55]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [23897.0, 24435.0, 24972.0, 25509.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.277
varray: [0.245, 0.282, 0.259, 0.322]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23897.0, 24435.0, 24972.0, 25509.0]
source_col: d(18_16)Mean
dD_mean | 11.36475
varray: [11.243, 11.35, 11.316, 11.55]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23897.0, 24435.0, 24972.0, 25509.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0336551135292
varray: [0.245, 0.282, 0.259, 0.322]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23897.0, 24435.0, 24972.0, 25509.0]
source_col: d(18_16)Mean
dD_stdev | 0.13131990202
varray: [11.243, 11.35, 11.316, 11.55]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23897.0, 24435.0, 24972.0, 25509.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.277
varray: [0.245, 0.282, 0.259, 0.322]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23897.0, 24435.0, 24972.0, 25509.0]
source_col: d(18_16)Mean
measured_dD_mean | 11.36475
varray: [11.243, 11.35, 11.316, 11.55]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23897.0, 24435.0, 24972.0, 25509.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0336551135292
varray: [0.245, 0.282, 0.259, 0.322]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23897.0, 24435.0, 24972.0, 25509.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.13131990202
varray: [11.243, 11.35, 11.316, 11.55]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23897.0, 24435.0, 24972.0, 25509.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.33403007105
actual_array: [ 1.93 -15.84]
varray: [0.245, 0.282, 0.259, 0.322]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.277
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [23897.0, 24435.0, 24972.0, 25509.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -6.69528259742
actual_array: [ 18.1 -119. ]
varray: [11.243, 11.35, 11.316, 11.55]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 11.36475
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [23897.0, 24435.0, 24972.0, 25509.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.06975
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1.38175
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 7 - Sample 7 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 7
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.82775
varray: [2.882, 2.771, 2.811, 2.847]
inst: mean_all
operation: mean_all(residuals)
tarray: [26100.0, 26637.0, 27175.0, 27712.0]
source_col: residuals
SAMP_residuals_STD | 0.0476611302146
varray: [2.882, 2.771, 2.811, 2.847]
inst: std_all
operation: std_all(residuals)
tarray: [26100.0, 26637.0, 27175.0, 27712.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.29575
varray: [0.373, 0.346, 0.09, 0.374]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [26100.0, 26637.0, 27175.0, 27712.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.137778505822
varray: [0.373, 0.346, 0.09, 0.374]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [26100.0, 26637.0, 27175.0, 27712.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1293.21975
varray: [1294.234, 1291.731, 1292.999, 1293.915]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [26100.0, 26637.0, 27175.0, 27712.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.1220788371
varray: [1294.234, 1291.731, 1292.999, 1293.915]
inst: std_all
operation: std_all(baseline_shift)
tarray: [26100.0, 26637.0, 27175.0, 27712.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.4255
varray: [1.12, 1.408, 1.559, 1.615]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [26100.0, 26637.0, 27175.0, 27712.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.221637090759
varray: [1.12, 1.408, 1.559, 1.615]
inst: std_all
operation: std_all(slope_shift)
tarray: [26100.0, 26637.0, 27175.0, 27712.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0364725924497
varray: [0.316, 0.311, 0.382, 0.37]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26100.0, 26637.0, 27175.0, 27712.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.211995872601
varray: [11.441, 11.281, 11.621, 11.768]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26100.0, 26637.0, 27175.0, 27712.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.42925
varray: [111.257, 108.009, 108.172, 110.279]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [26100.0, 26637.0, 27175.0, 27712.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.59797087479
varray: [111.257, 108.009, 108.172, 110.279]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [26100.0, 26637.0, 27175.0, 27712.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.82775
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.29575
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1293.21975
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.4255
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0267
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0267
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.953213973331
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.2741
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [0.316, 0.311, 0.382, 0.37]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [26100.0, 26637.0, 27175.0, 27712.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [11.441, 11.281, 11.621, 11.768]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [26100.0, 26637.0, 27175.0, 27712.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.34475
varray: [0.316, 0.311, 0.382, 0.37]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26100.0, 26637.0, 27175.0, 27712.0]
source_col: d(18_16)Mean
dD_mean | 11.52775
varray: [11.441, 11.281, 11.621, 11.768]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26100.0, 26637.0, 27175.0, 27712.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0364725924497
varray: [0.316, 0.311, 0.382, 0.37]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26100.0, 26637.0, 27175.0, 27712.0]
source_col: d(18_16)Mean
dD_stdev | 0.211995872601
varray: [11.441, 11.281, 11.621, 11.768]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26100.0, 26637.0, 27175.0, 27712.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.34475
varray: [0.316, 0.311, 0.382, 0.37]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26100.0, 26637.0, 27175.0, 27712.0]
source_col: d(18_16)Mean
measured_dD_mean | 11.52775
varray: [11.441, 11.281, 11.621, 11.768]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26100.0, 26637.0, 27175.0, 27712.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0364725924497
varray: [0.316, 0.311, 0.382, 0.37]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26100.0, 26637.0, 27175.0, 27712.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.211995872601
varray: [11.441, 11.281, 11.621, 11.768]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26100.0, 26637.0, 27175.0, 27712.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.26646182469
actual_array: [ 1.93 -15.84]
varray: [0.316, 0.311, 0.382, 0.37]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.34475
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [26100.0, 26637.0, 27175.0, 27712.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -6.54263429039
actual_array: [ 18.1 -119. ]
varray: [11.441, 11.281, 11.621, 11.768]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 11.52775
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [26100.0, 26637.0, 27175.0, 27712.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.067
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1.325
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 8 - Sample 8 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 8
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.854
varray: [2.832, 2.887, 2.914, 2.783]
inst: mean_all
operation: mean_all(residuals)
tarray: [28300.0, 28837.0, 29374.0, 29911.0]
source_col: residuals
SAMP_residuals_STD | 0.0583495215633
varray: [2.832, 2.887, 2.914, 2.783]
inst: std_all
operation: std_all(residuals)
tarray: [28300.0, 28837.0, 29374.0, 29911.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.21325
varray: [0.229, 0.295, 0.341, -0.012]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [28300.0, 28837.0, 29374.0, 29911.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.15704431434
varray: [0.229, 0.295, 0.341, -0.012]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [28300.0, 28837.0, 29374.0, 29911.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1293.935
varray: [1293.606, 1293.664, 1294.425, 1294.045]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [28300.0, 28837.0, 29374.0, 29911.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.38029900517
varray: [1293.606, 1293.664, 1294.425, 1294.045]
inst: std_all
operation: std_all(baseline_shift)
tarray: [28300.0, 28837.0, 29374.0, 29911.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.24125
varray: [0.871, 1.029, 1.692, 1.373]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [28300.0, 28837.0, 29374.0, 29911.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.366364458793
varray: [0.871, 1.029, 1.692, 1.373]
inst: std_all
operation: std_all(slope_shift)
tarray: [28300.0, 28837.0, 29374.0, 29911.0]
source_col: slope_shift
SAMP_d18O_STD | 0.010214368964
varray: [0.489, 0.497, 0.474, 0.494]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28300.0, 28837.0, 29374.0, 29911.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0702489620327
varray: [12.429, 12.508, 12.568, 12.584]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28300.0, 28837.0, 29374.0, 29911.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.1825
varray: [110.03, 110.654, 111.268, 108.778]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [28300.0, 28837.0, 29374.0, 29911.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.06403305087
varray: [110.03, 110.654, 111.268, 108.778]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [28300.0, 28837.0, 29374.0, 29911.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.854
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.21325
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1293.935
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.24125
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.05295
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.05295
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.89036254262
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 2.02735
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [0.489, 0.497, 0.474, 0.494]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [28300.0, 28837.0, 29374.0, 29911.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [12.429, 12.508, 12.568, 12.584]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [28300.0, 28837.0, 29374.0, 29911.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.4885
varray: [0.489, 0.497, 0.474, 0.494]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28300.0, 28837.0, 29374.0, 29911.0]
source_col: d(18_16)Mean
dD_mean | 12.52225
varray: [12.429, 12.508, 12.568, 12.584]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28300.0, 28837.0, 29374.0, 29911.0]
source_col: d(D_H)Mean
d18O_stdev | 0.010214368964
varray: [0.489, 0.497, 0.474, 0.494]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28300.0, 28837.0, 29374.0, 29911.0]
source_col: d(18_16)Mean
dD_stdev | 0.0702489620327
varray: [12.429, 12.508, 12.568, 12.584]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28300.0, 28837.0, 29374.0, 29911.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.4885
varray: [0.489, 0.497, 0.474, 0.494]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28300.0, 28837.0, 29374.0, 29911.0]
source_col: d(18_16)Mean
measured_dD_mean | 12.52225
varray: [12.429, 12.508, 12.568, 12.584]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28300.0, 28837.0, 29374.0, 29911.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.010214368964
varray: [0.489, 0.497, 0.474, 0.494]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28300.0, 28837.0, 29374.0, 29911.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0702489620327
varray: [12.429, 12.508, 12.568, 12.584]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28300.0, 28837.0, 29374.0, 29911.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.12309746433
actual_array: [ 1.93 -15.84]
varray: [0.489, 0.497, 0.474, 0.494]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.4885
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [28300.0, 28837.0, 29374.0, 29911.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -5.61129231892
actual_array: [ 18.1 -119. ]
varray: [12.429, 12.508, 12.568, 12.584]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 12.52225
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [28300.0, 28837.0, 29374.0, 29911.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.04075
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 2.04025
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 9 - Sample 9 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 9
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8295
varray: [2.771, 2.838, 2.847, 2.862]
inst: mean_all
operation: mean_all(residuals)
tarray: [30499.0, 31035.0, 31573.0, 32110.0]
source_col: residuals
SAMP_residuals_STD | 0.0402367990775
varray: [2.771, 2.838, 2.847, 2.862]
inst: std_all
operation: std_all(residuals)
tarray: [30499.0, 31035.0, 31573.0, 32110.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.23975
varray: [0.074, 0.312, 0.277, 0.296]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [30499.0, 31035.0, 31573.0, 32110.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.111422244936
varray: [0.074, 0.312, 0.277, 0.296]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [30499.0, 31035.0, 31573.0, 32110.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1293.351
varray: [1292.49, 1293.337, 1293.865, 1293.712]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [30499.0, 31035.0, 31573.0, 32110.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.615368182473
varray: [1292.49, 1293.337, 1293.865, 1293.712]
inst: std_all
operation: std_all(baseline_shift)
tarray: [30499.0, 31035.0, 31573.0, 32110.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.88025
varray: [1.112, 0.235, 0.944, 1.23]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [30499.0, 31035.0, 31573.0, 32110.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.445886663477
varray: [1.112, 0.235, 0.944, 1.23]
inst: std_all
operation: std_all(slope_shift)
tarray: [30499.0, 31035.0, 31573.0, 32110.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0145916642871
varray: [-0.071, -0.059, -0.048, -0.037]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30499.0, 31035.0, 31573.0, 32110.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.12598148012
varray: [9.294, 9.137, 9.114, 8.987]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30499.0, 31035.0, 31573.0, 32110.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.284
varray: [108.355, 111.149, 111.194, 110.438]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [30499.0, 31035.0, 31573.0, 32110.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.33180103619
varray: [108.355, 111.149, 111.194, 110.438]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [30499.0, 31035.0, 31573.0, 32110.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8295
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.23975
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1293.351
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.88025
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.02845
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.02845
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.01569054462
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 2.12885
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-0.071, -0.059, -0.048, -0.037]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [30499.0, 31035.0, 31573.0, 32110.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [9.294, 9.137, 9.114, 8.987]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [30499.0, 31035.0, 31573.0, 32110.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.05375
varray: [-0.071, -0.059, -0.048, -0.037]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30499.0, 31035.0, 31573.0, 32110.0]
source_col: d(18_16)Mean
dD_mean | 9.133
varray: [9.294, 9.137, 9.114, 8.987]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30499.0, 31035.0, 31573.0, 32110.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0145916642871
varray: [-0.071, -0.059, -0.048, -0.037]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30499.0, 31035.0, 31573.0, 32110.0]
source_col: d(18_16)Mean
dD_stdev | 0.12598148012
varray: [9.294, 9.137, 9.114, 8.987]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30499.0, 31035.0, 31573.0, 32110.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.05375
varray: [-0.071, -0.059, -0.048, -0.037]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30499.0, 31035.0, 31573.0, 32110.0]
source_col: d(18_16)Mean
measured_dD_mean | 9.133
varray: [9.294, 9.137, 9.114, 8.987]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30499.0, 31035.0, 31573.0, 32110.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0145916642871
varray: [-0.071, -0.059, -0.048, -0.037]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30499.0, 31035.0, 31573.0, 32110.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.12598148012
varray: [9.294, 9.137, 9.114, 8.987]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30499.0, 31035.0, 31573.0, 32110.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.66389276454
actual_array: [ 1.93 -15.84]
varray: [-0.071, -0.059, -0.048, -0.037]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -0.05375
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [30499.0, 31035.0, 31573.0, 32110.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -8.78530013873
actual_array: [ 18.1 -119. ]
varray: [9.294, 9.137, 9.114, 8.987]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 9.133
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [30499.0, 31035.0, 31573.0, 32110.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.06525
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1.45625
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 10 - Sample 10 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 10
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.782
varray: [2.811, 2.747, 2.773, 2.797]
inst: mean_all
operation: mean_all(residuals)
tarray: [32699.0, 33236.0, 33773.0, 34311.0]
source_col: residuals
SAMP_residuals_STD | 0.028118795612
varray: [2.811, 2.747, 2.773, 2.797]
inst: std_all
operation: std_all(residuals)
tarray: [32699.0, 33236.0, 33773.0, 34311.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.26
varray: [0.363, 0.095, 0.245, 0.337]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [32699.0, 33236.0, 33773.0, 34311.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.121089498581
varray: [0.363, 0.095, 0.245, 0.337]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [32699.0, 33236.0, 33773.0, 34311.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1290.8755
varray: [1292.621, 1290.366, 1289.449, 1291.066]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [32699.0, 33236.0, 33773.0, 34311.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.33884938162
varray: [1292.621, 1290.366, 1289.449, 1291.066]
inst: std_all
operation: std_all(baseline_shift)
tarray: [32699.0, 33236.0, 33773.0, 34311.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.3325
varray: [0.977, 1.217, 1.369, 1.767]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [32699.0, 33236.0, 33773.0, 34311.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.331583574181
varray: [0.977, 1.217, 1.369, 1.767]
inst: std_all
operation: std_all(slope_shift)
tarray: [32699.0, 33236.0, 33773.0, 34311.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0176351920885
varray: [0.194, 0.219, 0.177, 0.204]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32699.0, 33236.0, 33773.0, 34311.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0145258390463
varray: [9.902, 9.91, 9.877, 9.889]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32699.0, 33236.0, 33773.0, 34311.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 107.84325
varray: [109.897, 107.024, 106.399, 108.053]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [32699.0, 33236.0, 33773.0, 34311.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.5295873899
varray: [109.897, 107.024, 106.399, 108.053]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [32699.0, 33236.0, 33773.0, 34311.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.782
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.26
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1290.8755
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.3325
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.01905
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.01905
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.680102104568
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.3119
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [0.194, 0.219, 0.177, 0.204]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [32699.0, 33236.0, 33773.0, 34311.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [9.902, 9.91, 9.877, 9.889]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [32699.0, 33236.0, 33773.0, 34311.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.1985
varray: [0.194, 0.219, 0.177, 0.204]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32699.0, 33236.0, 33773.0, 34311.0]
source_col: d(18_16)Mean
dD_mean | 9.8945
varray: [9.902, 9.91, 9.877, 9.889]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32699.0, 33236.0, 33773.0, 34311.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0176351920885
varray: [0.194, 0.219, 0.177, 0.204]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32699.0, 33236.0, 33773.0, 34311.0]
source_col: d(18_16)Mean
dD_stdev | 0.0145258390463
varray: [9.902, 9.91, 9.877, 9.889]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32699.0, 33236.0, 33773.0, 34311.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.1985
varray: [0.194, 0.219, 0.177, 0.204]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32699.0, 33236.0, 33773.0, 34311.0]
source_col: d(18_16)Mean
measured_dD_mean | 9.8945
varray: [9.902, 9.91, 9.877, 9.889]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32699.0, 33236.0, 33773.0, 34311.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0176351920885
varray: [0.194, 0.219, 0.177, 0.204]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32699.0, 33236.0, 33773.0, 34311.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0145258390463
varray: [9.902, 9.91, 9.877, 9.889]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32699.0, 33236.0, 33773.0, 34311.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.41231947827
actual_array: [ 1.93 -15.84]
varray: [0.194, 0.219, 0.177, 0.204]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.1985
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [32699.0, 33236.0, 33773.0, 34311.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -8.07216096198
actual_array: [ 18.1 -119. ]
varray: [9.902, 9.91, 9.877, 9.889]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 9.8945
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [32699.0, 33236.0, 33773.0, 34311.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.11275
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.01925
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 11 - Sample 11 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 11
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8195
varray: [2.88, 2.758, 2.846, 2.794]
inst: mean_all
operation: mean_all(residuals)
tarray: [34900.0, 35438.0, 35975.0, 36512.0]
source_col: residuals
SAMP_residuals_STD | 0.0541448674083
varray: [2.88, 2.758, 2.846, 2.794]
inst: std_all
operation: std_all(residuals)
tarray: [34900.0, 35438.0, 35975.0, 36512.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.27975
varray: [0.468, -0.039, 0.557, 0.133]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [34900.0, 35438.0, 35975.0, 36512.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.280144456784
varray: [0.468, -0.039, 0.557, 0.133]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [34900.0, 35438.0, 35975.0, 36512.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1292.38475
varray: [1293.225, 1291.625, 1292.178, 1292.511]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [34900.0, 35438.0, 35975.0, 36512.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.668811071978
varray: [1293.225, 1291.625, 1292.178, 1292.511]
inst: std_all
operation: std_all(baseline_shift)
tarray: [34900.0, 35438.0, 35975.0, 36512.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.6445
varray: [1.315, 1.829, 1.728, 1.706]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [34900.0, 35438.0, 35975.0, 36512.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.226101009875
varray: [1.315, 1.829, 1.728, 1.706]
inst: std_all
operation: std_all(slope_shift)
tarray: [34900.0, 35438.0, 35975.0, 36512.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0410406302746
varray: [-0.152, -0.205, -0.227, -0.142]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34900.0, 35438.0, 35975.0, 36512.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.162011059293
varray: [8.947, 8.711, 8.696, 9.013]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34900.0, 35438.0, 35975.0, 36512.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.67675
varray: [110.75, 106.48, 109.842, 107.635]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [34900.0, 35438.0, 35975.0, 36512.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.96359591481
varray: [110.75, 106.48, 109.842, 107.635]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [34900.0, 35438.0, 35975.0, 36512.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8195
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.27975
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1292.38475
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.6445
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.01845
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.01845
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.658681565841
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.5216
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-0.152, -0.205, -0.227, -0.142]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [34900.0, 35438.0, 35975.0, 36512.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [8.947, 8.711, 8.696, 9.013]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [34900.0, 35438.0, 35975.0, 36512.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.1815
varray: [-0.152, -0.205, -0.227, -0.142]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [34900.0, 35438.0, 35975.0, 36512.0]
source_col: d(18_16)Mean
dD_mean | 8.84175
varray: [8.947, 8.711, 8.696, 9.013]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [34900.0, 35438.0, 35975.0, 36512.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0410406302746
varray: [-0.152, -0.205, -0.227, -0.142]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34900.0, 35438.0, 35975.0, 36512.0]
source_col: d(18_16)Mean
dD_stdev | 0.162011059293
varray: [8.947, 8.711, 8.696, 9.013]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34900.0, 35438.0, 35975.0, 36512.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.1815
varray: [-0.152, -0.205, -0.227, -0.142]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [34900.0, 35438.0, 35975.0, 36512.0]
source_col: d(18_16)Mean
measured_dD_mean | 8.84175
varray: [8.947, 8.711, 8.696, 9.013]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [34900.0, 35438.0, 35975.0, 36512.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0410406302746
varray: [-0.152, -0.205, -0.227, -0.142]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34900.0, 35438.0, 35975.0, 36512.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.162011059293
varray: [8.947, 8.711, 8.696, 9.013]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34900.0, 35438.0, 35975.0, 36512.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.79130004827
actual_array: [ 1.93 -15.84]
varray: [-0.152, -0.205, -0.227, -0.142]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -0.1815
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [34900.0, 35438.0, 35975.0, 36512.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -9.05805363214
actual_array: [ 18.1 -119. ]
varray: [8.947, 8.711, 8.696, 9.013]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 8.84175
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [34900.0, 35438.0, 35975.0, 36512.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.07525
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.49
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 12 - Sample 12 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 12
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.82325
varray: [2.883, 2.76, 2.824, 2.826]
inst: mean_all
operation: mean_all(residuals)
tarray: [37101.0, 37639.0, 38176.0, 38713.0]
source_col: residuals
SAMP_residuals_STD | 0.0502618145315
varray: [2.883, 2.76, 2.824, 2.826]
inst: std_all
operation: std_all(residuals)
tarray: [37101.0, 37639.0, 38176.0, 38713.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.21475
varray: [0.435, -0.136, 0.194, 0.366]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [37101.0, 37639.0, 38176.0, 38713.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.254848157406
varray: [0.435, -0.136, 0.194, 0.366]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [37101.0, 37639.0, 38176.0, 38713.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1292.738
varray: [1292.604, 1293.168, 1292.37, 1292.81]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [37101.0, 37639.0, 38176.0, 38713.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.338360754225
varray: [1292.604, 1293.168, 1292.37, 1292.81]
inst: std_all
operation: std_all(baseline_shift)
tarray: [37101.0, 37639.0, 38176.0, 38713.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.30925
varray: [1.292, 1.054, 1.608, 1.283]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [37101.0, 37639.0, 38176.0, 38713.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.22758935974
varray: [1.292, 1.054, 1.608, 1.283]
inst: std_all
operation: std_all(slope_shift)
tarray: [37101.0, 37639.0, 38176.0, 38713.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0261470202254
varray: [-0.15, -0.13, -0.096, -0.098]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37101.0, 37639.0, 38176.0, 38713.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.29494858309
varray: [8.747, 8.927, 9.325, 8.661]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37101.0, 37639.0, 38176.0, 38713.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.50925
varray: [109.767, 107.662, 107.526, 109.082]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [37101.0, 37639.0, 38176.0, 38713.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.09462181445
varray: [109.767, 107.662, 107.526, 109.082]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [37101.0, 37639.0, 38176.0, 38713.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.82325
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.21475
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1292.738
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.30925
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0222
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0222
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.792559932882
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.3541
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-0.15, -0.13, -0.096, -0.098]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [37101.0, 37639.0, 38176.0, 38713.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [8.747, 8.927, 9.325, 8.661]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [37101.0, 37639.0, 38176.0, 38713.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.1185
varray: [-0.15, -0.13, -0.096, -0.098]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37101.0, 37639.0, 38176.0, 38713.0]
source_col: d(18_16)Mean
dD_mean | 8.915
varray: [8.747, 8.927, 9.325, 8.661]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37101.0, 37639.0, 38176.0, 38713.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0261470202254
varray: [-0.15, -0.13, -0.096, -0.098]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37101.0, 37639.0, 38176.0, 38713.0]
source_col: d(18_16)Mean
dD_stdev | 0.29494858309
varray: [8.747, 8.927, 9.325, 8.661]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37101.0, 37639.0, 38176.0, 38713.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.1185
varray: [-0.15, -0.13, -0.096, -0.098]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37101.0, 37639.0, 38176.0, 38713.0]
source_col: d(18_16)Mean
measured_dD_mean | 8.915
varray: [8.747, 8.927, 9.325, 8.661]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37101.0, 37639.0, 38176.0, 38713.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0261470202254
varray: [-0.15, -0.13, -0.096, -0.098]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37101.0, 37639.0, 38176.0, 38713.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.29494858309
varray: [8.747, 8.927, 9.325, 8.661]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37101.0, 37639.0, 38176.0, 38713.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.72846905903
actual_array: [ 1.93 -15.84]
varray: [-0.15, -0.13, -0.096, -0.098]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -0.1185
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [37101.0, 37639.0, 38176.0, 38713.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -8.98945554324
actual_array: [ 18.1 -119. ]
varray: [8.747, 8.927, 9.325, 8.661]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 8.915
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [37101.0, 37639.0, 38176.0, 38713.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.0715
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.84325
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 13 - Sample 13 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 13
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.79625
varray: [2.782, 2.771, 2.818, 2.814]
inst: mean_all
operation: mean_all(residuals)
tarray: [39303.0, 39840.0, 40378.0, 40915.0]
source_col: residuals
SAMP_residuals_STD | 0.023300572239
varray: [2.782, 2.771, 2.818, 2.814]
inst: std_all
operation: std_all(residuals)
tarray: [39303.0, 39840.0, 40378.0, 40915.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.32325
varray: [0.552, 0.231, 0.34, 0.17]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [39303.0, 39840.0, 40378.0, 40915.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.167931285551
varray: [0.552, 0.231, 0.34, 0.17]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [39303.0, 39840.0, 40378.0, 40915.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1292.35875
varray: [1291.498, 1292.261, 1292.869, 1292.807]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [39303.0, 39840.0, 40378.0, 40915.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.635538813396
varray: [1291.498, 1292.261, 1292.869, 1292.807]
inst: std_all
operation: std_all(baseline_shift)
tarray: [39303.0, 39840.0, 40378.0, 40915.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.761
varray: [1.796, 1.771, 1.71, 1.767]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [39303.0, 39840.0, 40378.0, 40915.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0363409777891
varray: [1.796, 1.771, 1.71, 1.767]
inst: std_all
operation: std_all(slope_shift)
tarray: [39303.0, 39840.0, 40378.0, 40915.0]
source_col: slope_shift
SAMP_d18O_STD | 0.184434947881
varray: [-7.179, -7.438, -7.551, -7.587]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39303.0, 39840.0, 40378.0, 40915.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.41805470992
varray: [-51.202, -56.523, -57.919, -58.852]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39303.0, 39840.0, 40378.0, 40915.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.53025
varray: [108.772, 107.982, 109.192, 108.175]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [39303.0, 39840.0, 40378.0, 40915.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.554718171687
varray: [108.772, 107.982, 109.192, 108.175]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [39303.0, 39840.0, 40378.0, 40915.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.79625
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.32325
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1292.35875
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.761
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0048
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0048
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.171364309812
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.3751
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-7.179, -7.438, -7.551, -7.587]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [39303.0, 39840.0, 40378.0, 40915.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-51.202, -56.523, -57.919, -58.852]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [39303.0, 39840.0, 40378.0, 40915.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.43875
varray: [-7.179, -7.438, -7.551, -7.587]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39303.0, 39840.0, 40378.0, 40915.0]
source_col: d(18_16)Mean
dD_mean | -56.124
varray: [-51.202, -56.523, -57.919, -58.852]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39303.0, 39840.0, 40378.0, 40915.0]
source_col: d(D_H)Mean
d18O_stdev | 0.184434947881
varray: [-7.179, -7.438, -7.551, -7.587]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39303.0, 39840.0, 40378.0, 40915.0]
source_col: d(18_16)Mean
dD_stdev | 3.41805470992
varray: [-51.202, -56.523, -57.919, -58.852]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39303.0, 39840.0, 40378.0, 40915.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.43875
varray: [-7.179, -7.438, -7.551, -7.587]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39303.0, 39840.0, 40378.0, 40915.0]
source_col: d(18_16)Mean
measured_dD_mean | -56.124
varray: [-51.202, -56.523, -57.919, -58.852]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39303.0, 39840.0, 40378.0, 40915.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.184434947881
varray: [-7.179, -7.438, -7.551, -7.587]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39303.0, 39840.0, 40378.0, 40915.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.41805470992
varray: [-51.202, -56.523, -57.919, -58.852]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39303.0, 39840.0, 40378.0, 40915.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.0290809471
actual_array: [ 1.93 -15.84]
varray: [-7.179, -7.438, -7.551, -7.587]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.43875
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [39303.0, 39840.0, 40378.0, 40915.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -69.8980030369
actual_array: [ 18.1 -119. ]
varray: [-51.202, -56.523, -57.919, -58.852]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -56.124
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [39303.0, 39840.0, 40378.0, 40915.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.0985
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.464
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 14 - Sample 14 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 14
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.86475
varray: [2.802, 2.89, 2.901, 2.866]
inst: mean_all
operation: mean_all(residuals)
tarray: [41505.0, 42043.0, 42580.0, 43118.0]
source_col: residuals
SAMP_residuals_STD | 0.0443123383871
varray: [2.802, 2.89, 2.901, 2.866]
inst: std_all
operation: std_all(residuals)
tarray: [41505.0, 42043.0, 42580.0, 43118.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.44475
varray: [0.19, 0.456, 0.511, 0.622]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [41505.0, 42043.0, 42580.0, 43118.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.183330984833
varray: [0.19, 0.456, 0.511, 0.622]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [41505.0, 42043.0, 42580.0, 43118.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1292.604
varray: [1292.796, 1292.406, 1292.773, 1292.441]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [41505.0, 42043.0, 42580.0, 43118.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.209123567937
varray: [1292.796, 1292.406, 1292.773, 1292.441]
inst: std_all
operation: std_all(baseline_shift)
tarray: [41505.0, 42043.0, 42580.0, 43118.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.5375
varray: [1.516, 1.546, 1.692, 1.396]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [41505.0, 42043.0, 42580.0, 43118.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.121692234756
varray: [1.516, 1.546, 1.692, 1.396]
inst: std_all
operation: std_all(slope_shift)
tarray: [41505.0, 42043.0, 42580.0, 43118.0]
source_col: slope_shift
SAMP_d18O_STD | 0.21674158961
varray: [-0.008, 0.386, 0.446, 0.435]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41505.0, 42043.0, 42580.0, 43118.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.68878140447
varray: [4.49, 10.547, 11.963, 12.523]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41505.0, 42043.0, 42580.0, 43118.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.77875
varray: [107.724, 110.618, 109.775, 110.998]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [41505.0, 42043.0, 42580.0, 43118.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.46206734341
varray: [107.724, 110.618, 109.775, 110.998]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [41505.0, 42043.0, 42580.0, 43118.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.86475
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.44475
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1292.604
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.5375
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0637
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0637
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.2741471948
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.6236
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-0.008, 0.386, 0.446, 0.435]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [41505.0, 42043.0, 42580.0, 43118.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [4.49, 10.547, 11.963, 12.523]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [41505.0, 42043.0, 42580.0, 43118.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.31475
varray: [-0.008, 0.386, 0.446, 0.435]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [41505.0, 42043.0, 42580.0, 43118.0]
source_col: d(18_16)Mean
dD_mean | 9.88075
varray: [4.49, 10.547, 11.963, 12.523]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [41505.0, 42043.0, 42580.0, 43118.0]
source_col: d(D_H)Mean
d18O_stdev | 0.21674158961
varray: [-0.008, 0.386, 0.446, 0.435]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41505.0, 42043.0, 42580.0, 43118.0]
source_col: d(18_16)Mean
dD_stdev | 3.68878140447
varray: [4.49, 10.547, 11.963, 12.523]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41505.0, 42043.0, 42580.0, 43118.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.31475
varray: [-0.008, 0.386, 0.446, 0.435]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [41505.0, 42043.0, 42580.0, 43118.0]
source_col: d(18_16)Mean
measured_dD_mean | 9.88075
varray: [4.49, 10.547, 11.963, 12.523]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [41505.0, 42043.0, 42580.0, 43118.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.21674158961
varray: [-0.008, 0.386, 0.446, 0.435]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41505.0, 42043.0, 42580.0, 43118.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.68878140447
varray: [4.49, 10.547, 11.963, 12.523]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41505.0, 42043.0, 42580.0, 43118.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.29638134338
actual_array: [ 1.93 -15.84]
varray: [-0.008, 0.386, 0.446, 0.435]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.31475
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [41505.0, 42043.0, 42580.0, 43118.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -8.08503773635
actual_array: [ 18.1 -119. ]
varray: [4.49, 10.547, 11.963, 12.523]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 9.88075
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [41505.0, 42043.0, 42580.0, 43118.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.03
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.70925
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 15 - Sample 15 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 15
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.815
varray: [2.779, 2.832, 2.84, 2.809]
inst: mean_all
operation: mean_all(residuals)
tarray: [43705.0, 44242.0, 44779.0, 45316.0]
source_col: residuals
SAMP_residuals_STD | 0.0273617738216
varray: [2.779, 2.832, 2.84, 2.809]
inst: std_all
operation: std_all(residuals)
tarray: [43705.0, 44242.0, 44779.0, 45316.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.1975
varray: [0.059, 0.228, 0.268, 0.235]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [43705.0, 44242.0, 44779.0, 45316.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.093966306018
varray: [0.059, 0.228, 0.268, 0.235]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [43705.0, 44242.0, 44779.0, 45316.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1293.4805
varray: [1292.397, 1293.917, 1293.588, 1294.02]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [43705.0, 44242.0, 44779.0, 45316.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.745457577599
varray: [1292.397, 1293.917, 1293.588, 1294.02]
inst: std_all
operation: std_all(baseline_shift)
tarray: [43705.0, 44242.0, 44779.0, 45316.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.402
varray: [1.345, 1.043, 1.348, 1.872]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [43705.0, 44242.0, 44779.0, 45316.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.344454157569
varray: [1.345, 1.043, 1.348, 1.872]
inst: std_all
operation: std_all(slope_shift)
tarray: [43705.0, 44242.0, 44779.0, 45316.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0398026381035
varray: [0.19, 0.262, 0.227, 0.28]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43705.0, 44242.0, 44779.0, 45316.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.229735173624
varray: [9.879, 10.275, 10.34, 10.377]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43705.0, 44242.0, 44779.0, 45316.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.0745
varray: [107.377, 110.163, 109.36, 109.398]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [43705.0, 44242.0, 44779.0, 45316.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.19058822997
varray: [107.377, 110.163, 109.36, 109.398]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [43705.0, 44242.0, 44779.0, 45316.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.815
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.1975
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1293.4805
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.402
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.01395
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.01395
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.498027525392
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.91935
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [0.19, 0.262, 0.227, 0.28]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [43705.0, 44242.0, 44779.0, 45316.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [9.879, 10.275, 10.34, 10.377]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [43705.0, 44242.0, 44779.0, 45316.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.23975
varray: [0.19, 0.262, 0.227, 0.28]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [43705.0, 44242.0, 44779.0, 45316.0]
source_col: d(18_16)Mean
dD_mean | 10.21775
varray: [9.879, 10.275, 10.34, 10.377]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [43705.0, 44242.0, 44779.0, 45316.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0398026381035
varray: [0.19, 0.262, 0.227, 0.28]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43705.0, 44242.0, 44779.0, 45316.0]
source_col: d(18_16)Mean
dD_stdev | 0.229735173624
varray: [9.879, 10.275, 10.34, 10.377]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43705.0, 44242.0, 44779.0, 45316.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.23975
varray: [0.19, 0.262, 0.227, 0.28]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [43705.0, 44242.0, 44779.0, 45316.0]
source_col: d(18_16)Mean
measured_dD_mean | 10.21775
varray: [9.879, 10.275, 10.34, 10.377]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [43705.0, 44242.0, 44779.0, 45316.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0398026381035
varray: [0.19, 0.262, 0.227, 0.28]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43705.0, 44242.0, 44779.0, 45316.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.229735173624
varray: [9.879, 10.275, 10.34, 10.377]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43705.0, 44242.0, 44779.0, 45316.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.37118014008
actual_array: [ 1.93 -15.84]
varray: [0.19, 0.262, 0.227, 0.28]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.23975
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [43705.0, 44242.0, 44779.0, 45316.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -7.76943970278
actual_array: [ 18.1 -119. ]
varray: [9.879, 10.275, 10.34, 10.377]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 10.21775
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [43705.0, 44242.0, 44779.0, 45316.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.07975
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1.58575
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 16 - Sample 16 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 16
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.79633333333
varray: [2.806, 2.807, 2.776]
inst: mean_all
operation: mean_all(residuals)
tarray: [45903.0, 46441.0, 46978.0]
source_col: residuals
SAMP_residuals_STD | 0.017616280349
varray: [2.806, 2.807, 2.776]
inst: std_all
operation: std_all(residuals)
tarray: [45903.0, 46441.0, 46978.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.299
varray: [0.259, 0.395, 0.243]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [45903.0, 46441.0, 46978.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0835224520713
varray: [0.259, 0.395, 0.243]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [45903.0, 46441.0, 46978.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1291.03266667
varray: [1291.9, 1290.757, 1290.441]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [45903.0, 46441.0, 46978.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.767570409365
varray: [1291.9, 1290.757, 1290.441]
inst: std_all
operation: std_all(baseline_shift)
tarray: [45903.0, 46441.0, 46978.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.46566666667
varray: [1.155, 1.669, 1.573]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [45903.0, 46441.0, 46978.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.273293493031
varray: [1.155, 1.669, 1.573]
inst: std_all
operation: std_all(slope_shift)
tarray: [45903.0, 46441.0, 46978.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0180369990113
varray: [-0.008, -0.028, -0.044]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45903.0, 46441.0, 46978.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0428991064398
varray: [9.481, 9.546, 9.562]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45903.0, 46441.0, 46978.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 107.438333333
varray: [108.081, 107.683, 106.551]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [45903.0, 46441.0, 46978.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.793801822455
varray: [108.081, 107.683, 106.551]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [45903.0, 46441.0, 46978.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.79633333333
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.299
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1291.03266667
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.46566666667
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.00471666666667
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.00471666666667
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.168389234989
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.716816666667
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-0.008, -0.028, -0.044]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [45903.0, 46441.0, 46978.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [9.481, 9.546, 9.562]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [45903.0, 46441.0, 46978.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.0266666666667
varray: [-0.008, -0.028, -0.044]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45903.0, 46441.0, 46978.0]
source_col: d(18_16)Mean
dD_mean | 9.52966666667
varray: [9.481, 9.546, 9.562]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45903.0, 46441.0, 46978.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0180369990113
varray: [-0.008, -0.028, -0.044]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45903.0, 46441.0, 46978.0]
source_col: d(18_16)Mean
dD_stdev | 0.0428991064398
varray: [9.481, 9.546, 9.562]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45903.0, 46441.0, 46978.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.0266666666667
varray: [-0.008, -0.028, -0.044]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45903.0, 46441.0, 46978.0]
source_col: d(18_16)Mean
measured_dD_mean | 9.52966666667
varray: [9.481, 9.546, 9.562]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45903.0, 46441.0, 46978.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0180369990113
varray: [-0.008, -0.028, -0.044]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45903.0, 46441.0, 46978.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0428991064398
varray: [9.481, 9.546, 9.562]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45903.0, 46441.0, 46978.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.63688208795
actual_array: [ 1.93 -15.84]
varray: [-0.008, -0.028, -0.044]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -0.0266666666667
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [45903.0, 46441.0, 46978.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -8.41382470852
actual_array: [ 18.1 -119. ]
varray: [9.481, 9.546, 9.562]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 9.52966666667
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [45903.0, 46441.0, 46978.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.09841667
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.862083333333
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 17 - Sample 17 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 17
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.2664
varray: [0.0, 2.797, 2.807, 2.838, 2.89]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [47540.0, 48192.0, 48729.0, 49266.0, 49803.0]
source_col: residuals
SAMP_residuals_STD | 1.26747358947
varray: [0.0, 2.797, 2.807, 2.838, 2.89]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [47540.0, 48192.0, 48729.0, 49266.0, 49803.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.2036
varray: [0.0, 0.476, -0.05, 0.369, 0.223]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [47540.0, 48192.0, 48729.0, 49266.0, 49803.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.22787123557
varray: [0.0, 0.476, -0.05, 0.369, 0.223]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [47540.0, 48192.0, 48729.0, 49266.0, 49803.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1034.1732
varray: [0.0, 1290.514, 1294.095, 1293.358, 1292.899]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [47540.0, 48192.0, 48729.0, 49266.0, 49803.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 578.12194985
varray: [0.0, 1290.514, 1294.095, 1293.358, 1292.899]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [47540.0, 48192.0, 48729.0, 49266.0, 49803.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.0974
varray: [0.0, 1.212, 1.723, 1.293, 1.259]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [47540.0, 48192.0, 48729.0, 49266.0, 49803.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.646755208715
varray: [0.0, 1.212, 1.723, 1.293, 1.259]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [47540.0, 48192.0, 48729.0, 49266.0, 49803.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0555184653967
varray: [0.0, -0.049, 0.094, 0.015, -0.032]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [47540.0, 48192.0, 48729.0, 49266.0, 49803.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 4.45052646324
varray: [0.0, 9.78, 10.452, 9.88, 9.591]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [47540.0, 48192.0, 48729.0, 49266.0, 49803.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 87.3576
varray: [0.0, 108.038, 108.457, 110.558, 109.735]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [47540.0, 48192.0, 48729.0, 49266.0, 49803.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 48.8447041479
varray: [0.0, 108.038, 108.457, 110.558, 109.735]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [47540.0, 48192.0, 48729.0, 49266.0, 49803.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.2664
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.2036
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1034.1732
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.0974
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.53465
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.53465
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -19.0874850502
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -20.79755
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [0.0, -0.049, 0.094, 0.015, -0.032]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [47540.0, 48192.0, 48729.0, 49266.0, 49803.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [0.0, 9.78, 10.452, 9.88, 9.591]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [47540.0, 48192.0, 48729.0, 49266.0, 49803.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0056
varray: [0.0, -0.049, 0.094, 0.015, -0.032]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [47540.0, 48192.0, 48729.0, 49266.0, 49803.0]
source_col: d(18_16)Mean
dD_mean | 7.9406
varray: [0.0, 9.78, 10.452, 9.88, 9.591]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [47540.0, 48192.0, 48729.0, 49266.0, 49803.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0555184653967
varray: [0.0, -0.049, 0.094, 0.015, -0.032]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [47540.0, 48192.0, 48729.0, 49266.0, 49803.0]
source_col: d(18_16)Mean
dD_stdev | 4.45052646324
varray: [0.0, 9.78, 10.452, 9.88, 9.591]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [47540.0, 48192.0, 48729.0, 49266.0, 49803.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0056
varray: [0.0, -0.049, 0.094, 0.015, -0.032]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [47540.0, 48192.0, 48729.0, 49266.0, 49803.0]
source_col: d(18_16)Mean
measured_dD_mean | 7.9406
varray: [0.0, 9.78, 10.452, 9.88, 9.591]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [47540.0, 48192.0, 48729.0, 49266.0, 49803.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0555184653967
varray: [0.0, -0.049, 0.094, 0.015, -0.032]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [47540.0, 48192.0, 48729.0, 49266.0, 49803.0]
source_col: d(18_16)Mean
measured_dD_stdev | 4.45052646324
varray: [0.0, 9.78, 10.452, 9.88, 9.591]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [47540.0, 48192.0, 48729.0, 49266.0, 49803.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.60470198341
actual_array: [ 1.93 -15.84]
varray: [0.0, -0.049, 0.094, 0.015, -0.032]
fit_rsq: 1.0
fit_intercept: 2.61730844119
except: True
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.0056
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [47540.0, 48192.0, 48729.0, 49266.0, 49803.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -9.90197401182
actual_array: [ 18.1 -119. ]
varray: [0.0, 9.78, 10.452, 9.88, 9.591]
fit_rsq: 1.0
fit_intercept: 18.5140665208
except: True
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 7.9406
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [47540.0, 48192.0, 48729.0, 49266.0, 49803.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.62835
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -257.72155
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 18 - Sample 18 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 18
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.82475
varray: [2.802, 2.779, 2.83, 2.888]
inst: mean_all
operation: mean_all(residuals)
tarray: [50391.0, 50928.0, 51465.0, 52002.0]
source_col: residuals
SAMP_residuals_STD | 0.0470416482138
varray: [2.802, 2.779, 2.83, 2.888]
inst: std_all
operation: std_all(residuals)
tarray: [50391.0, 50928.0, 51465.0, 52002.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.22125
varray: [0.243, 0.204, 0.189, 0.249]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [50391.0, 50928.0, 51465.0, 52002.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0293300187521
varray: [0.243, 0.204, 0.189, 0.249]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [50391.0, 50928.0, 51465.0, 52002.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1292.3955
varray: [1291.716, 1291.716, 1291.789, 1294.361]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [50391.0, 50928.0, 51465.0, 52002.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.31078513368
varray: [1291.716, 1291.716, 1291.789, 1294.361]
inst: std_all
operation: std_all(baseline_shift)
tarray: [50391.0, 50928.0, 51465.0, 52002.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.51725
varray: [1.777, 1.636, 1.347, 1.309]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [50391.0, 50928.0, 51465.0, 52002.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.226513244646
varray: [1.777, 1.636, 1.347, 1.309]
inst: std_all
operation: std_all(slope_shift)
tarray: [50391.0, 50928.0, 51465.0, 52002.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0261661358757
varray: [1.586, 1.621, 1.619, 1.65]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50391.0, 50928.0, 51465.0, 52002.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.469549074468
varray: [16.181, 16.766, 16.827, 17.328]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50391.0, 50928.0, 51465.0, 52002.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.3935
varray: [107.813, 107.121, 108.001, 110.639]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [50391.0, 50928.0, 51465.0, 52002.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.54408149612
varray: [107.813, 107.121, 108.001, 110.639]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [50391.0, 50928.0, 51465.0, 52002.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.82475
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.22125
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1292.3955
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.51725
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0237
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0237
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.846111279699
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.23835
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [1.586, 1.621, 1.619, 1.65]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [50391.0, 50928.0, 51465.0, 52002.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [16.181, 16.766, 16.827, 17.328]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [50391.0, 50928.0, 51465.0, 52002.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.619
varray: [1.586, 1.621, 1.619, 1.65]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50391.0, 50928.0, 51465.0, 52002.0]
source_col: d(18_16)Mean
dD_mean | 16.7755
varray: [16.181, 16.766, 16.827, 17.328]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50391.0, 50928.0, 51465.0, 52002.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0261661358757
varray: [1.586, 1.621, 1.619, 1.65]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50391.0, 50928.0, 51465.0, 52002.0]
source_col: d(18_16)Mean
dD_stdev | 0.469549074468
varray: [16.181, 16.766, 16.827, 17.328]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50391.0, 50928.0, 51465.0, 52002.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.619
varray: [1.586, 1.621, 1.619, 1.65]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50391.0, 50928.0, 51465.0, 52002.0]
source_col: d(18_16)Mean
measured_dD_mean | 16.7755
varray: [16.181, 16.766, 16.827, 17.328]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50391.0, 50928.0, 51465.0, 52002.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0261661358757
varray: [1.586, 1.621, 1.619, 1.65]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50391.0, 50928.0, 51465.0, 52002.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.469549074468
varray: [16.181, 16.766, 16.827, 17.328]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50391.0, 50928.0, 51465.0, 52002.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.995630268608
actual_array: [ 1.93 -15.84]
varray: [1.586, 1.621, 1.619, 1.65]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 1.619
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [50391.0, 50928.0, 51465.0, 52002.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -1.62815482253
actual_array: [ 18.1 -119. ]
varray: [16.181, 16.766, 16.827, 17.328]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 16.7755
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [50391.0, 50928.0, 51465.0, 52002.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.07
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.50075
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 19 - Sample 19 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 19
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.817
varray: [2.837, 2.762, 2.825, 2.844]
inst: mean_all
operation: mean_all(residuals)
tarray: [52591.0, 53128.0, 53665.0, 54202.0]
source_col: residuals
SAMP_residuals_STD | 0.0374966665185
varray: [2.837, 2.762, 2.825, 2.844]
inst: std_all
operation: std_all(residuals)
tarray: [52591.0, 53128.0, 53665.0, 54202.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.25425
varray: [0.438, 0.272, 0.134, 0.173]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [52591.0, 53128.0, 53665.0, 54202.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.135573780651
varray: [0.438, 0.272, 0.134, 0.173]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [52591.0, 53128.0, 53665.0, 54202.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1292.02775
varray: [1291.86, 1290.353, 1292.794, 1293.104]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [52591.0, 53128.0, 53665.0, 54202.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.23536509448
varray: [1291.86, 1290.353, 1292.794, 1293.104]
inst: std_all
operation: std_all(baseline_shift)
tarray: [52591.0, 53128.0, 53665.0, 54202.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.41325
varray: [1.389, 1.604, 1.255, 1.405]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [52591.0, 53128.0, 53665.0, 54202.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.14385727881
varray: [1.389, 1.604, 1.255, 1.405]
inst: std_all
operation: std_all(slope_shift)
tarray: [52591.0, 53128.0, 53665.0, 54202.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0352230984819
varray: [1.645, 1.644, 1.625, 1.57]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52591.0, 53128.0, 53665.0, 54202.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.5397304111
varray: [23.472, 24.434, 24.724, 24.351]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52591.0, 53128.0, 53665.0, 54202.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.639
varray: [109.119, 106.922, 108.886, 109.629]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [52591.0, 53128.0, 53665.0, 54202.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.18597329931
varray: [109.119, 106.922, 108.886, 109.629]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [52591.0, 53128.0, 53665.0, 54202.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.817
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.25425
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1292.02775
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.41325
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.01595
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.01595
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.569429321147
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.48385
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [1.645, 1.644, 1.625, 1.57]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [52591.0, 53128.0, 53665.0, 54202.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [23.472, 24.434, 24.724, 24.351]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [52591.0, 53128.0, 53665.0, 54202.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.621
varray: [1.645, 1.644, 1.625, 1.57]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [52591.0, 53128.0, 53665.0, 54202.0]
source_col: d(18_16)Mean
dD_mean | 24.24525
varray: [23.472, 24.434, 24.724, 24.351]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [52591.0, 53128.0, 53665.0, 54202.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0352230984819
varray: [1.645, 1.644, 1.625, 1.57]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52591.0, 53128.0, 53665.0, 54202.0]
source_col: d(18_16)Mean
dD_stdev | 0.5397304111
varray: [23.472, 24.434, 24.724, 24.351]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52591.0, 53128.0, 53665.0, 54202.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.621
varray: [1.645, 1.644, 1.625, 1.57]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [52591.0, 53128.0, 53665.0, 54202.0]
source_col: d(18_16)Mean
measured_dD_mean | 24.24525
varray: [23.472, 24.434, 24.724, 24.351]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [52591.0, 53128.0, 53665.0, 54202.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0352230984819
varray: [1.645, 1.644, 1.625, 1.57]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52591.0, 53128.0, 53665.0, 54202.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.5397304111
varray: [23.472, 24.434, 24.724, 24.351]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52591.0, 53128.0, 53665.0, 54202.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.993635634029
actual_array: [ 1.93 -15.84]
varray: [1.645, 1.644, 1.625, 1.57]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 1.621
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [52591.0, 53128.0, 53665.0, 54202.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 5.367211383
actual_array: [ 18.1 -119. ]
varray: [23.472, 24.434, 24.724, 24.351]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 24.24525
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [52591.0, 53128.0, 53665.0, 54202.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.07775
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.133
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 20 - Sample 20 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 20
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.87425
varray: [2.872, 2.882, 2.883, 2.86]
inst: mean_all
operation: mean_all(residuals)
tarray: [54791.0, 55328.0, 55865.0, 56402.0]
source_col: residuals
SAMP_residuals_STD | 0.0107199191539
varray: [2.872, 2.882, 2.883, 2.86]
inst: std_all
operation: std_all(residuals)
tarray: [54791.0, 55328.0, 55865.0, 56402.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.26725
varray: [0.327, 0.191, 0.441, 0.11]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [54791.0, 55328.0, 55865.0, 56402.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.146402128855
varray: [0.327, 0.191, 0.441, 0.11]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [54791.0, 55328.0, 55865.0, 56402.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1294.1175
varray: [1293.742, 1294.175, 1294.317, 1294.236]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [54791.0, 55328.0, 55865.0, 56402.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.257001297014
varray: [1293.742, 1294.175, 1294.317, 1294.236]
inst: std_all
operation: std_all(baseline_shift)
tarray: [54791.0, 55328.0, 55865.0, 56402.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.18325
varray: [1.594, 0.738, 1.07, 1.331]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [54791.0, 55328.0, 55865.0, 56402.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.365886480574
varray: [1.594, 0.738, 1.07, 1.331]
inst: std_all
operation: std_all(slope_shift)
tarray: [54791.0, 55328.0, 55865.0, 56402.0]
source_col: slope_shift
SAMP_d18O_STD | 0.180717781822
varray: [7.155, 7.397, 7.525, 7.55]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54791.0, 55328.0, 55865.0, 56402.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.34729788218
varray: [46.787, 49.196, 49.521, 49.647]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54791.0, 55328.0, 55865.0, 56402.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.63175
varray: [109.789, 111.156, 111.828, 109.754]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [54791.0, 55328.0, 55865.0, 56402.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.03061870576
varray: [109.789, 111.156, 111.828, 109.754]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [54791.0, 55328.0, 55865.0, 56402.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.87425
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.26725
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1294.1175
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.18325
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0732
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0732
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.61330572464
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 2.4766
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [7.155, 7.397, 7.525, 7.55]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [54791.0, 55328.0, 55865.0, 56402.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [46.787, 49.196, 49.521, 49.647]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [54791.0, 55328.0, 55865.0, 56402.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 7.40675
varray: [7.155, 7.397, 7.525, 7.55]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [54791.0, 55328.0, 55865.0, 56402.0]
source_col: d(18_16)Mean
dD_mean | 48.78775
varray: [46.787, 49.196, 49.521, 49.647]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [54791.0, 55328.0, 55865.0, 56402.0]
source_col: d(D_H)Mean
d18O_stdev | 0.180717781822
varray: [7.155, 7.397, 7.525, 7.55]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54791.0, 55328.0, 55865.0, 56402.0]
source_col: d(18_16)Mean
dD_stdev | 1.34729788218
varray: [46.787, 49.196, 49.521, 49.647]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54791.0, 55328.0, 55865.0, 56402.0]
source_col: d(D_H)Mean
measured_d18O_mean | 7.40675
varray: [7.155, 7.397, 7.525, 7.55]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [54791.0, 55328.0, 55865.0, 56402.0]
source_col: d(18_16)Mean
measured_dD_mean | 48.78775
varray: [46.787, 49.196, 49.521, 49.647]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [54791.0, 55328.0, 55865.0, 56402.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.180717781822
varray: [7.155, 7.397, 7.525, 7.55]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54791.0, 55328.0, 55865.0, 56402.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.34729788218
varray: [46.787, 49.196, 49.521, 49.647]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54791.0, 55328.0, 55865.0, 56402.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 4.77659287342
actual_array: [ 1.93 -15.84]
varray: [7.155, 7.397, 7.525, 7.55]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 7.40675
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [54791.0, 55328.0, 55865.0, 56402.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 28.3510830118
actual_array: [ 18.1 -119. ]
varray: [46.787, 49.196, 49.521, 49.647]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 48.78775
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [54791.0, 55328.0, 55865.0, 56402.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.0205
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 2.22275
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 21 - Sample 21 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 21
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.809
varray: [2.787, 2.794, 2.81, 2.845]
inst: mean_all
operation: mean_all(residuals)
tarray: [56992.0, 57530.0, 58067.0, 58604.0]
source_col: residuals
SAMP_residuals_STD | 0.0258585898043
varray: [2.787, 2.794, 2.81, 2.845]
inst: std_all
operation: std_all(residuals)
tarray: [56992.0, 57530.0, 58067.0, 58604.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.2925
varray: [0.152, 0.282, 0.258, 0.478]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [56992.0, 57530.0, 58067.0, 58604.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.135954649301
varray: [0.152, 0.282, 0.258, 0.478]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [56992.0, 57530.0, 58067.0, 58604.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1291.2775
varray: [1290.999, 1291.274, 1291.507, 1291.33]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [56992.0, 57530.0, 58067.0, 58604.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.210555614189
varray: [1290.999, 1291.274, 1291.507, 1291.33]
inst: std_all
operation: std_all(baseline_shift)
tarray: [56992.0, 57530.0, 58067.0, 58604.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.35525
varray: [1.468, 1.564, 0.973, 1.416]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [56992.0, 57530.0, 58067.0, 58604.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.262103510087
varray: [1.468, 1.564, 0.973, 1.416]
inst: std_all
operation: std_all(slope_shift)
tarray: [56992.0, 57530.0, 58067.0, 58604.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0876318054894
varray: [4.17, 4.015, 3.987, 3.988]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56992.0, 57530.0, 58067.0, 58604.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.751347400785
varray: [35.066, 33.611, 33.406, 33.756]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56992.0, 57530.0, 58067.0, 58604.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.08925
varray: [107.158, 107.052, 108.818, 109.329]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [56992.0, 57530.0, 58067.0, 58604.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.15631177312
varray: [107.158, 107.052, 108.818, 109.329]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [56992.0, 57530.0, 58067.0, 58604.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.809
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.2925
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1291.2775
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.35525
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.00795
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.00795
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.283822138127
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.0659
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [4.17, 4.015, 3.987, 3.988]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [56992.0, 57530.0, 58067.0, 58604.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [35.066, 33.611, 33.406, 33.756]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [56992.0, 57530.0, 58067.0, 58604.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 4.04
varray: [4.17, 4.015, 3.987, 3.988]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56992.0, 57530.0, 58067.0, 58604.0]
source_col: d(18_16)Mean
dD_mean | 33.95975
varray: [35.066, 33.611, 33.406, 33.756]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56992.0, 57530.0, 58067.0, 58604.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0876318054894
varray: [4.17, 4.015, 3.987, 3.988]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56992.0, 57530.0, 58067.0, 58604.0]
source_col: d(18_16)Mean
dD_stdev | 0.751347400785
varray: [35.066, 33.611, 33.406, 33.756]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56992.0, 57530.0, 58067.0, 58604.0]
source_col: d(D_H)Mean
measured_d18O_mean | 4.04
varray: [4.17, 4.015, 3.987, 3.988]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56992.0, 57530.0, 58067.0, 58604.0]
source_col: d(18_16)Mean
measured_dD_mean | 33.95975
varray: [35.066, 33.611, 33.406, 33.756]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56992.0, 57530.0, 58067.0, 58604.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0876318054894
varray: [4.17, 4.015, 3.987, 3.988]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56992.0, 57530.0, 58067.0, 58604.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.751347400785
varray: [35.066, 33.611, 33.406, 33.756]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56992.0, 57530.0, 58067.0, 58604.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.41887488916
actual_array: [ 1.93 -15.84]
varray: [4.17, 4.015, 3.987, 3.988]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 4.04
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [56992.0, 57530.0, 58067.0, 58604.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 14.4647695347
actual_array: [ 18.1 -119. ]
varray: [35.066, 33.611, 33.406, 33.756]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 33.95975
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [56992.0, 57530.0, 58067.0, 58604.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.08575
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.61725
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 22 - Sample 22 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 22
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.79675
varray: [2.804, 2.778, 2.82, 2.785]
inst: mean_all
operation: mean_all(residuals)
tarray: [59189.0, 59725.0, 60262.0, 60798.0]
source_col: residuals
SAMP_residuals_STD | 0.018997806891
varray: [2.804, 2.778, 2.82, 2.785]
inst: std_all
operation: std_all(residuals)
tarray: [59189.0, 59725.0, 60262.0, 60798.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.214
varray: [0.385, 0.143, 0.184, 0.144]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [59189.0, 59725.0, 60262.0, 60798.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.115588350047
varray: [0.385, 0.143, 0.184, 0.144]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [59189.0, 59725.0, 60262.0, 60798.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1292.87375
varray: [1292.892, 1292.497, 1293.376, 1292.73]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [59189.0, 59725.0, 60262.0, 60798.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.37201825672
varray: [1292.892, 1292.497, 1293.376, 1292.73]
inst: std_all
operation: std_all(baseline_shift)
tarray: [59189.0, 59725.0, 60262.0, 60798.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.69475
varray: [1.819, 1.675, 1.784, 1.501]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [59189.0, 59725.0, 60262.0, 60798.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.142983390644
varray: [1.819, 1.675, 1.784, 1.501]
inst: std_all
operation: std_all(slope_shift)
tarray: [59189.0, 59725.0, 60262.0, 60798.0]
source_col: slope_shift
SAMP_d18O_STD | 0.316672806958
varray: [-6.802, -7.331, -7.429, -7.5]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59189.0, 59725.0, 60262.0, 60798.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 4.77877343921
varray: [-47.37, -55.247, -56.918, -57.87]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59189.0, 59725.0, 60262.0, 60798.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.61475
varray: [109.671, 107.615, 108.925, 108.248]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [59189.0, 59725.0, 60262.0, 60798.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.884293456194
varray: [109.671, 107.615, 108.925, 108.248]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [59189.0, 59725.0, 60262.0, 60798.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.79675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.214
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1292.87375
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.69475
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0043
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0043
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.153513860874
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.4596
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-6.802, -7.331, -7.429, -7.5]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [59189.0, 59725.0, 60262.0, 60798.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-47.37, -55.247, -56.918, -57.87]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [59189.0, 59725.0, 60262.0, 60798.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.2655
varray: [-6.802, -7.331, -7.429, -7.5]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [59189.0, 59725.0, 60262.0, 60798.0]
source_col: d(18_16)Mean
dD_mean | -54.35125
varray: [-47.37, -55.247, -56.918, -57.87]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [59189.0, 59725.0, 60262.0, 60798.0]
source_col: d(D_H)Mean
d18O_stdev | 0.316672806958
varray: [-6.802, -7.331, -7.429, -7.5]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59189.0, 59725.0, 60262.0, 60798.0]
source_col: d(18_16)Mean
dD_stdev | 4.77877343921
varray: [-47.37, -55.247, -56.918, -57.87]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59189.0, 59725.0, 60262.0, 60798.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.2655
varray: [-6.802, -7.331, -7.429, -7.5]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [59189.0, 59725.0, 60262.0, 60798.0]
source_col: d(18_16)Mean
measured_dD_mean | -54.35125
varray: [-47.37, -55.247, -56.918, -57.87]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [59189.0, 59725.0, 60262.0, 60798.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.316672806958
varray: [-6.802, -7.331, -7.429, -7.5]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59189.0, 59725.0, 60262.0, 60798.0]
source_col: d(18_16)Mean
measured_dD_stdev | 4.77877343921
varray: [-47.37, -55.247, -56.918, -57.87]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59189.0, 59725.0, 60262.0, 60798.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.85629572675
actual_array: [ 1.93 -15.84]
varray: [-6.802, -7.331, -7.429, -7.5]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.2655
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [59189.0, 59725.0, 60262.0, 60798.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -68.2378356363
actual_array: [ 18.1 -119. ]
varray: [-47.37, -55.247, -56.918, -57.87]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -54.35125
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [59189.0, 59725.0, 60262.0, 60798.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.098
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.979
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 23 - Sample 23 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 23
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.97225
varray: [2.926, 3.007, 3.064, 2.892]
inst: mean_all
operation: mean_all(residuals)
tarray: [61386.0, 61923.0, 62460.0, 62996.0]
source_col: residuals
SAMP_residuals_STD | 0.0778989730356
varray: [2.926, 3.007, 3.064, 2.892]
inst: std_all
operation: std_all(residuals)
tarray: [61386.0, 61923.0, 62460.0, 62996.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.36025
varray: [0.517, 0.329, 0.381, 0.214]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [61386.0, 61923.0, 62460.0, 62996.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.125653956033
varray: [0.517, 0.329, 0.381, 0.214]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [61386.0, 61923.0, 62460.0, 62996.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1297.4065
varray: [1294.691, 1298.909, 1299.7, 1296.326]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [61386.0, 61923.0, 62460.0, 62996.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.31366008739
varray: [1294.691, 1298.909, 1299.7, 1296.326]
inst: std_all
operation: std_all(baseline_shift)
tarray: [61386.0, 61923.0, 62460.0, 62996.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.5305
varray: [1.673, 1.658, 1.256, 1.535]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [61386.0, 61923.0, 62460.0, 62996.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.19316055498
varray: [1.673, 1.658, 1.256, 1.535]
inst: std_all
operation: std_all(slope_shift)
tarray: [61386.0, 61923.0, 62460.0, 62996.0]
source_col: slope_shift
SAMP_d18O_STD | 0.245092499545
varray: [0.437, 0.863, 0.943, 0.955]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61386.0, 61923.0, 62460.0, 62996.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.79260240468
varray: [5.506, 12.182, 13.342, 13.48]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61386.0, 61923.0, 62460.0, 62996.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.51925
varray: [112.114, 114.218, 116.448, 111.297]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [61386.0, 61923.0, 62460.0, 62996.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.30788147804
varray: [112.114, 114.218, 116.448, 111.297]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [61386.0, 61923.0, 62460.0, 62996.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.97225
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.36025
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1297.4065
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.5305
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.1712
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1712
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 6.11199371664
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 5.3641
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [0.437, 0.863, 0.943, 0.955]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [61386.0, 61923.0, 62460.0, 62996.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [5.506, 12.182, 13.342, 13.48]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [61386.0, 61923.0, 62460.0, 62996.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.7995
varray: [0.437, 0.863, 0.943, 0.955]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [61386.0, 61923.0, 62460.0, 62996.0]
source_col: d(18_16)Mean
dD_mean | 11.1275
varray: [5.506, 12.182, 13.342, 13.48]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [61386.0, 61923.0, 62460.0, 62996.0]
source_col: d(D_H)Mean
d18O_stdev | 0.245092499545
varray: [0.437, 0.863, 0.943, 0.955]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61386.0, 61923.0, 62460.0, 62996.0]
source_col: d(18_16)Mean
dD_stdev | 3.79260240468
varray: [5.506, 12.182, 13.342, 13.48]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61386.0, 61923.0, 62460.0, 62996.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.7995
varray: [0.437, 0.863, 0.943, 0.955]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [61386.0, 61923.0, 62460.0, 62996.0]
source_col: d(18_16)Mean
measured_dD_mean | 11.1275
varray: [5.506, 12.182, 13.342, 13.48]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [61386.0, 61923.0, 62460.0, 62996.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.245092499545
varray: [0.437, 0.863, 0.943, 0.955]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61386.0, 61923.0, 62460.0, 62996.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.79260240468
varray: [5.506, 12.182, 13.342, 13.48]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61386.0, 61923.0, 62460.0, 62996.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.81293178731
actual_array: [ 1.93 -15.84]
varray: [0.437, 0.863, 0.943, 0.955]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.7995
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [61386.0, 61923.0, 62460.0, 62996.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -6.91746548604
actual_array: [ 18.1 -119. ]
varray: [5.506, 12.182, 13.342, 13.48]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 11.1275
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [61386.0, 61923.0, 62460.0, 62996.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.9225
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 5.51175
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 24 - Sample 24 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 24
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.85625
varray: [2.92, 2.801, 2.924, 2.78]
inst: mean_all
operation: mean_all(residuals)
tarray: [63583.0, 64120.0, 64657.0, 65194.0]
source_col: residuals
SAMP_residuals_STD | 0.0764215283804
varray: [2.92, 2.801, 2.924, 2.78]
inst: std_all
operation: std_all(residuals)
tarray: [63583.0, 64120.0, 64657.0, 65194.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.26075
varray: [0.17, 0.29, 0.417, 0.166]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [63583.0, 64120.0, 64657.0, 65194.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.118999649859
varray: [0.17, 0.29, 0.417, 0.166]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [63583.0, 64120.0, 64657.0, 65194.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1295.393
varray: [1297.903, 1293.715, 1297.632, 1292.322]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [63583.0, 64120.0, 64657.0, 65194.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.80237672937
varray: [1297.903, 1293.715, 1297.632, 1292.322]
inst: std_all
operation: std_all(baseline_shift)
tarray: [63583.0, 64120.0, 64657.0, 65194.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.7705
varray: [1.811, 1.745, 1.563, 1.963]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [63583.0, 64120.0, 64657.0, 65194.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.165733722177
varray: [1.811, 1.745, 1.563, 1.963]
inst: std_all
operation: std_all(slope_shift)
tarray: [63583.0, 64120.0, 64657.0, 65194.0]
source_col: slope_shift
SAMP_d18O_STD | 0.221104801697
varray: [-6.019, -6.403, -6.444, -6.51]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63583.0, 64120.0, 64657.0, 65194.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.24002700606
varray: [-43.966, -49.623, -50.26, -51.109]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63583.0, 64120.0, 64657.0, 65194.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.57825
varray: [112.225, 109.321, 113.595, 107.172]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [63583.0, 64120.0, 64657.0, 65194.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.88650785264
varray: [112.225, 109.321, 113.595, 107.172]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [63583.0, 64120.0, 64657.0, 65194.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.85625
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.26075
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1295.393
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.7705
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0552
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0552
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.97068956284
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 2.4231
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-6.019, -6.403, -6.444, -6.51]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [63583.0, 64120.0, 64657.0, 65194.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-43.966, -49.623, -50.26, -51.109]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [63583.0, 64120.0, 64657.0, 65194.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.344
varray: [-6.019, -6.403, -6.444, -6.51]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63583.0, 64120.0, 64657.0, 65194.0]
source_col: d(18_16)Mean
dD_mean | -48.7395
varray: [-43.966, -49.623, -50.26, -51.109]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63583.0, 64120.0, 64657.0, 65194.0]
source_col: d(D_H)Mean
d18O_stdev | 0.221104801697
varray: [-6.019, -6.403, -6.444, -6.51]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63583.0, 64120.0, 64657.0, 65194.0]
source_col: d(18_16)Mean
dD_stdev | 3.24002700606
varray: [-43.966, -49.623, -50.26, -51.109]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63583.0, 64120.0, 64657.0, 65194.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.344
varray: [-6.019, -6.403, -6.444, -6.51]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63583.0, 64120.0, 64657.0, 65194.0]
source_col: d(18_16)Mean
measured_dD_mean | -48.7395
varray: [-43.966, -49.623, -50.26, -51.109]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63583.0, 64120.0, 64657.0, 65194.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.221104801697
varray: [-6.019, -6.403, -6.444, -6.51]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63583.0, 64120.0, 64657.0, 65194.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.24002700606
varray: [-43.966, -49.623, -50.26, -51.109]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63583.0, 64120.0, 64657.0, 65194.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.93726784452
actual_array: [ 1.93 -15.84]
varray: [-6.019, -6.403, -6.444, -6.51]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -6.344
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [63583.0, 64120.0, 64657.0, 65194.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -62.9824728325
actual_array: [ 18.1 -119. ]
varray: [-43.966, -49.623, -50.26, -51.109]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -48.7395
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [63583.0, 64120.0, 64657.0, 65194.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.0385
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 3.49825
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 25 - Sample 25 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 25
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.95625
varray: [2.98, 2.882, 2.943, 3.02]
inst: mean_all
operation: mean_all(residuals)
tarray: [65781.0, 66318.0, 66856.0, 67392.0]
source_col: residuals
SAMP_residuals_STD | 0.0586422771272
varray: [2.98, 2.882, 2.943, 3.02]
inst: std_all
operation: std_all(residuals)
tarray: [65781.0, 66318.0, 66856.0, 67392.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.2735
varray: [0.359, 0.369, 0.112, 0.254]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [65781.0, 66318.0, 66856.0, 67392.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.119572850876
varray: [0.359, 0.369, 0.112, 0.254]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [65781.0, 66318.0, 66856.0, 67392.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1297.855
varray: [1299.306, 1295.033, 1297.654, 1299.427]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [65781.0, 66318.0, 66856.0, 67392.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.04781753744
varray: [1299.306, 1295.033, 1297.654, 1299.427]
inst: std_all
operation: std_all(baseline_shift)
tarray: [65781.0, 66318.0, 66856.0, 67392.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.194
varray: [1.286, 1.55, 1.223, 0.717]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [65781.0, 66318.0, 66856.0, 67392.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.34812354129
varray: [1.286, 1.55, 1.223, 0.717]
inst: std_all
operation: std_all(slope_shift)
tarray: [65781.0, 66318.0, 66856.0, 67392.0]
source_col: slope_shift
SAMP_d18O_STD | 0.189198969342
varray: [0.524, 0.821, 0.928, 0.92]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65781.0, 66318.0, 66856.0, 67392.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.15245697829
varray: [6.335, 11.141, 12.619, 13.343]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65781.0, 66318.0, 66856.0, 67392.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.74275
varray: [115.409, 111.241, 112.611, 115.71]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [65781.0, 66318.0, 66856.0, 67392.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.17455533769
varray: [115.409, 111.241, 112.611, 115.71]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [65781.0, 66318.0, 66856.0, 67392.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.95625
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.2735
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1297.855
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.194
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.1552
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1552
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 5.5407793506
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 5.5876
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [0.524, 0.821, 0.928, 0.92]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [65781.0, 66318.0, 66856.0, 67392.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [6.335, 11.141, 12.619, 13.343]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [65781.0, 66318.0, 66856.0, 67392.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.79825
varray: [0.524, 0.821, 0.928, 0.92]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65781.0, 66318.0, 66856.0, 67392.0]
source_col: d(18_16)Mean
dD_mean | 10.8595
varray: [6.335, 11.141, 12.619, 13.343]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65781.0, 66318.0, 66856.0, 67392.0]
source_col: d(D_H)Mean
d18O_stdev | 0.189198969342
varray: [0.524, 0.821, 0.928, 0.92]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65781.0, 66318.0, 66856.0, 67392.0]
source_col: d(18_16)Mean
dD_stdev | 3.15245697829
varray: [6.335, 11.141, 12.619, 13.343]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65781.0, 66318.0, 66856.0, 67392.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.79825
varray: [0.524, 0.821, 0.928, 0.92]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65781.0, 66318.0, 66856.0, 67392.0]
source_col: d(18_16)Mean
measured_dD_mean | 10.8595
varray: [6.335, 11.141, 12.619, 13.343]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65781.0, 66318.0, 66856.0, 67392.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.189198969342
varray: [0.524, 0.821, 0.928, 0.92]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65781.0, 66318.0, 66856.0, 67392.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.15245697829
varray: [6.335, 11.141, 12.619, 13.343]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65781.0, 66318.0, 66856.0, 67392.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.81417843393
actual_array: [ 1.93 -15.84]
varray: [0.524, 0.821, 0.928, 0.92]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.79825
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [65781.0, 66318.0, 66856.0, 67392.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -7.16844552461
actual_array: [ 18.1 -119. ]
varray: [6.335, 11.141, 12.619, 13.343]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 10.8595
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [65781.0, 66318.0, 66856.0, 67392.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.9385
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 5.96025
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 26 - Sample 26 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 26
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.98375
varray: [3.065, 2.957, 2.92, 2.993]
inst: mean_all
operation: mean_all(residuals)
tarray: [88912.0, 89448.0, 89984.0, 90521.0]
source_col: residuals
SAMP_residuals_STD | 0.0618243479545
varray: [3.065, 2.957, 2.92, 2.993]
inst: std_all
operation: std_all(residuals)
tarray: [88912.0, 89448.0, 89984.0, 90521.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3
varray: [0.479, 0.219, 0.247, 0.255]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [88912.0, 89448.0, 89984.0, 90521.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.120327331337
varray: [0.479, 0.219, 0.247, 0.255]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [88912.0, 89448.0, 89984.0, 90521.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1298.63475
varray: [1300.688, 1298.562, 1296.754, 1298.535]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [88912.0, 89448.0, 89984.0, 90521.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.60917150008
varray: [1300.688, 1298.562, 1296.754, 1298.535]
inst: std_all
operation: std_all(baseline_shift)
tarray: [88912.0, 89448.0, 89984.0, 90521.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.61325
varray: [1.811, 1.306, 1.762, 1.574]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [88912.0, 89448.0, 89984.0, 90521.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.228892077917
varray: [1.811, 1.306, 1.762, 1.574]
inst: std_all
operation: std_all(slope_shift)
tarray: [88912.0, 89448.0, 89984.0, 90521.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0214223403639
varray: [1.258, 1.301, 1.304, 1.296]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [88912.0, 89448.0, 89984.0, 90521.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.228291297834
varray: [18.069, 18.164, 18.498, 18.514]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [88912.0, 89448.0, 89984.0, 90521.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.8945
varray: [116.252, 113.594, 112.264, 113.468]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [88912.0, 89448.0, 89984.0, 90521.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.68211563217
varray: [116.252, 113.594, 112.264, 113.468]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [88912.0, 89448.0, 89984.0, 90521.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.98375
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1298.63475
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.61325
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.1827
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1827
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 6.52255404223
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 5.73935
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [1.258, 1.301, 1.304, 1.296]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [88912.0, 89448.0, 89984.0, 90521.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [18.069, 18.164, 18.498, 18.514]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [88912.0, 89448.0, 89984.0, 90521.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.28975
varray: [1.258, 1.301, 1.304, 1.296]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [88912.0, 89448.0, 89984.0, 90521.0]
source_col: d(18_16)Mean
dD_mean | 18.31125
varray: [18.069, 18.164, 18.498, 18.514]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [88912.0, 89448.0, 89984.0, 90521.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0214223403639
varray: [1.258, 1.301, 1.304, 1.296]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [88912.0, 89448.0, 89984.0, 90521.0]
source_col: d(18_16)Mean
dD_stdev | 0.228291297834
varray: [18.069, 18.164, 18.498, 18.514]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [88912.0, 89448.0, 89984.0, 90521.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.28975
varray: [1.258, 1.301, 1.304, 1.296]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [88912.0, 89448.0, 89984.0, 90521.0]
source_col: d(18_16)Mean
measured_dD_mean | 18.31125
varray: [18.069, 18.164, 18.498, 18.514]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [88912.0, 89448.0, 89984.0, 90521.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0214223403639
varray: [1.258, 1.301, 1.304, 1.296]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [88912.0, 89448.0, 89984.0, 90521.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.228291297834
varray: [18.069, 18.164, 18.498, 18.514]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [88912.0, 89448.0, 89984.0, 90521.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.32399698616
actual_array: [ 1.93 -15.84]
varray: [1.258, 1.301, 1.304, 1.296]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 1.28975
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [88912.0, 89448.0, 89984.0, 90521.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -0.189936187348
actual_array: [ 18.1 -119. ]
varray: [18.069, 18.164, 18.498, 18.514]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 18.31125
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [88912.0, 89448.0, 89984.0, 90521.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.911
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 6.74
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 27 - Sample 27 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 27
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9665
varray: [3.017, 3.006, 2.899, 2.944]
inst: mean_all
operation: mean_all(residuals)
tarray: [91109.0, 91645.0, 92181.0, 92717.0]
source_col: residuals
SAMP_residuals_STD | 0.0552961722605
varray: [3.017, 3.006, 2.899, 2.944]
inst: std_all
operation: std_all(residuals)
tarray: [91109.0, 91645.0, 92181.0, 92717.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.293
varray: [0.481, 0.52, -0.039, 0.21]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [91109.0, 91645.0, 92181.0, 92717.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.26075914813
varray: [0.481, 0.52, -0.039, 0.21]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [91109.0, 91645.0, 92181.0, 92717.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1298.62275
varray: [1300.227, 1299.964, 1297.115, 1297.185]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [91109.0, 91645.0, 92181.0, 92717.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.70421093667
varray: [1300.227, 1299.964, 1297.115, 1297.185]
inst: std_all
operation: std_all(baseline_shift)
tarray: [91109.0, 91645.0, 92181.0, 92717.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.49575
varray: [1.649, 1.257, 1.774, 1.303]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [91109.0, 91645.0, 92181.0, 92717.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.254991993338
varray: [1.649, 1.257, 1.774, 1.303]
inst: std_all
operation: std_all(slope_shift)
tarray: [91109.0, 91645.0, 92181.0, 92717.0]
source_col: slope_shift
SAMP_d18O_STD | 0.119864854454
varray: [-2.231, -2.448, -2.486, -2.472]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [91109.0, 91645.0, 92181.0, 92717.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.95637017203
varray: [-11.241, -14.322, -15.273, -15.473]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [91109.0, 91645.0, 92181.0, 92717.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.9155
varray: [115.755, 115.949, 111.4, 112.558]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [91109.0, 91645.0, 92181.0, 92717.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.28687771135
varray: [115.755, 115.949, 111.4, 112.558]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [91109.0, 91645.0, 92181.0, 92717.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9665
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.293
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1298.62275
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.49575
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.16545
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.16545
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 5.90671355385
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 5.76035
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-2.231, -2.448, -2.486, -2.472]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [91109.0, 91645.0, 92181.0, 92717.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-11.241, -14.322, -15.273, -15.473]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [91109.0, 91645.0, 92181.0, 92717.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.40925
varray: [-2.231, -2.448, -2.486, -2.472]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [91109.0, 91645.0, 92181.0, 92717.0]
source_col: d(18_16)Mean
dD_mean | -14.07725
varray: [-11.241, -14.322, -15.273, -15.473]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [91109.0, 91645.0, 92181.0, 92717.0]
source_col: d(D_H)Mean
d18O_stdev | 0.119864854454
varray: [-2.231, -2.448, -2.486, -2.472]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [91109.0, 91645.0, 92181.0, 92717.0]
source_col: d(18_16)Mean
dD_stdev | 1.95637017203
varray: [-11.241, -14.322, -15.273, -15.473]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [91109.0, 91645.0, 92181.0, 92717.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.40925
varray: [-2.231, -2.448, -2.486, -2.472]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [91109.0, 91645.0, 92181.0, 92717.0]
source_col: d(18_16)Mean
measured_dD_mean | -14.07725
varray: [-11.241, -14.322, -15.273, -15.473]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [91109.0, 91645.0, 92181.0, 92717.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.119864854454
varray: [-2.231, -2.448, -2.486, -2.472]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [91109.0, 91645.0, 92181.0, 92717.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.95637017203
varray: [-11.241, -14.322, -15.273, -15.473]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [91109.0, 91645.0, 92181.0, 92717.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.01307363984
actual_array: [ 1.93 -15.84]
varray: [-2.231, -2.448, -2.486, -2.472]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.40925
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [91109.0, 91645.0, 92181.0, 92717.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -30.521529393
actual_array: [ 18.1 -119. ]
varray: [-11.241, -14.322, -15.273, -15.473]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -14.07725
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [91109.0, 91645.0, 92181.0, 92717.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.92825
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 6.728
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 28 - Sample 28 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 28
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.017
varray: [3.055, 2.997, 3.002, 3.014]
inst: mean_all
operation: mean_all(residuals)
tarray: [93306.0, 93842.0, 94378.0, 94914.0]
source_col: residuals
SAMP_residuals_STD | 0.0263185612575
varray: [3.055, 2.997, 3.002, 3.014]
inst: std_all
operation: std_all(residuals)
tarray: [93306.0, 93842.0, 94378.0, 94914.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.42625
varray: [0.36, 0.41, 0.462, 0.473]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [93306.0, 93842.0, 94378.0, 94914.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.052015222131
varray: [0.36, 0.41, 0.462, 0.473]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [93306.0, 93842.0, 94378.0, 94914.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1299.87125
varray: [1301.057, 1299.701, 1299.553, 1299.174]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [93306.0, 93842.0, 94378.0, 94914.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.821062066269
varray: [1301.057, 1299.701, 1299.553, 1299.174]
inst: std_all
operation: std_all(baseline_shift)
tarray: [93306.0, 93842.0, 94378.0, 94914.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.64575
varray: [1.526, 1.961, 1.599, 1.497]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [93306.0, 93842.0, 94378.0, 94914.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.214503107981
varray: [1.526, 1.961, 1.599, 1.497]
inst: std_all
operation: std_all(slope_shift)
tarray: [93306.0, 93842.0, 94378.0, 94914.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0464282600722
varray: [-2.816, -2.889, -2.876, -2.928]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [93306.0, 93842.0, 94378.0, 94914.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.298662574377
varray: [-18.346, -18.969, -18.699, -18.982]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [93306.0, 93842.0, 94378.0, 94914.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 115.32475
varray: [116.33, 114.81, 115.106, 115.053]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [93306.0, 93842.0, 94378.0, 94914.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.682445296953
varray: [116.33, 114.81, 115.106, 115.053]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [93306.0, 93842.0, 94378.0, 94914.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.017
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.42625
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1299.87125
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.64575
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.21595
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.21595
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 7.70960889666
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 7.1696
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-2.816, -2.889, -2.876, -2.928]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [93306.0, 93842.0, 94378.0, 94914.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-18.346, -18.969, -18.699, -18.982]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [93306.0, 93842.0, 94378.0, 94914.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.87725
varray: [-2.816, -2.889, -2.876, -2.928]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [93306.0, 93842.0, 94378.0, 94914.0]
source_col: d(18_16)Mean
dD_mean | -18.749
varray: [-18.346, -18.969, -18.699, -18.982]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [93306.0, 93842.0, 94378.0, 94914.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0464282600722
varray: [-2.816, -2.889, -2.876, -2.928]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [93306.0, 93842.0, 94378.0, 94914.0]
source_col: d(18_16)Mean
dD_stdev | 0.298662574377
varray: [-18.346, -18.969, -18.699, -18.982]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [93306.0, 93842.0, 94378.0, 94914.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.87725
varray: [-2.816, -2.889, -2.876, -2.928]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [93306.0, 93842.0, 94378.0, 94914.0]
source_col: d(18_16)Mean
measured_dD_mean | -18.749
varray: [-18.346, -18.969, -18.699, -18.982]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [93306.0, 93842.0, 94378.0, 94914.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0464282600722
varray: [-2.816, -2.889, -2.876, -2.928]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [93306.0, 93842.0, 94378.0, 94914.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.298662574377
varray: [-18.346, -18.969, -18.699, -18.982]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [93306.0, 93842.0, 94378.0, 94914.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.47981813131
actual_array: [ 1.93 -15.84]
varray: [-2.816, -2.889, -2.876, -2.928]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.87725
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [93306.0, 93842.0, 94378.0, 94914.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -34.8965890764
actual_array: [ 18.1 -119. ]
varray: [-18.346, -18.969, -18.699, -18.982]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -18.749
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [93306.0, 93842.0, 94378.0, 94914.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.87775
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 7.9765
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 29 - Sample 29 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 29
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.007
varray: [3.031, 3.003, 3.019, 2.975]
inst: mean_all
operation: mean_all(residuals)
tarray: [95499.0, 96035.0, 96571.0, 97107.0]
source_col: residuals
SAMP_residuals_STD | 0.0242212028328
varray: [3.031, 3.003, 3.019, 2.975]
inst: std_all
operation: std_all(residuals)
tarray: [95499.0, 96035.0, 96571.0, 97107.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4245
varray: [0.378, 0.381, 0.498, 0.441]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [95499.0, 96035.0, 96571.0, 97107.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0569473440996
varray: [0.378, 0.381, 0.498, 0.441]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [95499.0, 96035.0, 96571.0, 97107.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1299.8255
varray: [1301.297, 1300.33, 1299.653, 1298.022]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [95499.0, 96035.0, 96571.0, 97107.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.37867097839
varray: [1301.297, 1300.33, 1299.653, 1298.022]
inst: std_all
operation: std_all(baseline_shift)
tarray: [95499.0, 96035.0, 96571.0, 97107.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.55975
varray: [1.444, 1.415, 1.856, 1.524]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [95499.0, 96035.0, 96571.0, 97107.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.202807585328
varray: [1.444, 1.415, 1.856, 1.524]
inst: std_all
operation: std_all(slope_shift)
tarray: [95499.0, 96035.0, 96571.0, 97107.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0222111083319
varray: [-3.012, -3.044, -3.05, -3.006]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [95499.0, 96035.0, 96571.0, 97107.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0328468669231
varray: [-19.248, -19.266, -19.285, -19.208]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [95499.0, 96035.0, 96571.0, 97107.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 115.586
varray: [115.666, 115.977, 115.978, 114.723]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [95499.0, 96035.0, 96571.0, 97107.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.593777174817
varray: [115.666, 115.977, 115.978, 114.723]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [95499.0, 96035.0, 96571.0, 97107.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.007
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4245
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1299.8255
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.55975
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.20595
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.20595
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 7.35259991789
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 7.43085
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-3.012, -3.044, -3.05, -3.006]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [95499.0, 96035.0, 96571.0, 97107.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-19.248, -19.266, -19.285, -19.208]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [95499.0, 96035.0, 96571.0, 97107.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.028
varray: [-3.012, -3.044, -3.05, -3.006]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [95499.0, 96035.0, 96571.0, 97107.0]
source_col: d(18_16)Mean
dD_mean | -19.25175
varray: [-19.248, -19.266, -19.285, -19.208]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [95499.0, 96035.0, 96571.0, 97107.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0222111083319
varray: [-3.012, -3.044, -3.05, -3.006]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [95499.0, 96035.0, 96571.0, 97107.0]
source_col: d(18_16)Mean
dD_stdev | 0.0328468669231
varray: [-19.248, -19.266, -19.285, -19.208]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [95499.0, 96035.0, 96571.0, 97107.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.028
varray: [-3.012, -3.044, -3.05, -3.006]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [95499.0, 96035.0, 96571.0, 97107.0]
source_col: d(18_16)Mean
measured_dD_mean | -19.25175
varray: [-19.248, -19.266, -19.285, -19.208]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [95499.0, 96035.0, 96571.0, 97107.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0222111083319
varray: [-3.012, -3.044, -3.05, -3.006]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [95499.0, 96035.0, 96571.0, 97107.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0328468669231
varray: [-19.248, -19.266, -19.285, -19.208]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [95499.0, 96035.0, 96571.0, 97107.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.63016371269
actual_array: [ 1.93 -15.84]
varray: [-3.012, -3.044, -3.05, -3.006]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.028
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [95499.0, 96035.0, 96571.0, 97107.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -35.3674107719
actual_array: [ 18.1 -119. ]
varray: [-19.248, -19.266, -19.285, -19.208]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -19.25175
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [95499.0, 96035.0, 96571.0, 97107.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.88775
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 7.93075
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 30 - Sample 30 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 30
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9905
varray: [3.05, 2.956, 2.976, 2.98]
inst: mean_all
operation: mean_all(residuals)
tarray: [97692.0, 98228.0, 98764.0, 99301.0]
source_col: residuals
SAMP_residuals_STD | 0.0410325074382
varray: [3.05, 2.956, 2.976, 2.98]
inst: std_all
operation: std_all(residuals)
tarray: [97692.0, 98228.0, 98764.0, 99301.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.42175
varray: [0.54, 0.243, 0.362, 0.542]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [97692.0, 98228.0, 98764.0, 99301.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.146019119753
varray: [0.54, 0.243, 0.362, 0.542]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [97692.0, 98228.0, 98764.0, 99301.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1299.1045
varray: [1301.779, 1298.659, 1298.244, 1297.736]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [97692.0, 98228.0, 98764.0, 99301.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.82251410603
varray: [1301.779, 1298.659, 1298.244, 1297.736]
inst: std_all
operation: std_all(baseline_shift)
tarray: [97692.0, 98228.0, 98764.0, 99301.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.57025
varray: [1.652, 1.471, 1.029, 2.129]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [97692.0, 98228.0, 98764.0, 99301.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.455224029096
varray: [1.652, 1.471, 1.029, 2.129]
inst: std_all
operation: std_all(slope_shift)
tarray: [97692.0, 98228.0, 98764.0, 99301.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0616062767365
varray: [-2.727, -2.781, -2.792, -2.876]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [97692.0, 98228.0, 98764.0, 99301.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.28690706974
varray: [-18.218, -18.506, -18.352, -18.882]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [97692.0, 98228.0, 98764.0, 99301.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 115.1085
varray: [117.301, 114.149, 114.663, 114.321]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [97692.0, 98228.0, 98764.0, 99301.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.47719587959
varray: [117.301, 114.149, 114.663, 114.321]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [97692.0, 98228.0, 98764.0, 99301.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9905
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.42175
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1299.1045
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.57025
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.18945
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.18945
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 6.76353510291
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 6.95335
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-2.727, -2.781, -2.792, -2.876]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [97692.0, 98228.0, 98764.0, 99301.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-18.218, -18.506, -18.352, -18.882]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [97692.0, 98228.0, 98764.0, 99301.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.794
varray: [-2.727, -2.781, -2.792, -2.876]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [97692.0, 98228.0, 98764.0, 99301.0]
source_col: d(18_16)Mean
dD_mean | -18.4895
varray: [-18.218, -18.506, -18.352, -18.882]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [97692.0, 98228.0, 98764.0, 99301.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0616062767365
varray: [-2.727, -2.781, -2.792, -2.876]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [97692.0, 98228.0, 98764.0, 99301.0]
source_col: d(18_16)Mean
dD_stdev | 0.28690706974
varray: [-18.218, -18.506, -18.352, -18.882]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [97692.0, 98228.0, 98764.0, 99301.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.794
varray: [-2.727, -2.781, -2.792, -2.876]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [97692.0, 98228.0, 98764.0, 99301.0]
source_col: d(18_16)Mean
measured_dD_mean | -18.4895
varray: [-18.218, -18.506, -18.352, -18.882]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [97692.0, 98228.0, 98764.0, 99301.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0616062767365
varray: [-2.727, -2.781, -2.792, -2.876]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [97692.0, 98228.0, 98764.0, 99301.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.28690706974
varray: [-18.218, -18.506, -18.352, -18.882]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [97692.0, 98228.0, 98764.0, 99301.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.39679146696
actual_array: [ 1.93 -15.84]
varray: [-2.727, -2.781, -2.792, -2.876]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.794
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [97692.0, 98228.0, 98764.0, 99301.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -34.6535692257
actual_array: [ 18.1 -119. ]
varray: [-18.218, -18.506, -18.352, -18.882]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -18.4895
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [97692.0, 98228.0, 98764.0, 99301.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.90425
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 7.20975
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 31 - Sample 31 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 31
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9915
varray: [3.035, 3.033, 2.951, 2.947]
inst: mean_all
operation: mean_all(residuals)
tarray: [99886.0, 100422.0, 100958.0, 101494.0]
source_col: residuals
SAMP_residuals_STD | 0.0491087229183
varray: [3.035, 3.033, 2.951, 2.947]
inst: std_all
operation: std_all(residuals)
tarray: [99886.0, 100422.0, 100958.0, 101494.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4785
varray: [0.564, 0.602, 0.487, 0.261]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [99886.0, 100422.0, 100958.0, 101494.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.152688135317
varray: [0.564, 0.602, 0.487, 0.261]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [99886.0, 100422.0, 100958.0, 101494.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1299.1285
varray: [1301.183, 1298.78, 1298.293, 1298.258]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [99886.0, 100422.0, 100958.0, 101494.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.39023415294
varray: [1301.183, 1298.78, 1298.293, 1298.258]
inst: std_all
operation: std_all(baseline_shift)
tarray: [99886.0, 100422.0, 100958.0, 101494.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.6225
varray: [1.753, 1.324, 1.743, 1.67]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [99886.0, 100422.0, 100958.0, 101494.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.202409650626
varray: [1.753, 1.324, 1.743, 1.67]
inst: std_all
operation: std_all(slope_shift)
tarray: [99886.0, 100422.0, 100958.0, 101494.0]
source_col: slope_shift
SAMP_d18O_STD | 0.157625082606
varray: [-7.21, -7.467, -7.511, -7.566]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [99886.0, 100422.0, 100958.0, 101494.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.15910274574
varray: [-53.831, -57.29, -58.194, -58.561]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [99886.0, 100422.0, 100958.0, 101494.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 115.32025
varray: [117.464, 115.433, 114.685, 113.699]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [99886.0, 100422.0, 100958.0, 101494.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.5958666141
varray: [117.464, 115.433, 114.685, 113.699]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [99886.0, 100422.0, 100958.0, 101494.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9915
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4785
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1299.1285
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.6225
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.19045
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.19045
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 6.79923600079
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 7.1651
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-7.21, -7.467, -7.511, -7.566]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [99886.0, 100422.0, 100958.0, 101494.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-53.831, -57.29, -58.194, -58.561]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [99886.0, 100422.0, 100958.0, 101494.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.4385
varray: [-7.21, -7.467, -7.511, -7.566]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [99886.0, 100422.0, 100958.0, 101494.0]
source_col: d(18_16)Mean
dD_mean | -56.969
varray: [-53.831, -57.29, -58.194, -58.561]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [99886.0, 100422.0, 100958.0, 101494.0]
source_col: d(D_H)Mean
d18O_stdev | 0.157625082606
varray: [-7.21, -7.467, -7.511, -7.566]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [99886.0, 100422.0, 100958.0, 101494.0]
source_col: d(18_16)Mean
dD_stdev | 2.15910274574
varray: [-53.831, -57.29, -58.194, -58.561]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [99886.0, 100422.0, 100958.0, 101494.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.4385
varray: [-7.21, -7.467, -7.511, -7.566]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [99886.0, 100422.0, 100958.0, 101494.0]
source_col: d(18_16)Mean
measured_dD_mean | -56.969
varray: [-53.831, -57.29, -58.194, -58.561]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [99886.0, 100422.0, 100958.0, 101494.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.157625082606
varray: [-7.21, -7.467, -7.511, -7.566]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [99886.0, 100422.0, 100958.0, 101494.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.15910274574
varray: [-53.831, -57.29, -58.194, -58.561]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [99886.0, 100422.0, 100958.0, 101494.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.0288316178
actual_array: [ 1.93 -15.84]
varray: [-7.21, -7.467, -7.511, -7.566]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.4385
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [99886.0, 100422.0, 100958.0, 101494.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -70.6893393526
actual_array: [ 18.1 -119. ]
varray: [-53.831, -57.29, -58.194, -58.561]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -56.969
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [99886.0, 100422.0, 100958.0, 101494.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.90325
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 7.23375
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 32 - Sample 32 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 32
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9735
varray: [3.044, 2.932, 2.946, 2.972]
inst: mean_all
operation: mean_all(residuals)
tarray: [102080.0, 102616.0, 103152.0, 103688.0]
source_col: residuals
SAMP_residuals_STD | 0.0498363990138
varray: [3.044, 2.932, 2.946, 2.972]
inst: std_all
operation: std_all(residuals)
tarray: [102080.0, 102616.0, 103152.0, 103688.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.34775
varray: [0.495, 0.089, 0.409, 0.398]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [102080.0, 102616.0, 103152.0, 103688.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.177867694275
varray: [0.495, 0.089, 0.409, 0.398]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [102080.0, 102616.0, 103152.0, 103688.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1298.66725
varray: [1300.986, 1298.557, 1297.362, 1297.764]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [102080.0, 102616.0, 103152.0, 103688.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.6236065564
varray: [1300.986, 1298.557, 1297.362, 1297.764]
inst: std_all
operation: std_all(baseline_shift)
tarray: [102080.0, 102616.0, 103152.0, 103688.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.54875
varray: [1.712, 1.713, 1.311, 1.459]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [102080.0, 102616.0, 103152.0, 103688.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.198501679254
varray: [1.712, 1.713, 1.311, 1.459]
inst: std_all
operation: std_all(slope_shift)
tarray: [102080.0, 102616.0, 103152.0, 103688.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0901420175797
varray: [-3.001, -2.854, -2.811, -2.811]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [102080.0, 102616.0, 103152.0, 103688.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.57840053535
varray: [-22.316, -19.929, -18.807, -19.167]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [102080.0, 102616.0, 103152.0, 103688.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 114.8175
varray: [117.1, 112.781, 114.409, 114.98]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [102080.0, 102616.0, 103152.0, 103688.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.78422877831
varray: [117.1, 112.781, 114.409, 114.98]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [102080.0, 102616.0, 103152.0, 103688.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9735
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.34775
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1298.66725
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.54875
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.17245
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.17245
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 6.15661983899
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 6.66235
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-3.001, -2.854, -2.811, -2.811]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [102080.0, 102616.0, 103152.0, 103688.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-22.316, -19.929, -18.807, -19.167]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [102080.0, 102616.0, 103152.0, 103688.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.86925
varray: [-3.001, -2.854, -2.811, -2.811]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [102080.0, 102616.0, 103152.0, 103688.0]
source_col: d(18_16)Mean
dD_mean | -20.05475
varray: [-22.316, -19.929, -18.807, -19.167]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [102080.0, 102616.0, 103152.0, 103688.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0901420175797
varray: [-3.001, -2.854, -2.811, -2.811]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [102080.0, 102616.0, 103152.0, 103688.0]
source_col: d(18_16)Mean
dD_stdev | 1.57840053535
varray: [-22.316, -19.929, -18.807, -19.167]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [102080.0, 102616.0, 103152.0, 103688.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.86925
varray: [-3.001, -2.854, -2.811, -2.811]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [102080.0, 102616.0, 103152.0, 103688.0]
source_col: d(18_16)Mean
measured_dD_mean | -20.05475
varray: [-22.316, -19.929, -18.807, -19.167]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [102080.0, 102616.0, 103152.0, 103688.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0901420175797
varray: [-3.001, -2.854, -2.811, -2.811]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [102080.0, 102616.0, 103152.0, 103688.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.57840053535
varray: [-22.316, -19.929, -18.807, -19.167]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [102080.0, 102616.0, 103152.0, 103688.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.47183959299
actual_array: [ 1.93 -15.84]
varray: [-3.001, -2.854, -2.811, -2.811]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.86925
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [102080.0, 102616.0, 103152.0, 103688.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -36.1194143949
actual_array: [ 18.1 -119. ]
varray: [-22.316, -19.929, -18.807, -19.167]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -20.05475
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [102080.0, 102616.0, 103152.0, 103688.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.92125
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 6.7725
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 33 - Sample 33 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 33
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.979
varray: [2.987, 3.061, 2.926, 2.942]
inst: mean_all
operation: mean_all(residuals)
tarray: [104275.0, 104811.0, 105347.0, 105883.0]
source_col: residuals
SAMP_residuals_STD | 0.0604593527366
varray: [2.987, 3.061, 2.926, 2.942]
inst: std_all
operation: std_all(residuals)
tarray: [104275.0, 104811.0, 105347.0, 105883.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.271
varray: [0.221, 0.345, 0.402, 0.116]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [104275.0, 104811.0, 105347.0, 105883.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.128013020171
varray: [0.221, 0.345, 0.402, 0.116]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [104275.0, 104811.0, 105347.0, 105883.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1298.78425
varray: [1300.606, 1301.326, 1296.243, 1296.962]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [104275.0, 104811.0, 105347.0, 105883.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.55328629743
varray: [1300.606, 1301.326, 1296.243, 1296.962]
inst: std_all
operation: std_all(baseline_shift)
tarray: [104275.0, 104811.0, 105347.0, 105883.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.632
varray: [1.701, 1.497, 1.905, 1.425]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [104275.0, 104811.0, 105347.0, 105883.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.21630533974
varray: [1.701, 1.497, 1.905, 1.425]
inst: std_all
operation: std_all(slope_shift)
tarray: [104275.0, 104811.0, 105347.0, 105883.0]
source_col: slope_shift
SAMP_d18O_STD | 0.041161268202
varray: [-3.122, -3.134, -3.203, -3.194]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [104275.0, 104811.0, 105347.0, 105883.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.439603609327
varray: [-19.661, -19.322, -20.313, -20.076]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [104275.0, 104811.0, 105347.0, 105883.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 114.104
varray: [115.019, 117.179, 112.112, 112.106]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [104275.0, 104811.0, 105347.0, 105883.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.46663860344
varray: [115.019, 117.179, 112.112, 112.106]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [104275.0, 104811.0, 105347.0, 105883.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.979
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.271
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1298.78425
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.632
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.17795
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.17795
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 6.35297477732
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 5.94885
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-3.122, -3.134, -3.203, -3.194]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [104275.0, 104811.0, 105347.0, 105883.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-19.661, -19.322, -20.313, -20.076]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [104275.0, 104811.0, 105347.0, 105883.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.16325
varray: [-3.122, -3.134, -3.203, -3.194]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [104275.0, 104811.0, 105347.0, 105883.0]
source_col: d(18_16)Mean
dD_mean | -19.843
varray: [-19.661, -19.322, -20.313, -20.076]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [104275.0, 104811.0, 105347.0, 105883.0]
source_col: d(D_H)Mean
d18O_stdev | 0.041161268202
varray: [-3.122, -3.134, -3.203, -3.194]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [104275.0, 104811.0, 105347.0, 105883.0]
source_col: d(18_16)Mean
dD_stdev | 0.439603609327
varray: [-19.661, -19.322, -20.313, -20.076]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [104275.0, 104811.0, 105347.0, 105883.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.16325
varray: [-3.122, -3.134, -3.203, -3.194]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [104275.0, 104811.0, 105347.0, 105883.0]
source_col: d(18_16)Mean
measured_dD_mean | -19.843
varray: [-19.661, -19.322, -20.313, -20.076]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [104275.0, 104811.0, 105347.0, 105883.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.041161268202
varray: [-3.122, -3.134, -3.203, -3.194]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [104275.0, 104811.0, 105347.0, 105883.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.439603609327
varray: [-19.661, -19.322, -20.313, -20.076]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [104275.0, 104811.0, 105347.0, 105883.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.76505087609
actual_array: [ 1.93 -15.84]
varray: [-3.122, -3.134, -3.203, -3.194]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.16325
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [104275.0, 104811.0, 105347.0, 105883.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -35.9211120697
actual_array: [ 18.1 -119. ]
varray: [-19.661, -19.322, -20.313, -20.076]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -19.843
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [104275.0, 104811.0, 105347.0, 105883.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.91575
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 6.8895
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 34 - Sample 34 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 34
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.91675
varray: [2.957, 2.802, 2.991, 2.917]
inst: mean_all
operation: mean_all(residuals)
tarray: [106469.0, 107006.0, 107542.0, 108078.0]
source_col: residuals
SAMP_residuals_STD | 0.0822612707577
varray: [2.957, 2.802, 2.991, 2.917]
inst: std_all
operation: std_all(residuals)
tarray: [106469.0, 107006.0, 107542.0, 108078.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.30725
varray: [0.185, 0.092, 0.511, 0.441]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [106469.0, 107006.0, 107542.0, 108078.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.200566489391
varray: [0.185, 0.092, 0.511, 0.441]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [106469.0, 107006.0, 107542.0, 108078.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1296.85375
varray: [1300.029, 1292.544, 1299.278, 1295.564]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [106469.0, 107006.0, 107542.0, 108078.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.47354769412
varray: [1300.029, 1292.544, 1299.278, 1295.564]
inst: std_all
operation: std_all(baseline_shift)
tarray: [106469.0, 107006.0, 107542.0, 108078.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.46575
varray: [1.623, 1.951, 0.794, 1.495]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [106469.0, 107006.0, 107542.0, 108078.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.487270886332
varray: [1.623, 1.951, 0.794, 1.495]
inst: std_all
operation: std_all(slope_shift)
tarray: [106469.0, 107006.0, 107542.0, 108078.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0660353440717
varray: [-3.27, -3.364, -3.215, -3.331]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [106469.0, 107006.0, 107542.0, 108078.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.339734700416
varray: [-21.026, -21.532, -20.736, -21.268]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [106469.0, 107006.0, 107542.0, 108078.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 112.93725
varray: [114.649, 107.658, 116.724, 112.718]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [106469.0, 107006.0, 107542.0, 108078.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.88107007366
varray: [114.649, 107.658, 116.724, 112.718]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [106469.0, 107006.0, 107542.0, 108078.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.91675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.30725
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1296.85375
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.46575
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.1157
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1157
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 4.13059388444
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 4.7821
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-3.27, -3.364, -3.215, -3.331]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [106469.0, 107006.0, 107542.0, 108078.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-21.026, -21.532, -20.736, -21.268]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [106469.0, 107006.0, 107542.0, 108078.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.295
varray: [-3.27, -3.364, -3.215, -3.331]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [106469.0, 107006.0, 107542.0, 108078.0]
source_col: d(18_16)Mean
dD_mean | -21.1405
varray: [-21.026, -21.532, -20.736, -21.268]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [106469.0, 107006.0, 107542.0, 108078.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0660353440717
varray: [-3.27, -3.364, -3.215, -3.331]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [106469.0, 107006.0, 107542.0, 108078.0]
source_col: d(18_16)Mean
dD_stdev | 0.339734700416
varray: [-21.026, -21.532, -20.736, -21.268]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [106469.0, 107006.0, 107542.0, 108078.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.295
varray: [-3.27, -3.364, -3.215, -3.331]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [106469.0, 107006.0, 107542.0, 108078.0]
source_col: d(18_16)Mean
measured_dD_mean | -21.1405
varray: [-21.026, -21.532, -20.736, -21.268]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [106469.0, 107006.0, 107542.0, 108078.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0660353440717
varray: [-3.27, -3.364, -3.215, -3.331]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [106469.0, 107006.0, 107542.0, 108078.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.339734700416
varray: [-21.026, -21.532, -20.736, -21.268]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [106469.0, 107006.0, 107542.0, 108078.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.89644742898
actual_array: [ 1.93 -15.84]
varray: [-3.27, -3.364, -3.215, -3.331]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.295
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [106469.0, 107006.0, 107542.0, 108078.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -37.1362113236
actual_array: [ 18.1 -119. ]
varray: [-21.026, -21.532, -20.736, -21.268]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -21.1405
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [106469.0, 107006.0, 107542.0, 108078.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.978
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 4.959
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 35 - Sample 35 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 35
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.97725
varray: [2.994, 3.002, 2.975, 2.938]
inst: mean_all
operation: mean_all(residuals)
tarray: [108665.0, 109201.0, 109737.0, 110274.0]
source_col: residuals
SAMP_residuals_STD | 0.0285116935075
varray: [2.994, 3.002, 2.975, 2.938]
inst: std_all
operation: std_all(residuals)
tarray: [108665.0, 109201.0, 109737.0, 110274.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.31675
varray: [0.128, 0.274, 0.474, 0.391]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [108665.0, 109201.0, 109737.0, 110274.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.150216233033
varray: [0.128, 0.274, 0.474, 0.391]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [108665.0, 109201.0, 109737.0, 110274.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1298.4405
varray: [1300.674, 1298.696, 1297.482, 1296.91]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [108665.0, 109201.0, 109737.0, 110274.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.66482781892
varray: [1300.674, 1298.696, 1297.482, 1296.91]
inst: std_all
operation: std_all(baseline_shift)
tarray: [108665.0, 109201.0, 109737.0, 110274.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.697
varray: [2.037, 1.346, 2.071, 1.334]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [108665.0, 109201.0, 109737.0, 110274.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.412490807978
varray: [2.037, 1.346, 2.071, 1.334]
inst: std_all
operation: std_all(slope_shift)
tarray: [108665.0, 109201.0, 109737.0, 110274.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0290559804515
varray: [-2.909, -2.94, -2.975, -2.963]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [108665.0, 109201.0, 109737.0, 110274.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.199375692266
varray: [-19.506, -19.958, -19.852, -19.872]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [108665.0, 109201.0, 109737.0, 110274.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 114.13725
varray: [114.743, 114.674, 113.557, 113.575]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [108665.0, 109201.0, 109737.0, 110274.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.660264782745
varray: [114.743, 114.674, 113.557, 113.575]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [108665.0, 109201.0, 109737.0, 110274.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.97725
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.31675
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1298.4405
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.697
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.1762
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1762
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 6.29049820603
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 5.9821
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-2.909, -2.94, -2.975, -2.963]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [108665.0, 109201.0, 109737.0, 110274.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-19.506, -19.958, -19.852, -19.872]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [108665.0, 109201.0, 109737.0, 110274.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.94675
varray: [-2.909, -2.94, -2.975, -2.963]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [108665.0, 109201.0, 109737.0, 110274.0]
source_col: d(18_16)Mean
dD_mean | -19.797
varray: [-19.506, -19.958, -19.852, -19.872]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [108665.0, 109201.0, 109737.0, 110274.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0290559804515
varray: [-2.909, -2.94, -2.975, -2.963]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [108665.0, 109201.0, 109737.0, 110274.0]
source_col: d(18_16)Mean
dD_stdev | 0.199375692266
varray: [-19.506, -19.958, -19.852, -19.872]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [108665.0, 109201.0, 109737.0, 110274.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.94675
varray: [-2.909, -2.94, -2.975, -2.963]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [108665.0, 109201.0, 109737.0, 110274.0]
source_col: d(18_16)Mean
measured_dD_mean | -19.797
varray: [-19.506, -19.958, -19.852, -19.872]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [108665.0, 109201.0, 109737.0, 110274.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0290559804515
varray: [-2.909, -2.94, -2.975, -2.963]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [108665.0, 109201.0, 109737.0, 110274.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.199375692266
varray: [-19.506, -19.958, -19.852, -19.872]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [108665.0, 109201.0, 109737.0, 110274.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.54913168292
actual_array: [ 1.93 -15.84]
varray: [-2.909, -2.94, -2.975, -2.963]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.94675
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [108665.0, 109201.0, 109737.0, 110274.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -35.8780334064
actual_array: [ 18.1 -119. ]
varray: [-19.506, -19.958, -19.852, -19.872]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -19.797
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [108665.0, 109201.0, 109737.0, 110274.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.9175
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 6.54575
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 36 - Sample 36 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 36
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9565
varray: [3.0, 2.939, 2.932, 2.955]
inst: mean_all
operation: mean_all(residuals)
tarray: [110856.0, 111393.0, 111929.0, 112465.0]
source_col: residuals
SAMP_residuals_STD | 0.0305559595933
varray: [3.0, 2.939, 2.932, 2.955]
inst: std_all
operation: std_all(residuals)
tarray: [110856.0, 111393.0, 111929.0, 112465.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.2565
varray: [0.231, 0.134, 0.163, 0.498]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [110856.0, 111393.0, 111929.0, 112465.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.166053204325
varray: [0.231, 0.134, 0.163, 0.498]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [110856.0, 111393.0, 111929.0, 112465.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1298.549
varray: [1300.608, 1298.19, 1297.859, 1297.539]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [110856.0, 111393.0, 111929.0, 112465.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.39816093494
varray: [1300.608, 1298.19, 1297.859, 1297.539]
inst: std_all
operation: std_all(baseline_shift)
tarray: [110856.0, 111393.0, 111929.0, 112465.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.47625
varray: [1.698, 1.444, 1.354, 1.409]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [110856.0, 111393.0, 111929.0, 112465.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.152403794791
varray: [1.698, 1.444, 1.354, 1.409]
inst: std_all
operation: std_all(slope_shift)
tarray: [110856.0, 111393.0, 111929.0, 112465.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0475631159618
varray: [-3.212, -3.305, -3.29, -3.318]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [110856.0, 111393.0, 111929.0, 112465.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.225634475794
varray: [-20.785, -20.845, -21.203, -21.204]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [110856.0, 111393.0, 111929.0, 112465.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.815
varray: [115.071, 112.54, 113.025, 114.624]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [110856.0, 111393.0, 111929.0, 112465.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.22225774696
varray: [115.071, 112.54, 113.025, 114.624]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [110856.0, 111393.0, 111929.0, 112465.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9565
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.2565
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1298.549
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.47625
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.15545
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.15545
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 5.54970457507
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 5.65985
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-3.212, -3.305, -3.29, -3.318]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [110856.0, 111393.0, 111929.0, 112465.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-20.785, -20.845, -21.203, -21.204]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [110856.0, 111393.0, 111929.0, 112465.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.28125
varray: [-3.212, -3.305, -3.29, -3.318]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [110856.0, 111393.0, 111929.0, 112465.0]
source_col: d(18_16)Mean
dD_mean | -21.00925
varray: [-20.785, -20.845, -21.203, -21.204]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [110856.0, 111393.0, 111929.0, 112465.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0475631159618
varray: [-3.212, -3.305, -3.29, -3.318]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [110856.0, 111393.0, 111929.0, 112465.0]
source_col: d(18_16)Mean
dD_stdev | 0.225634475794
varray: [-20.785, -20.845, -21.203, -21.204]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [110856.0, 111393.0, 111929.0, 112465.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.28125
varray: [-3.212, -3.305, -3.29, -3.318]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [110856.0, 111393.0, 111929.0, 112465.0]
source_col: d(18_16)Mean
measured_dD_mean | -21.00925
varray: [-20.785, -20.845, -21.203, -21.204]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [110856.0, 111393.0, 111929.0, 112465.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0475631159618
varray: [-3.212, -3.305, -3.29, -3.318]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [110856.0, 111393.0, 111929.0, 112465.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.225634475794
varray: [-20.785, -20.845, -21.203, -21.204]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [110856.0, 111393.0, 111929.0, 112465.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.88273431625
actual_array: [ 1.93 -15.84]
varray: [-3.212, -3.305, -3.29, -3.318]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.28125
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [110856.0, 111393.0, 111929.0, 112465.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -37.0132966592
actual_array: [ 18.1 -119. ]
varray: [-20.785, -20.845, -21.203, -21.204]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -21.00925
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [110856.0, 111393.0, 111929.0, 112465.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.93825
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 6.65425
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 37 - Sample 37 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 37
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9765
varray: [3.008, 2.948, 2.903, 3.047]
inst: mean_all
operation: mean_all(residuals)
tarray: [113049.0, 113585.0, 114121.0, 114657.0]
source_col: residuals
SAMP_residuals_STD | 0.0637102817448
varray: [3.008, 2.948, 2.903, 3.047]
inst: std_all
operation: std_all(residuals)
tarray: [113049.0, 113585.0, 114121.0, 114657.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.42475
varray: [0.468, 0.304, -0.07, 0.997]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [113049.0, 113585.0, 114121.0, 114657.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.442981846882
varray: [0.468, 0.304, -0.07, 0.997]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [113049.0, 113585.0, 114121.0, 114657.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1297.79625
varray: [1300.09, 1297.071, 1297.591, 1296.433]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [113049.0, 113585.0, 114121.0, 114657.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.60081799403
varray: [1300.09, 1297.071, 1297.591, 1296.433]
inst: std_all
operation: std_all(baseline_shift)
tarray: [113049.0, 113585.0, 114121.0, 114657.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.37375
varray: [1.886, 1.147, 0.846, 1.616]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [113049.0, 113585.0, 114121.0, 114657.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.465840011878
varray: [1.886, 1.147, 0.846, 1.616]
inst: std_all
operation: std_all(slope_shift)
tarray: [113049.0, 113585.0, 114121.0, 114657.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0469068225315
varray: [-2.278, -2.212, -2.178, -2.179]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [113049.0, 113585.0, 114121.0, 114657.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.312412120977
varray: [-15.919, -15.693, -15.175, -15.645]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [113049.0, 113585.0, 114121.0, 114657.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 114.57175
varray: [115.891, 113.678, 112.215, 116.503]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [113049.0, 113585.0, 114121.0, 114657.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.98521256209
varray: [115.891, 113.678, 112.215, 116.503]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [113049.0, 113585.0, 114121.0, 114657.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9765
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.42475
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1297.79625
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.37375
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.17545
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.17545
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 6.26372253262
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 6.4166
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-2.278, -2.212, -2.178, -2.179]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [113049.0, 113585.0, 114121.0, 114657.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-15.919, -15.693, -15.175, -15.645]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [113049.0, 113585.0, 114121.0, 114657.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.21175
varray: [-2.278, -2.212, -2.178, -2.179]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [113049.0, 113585.0, 114121.0, 114657.0]
source_col: d(18_16)Mean
dD_mean | -15.608
varray: [-15.919, -15.693, -15.175, -15.645]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [113049.0, 113585.0, 114121.0, 114657.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0469068225315
varray: [-2.278, -2.212, -2.178, -2.179]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [113049.0, 113585.0, 114121.0, 114657.0]
source_col: d(18_16)Mean
dD_stdev | 0.312412120977
varray: [-15.919, -15.693, -15.175, -15.645]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [113049.0, 113585.0, 114121.0, 114657.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.21175
varray: [-2.278, -2.212, -2.178, -2.179]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [113049.0, 113585.0, 114121.0, 114657.0]
source_col: d(18_16)Mean
measured_dD_mean | -15.608
varray: [-15.919, -15.693, -15.175, -15.645]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [113049.0, 113585.0, 114121.0, 114657.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0469068225315
varray: [-2.278, -2.212, -2.178, -2.179]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [113049.0, 113585.0, 114121.0, 114657.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.312412120977
varray: [-15.919, -15.693, -15.175, -15.645]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [113049.0, 113585.0, 114121.0, 114657.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.81610347518
actual_array: [ 1.93 -15.84]
varray: [-2.278, -2.212, -2.178, -2.179]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.21175
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [113049.0, 113585.0, 114121.0, 114657.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -31.9550655647
actual_array: [ 18.1 -119. ]
varray: [-15.919, -15.693, -15.175, -15.645]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -15.608
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [113049.0, 113585.0, 114121.0, 114657.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.91825
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 5.9015
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 38 - Sample 38 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 38
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.987
varray: [3.076, 2.957, 2.957, 2.958]
inst: mean_all
operation: mean_all(residuals)
tarray: [115241.0, 115777.0, 116313.0, 116849.0]
source_col: residuals
SAMP_residuals_STD | 0.059335205963
varray: [3.076, 2.957, 2.957, 2.958]
inst: std_all
operation: std_all(residuals)
tarray: [115241.0, 115777.0, 116313.0, 116849.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.428
varray: [0.602, 0.362, 0.348, 0.4]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [115241.0, 115777.0, 116313.0, 116849.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.118062130536
varray: [0.602, 0.362, 0.348, 0.4]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [115241.0, 115777.0, 116313.0, 116849.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1298.161
varray: [1300.515, 1297.457, 1296.802, 1297.87]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [115241.0, 115777.0, 116313.0, 116849.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.62977442202
varray: [1300.515, 1297.457, 1296.802, 1297.87]
inst: std_all
operation: std_all(baseline_shift)
tarray: [115241.0, 115777.0, 116313.0, 116849.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.57625
varray: [1.657, 1.406, 1.617, 1.625]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [115241.0, 115777.0, 116313.0, 116849.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.114808173344
varray: [1.657, 1.406, 1.617, 1.625]
inst: std_all
operation: std_all(slope_shift)
tarray: [115241.0, 115777.0, 116313.0, 116849.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0333716446503
varray: [-1.47, -1.441, -1.396, -1.463]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [115241.0, 115777.0, 116313.0, 116849.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.38583156947
varray: [-9.086, -8.774, -8.171, -8.541]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [115241.0, 115777.0, 116313.0, 116849.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 114.83325
varray: [117.565, 114.42, 113.6, 113.748]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [115241.0, 115777.0, 116313.0, 116849.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.85579333889
varray: [117.565, 114.42, 113.6, 113.748]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [115241.0, 115777.0, 116313.0, 116849.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.987
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.428
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1298.161
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.57625
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.18595
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.18595
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 6.63858196034
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 6.6781
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-1.47, -1.441, -1.396, -1.463]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [115241.0, 115777.0, 116313.0, 116849.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-9.086, -8.774, -8.171, -8.541]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [115241.0, 115777.0, 116313.0, 116849.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.4425
varray: [-1.47, -1.441, -1.396, -1.463]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [115241.0, 115777.0, 116313.0, 116849.0]
source_col: d(18_16)Mean
dD_mean | -8.643
varray: [-9.086, -8.774, -8.171, -8.541]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [115241.0, 115777.0, 116313.0, 116849.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0333716446503
varray: [-1.47, -1.441, -1.396, -1.463]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [115241.0, 115777.0, 116313.0, 116849.0]
source_col: d(18_16)Mean
dD_stdev | 0.38583156947
varray: [-9.086, -8.774, -8.171, -8.541]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [115241.0, 115777.0, 116313.0, 116849.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.4425
varray: [-1.47, -1.441, -1.396, -1.463]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [115241.0, 115777.0, 116313.0, 116849.0]
source_col: d(18_16)Mean
measured_dD_mean | -8.643
varray: [-9.086, -8.774, -8.171, -8.541]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [115241.0, 115777.0, 116313.0, 116849.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0333716446503
varray: [-1.47, -1.441, -1.396, -1.463]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [115241.0, 115777.0, 116313.0, 116849.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.38583156947
varray: [-9.086, -8.774, -8.171, -8.541]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [115241.0, 115777.0, 116313.0, 116849.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.04891715027
actual_array: [ 1.93 -15.84]
varray: [-1.47, -1.441, -1.396, -1.463]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -1.4425
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [115241.0, 115777.0, 116313.0, 116849.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -25.4323940401
actual_array: [ 18.1 -119. ]
varray: [-9.086, -8.774, -8.171, -8.541]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -8.643
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [115241.0, 115777.0, 116313.0, 116849.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.90775
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 6.26625
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 39 - Sample 39 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 39
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9605
varray: [3.066, 2.9, 2.946, 2.93]
inst: mean_all
operation: mean_all(residuals)
tarray: [117433.0, 117970.0, 118505.0, 119041.0]
source_col: residuals
SAMP_residuals_STD | 0.0728720339591
varray: [3.066, 2.9, 2.946, 2.93]
inst: std_all
operation: std_all(residuals)
tarray: [117433.0, 117970.0, 118505.0, 119041.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3115
varray: [0.391, 0.409, 0.185, 0.261]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [117433.0, 117970.0, 118505.0, 119041.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.107049832633
varray: [0.391, 0.409, 0.185, 0.261]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [117433.0, 117970.0, 118505.0, 119041.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1298.18
varray: [1300.2, 1296.495, 1297.902, 1298.123]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [117433.0, 117970.0, 118505.0, 119041.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.52754247077
varray: [1300.2, 1296.495, 1297.902, 1298.123]
inst: std_all
operation: std_all(baseline_shift)
tarray: [117433.0, 117970.0, 118505.0, 119041.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.52425
varray: [1.137, 1.965, 1.687, 1.308]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [117433.0, 117970.0, 118505.0, 119041.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.373039207055
varray: [1.137, 1.965, 1.687, 1.308]
inst: std_all
operation: std_all(slope_shift)
tarray: [117433.0, 117970.0, 118505.0, 119041.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0672823651982
varray: [-1.851, -1.921, -2.015, -1.936]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [117433.0, 117970.0, 118505.0, 119041.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.602195081902
varray: [-11.004, -11.917, -12.434, -12.014]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [117433.0, 117970.0, 118505.0, 119041.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 114.413
varray: [117.132, 113.072, 113.366, 114.082]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [117433.0, 117970.0, 118505.0, 119041.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.86163118438
varray: [117.132, 113.072, 113.366, 114.082]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [117433.0, 117970.0, 118505.0, 119041.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9605
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3115
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1298.18
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.52425
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.15945
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.15945
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 5.69250816658
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 6.25785
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-1.851, -1.921, -2.015, -1.936]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [117433.0, 117970.0, 118505.0, 119041.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-11.004, -11.917, -12.434, -12.014]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [117433.0, 117970.0, 118505.0, 119041.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.93075
varray: [-1.851, -1.921, -2.015, -1.936]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [117433.0, 117970.0, 118505.0, 119041.0]
source_col: d(18_16)Mean
dD_mean | -11.84225
varray: [-11.004, -11.917, -12.434, -12.014]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [117433.0, 117970.0, 118505.0, 119041.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0672823651982
varray: [-1.851, -1.921, -2.015, -1.936]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [117433.0, 117970.0, 118505.0, 119041.0]
source_col: d(18_16)Mean
dD_stdev | 0.602195081902
varray: [-11.004, -11.917, -12.434, -12.014]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [117433.0, 117970.0, 118505.0, 119041.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.93075
varray: [-1.851, -1.921, -2.015, -1.936]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [117433.0, 117970.0, 118505.0, 119041.0]
source_col: d(18_16)Mean
measured_dD_mean | -11.84225
varray: [-11.004, -11.917, -12.434, -12.014]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [117433.0, 117970.0, 118505.0, 119041.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0672823651982
varray: [-1.851, -1.921, -2.015, -1.936]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [117433.0, 117970.0, 118505.0, 119041.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.602195081902
varray: [-11.004, -11.917, -12.434, -12.014]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [117433.0, 117970.0, 118505.0, 119041.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.53585731684
actual_array: [ 1.93 -15.84]
varray: [-1.851, -1.921, -2.015, -1.936]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -1.93075
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [117433.0, 117970.0, 118505.0, 119041.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -28.4284682504
actual_array: [ 18.1 -119. ]
varray: [-11.004, -11.917, -12.434, -12.014]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -11.84225
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [117433.0, 117970.0, 118505.0, 119041.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.93425
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 6.28525
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 40 - Sample 40 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 40
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.91125
varray: [2.946, 2.866, 2.931, 2.902]
inst: mean_all
operation: mean_all(residuals)
tarray: [119626.0, 120162.0, 120698.0, 121234.0]
source_col: residuals
SAMP_residuals_STD | 0.0352644769326
varray: [2.946, 2.866, 2.931, 2.902]
inst: std_all
operation: std_all(residuals)
tarray: [119626.0, 120162.0, 120698.0, 121234.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.39725
varray: [0.307, 0.424, 0.406, 0.452]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [119626.0, 120162.0, 120698.0, 121234.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.063073369975
varray: [0.307, 0.424, 0.406, 0.452]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [119626.0, 120162.0, 120698.0, 121234.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1295.909
varray: [1297.846, 1293.916, 1296.329, 1295.545]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [119626.0, 120162.0, 120698.0, 121234.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.63634898478
varray: [1297.846, 1293.916, 1296.329, 1295.545]
inst: std_all
operation: std_all(baseline_shift)
tarray: [119626.0, 120162.0, 120698.0, 121234.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.95875
varray: [1.967, 1.561, 2.556, 1.751]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [119626.0, 120162.0, 120698.0, 121234.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.431331562335
varray: [1.967, 1.561, 2.556, 1.751]
inst: std_all
operation: std_all(slope_shift)
tarray: [119626.0, 120162.0, 120698.0, 121234.0]
source_col: slope_shift
SAMP_d18O_STD | 0.161621986953
varray: [-7.223, -7.522, -7.469, -7.596]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [119626.0, 120162.0, 120698.0, 121234.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.29985840144
varray: [-53.526, -57.585, -57.94, -58.613]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [119626.0, 120162.0, 120698.0, 121234.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 112.151
varray: [113.43, 111.025, 111.956, 112.193]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [119626.0, 120162.0, 120698.0, 121234.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.990539583594
varray: [113.43, 111.025, 111.956, 112.193]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [119626.0, 120162.0, 120698.0, 121234.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.91125
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.39725
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1295.909
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.95875
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.1102
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1102
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 3.93423894611
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 3.99585
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-7.223, -7.522, -7.469, -7.596]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [119626.0, 120162.0, 120698.0, 121234.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-53.526, -57.585, -57.94, -58.613]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [119626.0, 120162.0, 120698.0, 121234.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.4525
varray: [-7.223, -7.522, -7.469, -7.596]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [119626.0, 120162.0, 120698.0, 121234.0]
source_col: d(18_16)Mean
dD_mean | -56.916
varray: [-53.526, -57.585, -57.94, -58.613]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [119626.0, 120162.0, 120698.0, 121234.0]
source_col: d(D_H)Mean
d18O_stdev | 0.161621986953
varray: [-7.223, -7.522, -7.469, -7.596]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [119626.0, 120162.0, 120698.0, 121234.0]
source_col: d(18_16)Mean
dD_stdev | 2.29985840144
varray: [-53.526, -57.585, -57.94, -58.613]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [119626.0, 120162.0, 120698.0, 121234.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.4525
varray: [-7.223, -7.522, -7.469, -7.596]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [119626.0, 120162.0, 120698.0, 121234.0]
source_col: d(18_16)Mean
measured_dD_mean | -56.916
varray: [-53.526, -57.585, -57.94, -58.613]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [119626.0, 120162.0, 120698.0, 121234.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.161621986953
varray: [-7.223, -7.522, -7.469, -7.596]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [119626.0, 120162.0, 120698.0, 121234.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.29985840144
varray: [-53.526, -57.585, -57.94, -58.613]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [119626.0, 120162.0, 120698.0, 121234.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.0427940599
actual_array: [ 1.93 -15.84]
varray: [-7.223, -7.522, -7.469, -7.596]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.4525
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [119626.0, 120162.0, 120698.0, 121234.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -70.6397052405
actual_array: [ 18.1 -119. ]
varray: [-53.526, -57.585, -57.94, -58.613]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -56.916
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [119626.0, 120162.0, 120698.0, 121234.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.9835
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 4.01425
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 41 - Sample 41 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 41
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.98325
varray: [2.959, 2.963, 3.054, 2.957]
inst: mean_all
operation: mean_all(residuals)
tarray: [121820.0, 122356.0, 122892.0, 123428.0]
source_col: residuals
SAMP_residuals_STD | 0.0472325805633
varray: [2.959, 2.963, 3.054, 2.957]
inst: std_all
operation: std_all(residuals)
tarray: [121820.0, 122356.0, 122892.0, 123428.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3605
varray: [0.258, 0.395, 0.616, 0.173]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [121820.0, 122356.0, 122892.0, 123428.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.193333045977
varray: [0.258, 0.395, 0.616, 0.173]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [121820.0, 122356.0, 122892.0, 123428.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1299.1105
varray: [1300.63, 1296.725, 1299.498, 1299.589]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [121820.0, 122356.0, 122892.0, 123428.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.67118810032
varray: [1300.63, 1296.725, 1299.498, 1299.589]
inst: std_all
operation: std_all(baseline_shift)
tarray: [121820.0, 122356.0, 122892.0, 123428.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.6875
varray: [1.667, 1.413, 1.908, 1.762]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [121820.0, 122356.0, 122892.0, 123428.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.20811935678
varray: [1.667, 1.413, 1.908, 1.762]
inst: std_all
operation: std_all(slope_shift)
tarray: [121820.0, 122356.0, 122892.0, 123428.0]
source_col: slope_shift
SAMP_d18O_STD | 0.120781828103
varray: [-2.92, -2.728, -2.716, -2.635]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [121820.0, 122356.0, 122892.0, 123428.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.99090850284
varray: [-22.067, -19.0, -17.844, -17.861]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [121820.0, 122356.0, 122892.0, 123428.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 114.68975
varray: [115.544, 113.327, 115.599, 114.289]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [121820.0, 122356.0, 122892.0, 123428.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.09150763473
varray: [115.544, 113.327, 115.599, 114.289]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [121820.0, 122356.0, 122892.0, 123428.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.98325
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3605
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1299.1105
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.6875
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.1822
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1822
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 6.5047035933
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 6.5346
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-2.92, -2.728, -2.716, -2.635]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [121820.0, 122356.0, 122892.0, 123428.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-22.067, -19.0, -17.844, -17.861]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [121820.0, 122356.0, 122892.0, 123428.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.74975
varray: [-2.92, -2.728, -2.716, -2.635]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [121820.0, 122356.0, 122892.0, 123428.0]
source_col: d(18_16)Mean
dD_mean | -19.193
varray: [-22.067, -19.0, -17.844, -17.861]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [121820.0, 122356.0, 122892.0, 123428.0]
source_col: d(D_H)Mean
d18O_stdev | 0.120781828103
varray: [-2.92, -2.728, -2.716, -2.635]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [121820.0, 122356.0, 122892.0, 123428.0]
source_col: d(18_16)Mean
dD_stdev | 1.99090850284
varray: [-22.067, -19.0, -17.844, -17.861]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [121820.0, 122356.0, 122892.0, 123428.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.74975
varray: [-2.92, -2.728, -2.716, -2.635]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [121820.0, 122356.0, 122892.0, 123428.0]
source_col: d(18_16)Mean
measured_dD_mean | -19.193
varray: [-22.067, -19.0, -17.844, -17.861]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [121820.0, 122356.0, 122892.0, 123428.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.120781828103
varray: [-2.92, -2.728, -2.716, -2.635]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [121820.0, 122356.0, 122892.0, 123428.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.99090850284
varray: [-22.067, -19.0, -17.844, -17.861]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [121820.0, 122356.0, 122892.0, 123428.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.3526601769
actual_array: [ 1.93 -15.84]
varray: [-2.92, -2.728, -2.716, -2.635]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.74975
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [121820.0, 122356.0, 122892.0, 123428.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -35.3123918269
actual_array: [ 18.1 -119. ]
varray: [-22.067, -19.0, -17.844, -17.861]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -19.193
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [121820.0, 122356.0, 122892.0, 123428.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.9115
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 7.21575
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 42 - Sample 42 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 42
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.01
varray: [3.083, 2.977, 3.014, 2.966]
inst: mean_all
operation: mean_all(residuals)
tarray: [124013.0, 124549.0, 125085.0, 125621.0]
source_col: residuals
SAMP_residuals_STD | 0.052820450585
varray: [3.083, 2.977, 3.014, 2.966]
inst: std_all
operation: std_all(residuals)
tarray: [124013.0, 124549.0, 125085.0, 125621.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.476
varray: [0.487, 0.371, 0.729, 0.317]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [124013.0, 124549.0, 125085.0, 125621.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.182971764671
varray: [0.487, 0.371, 0.729, 0.317]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [124013.0, 124549.0, 125085.0, 125621.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1298.282
varray: [1300.719, 1297.846, 1297.325, 1297.238]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [124013.0, 124549.0, 125085.0, 125621.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.64669871764
varray: [1300.719, 1297.846, 1297.325, 1297.238]
inst: std_all
operation: std_all(baseline_shift)
tarray: [124013.0, 124549.0, 125085.0, 125621.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.6235
varray: [1.42, 1.619, 1.538, 1.917]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [124013.0, 124549.0, 125085.0, 125621.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.212041662573
varray: [1.42, 1.619, 1.538, 1.917]
inst: std_all
operation: std_all(slope_shift)
tarray: [124013.0, 124549.0, 125085.0, 125621.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0307571129985
varray: [-2.396, -2.418, -2.421, -2.469]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [124013.0, 124549.0, 125085.0, 125621.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.19536717568
varray: [-16.701, -17.129, -16.988, -17.1]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [124013.0, 124549.0, 125085.0, 125621.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 114.7175
varray: [117.107, 113.408, 115.273, 113.082]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [124013.0, 124549.0, 125085.0, 125621.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.86260937039
varray: [117.107, 113.408, 115.273, 113.082]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [124013.0, 124549.0, 125085.0, 125621.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.01
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.476
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1298.282
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.6235
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.20895
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.20895
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 7.45970261152
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 6.56235
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-2.396, -2.418, -2.421, -2.469]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [124013.0, 124549.0, 125085.0, 125621.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-16.701, -17.129, -16.988, -17.1]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [124013.0, 124549.0, 125085.0, 125621.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.426
varray: [-2.396, -2.418, -2.421, -2.469]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [124013.0, 124549.0, 125085.0, 125621.0]
source_col: d(18_16)Mean
dD_mean | -16.9795
varray: [-16.701, -17.129, -16.988, -17.1]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [124013.0, 124549.0, 125085.0, 125621.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0307571129985
varray: [-2.396, -2.418, -2.421, -2.469]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [124013.0, 124549.0, 125085.0, 125621.0]
source_col: d(18_16)Mean
dD_stdev | 0.19536717568
varray: [-16.701, -17.129, -16.988, -17.1]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [124013.0, 124549.0, 125085.0, 125621.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.426
varray: [-2.396, -2.418, -2.421, -2.469]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [124013.0, 124549.0, 125085.0, 125621.0]
source_col: d(18_16)Mean
measured_dD_mean | -16.9795
varray: [-16.701, -17.129, -16.988, -17.1]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [124013.0, 124549.0, 125085.0, 125621.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0307571129985
varray: [-2.396, -2.418, -2.421, -2.469]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [124013.0, 124549.0, 125085.0, 125621.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.19536717568
varray: [-16.701, -17.129, -16.988, -17.1]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [124013.0, 124549.0, 125085.0, 125621.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.02977870444
actual_array: [ 1.93 -15.84]
varray: [-2.396, -2.418, -2.421, -2.469]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.426
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [124013.0, 124549.0, 125085.0, 125621.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -33.239465277
actual_array: [ 18.1 -119. ]
varray: [-16.701, -17.129, -16.988, -17.1]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -16.9795
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [124013.0, 124549.0, 125085.0, 125621.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.88475
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 6.38725
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 43 - Sample 43 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 43
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.96125
varray: [2.939, 3.021, 2.943, 2.942]
inst: mean_all
operation: mean_all(residuals)
tarray: [126205.0, 126741.0, 127277.0, 127812.0]
source_col: residuals
SAMP_residuals_STD | 0.0398695790464
varray: [2.939, 3.021, 2.943, 2.942]
inst: std_all
operation: std_all(residuals)
tarray: [126205.0, 126741.0, 127277.0, 127812.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3135
varray: [0.345, 0.275, 0.406, 0.228]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [126205.0, 126741.0, 127277.0, 127812.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0781899396768
varray: [0.345, 0.275, 0.406, 0.228]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [126205.0, 126741.0, 127277.0, 127812.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1298.67825
varray: [1298.109, 1301.067, 1297.318, 1298.219]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [126205.0, 126741.0, 127277.0, 127812.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.64229136169
varray: [1298.109, 1301.067, 1297.318, 1298.219]
inst: std_all
operation: std_all(baseline_shift)
tarray: [126205.0, 126741.0, 127277.0, 127812.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.61125
varray: [1.762, 1.432, 1.947, 1.304]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [126205.0, 126741.0, 127277.0, 127812.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.295514664047
varray: [1.762, 1.432, 1.947, 1.304]
inst: std_all
operation: std_all(slope_shift)
tarray: [126205.0, 126741.0, 127277.0, 127812.0]
source_col: slope_shift
SAMP_d18O_STD | 0.118265238623
varray: [-2.39, -2.116, -2.298, -2.332]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [126205.0, 126741.0, 127277.0, 127812.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.633521046743
varray: [-16.514, -15.251, -16.317, -16.639]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [126205.0, 126741.0, 127277.0, 127812.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 114.37775
varray: [114.166, 116.377, 113.5, 113.468]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [126205.0, 126741.0, 127277.0, 127812.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.37112225932
varray: [114.166, 116.377, 113.5, 113.468]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [126205.0, 126741.0, 127277.0, 127812.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.96125
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3135
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1298.67825
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.61125
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.1602
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1602
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 5.71928383999
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 6.2226
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-2.39, -2.116, -2.298, -2.332]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [126205.0, 126741.0, 127277.0, 127812.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-16.514, -15.251, -16.317, -16.639]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [126205.0, 126741.0, 127277.0, 127812.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.284
varray: [-2.39, -2.116, -2.298, -2.332]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [126205.0, 126741.0, 127277.0, 127812.0]
source_col: d(18_16)Mean
dD_mean | -16.18025
varray: [-16.514, -15.251, -16.317, -16.639]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [126205.0, 126741.0, 127277.0, 127812.0]
source_col: d(D_H)Mean
d18O_stdev | 0.118265238623
varray: [-2.39, -2.116, -2.298, -2.332]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [126205.0, 126741.0, 127277.0, 127812.0]
source_col: d(18_16)Mean
dD_stdev | 0.633521046743
varray: [-16.514, -15.251, -16.317, -16.639]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [126205.0, 126741.0, 127277.0, 127812.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.284
varray: [-2.39, -2.116, -2.298, -2.332]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [126205.0, 126741.0, 127277.0, 127812.0]
source_col: d(18_16)Mean
measured_dD_mean | -16.18025
varray: [-16.514, -15.251, -16.317, -16.639]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [126205.0, 126741.0, 127277.0, 127812.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.118265238623
varray: [-2.39, -2.116, -2.298, -2.332]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [126205.0, 126741.0, 127277.0, 127812.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.633521046743
varray: [-16.514, -15.251, -16.317, -16.639]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [126205.0, 126741.0, 127277.0, 127812.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.88815964934
actual_array: [ 1.93 -15.84]
varray: [-2.39, -2.116, -2.298, -2.332]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.284
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [126205.0, 126741.0, 127277.0, 127812.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -32.4909735016
actual_array: [ 18.1 -119. ]
varray: [-16.514, -15.251, -16.317, -16.639]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -16.18025
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [126205.0, 126741.0, 127277.0, 127812.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.9335
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 6.7835
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 44 - Sample 44 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 44
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.95625
varray: [2.975, 2.959, 2.908, 2.983]
inst: mean_all
operation: mean_all(residuals)
tarray: [128397.0, 128933.0, 129469.0, 130004.0]
source_col: residuals
SAMP_residuals_STD | 0.0336786282381
varray: [2.975, 2.959, 2.908, 2.983]
inst: std_all
operation: std_all(residuals)
tarray: [128397.0, 128933.0, 129469.0, 130004.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.36375
varray: [0.401, 0.299, 0.194, 0.561]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [128397.0, 128933.0, 129469.0, 130004.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.156314586651
varray: [0.401, 0.299, 0.194, 0.561]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [128397.0, 128933.0, 129469.0, 130004.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1298.1375
varray: [1300.107, 1298.009, 1297.225, 1297.209]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [128397.0, 128933.0, 129469.0, 130004.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.3650654441
varray: [1300.107, 1298.009, 1297.225, 1297.209]
inst: std_all
operation: std_all(baseline_shift)
tarray: [128397.0, 128933.0, 129469.0, 130004.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.5825
varray: [0.989, 2.258, 1.86, 1.223]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [128397.0, 128933.0, 129469.0, 130004.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.581603817044
varray: [0.989, 2.258, 1.86, 1.223]
inst: std_all
operation: std_all(slope_shift)
tarray: [128397.0, 128933.0, 129469.0, 130004.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0338661581327
varray: [-1.368, -1.292, -1.353, -1.354]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [128397.0, 128933.0, 129469.0, 130004.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.553908837265
varray: [-6.95, -5.717, -5.844, -6.143]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [128397.0, 128933.0, 129469.0, 130004.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 114.058
varray: [116.173, 112.952, 112.325, 114.782]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [128397.0, 128933.0, 129469.0, 130004.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.7534657111
varray: [116.173, 112.952, 112.325, 114.782]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [128397.0, 128933.0, 129469.0, 130004.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.95625
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.36375
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1298.1375
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.5825
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.1552
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1552
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 5.5407793506
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 5.90285
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-1.368, -1.292, -1.353, -1.354]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [128397.0, 128933.0, 129469.0, 130004.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-6.95, -5.717, -5.844, -6.143]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [128397.0, 128933.0, 129469.0, 130004.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.34175
varray: [-1.368, -1.292, -1.353, -1.354]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [128397.0, 128933.0, 129469.0, 130004.0]
source_col: d(18_16)Mean
dD_mean | -6.1635
varray: [-6.95, -5.717, -5.844, -6.143]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [128397.0, 128933.0, 129469.0, 130004.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0338661581327
varray: [-1.368, -1.292, -1.353, -1.354]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [128397.0, 128933.0, 129469.0, 130004.0]
source_col: d(18_16)Mean
dD_stdev | 0.553908837265
varray: [-6.95, -5.717, -5.844, -6.143]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [128397.0, 128933.0, 129469.0, 130004.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.34175
varray: [-1.368, -1.292, -1.353, -1.354]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [128397.0, 128933.0, 129469.0, 130004.0]
source_col: d(18_16)Mean
measured_dD_mean | -6.1635
varray: [-6.95, -5.717, -5.844, -6.143]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [128397.0, 128933.0, 129469.0, 130004.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0338661581327
varray: [-1.368, -1.292, -1.353, -1.354]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [128397.0, 128933.0, 129469.0, 130004.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.553908837265
varray: [-6.95, -5.717, -5.844, -6.143]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [128397.0, 128933.0, 129469.0, 130004.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.94843743335
actual_array: [ 1.93 -15.84]
varray: [-1.368, -1.292, -1.353, -1.354]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -1.34175
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [128397.0, 128933.0, 129469.0, 130004.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -23.110360437
actual_array: [ 18.1 -119. ]
varray: [-6.95, -5.717, -5.844, -6.143]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -6.1635
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [128397.0, 128933.0, 129469.0, 130004.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.9385
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 6.24275
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 45 - Sample 45 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 45
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.95975
varray: [3.023, 2.876, 2.972, 2.968]
inst: mean_all
operation: mean_all(residuals)
tarray: [130590.0, 131125.0, 131661.0, 132197.0]
source_col: residuals
SAMP_residuals_STD | 0.0611902770054
varray: [3.023, 2.876, 2.972, 2.968]
inst: std_all
operation: std_all(residuals)
tarray: [130590.0, 131125.0, 131661.0, 132197.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3915
varray: [0.279, 0.326, 0.599, 0.362]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [130590.0, 131125.0, 131661.0, 132197.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.142446481178
varray: [0.279, 0.326, 0.599, 0.362]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [130590.0, 131125.0, 131661.0, 132197.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1297.5555
varray: [1300.253, 1294.432, 1297.561, 1297.976]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [130590.0, 131125.0, 131661.0, 132197.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.3951067478
varray: [1300.253, 1294.432, 1297.561, 1297.976]
inst: std_all
operation: std_all(baseline_shift)
tarray: [130590.0, 131125.0, 131661.0, 132197.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.897
varray: [2.146, 1.939, 2.176, 1.327]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [130590.0, 131125.0, 131661.0, 132197.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.39433741897
varray: [2.146, 1.939, 2.176, 1.327]
inst: std_all
operation: std_all(slope_shift)
tarray: [130590.0, 131125.0, 131661.0, 132197.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0337577151675
varray: [-1.25, -1.326, -1.262, -1.269]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [130590.0, 131125.0, 131661.0, 132197.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.270625663971
varray: [-6.857, -7.499, -7.039, -7.1]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [130590.0, 131125.0, 131661.0, 132197.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.40125
varray: [115.509, 110.234, 113.622, 114.24]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [130590.0, 131125.0, 131661.0, 132197.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.25287333199
varray: [115.509, 110.234, 113.622, 114.24]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [130590.0, 131125.0, 131661.0, 132197.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.95975
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3915
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1297.5555
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.897
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.1587
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1587
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 5.66573249317
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 5.2461
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-1.25, -1.326, -1.262, -1.269]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [130590.0, 131125.0, 131661.0, 132197.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-6.857, -7.499, -7.039, -7.1]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [130590.0, 131125.0, 131661.0, 132197.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.27675
varray: [-1.25, -1.326, -1.262, -1.269]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [130590.0, 131125.0, 131661.0, 132197.0]
source_col: d(18_16)Mean
dD_mean | -7.12375
varray: [-6.857, -7.499, -7.039, -7.1]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [130590.0, 131125.0, 131661.0, 132197.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0337577151675
varray: [-1.25, -1.326, -1.262, -1.269]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [130590.0, 131125.0, 131661.0, 132197.0]
source_col: d(18_16)Mean
dD_stdev | 0.270625663971
varray: [-6.857, -7.499, -7.039, -7.1]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [130590.0, 131125.0, 131661.0, 132197.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.27675
varray: [-1.25, -1.326, -1.262, -1.269]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [130590.0, 131125.0, 131661.0, 132197.0]
source_col: d(18_16)Mean
measured_dD_mean | -7.12375
varray: [-6.857, -7.499, -7.039, -7.1]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [130590.0, 131125.0, 131661.0, 132197.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0337577151675
varray: [-1.25, -1.326, -1.262, -1.269]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [130590.0, 131125.0, 131661.0, 132197.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.270625663971
varray: [-6.857, -7.499, -7.039, -7.1]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [130590.0, 131125.0, 131661.0, 132197.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.88361180954
actual_array: [ 1.93 -15.84]
varray: [-1.25, -1.326, -1.262, -1.269]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -1.27675
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [130590.0, 131125.0, 131661.0, 132197.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -24.0096275341
actual_array: [ 18.1 -119. ]
varray: [-6.857, -7.499, -7.039, -7.1]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -7.12375
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [130590.0, 131125.0, 131661.0, 132197.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.935
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 5.66075
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 46 - Sample 46 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 46
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9115
varray: [2.94, 2.842, 2.939, 2.925]
inst: mean_all
operation: mean_all(residuals)
tarray: [132783.0, 133318.0, 133854.0, 134390.0]
source_col: residuals
SAMP_residuals_STD | 0.0468365953787
varray: [2.94, 2.842, 2.939, 2.925]
inst: std_all
operation: std_all(residuals)
tarray: [132783.0, 133318.0, 133854.0, 134390.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.336
varray: [0.163, 0.362, 0.334, 0.485]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [132783.0, 133318.0, 133854.0, 134390.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.132677553992
varray: [0.163, 0.362, 0.334, 0.485]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [132783.0, 133318.0, 133854.0, 134390.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1296.8055
varray: [1298.446, 1293.672, 1298.739, 1296.365]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [132783.0, 133318.0, 133854.0, 134390.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.34111974633
varray: [1298.446, 1293.672, 1298.739, 1296.365]
inst: std_all
operation: std_all(baseline_shift)
tarray: [132783.0, 133318.0, 133854.0, 134390.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.68525
varray: [1.986, 1.599, 1.914, 1.242]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [132783.0, 133318.0, 133854.0, 134390.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.339944480761
varray: [1.986, 1.599, 1.914, 1.242]
inst: std_all
operation: std_all(slope_shift)
tarray: [132783.0, 133318.0, 133854.0, 134390.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0591178202124
varray: [-2.593, -2.702, -2.678, -2.73]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [132783.0, 133318.0, 133854.0, 134390.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.59703789662
varray: [-16.354, -17.466, -17.32, -17.719]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [132783.0, 133318.0, 133854.0, 134390.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 112.68275
varray: [112.885, 110.961, 113.36, 113.525]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [132783.0, 133318.0, 133854.0, 134390.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.17946043455
varray: [112.885, 110.961, 113.36, 113.525]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [132783.0, 133318.0, 133854.0, 134390.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9115
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.336
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1296.8055
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.68525
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.11045
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.11045
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 3.94316417058
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 4.5276
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-2.593, -2.702, -2.678, -2.73]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [132783.0, 133318.0, 133854.0, 134390.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-16.354, -17.466, -17.32, -17.719]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [132783.0, 133318.0, 133854.0, 134390.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.67575
varray: [-2.593, -2.702, -2.678, -2.73]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [132783.0, 133318.0, 133854.0, 134390.0]
source_col: d(18_16)Mean
dD_mean | -17.21475
varray: [-16.354, -17.466, -17.32, -17.719]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [132783.0, 133318.0, 133854.0, 134390.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0591178202124
varray: [-2.593, -2.702, -2.678, -2.73]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [132783.0, 133318.0, 133854.0, 134390.0]
source_col: d(18_16)Mean
dD_stdev | 0.59703789662
varray: [-16.354, -17.466, -17.32, -17.719]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [132783.0, 133318.0, 133854.0, 134390.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.67575
varray: [-2.593, -2.702, -2.678, -2.73]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [132783.0, 133318.0, 133854.0, 134390.0]
source_col: d(18_16)Mean
measured_dD_mean | -17.21475
varray: [-16.354, -17.466, -17.32, -17.719]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [132783.0, 133318.0, 133854.0, 134390.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0591178202124
varray: [-2.593, -2.702, -2.678, -2.73]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [132783.0, 133318.0, 133854.0, 134390.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.59703789662
varray: [-16.354, -17.466, -17.32, -17.719]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [132783.0, 133318.0, 133854.0, 134390.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.27885869748
actual_array: [ 1.93 -15.84]
varray: [-2.593, -2.702, -2.678, -2.73]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.67575
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [132783.0, 133318.0, 133854.0, 134390.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -33.4597751803
actual_array: [ 18.1 -119. ]
varray: [-16.354, -17.466, -17.32, -17.719]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -17.21475
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [132783.0, 133318.0, 133854.0, 134390.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.98325
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 4.91075
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 47 - Sample 47 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 47
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.97825
varray: [3.066, 2.918, 2.954, 2.975]
inst: mean_all
operation: mean_all(residuals)
tarray: [134976.0, 135513.0, 136048.0, 136584.0]
source_col: residuals
SAMP_residuals_STD | 0.0630575134302
varray: [3.066, 2.918, 2.954, 2.975]
inst: std_all
operation: std_all(residuals)
tarray: [134976.0, 135513.0, 136048.0, 136584.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.50075
varray: [0.6, 0.621, 0.402, 0.38]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [134976.0, 135513.0, 136048.0, 136584.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.127335187596
varray: [0.6, 0.621, 0.402, 0.38]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [134976.0, 135513.0, 136048.0, 136584.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1297.96775
varray: [1301.139, 1295.001, 1298.009, 1297.722]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [134976.0, 135513.0, 136048.0, 136584.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.51134271855
varray: [1301.139, 1295.001, 1298.009, 1297.722]
inst: std_all
operation: std_all(baseline_shift)
tarray: [134976.0, 135513.0, 136048.0, 136584.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.28925
varray: [0.95, 0.888, 1.628, 1.691]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [134976.0, 135513.0, 136048.0, 136584.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.429048074229
varray: [0.95, 0.888, 1.628, 1.691]
inst: std_all
operation: std_all(slope_shift)
tarray: [134976.0, 135513.0, 136048.0, 136584.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0625159979525
varray: [-2.7, -2.838, -2.718, -2.727]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [134976.0, 135513.0, 136048.0, 136584.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.45733466958
varray: [-17.524, -18.625, -18.192, -18.241]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [134976.0, 135513.0, 136048.0, 136584.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 114.57275
varray: [118.207, 113.184, 113.184, 113.716]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [134976.0, 135513.0, 136048.0, 136584.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.43577822951
varray: [118.207, 113.184, 113.184, 113.716]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [134976.0, 135513.0, 136048.0, 136584.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.97825
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.50075
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1297.96775
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.28925
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.1772
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1772
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 6.32619910391
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 6.4176
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-2.7, -2.838, -2.718, -2.727]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [134976.0, 135513.0, 136048.0, 136584.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-17.524, -18.625, -18.192, -18.241]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [134976.0, 135513.0, 136048.0, 136584.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.74575
varray: [-2.7, -2.838, -2.718, -2.727]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [134976.0, 135513.0, 136048.0, 136584.0]
source_col: d(18_16)Mean
dD_mean | -18.1455
varray: [-17.524, -18.625, -18.192, -18.241]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [134976.0, 135513.0, 136048.0, 136584.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0625159979525
varray: [-2.7, -2.838, -2.718, -2.727]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [134976.0, 135513.0, 136048.0, 136584.0]
source_col: d(18_16)Mean
dD_stdev | 0.45733466958
varray: [-17.524, -18.625, -18.192, -18.241]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [134976.0, 135513.0, 136048.0, 136584.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.74575
varray: [-2.7, -2.838, -2.718, -2.727]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [134976.0, 135513.0, 136048.0, 136584.0]
source_col: d(18_16)Mean
measured_dD_mean | -18.1455
varray: [-17.524, -18.625, -18.192, -18.241]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [134976.0, 135513.0, 136048.0, 136584.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0625159979525
varray: [-2.7, -2.838, -2.718, -2.727]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [134976.0, 135513.0, 136048.0, 136584.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.45733466958
varray: [-17.524, -18.625, -18.192, -18.241]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [134976.0, 135513.0, 136048.0, 136584.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.34867090774
actual_array: [ 1.93 -15.84]
varray: [-2.7, -2.838, -2.718, -2.727]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.74575
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [134976.0, 135513.0, 136048.0, 136584.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -34.3314157433
actual_array: [ 18.1 -119. ]
varray: [-17.524, -18.625, -18.192, -18.241]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -18.1455
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [134976.0, 135513.0, 136048.0, 136584.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.9165
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 6.073
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 48 - Sample 48 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 48
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.961
varray: [2.988, 2.869, 2.965, 3.022]
inst: mean_all
operation: mean_all(residuals)
tarray: [137171.0, 137707.0, 138243.0, 138779.0]
source_col: residuals
SAMP_residuals_STD | 0.0656505902487
varray: [2.988, 2.869, 2.965, 3.022]
inst: std_all
operation: std_all(residuals)
tarray: [137171.0, 137707.0, 138243.0, 138779.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.33675
varray: [0.445, 0.188, 0.143, 0.571]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [137171.0, 137707.0, 138243.0, 138779.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.205147711012
varray: [0.445, 0.188, 0.143, 0.571]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [137171.0, 137707.0, 138243.0, 138779.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1297.84175
varray: [1299.028, 1294.794, 1299.415, 1298.13]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [137171.0, 137707.0, 138243.0, 138779.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.1019175967
varray: [1299.028, 1294.794, 1299.415, 1298.13]
inst: std_all
operation: std_all(baseline_shift)
tarray: [137171.0, 137707.0, 138243.0, 138779.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.70725
varray: [1.867, 1.636, 1.812, 1.514]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [137171.0, 137707.0, 138243.0, 138779.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.162187905427
varray: [1.867, 1.636, 1.812, 1.514]
inst: std_all
operation: std_all(slope_shift)
tarray: [137171.0, 137707.0, 138243.0, 138779.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0562101710844
varray: [-2.187, -2.131, -2.053, -2.1]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [137171.0, 137707.0, 138243.0, 138779.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.327402275496
varray: [-14.029, -13.561, -13.231, -13.606]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [137171.0, 137707.0, 138243.0, 138779.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.364
varray: [114.754, 109.884, 113.885, 114.933]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [137171.0, 137707.0, 138243.0, 138779.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.36471999752
varray: [114.754, 109.884, 113.885, 114.933]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [137171.0, 137707.0, 138243.0, 138779.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.961
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.33675
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1297.84175
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.70725
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.15995
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.15995
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 5.71035861552
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 5.20885
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-2.187, -2.131, -2.053, -2.1]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [137171.0, 137707.0, 138243.0, 138779.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-14.029, -13.561, -13.231, -13.606]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [137171.0, 137707.0, 138243.0, 138779.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.11775
varray: [-2.187, -2.131, -2.053, -2.1]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [137171.0, 137707.0, 138243.0, 138779.0]
source_col: d(18_16)Mean
dD_mean | -13.60675
varray: [-14.029, -13.561, -13.231, -13.606]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [137171.0, 137707.0, 138243.0, 138779.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0562101710844
varray: [-2.187, -2.131, -2.053, -2.1]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [137171.0, 137707.0, 138243.0, 138779.0]
source_col: d(18_16)Mean
dD_stdev | 0.327402275496
varray: [-14.029, -13.561, -13.231, -13.606]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [137171.0, 137707.0, 138243.0, 138779.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.11775
varray: [-2.187, -2.131, -2.053, -2.1]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [137171.0, 137707.0, 138243.0, 138779.0]
source_col: d(18_16)Mean
measured_dD_mean | -13.60675
varray: [-14.029, -13.561, -13.231, -13.606]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [137171.0, 137707.0, 138243.0, 138779.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0562101710844
varray: [-2.187, -2.131, -2.053, -2.1]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [137171.0, 137707.0, 138243.0, 138779.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.327402275496
varray: [-14.029, -13.561, -13.231, -13.606]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [137171.0, 137707.0, 138243.0, 138779.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.72235564997
actual_array: [ 1.93 -15.84]
varray: [-2.187, -2.131, -2.053, -2.1]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.11775
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [137171.0, 137707.0, 138243.0, 138779.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -30.0809095864
actual_array: [ 18.1 -119. ]
varray: [-14.029, -13.561, -13.231, -13.606]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -13.60675
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [137171.0, 137707.0, 138243.0, 138779.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.93375
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 5.947
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 49 - Sample 49 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 49
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.93
varray: [2.943, 2.879, 2.94, 2.958]
inst: mean_all
operation: mean_all(residuals)
tarray: [139366.0, 139902.0, 140438.0, 140974.0]
source_col: residuals
SAMP_residuals_STD | 0.0348998567332
varray: [2.943, 2.879, 2.94, 2.958]
inst: std_all
operation: std_all(residuals)
tarray: [139366.0, 139902.0, 140438.0, 140974.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.266
varray: [0.292, 0.247, 0.241, 0.284]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [139366.0, 139902.0, 140438.0, 140974.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0257293606605
varray: [0.292, 0.247, 0.241, 0.284]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [139366.0, 139902.0, 140438.0, 140974.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1297.09875
varray: [1298.209, 1294.997, 1297.648, 1297.541]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [139366.0, 139902.0, 140438.0, 140974.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.43146413973
varray: [1298.209, 1294.997, 1297.648, 1297.541]
inst: std_all
operation: std_all(baseline_shift)
tarray: [139366.0, 139902.0, 140438.0, 140974.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.70875
varray: [1.71, 1.679, 1.501, 1.945]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [139366.0, 139902.0, 140438.0, 140974.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.182447024275
varray: [1.71, 1.679, 1.501, 1.945]
inst: std_all
operation: std_all(slope_shift)
tarray: [139366.0, 139902.0, 140438.0, 140974.0]
source_col: slope_shift
SAMP_d18O_STD | 0.15868102176
varray: [-7.192, -7.476, -7.532, -7.51]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [139366.0, 139902.0, 140438.0, 140974.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.20942745223
varray: [-53.577, -57.22, -58.151, -58.302]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [139366.0, 139902.0, 140438.0, 140974.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 112.49025
varray: [112.899, 111.004, 113.056, 113.002]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [139366.0, 139902.0, 140438.0, 140974.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.992971424564
varray: [112.899, 111.004, 113.056, 113.002]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [139366.0, 139902.0, 140438.0, 140974.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.93
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.266
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1297.09875
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.70875
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.12895
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.12895
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 4.60363078131
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 4.3351
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-7.192, -7.476, -7.532, -7.51]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [139366.0, 139902.0, 140438.0, 140974.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-53.577, -57.22, -58.151, -58.302]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [139366.0, 139902.0, 140438.0, 140974.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.4275
varray: [-7.192, -7.476, -7.532, -7.51]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [139366.0, 139902.0, 140438.0, 140974.0]
source_col: d(18_16)Mean
dD_mean | -56.8125
varray: [-53.577, -57.22, -58.151, -58.302]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [139366.0, 139902.0, 140438.0, 140974.0]
source_col: d(D_H)Mean
d18O_stdev | 0.15868102176
varray: [-7.192, -7.476, -7.532, -7.51]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [139366.0, 139902.0, 140438.0, 140974.0]
source_col: d(18_16)Mean
dD_stdev | 2.20942745223
varray: [-53.577, -57.22, -58.151, -58.302]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [139366.0, 139902.0, 140438.0, 140974.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.4275
varray: [-7.192, -7.476, -7.532, -7.51]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [139366.0, 139902.0, 140438.0, 140974.0]
source_col: d(18_16)Mean
measured_dD_mean | -56.8125
varray: [-53.577, -57.22, -58.151, -58.302]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [139366.0, 139902.0, 140438.0, 140974.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.15868102176
varray: [-7.192, -7.476, -7.532, -7.51]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [139366.0, 139902.0, 140438.0, 140974.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.20942745223
varray: [-53.577, -57.22, -58.151, -58.302]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [139366.0, 139902.0, 140438.0, 140974.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.0178611276
actual_array: [ 1.93 -15.84]
varray: [-7.192, -7.476, -7.532, -7.51]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.4275
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [139366.0, 139902.0, 140438.0, 140974.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -70.542778248
actual_array: [ 18.1 -119. ]
varray: [-53.577, -57.22, -58.151, -58.302]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -56.8125
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [139366.0, 139902.0, 140438.0, 140974.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.96475
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 5.204
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 50 - Sample 50 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 50
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.86475
varray: [3.061, 2.759, 2.85, 2.789]
inst: mean_all
operation: mean_all(residuals)
tarray: [141558.0, 142094.0, 142630.0, 143166.0]
source_col: residuals
SAMP_residuals_STD | 0.136201749866
varray: [3.061, 2.759, 2.85, 2.789]
inst: std_all
operation: std_all(residuals)
tarray: [141558.0, 142094.0, 142630.0, 143166.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.27075
varray: [0.475, 0.248, 0.327, 0.033]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [141558.0, 142094.0, 142630.0, 143166.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.184322859859
varray: [0.475, 0.248, 0.327, 0.033]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [141558.0, 142094.0, 142630.0, 143166.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1293.9315
varray: [1300.724, 1289.995, 1292.703, 1292.304]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [141558.0, 142094.0, 142630.0, 143166.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 4.68302099789
varray: [1300.724, 1289.995, 1292.703, 1292.304]
inst: std_all
operation: std_all(baseline_shift)
tarray: [141558.0, 142094.0, 142630.0, 143166.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.50175
varray: [1.554, 1.408, 1.59, 1.455]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [141558.0, 142094.0, 142630.0, 143166.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.084641892701
varray: [1.554, 1.408, 1.59, 1.455]
inst: std_all
operation: std_all(slope_shift)
tarray: [141558.0, 142094.0, 142630.0, 143166.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0915691905974
varray: [-3.058, -2.984, -2.859, -2.886]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [141558.0, 142094.0, 142630.0, 143166.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.53729784579
varray: [-22.99, -20.835, -19.833, -19.634]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [141558.0, 142094.0, 142630.0, 143166.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.32875
varray: [116.755, 106.932, 109.432, 108.196]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [141558.0, 142094.0, 142630.0, 143166.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 4.40406565006
varray: [116.755, 106.932, 109.432, 108.196]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [141558.0, 142094.0, 142630.0, 143166.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.86475
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.27075
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1293.9315
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.50175
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0637
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0637
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.2741471948
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 2.1736
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-3.058, -2.984, -2.859, -2.886]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [141558.0, 142094.0, 142630.0, 143166.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-22.99, -20.835, -19.833, -19.634]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [141558.0, 142094.0, 142630.0, 143166.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.94675
varray: [-3.058, -2.984, -2.859, -2.886]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [141558.0, 142094.0, 142630.0, 143166.0]
source_col: d(18_16)Mean
dD_mean | -20.823
varray: [-22.99, -20.835, -19.833, -19.634]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [141558.0, 142094.0, 142630.0, 143166.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0915691905974
varray: [-3.058, -2.984, -2.859, -2.886]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [141558.0, 142094.0, 142630.0, 143166.0]
source_col: d(18_16)Mean
dD_stdev | 1.53729784579
varray: [-22.99, -20.835, -19.833, -19.634]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [141558.0, 142094.0, 142630.0, 143166.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.94675
varray: [-3.058, -2.984, -2.859, -2.886]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [141558.0, 142094.0, 142630.0, 143166.0]
source_col: d(18_16)Mean
measured_dD_mean | -20.823
varray: [-22.99, -20.835, -19.833, -19.634]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [141558.0, 142094.0, 142630.0, 143166.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0915691905974
varray: [-3.058, -2.984, -2.859, -2.886]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [141558.0, 142094.0, 142630.0, 143166.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.53729784579
varray: [-22.99, -20.835, -19.833, -19.634]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [141558.0, 142094.0, 142630.0, 143166.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.54913168292
actual_array: [ 1.93 -15.84]
varray: [-3.058, -2.984, -2.859, -2.886]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.94675
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [141558.0, 142094.0, 142630.0, 143166.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -36.8388748973
actual_array: [ 18.1 -119. ]
varray: [-22.99, -20.835, -19.833, -19.634]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -20.823
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [141558.0, 142094.0, 142630.0, 143166.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.03
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 2.03675
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 51 - Sample 51 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 51
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8475
varray: [2.867, 2.852, 2.868, 2.803]
inst: mean_all
operation: mean_all(residuals)
tarray: [143751.0, 144287.0, 144823.0, 145358.0]
source_col: residuals
SAMP_residuals_STD | 0.0305559595933
varray: [2.867, 2.852, 2.868, 2.803]
inst: std_all
operation: std_all(residuals)
tarray: [143751.0, 144287.0, 144823.0, 145358.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.38225
varray: [0.631, 0.222, 0.389, 0.287]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [143751.0, 144287.0, 144823.0, 145358.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.179512998601
varray: [0.631, 0.222, 0.389, 0.287]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [143751.0, 144287.0, 144823.0, 145358.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1292.56975
varray: [1292.684, 1293.102, 1293.646, 1290.847]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [143751.0, 144287.0, 144823.0, 145358.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.2141560512
varray: [1292.684, 1293.102, 1293.646, 1290.847]
inst: std_all
operation: std_all(baseline_shift)
tarray: [143751.0, 144287.0, 144823.0, 145358.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.53125
varray: [1.793, 1.624, 1.229, 1.479]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [143751.0, 144287.0, 144823.0, 145358.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.238886828156
varray: [1.793, 1.624, 1.229, 1.479]
inst: std_all
operation: std_all(slope_shift)
tarray: [143751.0, 144287.0, 144823.0, 145358.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0386210219095
varray: [-3.23, -3.311, -3.253, -3.301]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [143751.0, 144287.0, 144823.0, 145358.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.278717300025
varray: [-23.023, -23.466, -23.377, -23.694]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [143751.0, 144287.0, 144823.0, 145358.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.534
varray: [110.109, 109.857, 110.781, 107.389]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [143751.0, 144287.0, 144823.0, 145358.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.48222670331
varray: [110.109, 109.857, 110.781, 107.389]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [143751.0, 144287.0, 144823.0, 145358.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8475
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.38225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1292.56975
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.53125
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.04645
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.04645
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.65830670641
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.37885
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-3.23, -3.311, -3.253, -3.301]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [143751.0, 144287.0, 144823.0, 145358.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-23.023, -23.466, -23.377, -23.694]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [143751.0, 144287.0, 144823.0, 145358.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.27375
varray: [-3.23, -3.311, -3.253, -3.301]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [143751.0, 144287.0, 144823.0, 145358.0]
source_col: d(18_16)Mean
dD_mean | -23.39
varray: [-23.023, -23.466, -23.377, -23.694]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [143751.0, 144287.0, 144823.0, 145358.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0386210219095
varray: [-3.23, -3.311, -3.253, -3.301]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [143751.0, 144287.0, 144823.0, 145358.0]
source_col: d(18_16)Mean
dD_stdev | 0.278717300025
varray: [-23.023, -23.466, -23.377, -23.694]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [143751.0, 144287.0, 144823.0, 145358.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.27375
varray: [-3.23, -3.311, -3.253, -3.301]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [143751.0, 144287.0, 144823.0, 145358.0]
source_col: d(18_16)Mean
measured_dD_mean | -23.39
varray: [-23.023, -23.466, -23.377, -23.694]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [143751.0, 144287.0, 144823.0, 145358.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0386210219095
varray: [-3.23, -3.311, -3.253, -3.301]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [143751.0, 144287.0, 144823.0, 145358.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.278717300025
varray: [-23.023, -23.466, -23.377, -23.694]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [143751.0, 144287.0, 144823.0, 145358.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.87525443657
actual_array: [ 1.93 -15.84]
varray: [-3.23, -3.311, -3.253, -3.301]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.27375
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [143751.0, 144287.0, 144823.0, 145358.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -39.24285161
actual_array: [ 18.1 -119. ]
varray: [-23.023, -23.466, -23.377, -23.694]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -23.39
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [143751.0, 144287.0, 144823.0, 145358.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.04725
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.675
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 52 - Sample 52 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 52
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.908
varray: [3.064, 2.863, 2.878, 2.827]
inst: mean_all
operation: mean_all(residuals)
tarray: [145944.0, 146480.0, 147016.0, 147552.0]
source_col: residuals
SAMP_residuals_STD | 0.106179093987
varray: [3.064, 2.863, 2.878, 2.827]
inst: std_all
operation: std_all(residuals)
tarray: [145944.0, 146480.0, 147016.0, 147552.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.41575
varray: [0.671, 0.457, 0.216, 0.319]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [145944.0, 146480.0, 147016.0, 147552.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.196735651743
varray: [0.671, 0.457, 0.216, 0.319]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [145944.0, 146480.0, 147016.0, 147552.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1294.8855
varray: [1298.581, 1292.253, 1295.506, 1293.202]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [145944.0, 146480.0, 147016.0, 147552.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.81697053588
varray: [1298.581, 1292.253, 1295.506, 1293.202]
inst: std_all
operation: std_all(baseline_shift)
tarray: [145944.0, 146480.0, 147016.0, 147552.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.5
varray: [1.757, 1.254, 1.493, 1.496]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [145944.0, 146480.0, 147016.0, 147552.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.205450724019
varray: [1.757, 1.254, 1.493, 1.496]
inst: std_all
operation: std_all(slope_shift)
tarray: [145944.0, 146480.0, 147016.0, 147552.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0908772065298
varray: [-4.134, -4.284, -4.226, -4.348]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [145944.0, 146480.0, 147016.0, 147552.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.95833849274
varray: [-34.299, -36.064, -35.417, -36.504]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [145944.0, 146480.0, 147016.0, 147552.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.95575
varray: [116.398, 110.274, 111.723, 109.428]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [145944.0, 146480.0, 147016.0, 147552.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.10942442423
varray: [116.398, 110.274, 111.723, 109.428]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [145944.0, 146480.0, 147016.0, 147552.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.908
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.41575
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1294.8855
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.5
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.10695
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.10695
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 3.81821102801
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 3.8006
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-4.134, -4.284, -4.226, -4.348]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [145944.0, 146480.0, 147016.0, 147552.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-34.299, -36.064, -35.417, -36.504]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [145944.0, 146480.0, 147016.0, 147552.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.248
varray: [-4.134, -4.284, -4.226, -4.348]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [145944.0, 146480.0, 147016.0, 147552.0]
source_col: d(18_16)Mean
dD_mean | -35.571
varray: [-34.299, -36.064, -35.417, -36.504]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [145944.0, 146480.0, 147016.0, 147552.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0908772065298
varray: [-4.134, -4.284, -4.226, -4.348]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [145944.0, 146480.0, 147016.0, 147552.0]
source_col: d(18_16)Mean
dD_stdev | 0.95833849274
varray: [-34.299, -36.064, -35.417, -36.504]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [145944.0, 146480.0, 147016.0, 147552.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.248
varray: [-4.134, -4.284, -4.226, -4.348]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [145944.0, 146480.0, 147016.0, 147552.0]
source_col: d(18_16)Mean
measured_dD_mean | -35.571
varray: [-34.299, -36.064, -35.417, -36.504]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [145944.0, 146480.0, 147016.0, 147552.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0908772065298
varray: [-4.134, -4.284, -4.226, -4.348]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [145944.0, 146480.0, 147016.0, 147552.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.95833849274
varray: [-34.299, -36.064, -35.417, -36.504]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [145944.0, 146480.0, 147016.0, 147552.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.84689080582
actual_array: [ 1.93 -15.84]
varray: [-4.134, -4.284, -4.226, -4.348]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -4.248
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [145944.0, 146480.0, 147016.0, 147552.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -50.6502689599
actual_array: [ 18.1 -119. ]
varray: [-34.299, -36.064, -35.417, -36.504]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -35.571
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [145944.0, 146480.0, 147016.0, 147552.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.98675
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 2.99075
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 53 - Sample 53 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 53
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.91475
varray: [2.985, 2.842, 2.958, 2.874]
inst: mean_all
operation: mean_all(residuals)
tarray: [148139.0, 148674.0, 149210.0, 149746.0]
source_col: residuals
SAMP_residuals_STD | 0.0677218576237
varray: [2.985, 2.842, 2.958, 2.874]
inst: std_all
operation: std_all(residuals)
tarray: [148139.0, 148674.0, 149210.0, 149746.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.22975
varray: [0.168, 0.3, 0.267, 0.184]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [148139.0, 148674.0, 149210.0, 149746.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0638455166789
varray: [0.168, 0.3, 0.267, 0.184]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [148139.0, 148674.0, 149210.0, 149746.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1297.4095
varray: [1300.183, 1294.881, 1297.651, 1296.923]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [148139.0, 148674.0, 149210.0, 149746.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.18941537098
varray: [1300.183, 1294.881, 1297.651, 1296.923]
inst: std_all
operation: std_all(baseline_shift)
tarray: [148139.0, 148674.0, 149210.0, 149746.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.506
varray: [1.328, 1.499, 1.662, 1.535]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [148139.0, 148674.0, 149210.0, 149746.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.137731623094
varray: [1.328, 1.499, 1.662, 1.535]
inst: std_all
operation: std_all(slope_shift)
tarray: [148139.0, 148674.0, 149210.0, 149746.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0629417720331
varray: [-1.593, -1.612, -1.499, -1.49]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [148139.0, 148674.0, 149210.0, 149746.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.3684535067
varray: [-9.214, -7.29, -6.402, -6.228]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [148139.0, 148674.0, 149210.0, 149746.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 112.69725
varray: [115.028, 110.822, 113.389, 111.55]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [148139.0, 148674.0, 149210.0, 149746.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.89241105031
varray: [115.028, 110.822, 113.389, 111.55]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [148139.0, 148674.0, 149210.0, 149746.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.91475
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.22975
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1297.4095
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.506
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.1137
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1137
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 4.05919208868
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 4.5421
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-1.593, -1.612, -1.499, -1.49]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [148139.0, 148674.0, 149210.0, 149746.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-9.214, -7.29, -6.402, -6.228]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [148139.0, 148674.0, 149210.0, 149746.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.5485
varray: [-1.593, -1.612, -1.499, -1.49]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [148139.0, 148674.0, 149210.0, 149746.0]
source_col: d(18_16)Mean
dD_mean | -7.2835
varray: [-9.214, -7.29, -6.402, -6.228]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [148139.0, 148674.0, 149210.0, 149746.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0629417720331
varray: [-1.593, -1.612, -1.499, -1.49]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [148139.0, 148674.0, 149210.0, 149746.0]
source_col: d(18_16)Mean
dD_stdev | 1.3684535067
varray: [-9.214, -7.29, -6.402, -6.228]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [148139.0, 148674.0, 149210.0, 149746.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.5485
varray: [-1.593, -1.612, -1.499, -1.49]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [148139.0, 148674.0, 149210.0, 149746.0]
source_col: d(18_16)Mean
measured_dD_mean | -7.2835
varray: [-9.214, -7.29, -6.402, -6.228]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [148139.0, 148674.0, 149210.0, 149746.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0629417720331
varray: [-1.593, -1.612, -1.499, -1.49]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [148139.0, 148674.0, 149210.0, 149746.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.3684535067
varray: [-9.214, -7.29, -6.402, -6.228]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [148139.0, 148674.0, 149210.0, 149746.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.15463278295
actual_array: [ 1.93 -15.84]
varray: [-1.593, -1.612, -1.499, -1.49]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -1.5485
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [148139.0, 148674.0, 149210.0, 149746.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -24.1592322399
actual_array: [ 18.1 -119. ]
varray: [-9.214, -7.29, -6.402, -6.228]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -7.2835
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [148139.0, 148674.0, 149210.0, 149746.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.98
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 5.51475
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 54 - Sample 54 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 54
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8695
varray: [2.997, 2.795, 2.879, 2.807]
inst: mean_all
operation: mean_all(residuals)
tarray: [150333.0, 150869.0, 151405.0, 151941.0]
source_col: residuals
SAMP_residuals_STD | 0.0927415764369
varray: [2.997, 2.795, 2.879, 2.807]
inst: std_all
operation: std_all(residuals)
tarray: [150333.0, 150869.0, 151405.0, 151941.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3515
varray: [0.55, 0.128, 0.419, 0.309]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [150333.0, 150869.0, 151405.0, 151941.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.178621573912
varray: [0.55, 0.128, 0.419, 0.309]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [150333.0, 150869.0, 151405.0, 151941.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1294.3255
varray: [1298.255, 1292.528, 1294.766, 1291.753]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [150333.0, 150869.0, 151405.0, 151941.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.91454816395
varray: [1298.255, 1292.528, 1294.766, 1291.753]
inst: std_all
operation: std_all(baseline_shift)
tarray: [150333.0, 150869.0, 151405.0, 151941.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.4
varray: [1.76, 1.086, 1.262, 1.492]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [150333.0, 150869.0, 151405.0, 151941.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.291949767369
varray: [1.76, 1.086, 1.262, 1.492]
inst: std_all
operation: std_all(slope_shift)
tarray: [150333.0, 150869.0, 151405.0, 151941.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0417372735094
varray: [0.476, 0.464, 0.557, 0.511]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [150333.0, 150869.0, 151405.0, 151941.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.975867135765
varray: [14.768, 16.122, 16.899, 16.771]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [150333.0, 150869.0, 151405.0, 151941.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.11125
varray: [114.856, 109.213, 112.027, 108.349]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [150333.0, 150869.0, 151405.0, 151941.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.9493009765
varray: [114.856, 109.213, 112.027, 108.349]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [150333.0, 150869.0, 151405.0, 151941.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8695
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3515
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1294.3255
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.4
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.06845
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.06845
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.44372645972
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 2.9561
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [0.476, 0.464, 0.557, 0.511]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [150333.0, 150869.0, 151405.0, 151941.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [14.768, 16.122, 16.899, 16.771]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [150333.0, 150869.0, 151405.0, 151941.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.502
varray: [0.476, 0.464, 0.557, 0.511]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [150333.0, 150869.0, 151405.0, 151941.0]
source_col: d(18_16)Mean
dD_mean | 16.14
varray: [14.768, 16.122, 16.899, 16.771]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [150333.0, 150869.0, 151405.0, 151941.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0417372735094
varray: [0.476, 0.464, 0.557, 0.511]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [150333.0, 150869.0, 151405.0, 151941.0]
source_col: d(18_16)Mean
dD_stdev | 0.975867135765
varray: [14.768, 16.122, 16.899, 16.771]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [150333.0, 150869.0, 151405.0, 151941.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.502
varray: [0.476, 0.464, 0.557, 0.511]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [150333.0, 150869.0, 151405.0, 151941.0]
source_col: d(18_16)Mean
measured_dD_mean | 16.14
varray: [14.768, 16.122, 16.899, 16.771]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [150333.0, 150869.0, 151405.0, 151941.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0417372735094
varray: [0.476, 0.464, 0.557, 0.511]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [150333.0, 150869.0, 151405.0, 151941.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.975867135765
varray: [14.768, 16.122, 16.899, 16.771]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [150333.0, 150869.0, 151405.0, 151941.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.10963368093
actual_array: [ 1.93 -15.84]
varray: [0.476, 0.464, 0.557, 0.511]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.502
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [150333.0, 150869.0, 151405.0, 151941.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -2.22329592144
actual_array: [ 18.1 -119. ]
varray: [14.768, 16.122, 16.899, 16.771]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 16.14
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [150333.0, 150869.0, 151405.0, 151941.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.02525
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 2.43075
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 55 - Sample 55 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 55
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.90425
varray: [2.95, 2.839, 2.847, 2.981]
inst: mean_all
operation: mean_all(residuals)
tarray: [152528.0, 153064.0, 153600.0, 154136.0]
source_col: residuals
SAMP_residuals_STD | 0.0719229912244
varray: [2.95, 2.839, 2.847, 2.981]
inst: std_all
operation: std_all(residuals)
tarray: [152528.0, 153064.0, 153600.0, 154136.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.2345
varray: [0.341, 0.131, 0.114, 0.352]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [152528.0, 153064.0, 153600.0, 154136.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.129590380301
varray: [0.341, 0.131, 0.114, 0.352]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [152528.0, 153064.0, 153600.0, 154136.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1295.9815
varray: [1298.412, 1293.421, 1295.507, 1296.586]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [152528.0, 153064.0, 153600.0, 154136.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.08598985296
varray: [1298.412, 1293.421, 1295.507, 1296.586]
inst: std_all
operation: std_all(baseline_shift)
tarray: [152528.0, 153064.0, 153600.0, 154136.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.40625
varray: [1.458, 0.861, 1.768, 1.538]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [152528.0, 153064.0, 153600.0, 154136.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.386521560416
varray: [1.458, 0.861, 1.768, 1.538]
inst: std_all
operation: std_all(slope_shift)
tarray: [152528.0, 153064.0, 153600.0, 154136.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0355844066973
varray: [1.708, 1.712, 1.669, 1.756]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [152528.0, 153064.0, 153600.0, 154136.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.346330646445
varray: [23.244, 23.419, 23.87, 23.96]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [152528.0, 153064.0, 153600.0, 154136.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.6825
varray: [114.134, 109.6, 110.051, 112.945]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [152528.0, 153064.0, 153600.0, 154136.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.20623034458
varray: [114.134, 109.6, 110.051, 112.945]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [152528.0, 153064.0, 153600.0, 154136.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.90425
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.2345
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1295.9815
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.40625
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.1032
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1032
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 3.68433266097
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 3.52735
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [1.708, 1.712, 1.669, 1.756]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [152528.0, 153064.0, 153600.0, 154136.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [23.244, 23.419, 23.87, 23.96]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [152528.0, 153064.0, 153600.0, 154136.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.71125
varray: [1.708, 1.712, 1.669, 1.756]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [152528.0, 153064.0, 153600.0, 154136.0]
source_col: d(18_16)Mean
dD_mean | 23.62325
varray: [23.244, 23.419, 23.87, 23.96]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [152528.0, 153064.0, 153600.0, 154136.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0355844066973
varray: [1.708, 1.712, 1.669, 1.756]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [152528.0, 153064.0, 153600.0, 154136.0]
source_col: d(18_16)Mean
dD_stdev | 0.346330646445
varray: [23.244, 23.419, 23.87, 23.96]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [152528.0, 153064.0, 153600.0, 154136.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.71125
varray: [1.708, 1.712, 1.669, 1.756]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [152528.0, 153064.0, 153600.0, 154136.0]
source_col: d(18_16)Mean
measured_dD_mean | 23.62325
varray: [23.244, 23.419, 23.87, 23.96]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [152528.0, 153064.0, 153600.0, 154136.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0355844066973
varray: [1.708, 1.712, 1.669, 1.756]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [152528.0, 153064.0, 153600.0, 154136.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.346330646445
varray: [23.244, 23.419, 23.87, 23.96]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [152528.0, 153064.0, 153600.0, 154136.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.903627748656
actual_array: [ 1.93 -15.84]
varray: [1.708, 1.712, 1.669, 1.756]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 1.71125
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [152528.0, 153064.0, 153600.0, 154136.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 4.78471293528
actual_array: [ 18.1 -119. ]
varray: [23.244, 23.419, 23.87, 23.96]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 23.62325
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [152528.0, 153064.0, 153600.0, 154136.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.9905
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 4.08675
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 56 - Sample 56 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 56
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.88925
varray: [2.899, 2.833, 2.875, 2.95]
inst: mean_all
operation: mean_all(residuals)
tarray: [154723.0, 155259.0, 155796.0, 156331.0]
source_col: residuals
SAMP_residuals_STD | 0.0488287824956
varray: [2.899, 2.833, 2.875, 2.95]
inst: std_all
operation: std_all(residuals)
tarray: [154723.0, 155259.0, 155796.0, 156331.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.33625
varray: [0.153, 0.422, 0.375, 0.395]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [154723.0, 155259.0, 155796.0, 156331.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.123675314163
varray: [0.153, 0.422, 0.375, 0.395]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [154723.0, 155259.0, 155796.0, 156331.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1295.50625
varray: [1297.239, 1292.055, 1295.649, 1297.082]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [154723.0, 155259.0, 155796.0, 156331.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.40948920936
varray: [1297.239, 1292.055, 1295.649, 1297.082]
inst: std_all
operation: std_all(baseline_shift)
tarray: [154723.0, 155259.0, 155796.0, 156331.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.56775
varray: [1.775, 1.714, 1.335, 1.447]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [154723.0, 155259.0, 155796.0, 156331.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.21062980954
varray: [1.775, 1.714, 1.335, 1.447]
inst: std_all
operation: std_all(slope_shift)
tarray: [154723.0, 155259.0, 155796.0, 156331.0]
source_col: slope_shift
SAMP_d18O_STD | 0.175290948616
varray: [-2.767, -3.108, -3.152, -3.076]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [154723.0, 155259.0, 155796.0, 156331.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.18981344335
varray: [-12.421, -16.323, -16.979, -16.97]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [154723.0, 155259.0, 155796.0, 156331.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.69075
varray: [111.83, 109.028, 112.574, 113.331]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [154723.0, 155259.0, 155796.0, 156331.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.87795800006
varray: [111.83, 109.028, 112.574, 113.331]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [154723.0, 155259.0, 155796.0, 156331.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.88925
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.33625
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1295.50625
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.56775
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0882
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0882
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 3.1488191928
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 3.5356
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-2.767, -3.108, -3.152, -3.076]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [154723.0, 155259.0, 155796.0, 156331.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-12.421, -16.323, -16.979, -16.97]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [154723.0, 155259.0, 155796.0, 156331.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.02575
varray: [-2.767, -3.108, -3.152, -3.076]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [154723.0, 155259.0, 155796.0, 156331.0]
source_col: d(18_16)Mean
dD_mean | -15.67325
varray: [-12.421, -16.323, -16.979, -16.97]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [154723.0, 155259.0, 155796.0, 156331.0]
source_col: d(D_H)Mean
d18O_stdev | 0.175290948616
varray: [-2.767, -3.108, -3.152, -3.076]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [154723.0, 155259.0, 155796.0, 156331.0]
source_col: d(18_16)Mean
dD_stdev | 2.18981344335
varray: [-12.421, -16.323, -16.979, -16.97]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [154723.0, 155259.0, 155796.0, 156331.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.02575
varray: [-2.767, -3.108, -3.152, -3.076]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [154723.0, 155259.0, 155796.0, 156331.0]
source_col: d(18_16)Mean
measured_dD_mean | -15.67325
varray: [-12.421, -16.323, -16.979, -16.97]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [154723.0, 155259.0, 155796.0, 156331.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.175290948616
varray: [-2.767, -3.108, -3.152, -3.076]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [154723.0, 155259.0, 155796.0, 156331.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.18981344335
varray: [-12.421, -16.323, -16.979, -16.97]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [154723.0, 155259.0, 155796.0, 156331.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.62791974879
actual_array: [ 1.93 -15.84]
varray: [-2.767, -3.108, -3.152, -3.076]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.02575
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [154723.0, 155259.0, 155796.0, 156331.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -32.0161717122
actual_array: [ 18.1 -119. ]
varray: [-12.421, -16.323, -16.979, -16.97]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -15.67325
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [154723.0, 155259.0, 155796.0, 156331.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.0055
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 3.6115
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 57 - Sample 57 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 57
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.89975
varray: [2.948, 2.855, 2.902, 2.894]
inst: mean_all
operation: mean_all(residuals)
tarray: [156917.0, 157453.0, 157989.0, 158525.0]
source_col: residuals
SAMP_residuals_STD | 0.0381608438062
varray: [2.948, 2.855, 2.902, 2.894]
inst: std_all
operation: std_all(residuals)
tarray: [156917.0, 157453.0, 157989.0, 158525.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.34275
varray: [0.348, 0.35, 0.275, 0.398]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [156917.0, 157453.0, 157989.0, 158525.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0507370673177
varray: [0.348, 0.35, 0.275, 0.398]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [156917.0, 157453.0, 157989.0, 158525.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1295.42575
varray: [1297.685, 1292.476, 1296.2, 1295.342]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [156917.0, 157453.0, 157989.0, 158525.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.19178106799
varray: [1297.685, 1292.476, 1296.2, 1295.342]
inst: std_all
operation: std_all(baseline_shift)
tarray: [156917.0, 157453.0, 157989.0, 158525.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.3545
varray: [0.926, 1.952, 1.305, 1.235]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [156917.0, 157453.0, 157989.0, 158525.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.431025521286
varray: [0.926, 1.952, 1.305, 1.235]
inst: std_all
operation: std_all(slope_shift)
tarray: [156917.0, 157453.0, 157989.0, 158525.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0662841107154
varray: [-0.101, -0.045, 0.0, 0.055]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [156917.0, 157453.0, 157989.0, 158525.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.35788094716
varray: [7.477, 9.222, 10.052, 10.581]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [156917.0, 157453.0, 157989.0, 158525.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 112.1355
varray: [114.552, 109.066, 112.346, 112.578]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [156917.0, 157453.0, 157989.0, 158525.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.27313432658
varray: [114.552, 109.066, 112.346, 112.578]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [156917.0, 157453.0, 157989.0, 158525.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.89975
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.34275
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1295.42575
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.3545
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0987
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0987
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 3.52367862052
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 3.98035
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-0.101, -0.045, 0.0, 0.055]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [156917.0, 157453.0, 157989.0, 158525.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [7.477, 9.222, 10.052, 10.581]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [156917.0, 157453.0, 157989.0, 158525.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.02275
varray: [-0.101, -0.045, 0.0, 0.055]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [156917.0, 157453.0, 157989.0, 158525.0]
source_col: d(18_16)Mean
dD_mean | 9.333
varray: [7.477, 9.222, 10.052, 10.581]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [156917.0, 157453.0, 157989.0, 158525.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0662841107154
varray: [-0.101, -0.045, 0.0, 0.055]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [156917.0, 157453.0, 157989.0, 158525.0]
source_col: d(18_16)Mean
dD_stdev | 1.35788094716
varray: [7.477, 9.222, 10.052, 10.581]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [156917.0, 157453.0, 157989.0, 158525.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.02275
varray: [-0.101, -0.045, 0.0, 0.055]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [156917.0, 157453.0, 157989.0, 158525.0]
source_col: d(18_16)Mean
measured_dD_mean | 9.333
varray: [7.477, 9.222, 10.052, 10.581]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [156917.0, 157453.0, 157989.0, 158525.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0662841107154
varray: [-0.101, -0.045, 0.0, 0.055]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [156917.0, 157453.0, 157989.0, 158525.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.35788094716
varray: [7.477, 9.222, 10.052, 10.581]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [156917.0, 157453.0, 157989.0, 158525.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.63297592857
actual_array: [ 1.93 -15.84]
varray: [-0.101, -0.045, 0.0, 0.055]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -0.02275
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [156917.0, 157453.0, 157989.0, 158525.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -8.59800160249
actual_array: [ 18.1 -119. ]
varray: [7.477, 9.222, 10.052, 10.581]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 9.333
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [156917.0, 157453.0, 157989.0, 158525.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.995
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 3.531
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 58 - Sample 58 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 58
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.87275
varray: [2.92, 2.788, 2.87, 2.913]
inst: mean_all
operation: mean_all(residuals)
tarray: [159111.0, 159647.0, 160183.0, 160718.0]
source_col: residuals
SAMP_residuals_STD | 0.0606705584832
varray: [2.92, 2.788, 2.87, 2.913]
inst: std_all
operation: std_all(residuals)
tarray: [159111.0, 159647.0, 160183.0, 160718.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.28775
varray: [0.408, 0.106, 0.37, 0.267]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [159111.0, 159647.0, 160183.0, 160718.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.135016974242
varray: [0.408, 0.106, 0.37, 0.267]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [159111.0, 159647.0, 160183.0, 160718.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1295.65775
varray: [1297.843, 1293.242, 1295.208, 1296.338]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [159111.0, 159647.0, 160183.0, 160718.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.93874364388
varray: [1297.843, 1293.242, 1295.208, 1296.338]
inst: std_all
operation: std_all(baseline_shift)
tarray: [159111.0, 159647.0, 160183.0, 160718.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.99425
varray: [2.122, 1.972, 2.173, 1.71]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [159111.0, 159647.0, 160183.0, 160718.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.207817828879
varray: [2.122, 1.972, 2.173, 1.71]
inst: std_all
operation: std_all(slope_shift)
tarray: [159111.0, 159647.0, 160183.0, 160718.0]
source_col: slope_shift
SAMP_d18O_STD | 0.233923064275
varray: [-7.052, -7.476, -7.548, -7.524]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [159111.0, 159647.0, 160183.0, 160718.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.72829446038
varray: [-50.308, -56.79, -57.936, -58.249]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [159111.0, 159647.0, 160183.0, 160718.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.14375
varray: [112.848, 109.096, 110.258, 112.373]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [159111.0, 159647.0, 160183.0, 160718.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.76949679759
varray: [112.848, 109.096, 110.258, 112.373]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [159111.0, 159647.0, 160183.0, 160718.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.87275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.28775
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1295.65775
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.99425
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0717
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0717
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.55975437782
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 2.9886
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-7.052, -7.476, -7.548, -7.524]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [159111.0, 159647.0, 160183.0, 160718.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-50.308, -56.79, -57.936, -58.249]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [159111.0, 159647.0, 160183.0, 160718.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.4
varray: [-7.052, -7.476, -7.548, -7.524]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [159111.0, 159647.0, 160183.0, 160718.0]
source_col: d(18_16)Mean
dD_mean | -55.82075
varray: [-50.308, -56.79, -57.936, -58.249]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [159111.0, 159647.0, 160183.0, 160718.0]
source_col: d(D_H)Mean
d18O_stdev | 0.233923064275
varray: [-7.052, -7.476, -7.548, -7.524]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [159111.0, 159647.0, 160183.0, 160718.0]
source_col: d(18_16)Mean
dD_stdev | 3.72829446038
varray: [-50.308, -56.79, -57.936, -58.249]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [159111.0, 159647.0, 160183.0, 160718.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.4
varray: [-7.052, -7.476, -7.548, -7.524]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [159111.0, 159647.0, 160183.0, 160718.0]
source_col: d(18_16)Mean
measured_dD_mean | -55.82075
varray: [-50.308, -56.79, -57.936, -58.249]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [159111.0, 159647.0, 160183.0, 160718.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.233923064275
varray: [-7.052, -7.476, -7.548, -7.524]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [159111.0, 159647.0, 160183.0, 160718.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.72829446038
varray: [-50.308, -56.79, -57.936, -58.249]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [159111.0, 159647.0, 160183.0, 160718.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.99043490218
actual_array: [ 1.93 -15.84]
varray: [-7.052, -7.476, -7.548, -7.524]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.4
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [159111.0, 159647.0, 160183.0, 160718.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -69.6140116314
actual_array: [ 18.1 -119. ]
varray: [-50.308, -56.79, -57.936, -58.249]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -55.82075
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [159111.0, 159647.0, 160183.0, 160718.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.022
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 3.763
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 59 - Sample 59 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 59
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.89975
varray: [2.955, 2.861, 2.867, 2.916]
inst: mean_all
operation: mean_all(residuals)
tarray: [161305.0, 161841.0, 162377.0, 162913.0]
source_col: residuals
SAMP_residuals_STD | 0.0443123383871
varray: [2.955, 2.861, 2.867, 2.916]
inst: std_all
operation: std_all(residuals)
tarray: [161305.0, 161841.0, 162377.0, 162913.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.10675
varray: [-0.107, 0.258, 0.198, 0.078]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [161305.0, 161841.0, 162377.0, 162913.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.160954186028
varray: [-0.107, 0.258, 0.198, 0.078]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [161305.0, 161841.0, 162377.0, 162913.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1296.374
varray: [1300.231, 1292.902, 1294.953, 1297.41]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [161305.0, 161841.0, 162377.0, 162913.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.16353020111
varray: [1300.231, 1292.902, 1294.953, 1297.41]
inst: std_all
operation: std_all(baseline_shift)
tarray: [161305.0, 161841.0, 162377.0, 162913.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.185
varray: [0.963, 1.572, 1.217, 0.988]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [161305.0, 161841.0, 162377.0, 162913.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.282185518173
varray: [0.963, 1.572, 1.217, 0.988]
inst: std_all
operation: std_all(slope_shift)
tarray: [161305.0, 161841.0, 162377.0, 162913.0]
source_col: slope_shift
SAMP_d18O_STD | 0.162835295109
varray: [-0.902, -0.775, -0.689, -0.514]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [161305.0, 161841.0, 162377.0, 162913.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.41794959625
varray: [-5.226, -1.791, -0.726, 0.358]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [161305.0, 161841.0, 162377.0, 162913.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.7845
varray: [114.412, 108.866, 111.115, 112.745]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [161305.0, 161841.0, 162377.0, 162913.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.36588172429
varray: [114.412, 108.866, 111.115, 112.745]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [161305.0, 161841.0, 162377.0, 162913.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.89975
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.10675
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1296.374
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.185
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0987
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0987
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 3.52367862052
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 3.62935
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-0.902, -0.775, -0.689, -0.514]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [161305.0, 161841.0, 162377.0, 162913.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-5.226, -1.791, -0.726, 0.358]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [161305.0, 161841.0, 162377.0, 162913.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.72
varray: [-0.902, -0.775, -0.689, -0.514]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [161305.0, 161841.0, 162377.0, 162913.0]
source_col: d(18_16)Mean
dD_mean | -1.84625
varray: [-5.226, -1.791, -0.726, 0.358]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [161305.0, 161841.0, 162377.0, 162913.0]
source_col: d(D_H)Mean
d18O_stdev | 0.162835295109
varray: [-0.902, -0.775, -0.689, -0.514]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [161305.0, 161841.0, 162377.0, 162913.0]
source_col: d(18_16)Mean
dD_stdev | 2.41794959625
varray: [-5.226, -1.791, -0.726, 0.358]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [161305.0, 161841.0, 162377.0, 162913.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.72
varray: [-0.902, -0.775, -0.689, -0.514]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [161305.0, 161841.0, 162377.0, 162913.0]
source_col: d(18_16)Mean
measured_dD_mean | -1.84625
varray: [-5.226, -1.791, -0.726, 0.358]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [161305.0, 161841.0, 162377.0, 162913.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.162835295109
varray: [-0.902, -0.775, -0.689, -0.514]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [161305.0, 161841.0, 162377.0, 162913.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.41794959625
varray: [-5.226, -1.791, -0.726, 0.358]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [161305.0, 161841.0, 162377.0, 162913.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.32835540864
actual_array: [ 1.93 -15.84]
varray: [-0.902, -0.775, -0.689, -0.514]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -0.72
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [161305.0, 161841.0, 162377.0, 162913.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -19.067287409
actual_array: [ 18.1 -119. ]
varray: [-5.226, -1.791, -0.726, 0.358]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -1.84625
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [161305.0, 161841.0, 162377.0, 162913.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.995
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 4.47925
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 60 - Sample 60 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 60
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.797
varray: [2.747, 2.758, 2.873, 2.81]
inst: mean_all
operation: mean_all(residuals)
tarray: [163500.0, 164035.0, 164571.0, 165107.0]
source_col: residuals
SAMP_residuals_STD | 0.057636793804
varray: [2.747, 2.758, 2.873, 2.81]
inst: std_all
operation: std_all(residuals)
tarray: [163500.0, 164035.0, 164571.0, 165107.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.807
varray: [-0.866, -0.82, -0.686, -0.856]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [163500.0, 164035.0, 164571.0, 165107.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0830501856309
varray: [-0.866, -0.82, -0.686, -0.856]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [163500.0, 164035.0, 164571.0, 165107.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 2099.81125
varray: [2078.645, 2087.595, 2115.845, 2117.16]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [163500.0, 164035.0, 164571.0, 165107.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 19.6240258251
varray: [2078.645, 2087.595, 2115.845, 2117.16]
inst: std_all
operation: std_all(baseline_shift)
tarray: [163500.0, 164035.0, 164571.0, 165107.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -28.5245
varray: [-27.628, -28.461, -29.259, -28.75]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [163500.0, 164035.0, 164571.0, 165107.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.682662679025
varray: [-27.628, -28.461, -29.259, -28.75]
inst: std_all
operation: std_all(slope_shift)
tarray: [163500.0, 164035.0, 164571.0, 165107.0]
source_col: slope_shift
SAMP_d18O_STD | 0.102457714855
varray: [0.791, 0.875, 0.945, 1.032]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [163500.0, 164035.0, 164571.0, 165107.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.222754723107
varray: [-2.923, -2.977, -2.709, -2.489]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [163500.0, 164035.0, 164571.0, 165107.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 102.3195
varray: [100.566, 100.423, 104.907, 103.382]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [163500.0, 164035.0, 164571.0, 165107.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.19814596118
varray: [100.566, 100.423, 104.907, 103.382]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [163500.0, 164035.0, 164571.0, 165107.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.797
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.807
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 2099.81125
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.5245
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.00405
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.00405
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.144588636404
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.83565
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [0.791, 0.875, 0.945, 1.032]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [163500.0, 164035.0, 164571.0, 165107.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-2.923, -2.977, -2.709, -2.489]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [163500.0, 164035.0, 164571.0, 165107.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.91075
varray: [0.791, 0.875, 0.945, 1.032]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [163500.0, 164035.0, 164571.0, 165107.0]
source_col: d(18_16)Mean
dD_mean | -2.7745
varray: [-2.923, -2.977, -2.709, -2.489]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [163500.0, 164035.0, 164571.0, 165107.0]
source_col: d(D_H)Mean
d18O_stdev | 0.102457714855
varray: [0.791, 0.875, 0.945, 1.032]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [163500.0, 164035.0, 164571.0, 165107.0]
source_col: d(18_16)Mean
dD_stdev | 0.222754723107
varray: [-2.923, -2.977, -2.709, -2.489]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [163500.0, 164035.0, 164571.0, 165107.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.91075
varray: [0.791, 0.875, 0.945, 1.032]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [163500.0, 164035.0, 164571.0, 165107.0]
source_col: d(18_16)Mean
measured_dD_mean | -2.7745
varray: [-2.923, -2.977, -2.709, -2.489]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [163500.0, 164035.0, 164571.0, 165107.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.102457714855
varray: [0.791, 0.875, 0.945, 1.032]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [163500.0, 164035.0, 164571.0, 165107.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.222754723107
varray: [-2.923, -2.977, -2.709, -2.489]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [163500.0, 164035.0, 164571.0, 165107.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.70198023886
actual_array: [ 1.93 -15.84]
varray: [0.791, 0.875, 0.945, 1.032]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.91075
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [163500.0, 164035.0, 164571.0, 165107.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -19.9365867403
actual_array: [ 18.1 -119. ]
varray: [-2.923, -2.977, -2.709, -2.489]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -2.7745
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [163500.0, 164035.0, 164571.0, 165107.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.09775
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 807.9165
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 61 - Sample 61 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 61
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.85375
varray: [2.742, 2.893, 2.839, 2.941]
inst: mean_all
operation: mean_all(residuals)
tarray: [165695.0, 166231.0, 166766.0, 167302.0]
source_col: residuals
SAMP_residuals_STD | 0.0853595337382
varray: [2.742, 2.893, 2.839, 2.941]
inst: std_all
operation: std_all(residuals)
tarray: [165695.0, 166231.0, 166766.0, 167302.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.33975
varray: [0.262, 0.474, 0.26, 0.363]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [165695.0, 166231.0, 166766.0, 167302.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.101601755234
varray: [0.262, 0.474, 0.26, 0.363]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [165695.0, 166231.0, 166766.0, 167302.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1299.96075
varray: [1303.878, 1298.745, 1297.097, 1300.123]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [165695.0, 166231.0, 166766.0, 167302.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.88965307895
varray: [1303.878, 1298.745, 1297.097, 1300.123]
inst: std_all
operation: std_all(baseline_shift)
tarray: [165695.0, 166231.0, 166766.0, 167302.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.58425
varray: [1.542, 1.523, 1.284, 1.988]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [165695.0, 166231.0, 166766.0, 167302.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.293655552192
varray: [1.542, 1.523, 1.284, 1.988]
inst: std_all
operation: std_all(slope_shift)
tarray: [165695.0, 166231.0, 166766.0, 167302.0]
source_col: slope_shift
SAMP_d18O_STD | 0.149709440807
varray: [-6.158, -6.399, -6.505, -6.427]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [165695.0, 166231.0, 166766.0, 167302.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.81870439233
varray: [-37.606, -40.677, -41.426, -41.421]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [165695.0, 166231.0, 166766.0, 167302.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.21425
varray: [107.536, 112.127, 110.876, 114.318]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [165695.0, 166231.0, 166766.0, 167302.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.83491991362
varray: [107.536, 112.127, 110.876, 114.318]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [165695.0, 166231.0, 166766.0, 167302.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.85375
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.33975
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1299.96075
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.58425
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0527
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0527
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.88143731815
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 3.0591
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-6.158, -6.399, -6.505, -6.427]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [165695.0, 166231.0, 166766.0, 167302.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-37.606, -40.677, -41.426, -41.421]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [165695.0, 166231.0, 166766.0, 167302.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.37225
varray: [-6.158, -6.399, -6.505, -6.427]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [165695.0, 166231.0, 166766.0, 167302.0]
source_col: d(18_16)Mean
dD_mean | -40.2825
varray: [-37.606, -40.677, -41.426, -41.421]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [165695.0, 166231.0, 166766.0, 167302.0]
source_col: d(D_H)Mean
d18O_stdev | 0.149709440807
varray: [-6.158, -6.399, -6.505, -6.427]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [165695.0, 166231.0, 166766.0, 167302.0]
source_col: d(18_16)Mean
dD_stdev | 1.81870439233
varray: [-37.606, -40.677, -41.426, -41.421]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [165695.0, 166231.0, 166766.0, 167302.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.37225
varray: [-6.158, -6.399, -6.505, -6.427]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [165695.0, 166231.0, 166766.0, 167302.0]
source_col: d(18_16)Mean
measured_dD_mean | -40.2825
varray: [-37.606, -40.677, -41.426, -41.421]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [165695.0, 166231.0, 166766.0, 167302.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.149709440807
varray: [-6.158, -6.399, -6.505, -6.427]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [165695.0, 166231.0, 166766.0, 167302.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.81870439233
varray: [-37.606, -40.677, -41.426, -41.421]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [165695.0, 166231.0, 166766.0, 167302.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.96544205794
actual_array: [ 1.93 -15.84]
varray: [-6.158, -6.399, -6.505, -6.427]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -6.37225
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [165695.0, 166231.0, 166766.0, 167302.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -55.0625542274
actual_array: [ 18.1 -119. ]
varray: [-37.606, -40.677, -41.426, -41.421]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -40.2825
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [165695.0, 166231.0, 166766.0, 167302.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.041
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 8.066
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 62 - Sample 62 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 62
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.867
varray: [2.785, 2.867, 2.914, 2.902]
inst: mean_all
operation: mean_all(residuals)
tarray: [167890.0, 168426.0, 168962.0, 169497.0]
source_col: residuals
SAMP_residuals_STD | 0.0581893461039
varray: [2.785, 2.867, 2.914, 2.902]
inst: std_all
operation: std_all(residuals)
tarray: [167890.0, 168426.0, 168962.0, 169497.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.2235
varray: [0.308, 0.132, 0.4, 0.054]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [167890.0, 168426.0, 168962.0, 169497.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.15852970279
varray: [0.308, 0.132, 0.4, 0.054]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [167890.0, 168426.0, 168962.0, 169497.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1297.705
varray: [1293.456, 1298.044, 1298.807, 1300.513]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [167890.0, 168426.0, 168962.0, 169497.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.01486262816
varray: [1293.456, 1298.044, 1298.807, 1300.513]
inst: std_all
operation: std_all(baseline_shift)
tarray: [167890.0, 168426.0, 168962.0, 169497.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.98075
varray: [1.537, 2.408, 2.434, 1.544]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [167890.0, 168426.0, 168962.0, 169497.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.508475745079
varray: [1.537, 2.408, 2.434, 1.544]
inst: std_all
operation: std_all(slope_shift)
tarray: [167890.0, 168426.0, 168962.0, 169497.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0474552420708
varray: [-8.18, -8.228, -8.166, -8.27]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [167890.0, 168426.0, 168962.0, 169497.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.915715157677
varray: [-58.646, -60.001, -59.998, -60.862]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [167890.0, 168426.0, 168962.0, 169497.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.5245
varray: [108.859, 111.308, 112.698, 113.233]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [167890.0, 168426.0, 168962.0, 169497.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.95344456452
varray: [108.859, 111.308, 112.698, 113.233]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [167890.0, 168426.0, 168962.0, 169497.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.867
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.2235
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1297.705
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.98075
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.06595
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.06595
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.35447421503
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 3.36935
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-8.18, -8.228, -8.166, -8.27]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [167890.0, 168426.0, 168962.0, 169497.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-58.646, -60.001, -59.998, -60.862]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [167890.0, 168426.0, 168962.0, 169497.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.211
varray: [-8.18, -8.228, -8.166, -8.27]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [167890.0, 168426.0, 168962.0, 169497.0]
source_col: d(18_16)Mean
dD_mean | -59.87675
varray: [-58.646, -60.001, -59.998, -60.862]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [167890.0, 168426.0, 168962.0, 169497.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0474552420708
varray: [-8.18, -8.228, -8.166, -8.27]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [167890.0, 168426.0, 168962.0, 169497.0]
source_col: d(18_16)Mean
dD_stdev | 0.915715157677
varray: [-58.646, -60.001, -59.998, -60.862]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [167890.0, 168426.0, 168962.0, 169497.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.211
varray: [-8.18, -8.228, -8.166, -8.27]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [167890.0, 168426.0, 168962.0, 169497.0]
source_col: d(18_16)Mean
measured_dD_mean | -59.87675
varray: [-58.646, -60.001, -59.998, -60.862]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [167890.0, 168426.0, 168962.0, 169497.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0474552420708
varray: [-8.18, -8.228, -8.166, -8.27]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [167890.0, 168426.0, 168962.0, 169497.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.915715157677
varray: [-58.646, -60.001, -59.998, -60.862]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [167890.0, 168426.0, 168962.0, 169497.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.7992592239
actual_array: [ 1.93 -15.84]
varray: [-8.18, -8.228, -8.166, -8.27]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -8.211
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [167890.0, 168426.0, 168962.0, 169497.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -73.4124259464
actual_array: [ 18.1 -119. ]
varray: [-58.646, -60.001, -59.998, -60.862]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -59.87675
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [167890.0, 168426.0, 168962.0, 169497.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.02775
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 5.81025
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 63 - Sample 63 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 63
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.962
varray: [2.879, 3.06, 2.936, 2.973]
inst: mean_all
operation: mean_all(residuals)
tarray: [170085.0, 170621.0, 171157.0, 171693.0]
source_col: residuals
SAMP_residuals_STD | 0.0759166209294
varray: [2.879, 3.06, 2.936, 2.973]
inst: std_all
operation: std_all(residuals)
tarray: [170085.0, 170621.0, 171157.0, 171693.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.42875
varray: [0.409, 0.391, 0.333, 0.582]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [170085.0, 170621.0, 171157.0, 171693.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.107189473986
varray: [0.409, 0.391, 0.333, 0.582]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [170085.0, 170621.0, 171157.0, 171693.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1297.7975
varray: [1294.448, 1301.469, 1297.648, 1297.625]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [170085.0, 170621.0, 171157.0, 171693.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.87234915473
varray: [1294.448, 1301.469, 1297.648, 1297.625]
inst: std_all
operation: std_all(baseline_shift)
tarray: [170085.0, 170621.0, 171157.0, 171693.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.4845
varray: [1.504, 1.76, 1.5, 1.174]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [170085.0, 170621.0, 171157.0, 171693.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.240090954987
varray: [1.504, 1.76, 1.5, 1.174]
inst: std_all
operation: std_all(slope_shift)
tarray: [170085.0, 170621.0, 171157.0, 171693.0]
source_col: slope_shift
SAMP_d18O_STD | 0.339830717662
varray: [-4.752, -3.922, -4.385, -4.378]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [170085.0, 170621.0, 171157.0, 171693.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.49860447717
varray: [-27.864, -22.421, -23.321, -23.028]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [170085.0, 170621.0, 171157.0, 171693.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.6125
varray: [110.469, 116.809, 112.52, 114.652]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [170085.0, 170621.0, 171157.0, 171693.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.73089246706
varray: [110.469, 116.809, 112.52, 114.652]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [170085.0, 170621.0, 171157.0, 171693.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.962
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.42875
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1297.7975
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.4845
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.16095
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.16095
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 5.7460595134
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 5.45735
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-4.752, -3.922, -4.385, -4.378]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [170085.0, 170621.0, 171157.0, 171693.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-27.864, -22.421, -23.321, -23.028]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [170085.0, 170621.0, 171157.0, 171693.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.35925
varray: [-4.752, -3.922, -4.385, -4.378]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [170085.0, 170621.0, 171157.0, 171693.0]
source_col: d(18_16)Mean
dD_mean | -24.1585
varray: [-27.864, -22.421, -23.321, -23.028]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [170085.0, 170621.0, 171157.0, 171693.0]
source_col: d(D_H)Mean
d18O_stdev | 0.339830717662
varray: [-4.752, -3.922, -4.385, -4.378]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [170085.0, 170621.0, 171157.0, 171693.0]
source_col: d(18_16)Mean
dD_stdev | 2.49860447717
varray: [-27.864, -22.421, -23.321, -23.028]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [170085.0, 170621.0, 171157.0, 171693.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.35925
varray: [-4.752, -3.922, -4.385, -4.378]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [170085.0, 170621.0, 171157.0, 171693.0]
source_col: d(18_16)Mean
measured_dD_mean | -24.1585
varray: [-27.864, -22.421, -23.321, -23.028]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [170085.0, 170621.0, 171157.0, 171693.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.339830717662
varray: [-4.752, -3.922, -4.385, -4.378]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [170085.0, 170621.0, 171157.0, 171693.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.49860447717
varray: [-27.864, -22.421, -23.321, -23.028]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [170085.0, 170621.0, 171157.0, 171693.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.95784235427
actual_array: [ 1.93 -15.84]
varray: [-4.752, -3.922, -4.385, -4.378]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -4.35925
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [170085.0, 170621.0, 171157.0, 171693.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -39.9625462355
actual_array: [ 18.1 -119. ]
varray: [-27.864, -22.421, -23.321, -23.028]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -24.1585
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [170085.0, 170621.0, 171157.0, 171693.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.93275
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 5.90275
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 64 - Sample 64 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 64
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.023
varray: [2.987, 3.202, 2.916, 2.987]
inst: mean_all
operation: mean_all(residuals)
tarray: [172279.0, 172814.0, 173350.0, 173886.0]
source_col: residuals
SAMP_residuals_STD | 0.123938156621
varray: [2.987, 3.202, 2.916, 2.987]
inst: std_all
operation: std_all(residuals)
tarray: [172279.0, 172814.0, 173350.0, 173886.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.22025
varray: [0.119, 0.368, 0.275, 0.119]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [172279.0, 172814.0, 173350.0, 173886.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.122923756858
varray: [0.119, 0.368, 0.275, 0.119]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [172279.0, 172814.0, 173350.0, 173886.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1301.17725
varray: [1298.982, 1309.888, 1296.857, 1298.982]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [172279.0, 172814.0, 173350.0, 173886.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 5.89293279418
varray: [1298.982, 1309.888, 1296.857, 1298.982]
inst: std_all
operation: std_all(baseline_shift)
tarray: [172279.0, 172814.0, 173350.0, 173886.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.244
varray: [0.99, 1.745, 1.251, 0.99]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [172279.0, 172814.0, 173350.0, 173886.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.355941006348
varray: [0.99, 1.745, 1.251, 0.99]
inst: std_all
operation: std_all(slope_shift)
tarray: [172279.0, 172814.0, 173350.0, 173886.0]
source_col: slope_shift
SAMP_d18O_STD | 0.124601698758
varray: [1.3, 1.527, 1.248, 1.3]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [172279.0, 172814.0, 173350.0, 173886.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.380062604492
varray: [25.258, 25.23, 24.489, 25.258]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [172279.0, 172814.0, 173350.0, 173886.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 115.64475
varray: [114.277, 121.983, 112.042, 114.277]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [172279.0, 172814.0, 173350.0, 173886.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 4.35487086491
varray: [114.277, 121.983, 112.042, 114.277]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [172279.0, 172814.0, 173350.0, 173886.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.023
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.22025
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1301.17725
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.244
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.22195
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.22195
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 7.92381428393
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 7.4896
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [1.3, 1.527, 1.248, 1.3]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [172279.0, 172814.0, 173350.0, 173886.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [25.258, 25.23, 24.489, 25.258]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [172279.0, 172814.0, 173350.0, 173886.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.34375
varray: [1.3, 1.527, 1.248, 1.3]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [172279.0, 172814.0, 173350.0, 173886.0]
source_col: d(18_16)Mean
dD_mean | 25.05875
varray: [25.258, 25.23, 24.489, 25.258]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [172279.0, 172814.0, 173350.0, 173886.0]
source_col: d(D_H)Mean
d18O_stdev | 0.124601698758
varray: [1.3, 1.527, 1.248, 1.3]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [172279.0, 172814.0, 173350.0, 173886.0]
source_col: d(18_16)Mean
dD_stdev | 0.380062604492
varray: [25.258, 25.23, 24.489, 25.258]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [172279.0, 172814.0, 173350.0, 173886.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.34375
varray: [1.3, 1.527, 1.248, 1.3]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [172279.0, 172814.0, 173350.0, 173886.0]
source_col: d(18_16)Mean
measured_dD_mean | 25.05875
varray: [25.258, 25.23, 24.489, 25.258]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [172279.0, 172814.0, 173350.0, 173886.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.124601698758
varray: [1.3, 1.527, 1.248, 1.3]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [172279.0, 172814.0, 173350.0, 173886.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.380062604492
varray: [25.258, 25.23, 24.489, 25.258]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [172279.0, 172814.0, 173350.0, 173886.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.27014185253
actual_array: [ 1.93 -15.84]
varray: [1.3, 1.527, 1.248, 1.3]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 1.34375
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [172279.0, 172814.0, 173350.0, 173886.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 6.12904817917
actual_array: [ 18.1 -119. ]
varray: [25.258, 25.23, 24.489, 25.258]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 25.05875
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [172279.0, 172814.0, 173350.0, 173886.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1288.87175
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 9.2825
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 65 - Sample 65 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Sample Variables
Sample 65
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.871
varray: [2.886, 2.857, 2.874, 2.867]
inst: mean_all
operation: mean_all(residuals)
tarray: [174473.0, 175008.0, 175544.0, 176079.0]
source_col: residuals
SAMP_residuals_STD | 0.0121928941054
varray: [2.886, 2.857, 2.874, 2.867]
inst: std_all
operation: std_all(residuals)
tarray: [174473.0, 175008.0, 175544.0, 176079.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.38275
varray: [0.202, 0.545, 0.466, 0.318]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [174473.0, 175008.0, 175544.0, 176079.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.152882035134
varray: [0.202, 0.545, 0.466, 0.318]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [174473.0, 175008.0, 175544.0, 176079.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1295.395
varray: [1297.92, 1294.542, 1293.968, 1295.15]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [174473.0, 175008.0, 175544.0, 176079.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.75115085967
varray: [1297.92, 1294.542, 1293.968, 1295.15]
inst: std_all
operation: std_all(baseline_shift)
tarray: [174473.0, 175008.0, 175544.0, 176079.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.7975
varray: [1.928, 1.856, 1.84, 1.566]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [174473.0, 175008.0, 175544.0, 176079.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.159008385523
varray: [1.928, 1.856, 1.84, 1.566]
inst: std_all
operation: std_all(slope_shift)
tarray: [174473.0, 175008.0, 175544.0, 176079.0]
source_col: slope_shift
SAMP_d18O_STD | 0.251810113644
varray: [-6.972, -7.373, -7.481, -7.524]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [174473.0, 175008.0, 175544.0, 176079.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 4.32805475936
varray: [-48.944, -56.126, -57.735, -58.331]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [174473.0, 175008.0, 175544.0, 176079.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.38625
varray: [112.343, 111.346, 110.45, 111.406]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [174473.0, 175008.0, 175544.0, 176079.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.773292689306
varray: [112.343, 111.346, 110.45, 111.406]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [174473.0, 175008.0, 175544.0, 176079.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.871
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.38275
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1295.395
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.7975
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.06995
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.06995
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.49727780654
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 3.2311
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2653
varray: [-6.972, -7.373, -7.481, -7.524]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [174473.0, 175008.0, 175544.0, 176079.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.5558
varray: [-48.944, -56.126, -57.735, -58.331]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [174473.0, 175008.0, 175544.0, 176079.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.3375
varray: [-6.972, -7.373, -7.481, -7.524]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [174473.0, 175008.0, 175544.0, 176079.0]
source_col: d(18_16)Mean
dD_mean | -55.284
varray: [-48.944, -56.126, -57.735, -58.331]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [174473.0, 175008.0, 175544.0, 176079.0]
source_col: d(D_H)Mean
d18O_stdev | 0.251810113644
varray: [-6.972, -7.373, -7.481, -7.524]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [174473.0, 175008.0, 175544.0, 176079.0]
source_col: d(18_16)Mean
dD_stdev | 4.32805475936
varray: [-48.944, -56.126, -57.735, -58.331]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [174473.0, 175008.0, 175544.0, 176079.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.3375
varray: [-6.972, -7.373, -7.481, -7.524]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [174473.0, 175008.0, 175544.0, 176079.0]
source_col: d(18_16)Mean
measured_dD_mean | -55.284
varray: [-48.944, -56.126, -57.735, -58.331]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [174473.0, 175008.0, 175544.0, 176079.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.251810113644
varray: [-6.972, -7.373, -7.481, -7.524]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [174473.0, 175008.0, 175544.0, 176079.0]
source_col: d(18_16)Mean
measured_dD_stdev | 4.32805475936
varray: [-48.944, -56.126, -57.735, -58.331]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [174473.0, 175008.0, 175544.0, 176079.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.92810257159
actual_array: [ 1.93 -15.84]
varray: [-6.972, -7.373, -7.481, -7.524]
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.3375
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: [174473.0, 175008.0, 175544.0, 176079.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -69.1113491847
actual_array: [ 18.1 -119. ]
varray: [-48.944, -56.126, -57.735, -58.331]
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -55.284
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: [174473.0, 175008.0, 175544.0, 176079.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1289.02375
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 3.50025
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 1 - Sample 1 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 1 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.089
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.076
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.6775865146
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -8.089
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.2814855192
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -65.076
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 1 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.945
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.796
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5339728249
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.945
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.0827748872
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -63.796
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 1 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.774
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.106
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3634315684
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.774
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.500102256
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -62.106
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 1 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.704
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.064
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2936193582
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.704
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.5242768822
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -61.064
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 1 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.622
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.397
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2118393404
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.622
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.8996362638
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -60.397
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 1 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.559
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.071
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1490083512
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.559
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.5943396497
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -60.071
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 1 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.539
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.778
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1290620054
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.539
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.3199472941
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -59.778
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 1 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.616
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.828
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2058554367
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.616
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.3667719282
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -59.828
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 1 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.553
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.984
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1430244475
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.553
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.5128647864
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -59.984
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 1 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.566
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.99
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1559895722
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.566
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.5184837425
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -59.99
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 2 - Sample 2 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 2 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.69
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.328
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.5525
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 37.8415
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.875
operation: col(d(18_16)Mean)
inst: col
dD_actual | 27.309
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.2543175364
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 3.875
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 8.23639083508
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 27.309
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 2 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.69
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.328
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.5525
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 37.8415
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.425
operation: col(d(18_16)Mean)
inst: col
dD_actual | 35.67
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.80284204559
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 4.425
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 16.0664061427
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 35.67
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 2 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.69
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.328
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.5525
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 37.8415
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.604
operation: col(d(18_16)Mean)
inst: col
dD_actual | 37.859
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.98136184041
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 4.604
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 18.1163886219
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 37.859
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 2 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.69
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.328
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.5525
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 37.8415
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.634
operation: col(d(18_16)Mean)
inst: col
dD_actual | 38.931
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.01128135909
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 4.634
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.1203087762
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 38.931
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 2 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.69
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.328
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.5525
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 37.8415
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.566
operation: col(d(18_16)Mean)
inst: col
dD_actual | 39.076
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.94346378341
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 4.566
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.256100215
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 39.076
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 2 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.69
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.328
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.5525
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 37.8415
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.657
operation: col(d(18_16)Mean)
inst: col
dD_actual | 39.468
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.03421965675
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 4.657
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.623205346
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 39.468
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 2 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.69
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.328
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.5525
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 37.8415
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.647
operation: col(d(18_16)Mean)
inst: col
dD_actual | 39.646
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.02424648385
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 4.647
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.7899010433
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 39.646
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 2 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.69
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.328
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.5525
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 37.8415
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.683
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.008
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.06014990627
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 4.683
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.1289113939
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 40.008
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 2 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.69
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.328
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.5525
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 37.8415
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.744
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.12
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.12098626093
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 4.744
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.2337985742
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 40.12
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 2 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.69
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.328
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.5525
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 37.8415
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.69
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.328
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.0671311273
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 4.69
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.4285890518
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 40.328
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 3 - Sample 3 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 3 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.217
operation: col(d(18_16)Mean)
inst: col
dD_actual | -92.943
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.7945122855
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -12.217
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -104.378727067
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -92.943
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 3 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.128
operation: col(d(18_16)Mean)
inst: col
dD_actual | -105.802
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.7030683362
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -13.128
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -116.421086454
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -105.802
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 3 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.347
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.88
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9214808225
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -13.347
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -119.303610927
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -108.88
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 3 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.31
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.777
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.8845800828
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -13.31
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.143644862
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -109.777
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 3 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.406
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.581
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9803225426
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -13.406
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.896584978
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -110.581
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 3 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.456
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.82
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0301884071
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -13.456
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.120406729
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -110.82
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 3 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.424
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.487
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9982742538
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -13.424
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.745047347
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -111.487
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 3 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.422
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.33
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9962796193
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -13.422
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.598017996
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -111.33
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 3 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.464
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.891
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0381669454
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -13.464
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.12339039
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -111.891
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 3 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.479
operation: col(d(18_16)Mean)
inst: col
dD_actual | -112.047
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0531267048
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -13.479
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.269483249
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -112.047
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 4 - Sample 4 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 4 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.031
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.537
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.6197421118
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -8.031
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.6497013265
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -66.537
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 4 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.811
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.33
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4003323081
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.811
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.7098766166
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -62.33
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 4 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.75
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.365
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3394959535
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.75
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.8061611792
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -61.365
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 4 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.768
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.864
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3574476647
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.768
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3369783459
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -60.864
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 4 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.769
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.764
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.358444982
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.769
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.2433290778
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -60.764
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 4 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.691
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.768
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2806542334
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.691
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.2470750485
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -60.768
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 4 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.695
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.403
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2846435026
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.695
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.9052552199
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -60.403
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 4 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.732
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.553
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3215442423
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.732
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.0457291221
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -60.553
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 4 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.788
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.877
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3773940105
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.788
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3491527508
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -60.877
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 4 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.739
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.414
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3285254633
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.739
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.9155566394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -60.414
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 5 - Sample 5 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 5 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.189
operation: col(d(18_16)Mean)
inst: col
dD_actual | 5.349
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.42179399252
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.189
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -12.3289884445
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 5.349
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 5 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.535
operation: col(d(18_16)Mean)
inst: col
dD_actual | 11.125
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.07672221037
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.535
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -6.91980671775
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 11.125
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 5 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.624
operation: col(d(18_16)Mean)
inst: col
dD_actual | 12.932
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.98796097161
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.624
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -5.22756444279
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 12.932
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 5 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.677
operation: col(d(18_16)Mean)
inst: col
dD_actual | 13.665
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.93510315527
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.677
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -4.54111530745
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 13.665
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 6 - Sample 6 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 6 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.245
operation: col(d(18_16)Mean)
inst: col
dD_actual | 11.243
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.36594422432
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.245
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -6.80930058136
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 11.243
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 6 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.282
operation: col(d(18_16)Mean)
inst: col
dD_actual | 11.35
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.32904348461
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.282
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -6.70909586447
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 11.35
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 6 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.259
operation: col(d(18_16)Mean)
inst: col
dD_actual | 11.316
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.35198178226
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.259
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -6.74093661563
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 11.316
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 6 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.322
operation: col(d(18_16)Mean)
inst: col
dD_actual | 11.55
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.28915079303
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.322
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -6.52179732823
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 11.55
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 7 - Sample 7 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 7 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.316
operation: col(d(18_16)Mean)
inst: col
dD_actual | 11.441
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.29513469676
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.316
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -6.62387503048
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 11.441
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 7 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.311
operation: col(d(18_16)Mean)
inst: col
dD_actual | 11.281
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.30012128321
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.311
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -6.77371385948
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 11.281
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 7 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.382
operation: col(d(18_16)Mean)
inst: col
dD_actual | 11.621
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.22931175566
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.382
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -6.45530634786
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 11.621
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 7 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.37
operation: col(d(18_16)Mean)
inst: col
dD_actual | 11.768
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.24127956313
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.37
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -6.31764192372
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 11.768
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 8 - Sample 8 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 8 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.489
operation: col(d(18_16)Mean)
inst: col
dD_actual | 12.429
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.12259880569
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.489
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -5.69862026144
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 12.429
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 8 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.497
operation: col(d(18_16)Mean)
inst: col
dD_actual | 12.508
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.11462026737
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.497
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -5.62463733962
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 12.508
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 8 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.474
operation: col(d(18_16)Mean)
inst: col
dD_actual | 12.568
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.13755856503
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.474
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -5.56844777875
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 12.568
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 8 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.494
operation: col(d(18_16)Mean)
inst: col
dD_actual | 12.584
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.11761221924
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.494
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -5.55346389585
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 12.584
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 9 - Sample 9 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 9 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.071
operation: col(d(18_16)Mean)
inst: col
dD_actual | 9.294
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.68109648778
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -0.071
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.63452481706
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 9.294
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 9 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.059
operation: col(d(18_16)Mean)
inst: col
dD_actual | 9.137
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.66912868031
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -0.059
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.78155416801
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 9.137
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 9 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.048
operation: col(d(18_16)Mean)
inst: col
dD_actual | 9.114
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.65815819012
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -0.048
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.80309349968
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 9.114
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 9 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.037
operation: col(d(18_16)Mean)
inst: col
dD_actual | 8.987
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.64718769994
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -0.037
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.92202807019
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 8.987
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 10 - Sample 10 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 10 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.194
operation: col(d(18_16)Mean)
inst: col
dD_actual | 9.902
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.41680740608
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.194
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.06513726688
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 9.902
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 10 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.219
operation: col(d(18_16)Mean)
inst: col
dD_actual | 9.91
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.39187447384
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.219
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.05764532543
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 9.91
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 10 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.177
operation: col(d(18_16)Mean)
inst: col
dD_actual | 9.877
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.4337618
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.177
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.08854958391
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 9.877
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 10 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.204
operation: col(d(18_16)Mean)
inst: col
dD_actual | 9.889
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.40683423318
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.204
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.07731167173
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 9.889
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 11 - Sample 11 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 11 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.152
operation: col(d(18_16)Mean)
inst: col
dD_actual | 8.947
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.76187918823
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -0.152
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.95948777744
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 8.947
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 11 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.205
operation: col(d(18_16)Mean)
inst: col
dD_actual | 8.711
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.81473700457
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -0.205
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -9.18050005021
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 8.711
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 11 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.227
operation: col(d(18_16)Mean)
inst: col
dD_actual | 8.696
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.83667798494
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -0.227
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -9.19454744042
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 8.696
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 11 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.142
operation: col(d(18_16)Mean)
inst: col
dD_actual | 9.013
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.75190601533
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -0.142
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.89767926048
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 9.013
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 12 - Sample 12 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 12 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.15
operation: col(d(18_16)Mean)
inst: col
dD_actual | 8.747
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.75988455365
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -0.15
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -9.14678631368
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 8.747
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 12 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.13
operation: col(d(18_16)Mean)
inst: col
dD_actual | 8.927
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.73993820786
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -0.13
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.97821763106
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 8.927
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 12 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.096
operation: col(d(18_16)Mean)
inst: col
dD_actual | 9.325
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.70602942002
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -0.096
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.60549354394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 9.325
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 12 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.098
operation: col(d(18_16)Mean)
inst: col
dD_actual | 8.661
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.7080240546
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -0.098
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -9.22732468427
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 8.661
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 13 - Sample 13 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 13 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.179
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.202
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.77002778121
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.179
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -65.28858606
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -51.202
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 13 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.438
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.523
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0283329592
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.438
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.2716636167
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -56.523
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 13 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.551
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.919
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1410298129
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.551
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.5790073997
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -57.919
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 13 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.587
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.852
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1769332353
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.587
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.4527550713
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -58.852
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 14 - Sample 14 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 14 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.008
operation: col(d(18_16)Mean)
inst: col
dD_actual | 4.49
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.61826549855
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -0.008
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -13.1334356576
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 4.49
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 14 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.386
operation: col(d(18_16)Mean)
inst: col
dD_actual | 10.547
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.2253224865
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.386
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -7.46109948749
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 10.547
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 14 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.446
operation: col(d(18_16)Mean)
inst: col
dD_actual | 11.963
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.16548344914
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.446
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -6.13502585089
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 11.963
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 14 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.435
operation: col(d(18_16)Mean)
inst: col
dD_actual | 12.523
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.17645393932
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.435
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -5.6105899494
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 12.523
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 15 - Sample 15 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 15 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.19
operation: col(d(18_16)Mean)
inst: col
dD_actual | 9.879
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.42079667523
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.19
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.08667659854
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 9.879
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 15 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.262
operation: col(d(18_16)Mean)
inst: col
dD_actual | 10.275
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.34898983039
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.262
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -7.71582549678
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 10.275
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 15 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.227
operation: col(d(18_16)Mean)
inst: col
dD_actual | 10.34
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.38389593553
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.227
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -7.6549534725
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 10.34
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 15 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.28
operation: col(d(18_16)Mean)
inst: col
dD_actual | 10.377
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.33103811918
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.28
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -7.6203032433
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 10.377
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 16 - Sample 16 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 16 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.008
operation: col(d(18_16)Mean)
inst: col
dD_actual | 9.481
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.61826549855
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -0.008
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.45940068567
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 9.481
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 16 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.028
operation: col(d(18_16)Mean)
inst: col
dD_actual | 9.546
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.63821184434
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -0.028
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.39852866139
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 9.546
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 16 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.044
operation: col(d(18_16)Mean)
inst: col
dD_actual | 9.562
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.65416892097
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -0.044
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.38354477849
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 9.562
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 17 - Sample 17 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 17 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.61028696023
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.3382877963
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 17 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.049
operation: col(d(18_16)Mean)
inst: col
dD_actual | 9.78
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.65915550741
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -0.049
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.17938937398
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 9.78
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 17 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.094
operation: col(d(18_16)Mean)
inst: col
dD_actual | 10.452
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.51653913502
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.094
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -7.55006629221
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 10.452
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 17 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.015
operation: col(d(18_16)Mean)
inst: col
dD_actual | 9.88
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.59532720089
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.015
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.08574010586
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 9.88
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 17 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.032
operation: col(d(18_16)Mean)
inst: col
dD_actual | 9.591
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.64220111349
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -0.032
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.35638649073
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 9.591
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 18 - Sample 18 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 18 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.586
operation: col(d(18_16)Mean)
inst: col
dD_actual | 16.181
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.02854173916
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 1.586
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -2.18489972151
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 16.181
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 18 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.621
operation: col(d(18_16)Mean)
inst: col
dD_actual | 16.766
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.993635634029
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 1.621
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -1.637051503
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 16.766
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 18 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.619
operation: col(d(18_16)Mean)
inst: col
dD_actual | 16.827
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.995630268608
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 1.619
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -1.57992544944
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 16.827
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 18 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.65
operation: col(d(18_16)Mean)
inst: col
dD_actual | 17.328
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.964713432635
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 1.65
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -1.11074261615
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 17.328
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 19 - Sample 19 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 19 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.645
operation: col(d(18_16)Mean)
inst: col
dD_actual | 23.472
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.969700019082
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 1.645
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 4.64306841725
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 23.472
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 19 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.644
operation: col(d(18_16)Mean)
inst: col
dD_actual | 24.434
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.970697336371
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 1.644
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 5.54397437658
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 24.434
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 19 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.625
operation: col(d(18_16)Mean)
inst: col
dD_actual | 24.724
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.989646364871
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 1.625
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 5.81555725413
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 24.724
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 19 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.57
operation: col(d(18_16)Mean)
inst: col
dD_actual | 24.351
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.04449881579
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 1.57
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 5.46624548404
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 24.351
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 20 - Sample 20 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 20 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 7.155
operation: col(d(18_16)Mean)
inst: col
dD_actual | 46.787
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 4.5255182458
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 7.155
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 26.4773952798
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 46.787
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 20 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 7.397
operation: col(d(18_16)Mean)
inst: col
dD_actual | 49.196
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 4.76686902985
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 7.397
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 28.7334061489
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 49.196
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 20 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 7.525
operation: col(d(18_16)Mean)
inst: col
dD_actual | 49.521
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 4.8945256429
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 7.525
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 29.0377662703
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 49.521
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 20 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 7.55
operation: col(d(18_16)Mean)
inst: col
dD_actual | 49.647
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 4.91945857513
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 7.55
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 29.1557643481
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 49.647
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 21 - Sample 21 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 21 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.17
operation: col(d(18_16)Mean)
inst: col
dD_actual | 35.066
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.54852613678
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 4.17
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 15.5007645633
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 35.066
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 21 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.015
operation: col(d(18_16)Mean)
inst: col
dD_actual | 33.611
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.39394195692
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 4.015
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 14.1381677121
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 33.611
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 21 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.987
operation: col(d(18_16)Mean)
inst: col
dD_actual | 33.406
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.36601707281
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 3.987
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 13.9461867125
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 33.406
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 21 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.988
operation: col(d(18_16)Mean)
inst: col
dD_actual | 33.756
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.3670143901
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 3.988
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 14.2739591509
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 33.756
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 22 - Sample 22 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 22 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.802
operation: col(d(18_16)Mean)
inst: col
dD_actual | -47.37
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.39403916308
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -6.802
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.6999461056
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -47.37
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 22 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.331
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.247
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.9216200092
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.331
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.0766989555
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -55.247
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 22 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.429
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.918
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0193571036
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.429
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.6415782258
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -56.918
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 22 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.5
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.87
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0901666311
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.5
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.5331192583
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -57.87
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 23 - Sample 23 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 23 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.437
operation: col(d(18_16)Mean)
inst: col
dD_actual | 5.506
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.17445930474
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.437
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -12.1819590935
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 5.506
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 23 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.863
operation: col(d(18_16)Mean)
inst: col
dD_actual | 12.182
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.74960213943
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.863
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -5.9299339537
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 12.182
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 23 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.943
operation: col(d(18_16)Mean)
inst: col
dD_actual | 13.342
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.66981675628
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.943
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -4.84360244349
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 13.342
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 23 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.955
operation: col(d(18_16)Mean)
inst: col
dD_actual | 13.48
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.6578489488
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.955
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -4.71436645348
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 13.48
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 24 - Sample 24 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 24 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.019
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.966
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.61313972544
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -6.019
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.5121250187
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -43.966
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 24 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.403
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.623
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.99610956459
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -6.403
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -63.8098641163
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -49.623
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 24 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.444
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.26
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.03699957346
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -6.444
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.4064099543
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -50.26
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 24 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.51
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.109
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.10282251456
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -6.51
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -65.2014922406
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -51.109
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 25 - Sample 25 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 25 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.524
operation: col(d(18_16)Mean)
inst: col
dD_actual | 6.335
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.08769270056
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.524
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -11.4056066608
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 6.335
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 25 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.821
operation: col(d(18_16)Mean)
inst: col
dD_actual | 11.141
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.79148946559
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.821
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -6.90482283485
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 11.141
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 25 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.928
operation: col(d(18_16)Mean)
inst: col
dD_actual | 12.619
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.68477651562
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.928
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -5.52068665201
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 12.619
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 25 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.92
operation: col(d(18_16)Mean)
inst: col
dD_actual | 13.343
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.69275505393
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.92
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -4.84266595081
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 13.343
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 26 - Sample 26 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 26 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.258
operation: col(d(18_16)Mean)
inst: col
dD_actual | 18.069
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.3556618101
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 1.258
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.416801539373
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 18.069
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 26 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.301
operation: col(d(18_16)Mean)
inst: col
dD_actual | 18.164
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.31277716665
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 1.301
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.327834734657
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 18.164
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 26 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.304
operation: col(d(18_16)Mean)
inst: col
dD_actual | 18.498
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.30978521479
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 1.304
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.0150461791304
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 18.498
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 26 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.296
operation: col(d(18_16)Mean)
inst: col
dD_actual | 18.514
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.3177637531
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 1.296
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -6.22962308843e-05
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 18.514
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 27 - Sample 27 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 27 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.231
operation: col(d(18_16)Mean)
inst: col
dD_actual | -11.241
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.835301833
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.231
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.8654020259
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -11.241
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 27 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.448
operation: col(d(18_16)Mean)
inst: col
dD_actual | -14.322
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.05171968481
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.448
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -30.7507359767
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -14.322
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 27 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.486
operation: col(d(18_16)Mean)
inst: col
dD_actual | -15.273
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.08961774181
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.486
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -31.6413405165
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -15.273
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 27 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.472
operation: col(d(18_16)Mean)
inst: col
dD_actual | -15.473
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.07565529976
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.472
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -31.8286390528
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -15.473
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 28 - Sample 28 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 28 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.816
operation: col(d(18_16)Mean)
inst: col
dD_actual | -18.346
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.41873244733
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.816
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -34.5191825259
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -18.346
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 28 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.889
operation: col(d(18_16)Mean)
inst: col
dD_actual | -18.969
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.49153660946
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.889
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -35.1026174663
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -18.969
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 28 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.876
operation: col(d(18_16)Mean)
inst: col
dD_actual | -18.699
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.4785714847
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.876
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -34.8497644424
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -18.699
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 28 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.928
operation: col(d(18_16)Mean)
inst: col
dD_actual | -18.982
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.53043198375
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.928
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -35.1147918712
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -18.982
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 29 - Sample 29 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 29 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.012
operation: col(d(18_16)Mean)
inst: col
dD_actual | -19.248
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.61420663606
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.012
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -35.3638989244
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -19.248
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 29 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.044
operation: col(d(18_16)Mean)
inst: col
dD_actual | -19.266
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.64612078932
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.044
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -35.3807557926
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -19.266
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 29 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.05
operation: col(d(18_16)Mean)
inst: col
dD_actual | -19.285
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.65210469306
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.05
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -35.3985491536
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -19.285
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 29 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.006
operation: col(d(18_16)Mean)
inst: col
dD_actual | -19.208
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.60822273232
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.006
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -35.3264392171
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -19.208
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 30 - Sample 30 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 30 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.727
operation: col(d(18_16)Mean)
inst: col
dD_actual | -18.218
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.32997120857
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.727
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -34.3993114627
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -18.218
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 30 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.781
operation: col(d(18_16)Mean)
inst: col
dD_actual | -18.506
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.3838263422
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.781
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -34.6690213549
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -18.506
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 30 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.792
operation: col(d(18_16)Mean)
inst: col
dD_actual | -18.352
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.39479683238
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.792
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -34.524801482
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -18.352
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 30 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.876
operation: col(d(18_16)Mean)
inst: col
dD_actual | -18.882
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.4785714847
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.876
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -35.021142603
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -18.882
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 31 - Sample 31 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 31 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.21
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.831
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.80094461718
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.21
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.7506253189
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -53.831
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 31 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.467
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.29
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0572551606
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.467
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.9899535032
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -57.29
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 31 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.511
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.194
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1011371213
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.511
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.8365428871
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -58.194
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 31 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.566
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.561
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1559895722
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.566
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.1802357011
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -58.561
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 32 - Sample 32 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 32 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.001
operation: col(d(18_16)Mean)
inst: col
dD_actual | -22.316
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.60323614588
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.001
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -38.2370584703
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -22.316
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 32 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.854
operation: col(d(18_16)Mean)
inst: col
dD_actual | -19.929
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.45663050433
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.854
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -36.0016504403
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -19.929
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 32 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.811
operation: col(d(18_16)Mean)
inst: col
dD_actual | -18.807
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.41374586088
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.811
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -34.9509056519
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -18.807
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 32 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.811
operation: col(d(18_16)Mean)
inst: col
dD_actual | -19.167
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.41374586088
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.811
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -35.2880430172
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -19.167
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 33 - Sample 33 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 33 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.122
operation: col(d(18_16)Mean)
inst: col
dD_actual | -19.661
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.7239115379
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.122
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -35.7506704017
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -19.661
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 33 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.134
operation: col(d(18_16)Mean)
inst: col
dD_actual | -19.322
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.73587934537
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.134
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -35.4331993828
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -19.322
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 33 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.203
operation: col(d(18_16)Mean)
inst: col
dD_actual | -20.313
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.80469423835
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.203
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -36.3612636299
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -20.313
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 33 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.194
operation: col(d(18_16)Mean)
inst: col
dD_actual | -20.076
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.79571838274
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.194
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -36.1393148644
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -20.076
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 34 - Sample 34 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 34 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.27
operation: col(d(18_16)Mean)
inst: col
dD_actual | -21.026
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.87151449674
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.27
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -37.0289829116
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -21.026
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 34 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.364
operation: col(d(18_16)Mean)
inst: col
dD_actual | -21.532
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.96526232195
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.364
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -37.5028482083
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -21.532
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 34 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.215
operation: col(d(18_16)Mean)
inst: col
dD_actual | -20.736
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.81666204582
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.215
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -36.757400034
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -20.736
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 34 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.331
operation: col(d(18_16)Mean)
inst: col
dD_actual | -21.268
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.9323508514
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.331
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -37.2556141404
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -21.268
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 35 - Sample 35 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 35 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.909
operation: col(d(18_16)Mean)
inst: col
dD_actual | -19.506
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.51148295525
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.909
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -35.6055140361
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -19.506
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 35 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.94
operation: col(d(18_16)Mean)
inst: col
dD_actual | -19.958
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.54239979122
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.94
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -36.028808728
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -19.958
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 35 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.975
operation: col(d(18_16)Mean)
inst: col
dD_actual | -19.852
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.57730589635
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.975
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -35.9295405038
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -19.852
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 35 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.963
operation: col(d(18_16)Mean)
inst: col
dD_actual | -19.872
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.56533808888
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.963
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -35.9482703574
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -19.872
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 36 - Sample 36 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 36 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.212
operation: col(d(18_16)Mean)
inst: col
dD_actual | -20.785
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.81367009395
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.212
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -36.8032881754
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -20.785
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 36 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.305
operation: col(d(18_16)Mean)
inst: col
dD_actual | -20.845
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.90642060187
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.305
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -36.8594777363
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -20.845
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 36 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.29
operation: col(d(18_16)Mean)
inst: col
dD_actual | -21.203
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.89146084253
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.29
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -37.1947421161
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -21.203
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 36 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.318
operation: col(d(18_16)Mean)
inst: col
dD_actual | -21.204
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.91938572663
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.318
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -37.1956786088
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -21.204
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 37 - Sample 37 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 37 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.278
operation: col(d(18_16)Mean)
inst: col
dD_actual | -15.919
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.8821757456
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.278
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -32.2463147886
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -15.919
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 37 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.212
operation: col(d(18_16)Mean)
inst: col
dD_actual | -15.693
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.8163528045
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.212
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -32.0346674426
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -15.693
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 37 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.178
operation: col(d(18_16)Mean)
inst: col
dD_actual | -15.175
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.78244401666
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.178
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -31.5495642338
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -15.175
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 37 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.179
operation: col(d(18_16)Mean)
inst: col
dD_actual | -15.645
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.78344133395
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.179
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -31.9897157939
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -15.645
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 38 - Sample 38 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 38 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.47
operation: col(d(18_16)Mean)
inst: col
dD_actual | -9.086
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.07634337573
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -1.47
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.8472602978
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -9.086
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 38 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.441
operation: col(d(18_16)Mean)
inst: col
dD_actual | -8.774
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.04742117433
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -1.441
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.5550745813
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -8.774
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 38 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.396
operation: col(d(18_16)Mean)
inst: col
dD_actual | -8.171
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.00254189631
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -1.396
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -24.9903694945
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -8.171
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 38 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.463
operation: col(d(18_16)Mean)
inst: col
dD_actual | -8.541
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.0693621547
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -1.463
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.3368717866
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -8.541
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 39 - Sample 39 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 39 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.851
operation: col(d(18_16)Mean)
inst: col
dD_actual | -11.004
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.45632126301
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -1.851
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.6434532604
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -11.004
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 39 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.921
operation: col(d(18_16)Mean)
inst: col
dD_actual | -11.917
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.52613347327
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -1.921
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.4984710784
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -11.917
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 39 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.015
operation: col(d(18_16)Mean)
inst: col
dD_actual | -12.434
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.61988129848
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.015
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.9826377945
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -12.434
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 39 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.936
operation: col(d(18_16)Mean)
inst: col
dD_actual | -12.014
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.54109323261
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -1.936
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.5893108684
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -12.014
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 40 - Sample 40 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 40 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.223
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.526
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.81390974194
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.223
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.4649950511
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -53.526
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 40 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.522
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.585
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1121076115
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.522
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.2662188442
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -57.585
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 40 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.469
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.94
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0592497951
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.469
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.598673746
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -57.94
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 40 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.596
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.613
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1859090909
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.596
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.2289333205
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -58.613
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 41 - Sample 41 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 41 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.92
operation: col(d(18_16)Mean)
inst: col
dD_actual | -22.067
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.52245344543
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.92
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -38.0038717927
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -22.067
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 41 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.728
operation: col(d(18_16)Mean)
inst: col
dD_actual | -19.0
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.33096852586
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.728
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -35.1316487394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -19.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 41 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.716
operation: col(d(18_16)Mean)
inst: col
dD_actual | -17.844
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.31900071838
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.716
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -34.0490631999
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -17.844
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 41 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.635
operation: col(d(18_16)Mean)
inst: col
dD_actual | -17.861
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.23821801794
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.635
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -34.0649835755
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -17.861
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 42 - Sample 42 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 42 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.396
operation: col(d(18_16)Mean)
inst: col
dD_actual | -16.701
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.99985918576
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.396
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -32.9786520653
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -16.701
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 42 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.418
operation: col(d(18_16)Mean)
inst: col
dD_actual | -17.129
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.02180016613
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.418
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -33.3794709329
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -17.129
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 42 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.421
operation: col(d(18_16)Mean)
inst: col
dD_actual | -16.988
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.02479211799
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.421
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -33.2474254648
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -16.988
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 42 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.469
operation: col(d(18_16)Mean)
inst: col
dD_actual | -17.1
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.07266334789
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.469
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -33.3523126451
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -17.1
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 43 - Sample 43 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 43 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.39
operation: col(d(18_16)Mean)
inst: col
dD_actual | -16.514
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.99387528202
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.39
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -32.8035279339
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -16.514
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 43 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.116
operation: col(d(18_16)Mean)
inst: col
dD_actual | -15.251
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.72061034471
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.116
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -31.6207376775
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -15.251
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 43 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.298
operation: col(d(18_16)Mean)
inst: col
dD_actual | -16.317
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.90212209139
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.298
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -32.6190388757
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -16.317
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 43 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.332
operation: col(d(18_16)Mean)
inst: col
dD_actual | -16.639
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.93603087923
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.332
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -32.9205895191
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -16.639
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 44 - Sample 44 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 44 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.368
operation: col(d(18_16)Mean)
inst: col
dD_actual | -6.95
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.9746170122
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -1.368
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -23.8469119308
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -6.95
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 44 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.292
operation: col(d(18_16)Mean)
inst: col
dD_actual | -5.717
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.8988208982
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -1.292
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -22.6922164548
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -5.717
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 44 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.353
operation: col(d(18_16)Mean)
inst: col
dD_actual | -5.844
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.95965725286
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -1.353
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -22.8111510253
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -5.844
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 44 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.354
operation: col(d(18_16)Mean)
inst: col
dD_actual | -6.143
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.96065457015
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -1.354
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -23.091162337
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -6.143
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 45 - Sample 45 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 45 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.25
operation: col(d(18_16)Mean)
inst: col
dD_actual | -6.857
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.85693357205
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -1.25
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -23.7598181114
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -6.857
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 45 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.326
operation: col(d(18_16)Mean)
inst: col
dD_actual | -7.499
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.93272968604
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -1.326
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -24.3610464127
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -7.499
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 45 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.262
operation: col(d(18_16)Mean)
inst: col
dD_actual | -7.039
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.86890137952
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -1.262
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -23.9302597794
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -7.039
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 45 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.269
operation: col(d(18_16)Mean)
inst: col
dD_actual | -7.1
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.87588260055
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -1.269
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -23.9873858329
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -7.1
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 46 - Sample 46 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 46 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.593
operation: col(d(18_16)Mean)
inst: col
dD_actual | -16.354
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.19633069178
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.593
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -32.6536891049
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -16.354
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 46 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.702
operation: col(d(18_16)Mean)
inst: col
dD_actual | -17.466
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.30503827633
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.702
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -33.6950689664
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -17.466
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 46 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.678
operation: col(d(18_16)Mean)
inst: col
dD_actual | -17.32
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.28110266138
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.678
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -33.558341035
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -17.32
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 46 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.73
operation: col(d(18_16)Mean)
inst: col
dD_actual | -17.719
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.33296316044
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.73
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -33.9320016148
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -17.719
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 47 - Sample 47 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 47 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.7
operation: col(d(18_16)Mean)
inst: col
dD_actual | -17.524
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.30304364175
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.7
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -33.7493855419
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -17.524
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 47 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.838
operation: col(d(18_16)Mean)
inst: col
dD_actual | -18.625
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.4406734277
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.838
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -34.780463984
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -18.625
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 47 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.718
operation: col(d(18_16)Mean)
inst: col
dD_actual | -18.192
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.32099535296
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.718
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -34.374962653
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -18.192
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 47 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.727
operation: col(d(18_16)Mean)
inst: col
dD_actual | -18.241
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.32997120857
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.727
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -34.4208507944
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -18.241
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 48 - Sample 48 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 48 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.187
operation: col(d(18_16)Mean)
inst: col
dD_actual | -14.029
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.79141987226
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.187
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -30.4763436211
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -14.029
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 48 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.131
operation: col(d(18_16)Mean)
inst: col
dD_actual | -13.561
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.73557010405
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.131
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -30.0380650463
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -13.561
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 48 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.053
operation: col(d(18_16)Mean)
inst: col
dD_actual | -13.231
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.65777935548
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.053
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -29.7290224615
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -13.231
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 48 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.1
operation: col(d(18_16)Mean)
inst: col
dD_actual | -13.606
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.70465326808
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.1
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -30.0802072169
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -13.606
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 49 - Sample 49 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 49 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.192
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.577
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.78299290597
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.192
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.5127561779
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -53.577
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 49 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.476
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.22
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0662310162
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.476
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.9243990156
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -57.22
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 49 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.532
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.151
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1220807844
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.532
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.7962737018
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -58.151
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 49 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.51
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.302
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.100139804
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.51
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.9376840966
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -58.302
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 50 - Sample 50 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 50 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.058
operation: col(d(18_16)Mean)
inst: col
dD_actual | -22.99
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.66008323138
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.058
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -38.8682545375
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -22.99
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 50 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.984
operation: col(d(18_16)Mean)
inst: col
dD_actual | -20.835
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.58628175196
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.984
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -36.8501128095
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -20.835
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 50 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.859
operation: col(d(18_16)Mean)
inst: col
dD_actual | -19.833
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.46161709077
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.859
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -35.9117471429
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -19.833
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 50 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.886
operation: col(d(18_16)Mean)
inst: col
dD_actual | -19.634
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.48854465759
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.886
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -35.7253850993
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -19.634
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 51 - Sample 51 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 51 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.23
operation: col(d(18_16)Mean)
inst: col
dD_actual | -23.023
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.83162180516
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.23
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -38.899158796
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -23.023
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 51 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.311
operation: col(d(18_16)Mean)
inst: col
dD_actual | -23.466
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.91240450561
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.311
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -39.3140250537
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -23.466
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 51 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.253
operation: col(d(18_16)Mean)
inst: col
dD_actual | -23.377
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.85456010282
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.253
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -39.2306772051
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -23.377
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 51 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.301
operation: col(d(18_16)Mean)
inst: col
dD_actual | -23.694
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.90243133271
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.301
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -39.5275453851
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -23.694
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 52 - Sample 52 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 52 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.134
operation: col(d(18_16)Mean)
inst: col
dD_actual | -34.299
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.73319663483
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -4.134
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -49.4590502694
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -34.299
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 52 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.284
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.064
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.88279422824
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -4.284
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.1119598517
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -36.064
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 52 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.226
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.417
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.82494982546
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -4.226
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -50.506049087
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -35.417
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 52 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.348
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.504
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.94662253477
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -4.348
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.5240166315
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -36.504
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 53 - Sample 53 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 53 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.593
operation: col(d(18_16)Mean)
inst: col
dD_actual | -9.214
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.19901340233
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -1.593
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.967131361
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -9.214
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 53 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.612
operation: col(d(18_16)Mean)
inst: col
dD_actual | -7.29
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.21796243083
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -1.612
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -24.1653194424
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -7.29
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 53 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.499
operation: col(d(18_16)Mean)
inst: col
dD_actual | -6.402
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.10526557712
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -1.499
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -23.3337139414
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -6.402
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 53 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.49
operation: col(d(18_16)Mean)
inst: col
dD_actual | -6.228
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.09628972151
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -1.49
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -23.1707642149
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -6.228
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 54 - Sample 54 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 54 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.476
operation: col(d(18_16)Mean)
inst: col
dD_actual | 14.768
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.13556393045
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.476
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -3.50816388007
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 14.768
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 54 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.464
operation: col(d(18_16)Mean)
inst: col
dD_actual | 16.122
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.14753173792
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.464
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -2.2401527897
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 16.122
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 54 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.557
operation: col(d(18_16)Mean)
inst: col
dD_actual | 16.899
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.05478123001
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.557
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -1.5124979764
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 16.899
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 54 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.511
operation: col(d(18_16)Mean)
inst: col
dD_actual | 16.771
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.10065782532
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.511
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -1.63236903959
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 16.771
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 55 - Sample 55 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 55 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.708
operation: col(d(18_16)Mean)
inst: col
dD_actual | 23.244
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.906869029847
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 1.708
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 4.42954808593
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 23.244
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 55 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.712
operation: col(d(18_16)Mean)
inst: col
dD_actual | 23.419
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.902879760689
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 1.712
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 4.59343430514
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 23.419
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 55 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.669
operation: col(d(18_16)Mean)
inst: col
dD_actual | 23.87
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.945764404135
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 1.669
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 5.01579250437
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 23.87
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 55 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.756
operation: col(d(18_16)Mean)
inst: col
dD_actual | 23.96
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.858997799953
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 1.756
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 5.10007684568
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 23.96
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 56 - Sample 56 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 56 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.767
operation: col(d(18_16)Mean)
inst: col
dD_actual | -12.421
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.36986390014
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -2.767
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.9704633897
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -12.421
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 56 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.108
operation: col(d(18_16)Mean)
inst: col
dD_actual | -16.323
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.70994909585
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.108
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -32.6246578318
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -16.323
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 56 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.152
operation: col(d(18_16)Mean)
inst: col
dD_actual | -16.979
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.75383105658
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.152
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -33.2389970307
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -16.979
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 56 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.076
operation: col(d(18_16)Mean)
inst: col
dD_actual | -16.97
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.67803494259
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.076
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -33.2305685966
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -16.97
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 57 - Sample 57 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 57 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.101
operation: col(d(18_16)Mean)
inst: col
dD_actual | 7.477
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.71101600647
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -0.101
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -10.3361320188
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 7.477
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 57 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.045
operation: col(d(18_16)Mean)
inst: col
dD_actual | 9.222
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.65516623826
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -0.045
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.7019522901
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 9.222
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 57 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
dD_actual | 10.052
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.61028696023
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -7.92466336469
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 10.052
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 57 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.055
operation: col(d(18_16)Mean)
inst: col
dD_actual | 10.581
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.55543450931
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.055
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -7.42925873633
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 10.581
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 58 - Sample 58 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 58 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.052
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.308
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.64336848545
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.052
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.451361603
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -50.308
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 58 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.476
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.79
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0662310162
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.476
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.5217071626
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -56.79
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 58 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.548
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.936
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.138037861
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.548
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.5949277753
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -57.936
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 58 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.524
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.249
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1141022461
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.524
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.8880499845
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -58.249
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 59 - Sample 59 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 59 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.902
operation: col(d(18_16)Mean)
inst: col
dD_actual | -5.226
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.50986715532
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -0.902
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -22.2323985483
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -5.226
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 59 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.775
operation: col(d(18_16)Mean)
inst: col
dD_actual | -1.791
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.38320785956
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -0.775
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -19.0155461884
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -1.791
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 59 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.689
operation: col(d(18_16)Mean)
inst: col
dD_actual | -0.726
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.29743857266
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -0.689
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.0181814829
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -0.726
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 59 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.514
operation: col(d(18_16)Mean)
inst: col
dD_actual | 0.358
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.12290804701
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -0.514
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.0030234164
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 0.358
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 60 - Sample 60 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 60 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.791
operation: col(d(18_16)Mean)
inst: col
dD_actual | -2.923
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.82140898427
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.791
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -20.0756559035
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -2.923
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 60 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.875
operation: col(d(18_16)Mean)
inst: col
dD_actual | -2.977
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.73763433196
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.875
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -20.1262265083
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -2.977
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 60 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.945
operation: col(d(18_16)Mean)
inst: col
dD_actual | -2.709
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.6678221217
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 0.945
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -19.8752464697
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -2.709
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 60 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.032
operation: col(d(18_16)Mean)
inst: col
dD_actual | -2.489
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.58105551752
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 1.032
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -19.6692180798
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -2.489
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 61 - Sample 61 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 61 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.158
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.606
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.75176682868
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -6.158
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.5560315662
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -37.606
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 61 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.399
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.677
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.99212029543
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -6.399
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.4320005902
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -40.677
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 61 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.505
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.426
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.09783592812
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -6.505
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.1334336084
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -41.426
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 61 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.427
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.421
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.02004517954
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -6.427
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.128751145
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -41.421
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 62 - Sample 62 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 62 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.18
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.646
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.7683423879
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -8.18
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.259837579
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -58.646
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 62 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.228
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.001
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.8162136178
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -8.228
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.528785162
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -60.001
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 62 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.166
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.998
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.7543799459
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -8.166
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.525975684
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -59.998
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 62 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.27
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.862
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.858100944
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -8.27
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3351053605
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -60.862
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 63 - Sample 63 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 63 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.752
operation: col(d(18_16)Mean)
inst: col
dD_actual | -27.864
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.34953871971
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -4.752
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -43.4327198657
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -27.864
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 63 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.922
operation: col(d(18_16)Mean)
inst: col
dD_actual | -22.421
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.52176536946
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -3.922
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -38.3353902019
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -22.421
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 63 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.385
operation: col(d(18_16)Mean)
inst: col
dD_actual | -23.321
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.98352327448
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -4.385
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -39.178233615
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -23.321
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 63 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.378
operation: col(d(18_16)Mean)
inst: col
dD_actual | -23.028
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.97654205345
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -4.378
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -38.9038412594
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -23.028
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 64 - Sample 64 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 64 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.3
operation: col(d(18_16)Mean)
inst: col
dD_actual | 25.258
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.31377448394
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 1.3
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 6.3156443459
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 25.258
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 64 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.527
operation: col(d(18_16)Mean)
inst: col
dD_actual | 25.23
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.08738345924
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 1.527
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 6.28942255083
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 25.23
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 64 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.248
operation: col(d(18_16)Mean)
inst: col
dD_actual | 24.489
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.36563498299
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 1.248
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 5.59548147404
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 24.489
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 64 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.3
operation: col(d(18_16)Mean)
inst: col
dD_actual | 25.258
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.31377448394
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: 1.3
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 6.3156443459
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: 25.258
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean

## 65 - Sample 65 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:28:20
Detail Variables
Detail 65 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.972
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.944
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.56358310229
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -6.972
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -63.1739855858
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -48.944
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 65 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.373
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.126
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.96350733536
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.373
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.8998760223
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -56.126
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 65 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.481
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.735
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0712176026
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.481
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.4066927464
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -57.735
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean
Detail 65 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.5525
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 37.8415
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.479
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -112.047
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2653
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.5558
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.524
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.331
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1141022461
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.61730844119
measure_array: [ 4.5525 -13.2653]
inst: fit_type1
actual_mean: -7.524
operation: fit_type1(d(18_16)Mean)
fit_slope: 1.00268992684
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.9648423844
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 18.5140665208
measure_array: [ 37.8415 -108.5558]
inst: fit_type1
actual_mean: -58.331
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06781400438
tarray: []
source_col: d(D_H)Mean