## Summary
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
221118_22-158_22-258_HIDS2052
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -9.2137 | -67.218 |  | 47.446 | 4.3007 | 0.40168 | 1.4555 |  | 4.6048
2 |  | PZ | 1.93 | 18.1 |  |  |  | 5.0737 | 39.708
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.1513 | 2.035
4 |  | EV | -9.8517 | -70.2764 |  |  |  | 0.11367 | 1.9486
5 |  | WOOD5-D-2 | -8.8786 | -94.1872 |  | 30.8194 |  | 0.034766 | 1.5093 |  | -3.6729
6 |  | WOOD5-D-3 | -9.7024 | -96.8446 |  | 19.4284 |  | 0.021093 | 0.1241
7 |  | WOOD5-B-5 | -2.3948 | -93.9961 | -6.4925 | 636.4344 | 24.888 | 0.037327 | 0.58344 | -39.5637 | -1.8467
8 |  | WOOD5-C-2 | -1.1826 | -53.1134 | -2.326 | 407.4334 |  | 0.15744 | 2.8894 | -19.9618 | -0.25515
9 |  | WOOD5-B-2 | -0.069196 | -88.2585 | -8.554 | 1435.0586 | 45.8849 | 0.20502 | 2.7166 | -72.586 | -3.1222
10 |  | WOOD5-010-1 | -6.808 | -80.3743 |  | 9.5889 |  | 0.076151 | 1.3805
11 |  | WOOD5-1020-3 | -10.2643 | -94.9168 |  |  |  | 0.1076 | 0.81593
12 |  | WOOD5-D-5 | -1.5671 | -101.2636 | -8.1065 | 563.4469 | 49.7806 | 0.067079 | 1.3862 | -41.5102 | 10.9718
13 |  | EV | -9.8082 | -72.0607 |  |  |  | 0.052284 | 2.5547
14 |  | WOOD5-D-1 | -9.7568 | -102.5235 |  | 8.5136 |  | 0.013856 | 1.7241
15 |  | WOOD5-B-3 | 3.2514 | -78.5197 | -9.2572 | 1005.7589 | 54.7435 | 0.05526 | 0.37352 | -60.0472 | -0.48215
16 |  | WOOD5-B-4 | -1.4081 | -81.0506 | -5.9982 | 563.2294 | 16.6108 | 0.042555 | 0.58021 | -38.5312 | -5.3537
17 |  | WOOD5-B-1 | -7.8246 | -91.6904 | -1.0305 | 29.8851 |  | 0.018484 | 0.17665 |  | 0.7311
18 |  | WOOD5-C-1 | -6.2778 | -71.8974 | -0.962 | 585.3554 |  | 0.084306 | 1.3638 | -21.9902 | -8.4907
19 |  | WOOD5-A-4 | -3.0997 | -74.7198 | -0.90725 | 95.2529 |  | 0.11486 | 0.12144 | -5.579 | -3.0447
20 |  | WOOD5-1020-2 | -6.44 | -78.9583 |  |  |  | 0.088977 | 0.4055
21 |  | WOOD5-E-5 | -2.4407 | -58.518 |  | 72.0639 |  | 0.097141 | 1.3926 |  | -5.1752
22 |  | EV | -9.5589 | -68.8998 |  |  |  | 0.15184 | 0.49448
23 |  | WOOD5-2030-2 | -9.1733 | -87.2851 |  |  |  | 0.066646 | 0.73148
24 |  | WOOD5-3040-4 | -10.6931 | -93.7959 |  | 4.4081 |  | 0.059601 | 0.26679
25 |  | WOOD5-E-1 | -3.6702 | -61.6879 |  | 57.8456 |  | 0.18546 | 1.7773 |  | -4.5444
26 |  | WOOD5-1020-1 | -7.4739 | -82.3185 |  | 35.2569 |  | 0.092056 | 1.1962 |  | -6.3864
27 |  | WOOD5-C-5 | -7.5726 | -77.244 |  | 5.7164 |  | 0.04588 | 0.39461
28 |  | WOOD5-E-3 | -3.1413 | -59.1648 |  | 39.8064 |  | 0.16241 | 1.3488 |  | -4.7372
29 |  | WOOD5-E-2 | -3.2294 | -55.0734 |  | 7.0296 |  | 0.13194 | 0.69356
30 |  | WOOD5-A-5 | -4.8147 | -70.9398 |  | 46.3786 |  | 0.070088 | 0.98774 |  | -8.6097
31 |  | EV | -9.6267 | -69.4985 |  |  |  | 0.12315 | 0.13231
32 |  | WOOD5-A-3 | -5.5269 | -67.9635 |  | 28.9446 |  | 0.11163 | 0.18866 |  | -4.7754
33 |  | WOOD5-deep-3 | -14.6735 | -112.2576 |  |  |  | 0.21541 | 2.626
34 |  | WOOD5-010-2 | -0.64728 | -59.9723 |  |  |  | 0.38647 | 3.061
35 |  | WOOD5-E-4 | -2.7998 | -58.6706 |  | 106.8649 |  | 0.065546 | 0.22901 |  | -3.4427
36 |  | WOOD5-A-1 | -4.2431 | -62.1202 |  | 42.6289 |  | 0.030692 | 0.23779 |  | -2.9924
37 |  | WOOD5-deep-2 | -6.941 | -85.9924 | -0.852 | 162.7306 |  | 0.075796 | 1.3128 | -7.2255 | -4.8097
38 |  | WOOD5-deep-1 | -17.6185 | -126.5264 |  |  |  | 0.26501 | 2.2918
39 |  | WOOD5-3040-2 | -6.6314 | -84.3414 |  | 102.1989 |  | 0.30547 | 2.4565 |  | -2.7622
40 |  | EV | -9.7103 | -70.4553 |  |  |  | 0.049383 | 0.91158
41 |  | WOOD5-A-2 | -4.8747 | -66.3602 |  | 17.2116 |  | 0.28851 | 1.0041
42 |  | WOOD5-C-4 | -2.6173 | -59.6638 |  |  |  | 0.018892 | 0.30483
43 |  | WOOD5-2030-3 | -13.3913 | -107.404 |  |  |  | 0.29166 | 2.6985
44 | * | WOOD5-3040-3 | -7.5767 | -59.8927 |  |  |  | 6.4933 | 58.8708
45 | * | WOOD5-010-3 | -1.4949 | -8.8108 |  |  |  | 0 | 0
46 | * | WOOD5-C-5 | -3.853 | -35.0835 |  |  |  | 3.7764 | 45.4178
47 | * | SP-21-775 | 0.92266 | -3.9948 | -0.885 |  |  | 1.9363 | 4.1942
48 |  | SP-21-787 | 1.4133 | 12.4823 |  |  |  | 0.43451 | 3.481
49 | * | EV | -1.4949 | -8.8108 |  |  |  | 0 | 0
50 | * | SP-21-625 | -0.36581 | -11.871 |  |  |  | 0.90424 | 2.6477
51 | * | SP-21-623 | -0.074513 | -9.5212 |  |  |  | 1.1452 | 1.4554
52 | * | SP-21-771 | -0.31137 | -8.4003 |  |  |  | 1.8954 | 0.70956
53 | * | SP-21-789 | 0.76909 | -5.3575 |  |  |  | 1.822 | 3.2995
54 | * | SP-21-778 | 1.179 | -4.1652 | -1.1297 |  |  | 2.1503 | 4.271
55 | * | SP-21-839 | -0.82618 | -3.5229 |  |  |  | 1.071 | 9.1412
56 |  | SP-21-773 | 3.801 | 10.9133 | -2.1943 | 231.27 | 71.9546 | 0.065041 | 0.7254 | -14.277 | 73.7091
57 |  | SP-21-772 | 2.4325 | -1.5912 | -1.254 |  |  | 0.49539 | 3.7068
58 | * | EV | -1.4949 | -8.8108 |  |  |  | 0 | 0
59 | * | SP-21-626 | -0.43899 | -14.1391 |  |  |  | 0.86275 | 4.8189
60 | * | SP-21-622 | -0.042428 | -11.9338 |  |  |  | 1.168 | 2.8511
61 | * | SP-21-624 | -0.64701 | -8.3686 |  |  |  | 1.3579 | 0.76441
62 |  | SP-21-791 | 0.86318 | -7.0507 |  |  |  | 0.30275 | 2.5946
63 |  | SP-21-774 | -0.052883 | -8.382 |  |  |  | 0.25537 | 2.3241
64 | * | EV | -5.4409 | -35.7513 |  |  |  | 4.2302 | 31.2601

## Detail
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
1 |  | EV | -6.107 | -54.516 | -8.0999 | -63.4323 | 0.609 | 1529.77 | 2022/11/18 12:44:28
1 |  | EV | -6.837 | -57.248 | -8.8894 | -66.1696 | 0.33 | 1385.05 | 2022/11/18 12:53:25
1 |  | EV | -7.245 | -58.741 | -9.3307 | -67.6655 | 0.443 | 1331.37 | 2022/11/18 13:04:11
1 |  | EV | -7.129 | -58.353 | -9.2052 | -67.2767 | 0.41 | 1340.422 | 2022/11/18 13:16:01
1 |  | EV | -7.203 | -58.641 | -9.2852 | -67.5653 | 0.478 | 1331.35 | 2022/11/18 13:26:08
1 |  | EV | -7.273 | -58.81 | -9.3609 | -67.7346 | 0.316 | 1307.86 | 2022/11/18 13:35:05
1 |  | EV | -7.395 | -59 | -9.4929 | -67.925 | 0.428 | 1301.865 | 2022/11/18 13:44:03
1 |  | EV | -7.349 | -58.974 | -9.4431 | -67.8989 | 0.675 | 1300.948 | 2022/11/18 13:53:00
1 |  | EV | -7.42 | -59.346 | -9.5199 | -68.2716 | 0.519 | 1298.307 | 2022/11/18 14:01:57
1 |  | EV | -7.411 | -59.315 | -9.5102 | -68.2406 | 0.353 | 1297.452 | 2022/11/18 14:10:54
Sample 1 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -9.2137 | -67.218 |  | 47.446 | 4.3007 | 0.40168 | 1.4555 |  | 4.6048
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
2 |  | PZ | 4.01 | 28.804 | 2.842 | 20.0489 | 0.457 | 1298.261 | 2022/11/18 14:21:41
2 |  | PZ | 4.535 | 36.832 | 3.4098 | 28.0925 | 0.439 | 1295.929 | 2022/11/18 14:36:19
2 |  | PZ | 4.691 | 38.754 | 3.5786 | 30.0182 | 0.615 | 1295.972 | 2022/11/18 14:46:22
2 |  | PZ | 4.763 | 39.807 | 3.6564 | 31.0732 | 0.487 | 1295.894 | 2022/11/18 14:55:19
2 |  | PZ | 4.819 | 39.876 | 3.717 | 31.1424 | 0.191 | 1295.682 | 2022/11/18 15:04:15
2 |  | PZ | 4.777 | 40.238 | 3.6716 | 31.5051 | 0.216 | 1295.602 | 2022/11/18 15:13:17
2 |  | PZ | 4.806 | 40.639 | 3.7029 | 31.9068 | 0.573 | 1294.809 | 2022/11/18 15:22:13
2 |  | PZ | 4.84 | 40.948 | 3.7397 | 32.2164 | 0.504 | 1295.128 | 2022/11/18 15:31:10
2 |  | PZ | 4.859 | 40.953 | 3.7603 | 32.2215 | 0.444 | 1294.939 | 2022/11/18 15:40:08
2 |  | PZ | 4.847 | 40.984 | 3.7473 | 32.2525 | 0.214 | 1294.768 | 2022/11/18 15:49:32
2 |  | PZ | -12.113 | -92.388 | -14.5956 | -101.3776 | 0.384 | 1298.378 | 2022/11/18 15:59:34
Sample 2 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
2 |  | PZ | 1.93 | 18.1 |  |  |  | 5.0737 | 39.708
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
3 |  | UT2 | -12.906 | -105.184 | -15.4532 | -114.1983 | 0.389 | 1294.586 | 2022/11/18 16:09:23
3 |  | UT2 | -13.147 | -107.874 | -15.7139 | -116.8935 | 0.415 | 1294.372 | 2022/11/18 16:18:21
3 |  | UT2 | -13.225 | -109.468 | -15.7983 | -118.4906 | 0.564 | 1294.343 | 2022/11/18 16:27:18
3 |  | UT2 | -13.24 | -110.097 | -15.8145 | -119.1208 | 0.27 | 1293.844 | 2022/11/18 16:36:15
3 |  | UT2 | -13.302 | -110.614 | -15.8815 | -119.6388 | 0.056 | 1294.249 | 2022/11/18 16:45:12
3 |  | UT2 | -13.359 | -111.064 | -15.9432 | -120.0897 | 0.37 | 1293.664 | 2022/11/18 16:54:09
3 |  | UT2 | -13.35 | -111.175 | -15.9334 | -120.2009 | 0.472 | 1293.094 | 2022/11/18 17:03:06
3 |  | UT2 | -13.305 | -111.049 | -15.8848 | -120.0747 | 0.281 | 1293.282 | 2022/11/18 17:12:03
3 |  | UT2 | -13.351 | -111.485 | -15.9345 | -120.5115 | 0.539 | 1293.696 | 2022/11/18 17:21:00
3 |  | UT2 | -13.451 | -111.754 | -16.0427 | -120.781 | 0.345 | 1294.369 | 2022/11/18 17:29:57
Sample 3 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.1513 | 2.035
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
4 |  | EV | -8.017 | -66.554 | -10.1656 | -75.4936 | 0.088 | 1294.797 | 2022/11/18 17:39:48
4 |  | EV | -7.77 | -62.362 | -9.8985 | -71.2935 | 0.371 | 1292.458 | 2022/11/18 17:48:45
4 |  | EV | -7.704 | -61.345 | -9.8271 | -70.2745 | 0.343 | 1293.273 | 2022/11/18 17:57:42
4 |  | EV | -7.772 | -61.072 | -9.9006 | -70.001 | 0.486 | 1292.024 | 2022/11/18 18:06:39
4 |  | EV | -7.641 | -60.439 | -9.759 | -69.3667 | 0.565 | 1296.447 | 2022/11/18 18:15:36
4 |  | EV | -7.676 | -60.479 | -9.7968 | -69.4068 | 0.217 | 1293.239 | 2022/11/18 18:24:33
4 |  | EV | -7.661 | -60.432 | -9.7806 | -69.3597 | 0.326 | 1293.214 | 2022/11/18 18:33:31
4 |  | EV | -7.73 | -60.284 | -9.8552 | -69.2114 | 0.389 | 1293.26 | 2022/11/18 18:42:29
4 |  | EV | -7.625 | -60.196 | -9.7416 | -69.1233 | 0.522 | 1292.993 | 2022/11/18 18:51:26
4 |  | EV | -7.672 | -60.306 | -9.7925 | -69.2335 | 0.327 | 1293.316 | 2022/11/18 19:00:23
Sample 4 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
4 |  | EV | -9.8517 | -70.2764 |  |  |  | 0.11367 | 1.9486
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
5 |  | WOOD5-D-2 | -6.791 | -82.99 | -8.8396 | -91.9614 | -0.289 | 1326.754 | 2022/11/18 19:10:13
5 |  | WOOD5-D-2 | -6.874 | -85.569 | -8.9294 | -94.5454 | -0.478 | 1327.776 | 2022/11/18 19:19:09
5 |  | WOOD5-D-2 | -6.827 | -86.014 | -8.8786 | -94.9912 | -0.467 | 1323.925 | 2022/11/18 19:28:06
5 |  | WOOD5-D-2 | -6.816 | -86.273 | -8.8667 | -95.2507 | -0.498 | 1324.796 | 2022/11/18 19:37:04
Sample 5 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
5 |  | WOOD5-D-2 | -8.8786 | -94.1872 |  | 30.8194 |  | 0.034766 | 1.5093 |  | -3.6729
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
6 |  | WOOD5-D-3 | -7.587 | -87.678 | -9.7005 | -96.6585 | -0.313 | 1316.046 | 2022/11/18 19:46:54
6 |  | WOOD5-D-3 | -7.618 | -87.937 | -9.7341 | -96.918 | -0.463 | 1312.006 | 2022/11/18 19:55:51
6 |  | WOOD5-D-3 | -7.582 | -87.918 | -9.6951 | -96.8989 | -0.365 | 1315.475 | 2022/11/18 20:04:48
6 |  | WOOD5-D-3 | -7.568 | -87.922 | -9.68 | -96.9029 | -0.28 | 1314.16 | 2022/11/18 20:13:45
Sample 6 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
6 |  | WOOD5-D-3 | -9.7024 | -96.8446 |  | 19.4284 |  | 0.021093 | 0.1241
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
7 |  | WOOD5-B-5 | -0.831 | -84.193 | -2.3937 | -93.1667 | -6.537 | 1913.458 | 2022/11/18 20:23:36
7 |  | WOOD5-B-5 | -0.885 | -85.44 | -2.4521 | -94.4161 | -6.6 | 1920.035 | 2022/11/18 20:32:33
7 |  | WOOD5-B-5 | -0.809 | -85.029 | -2.3699 | -94.0043 | -6.584 | 1944.877 | 2022/11/18 20:41:30
7 |  | WOOD5-B-5 | -0.803 | -85.421 | -2.3634 | -94.3971 | -6.249 | 1947.341 | 2022/11/18 20:50:27
Sample 7 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
7 |  | WOOD5-B-5 | -2.3948 | -93.9961 | -6.4925 | 636.4344 | 24.888 | 0.037327 | 0.58344 | -39.5637 | -1.8467
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
8 |  | WOOD5-C-2 | 0.058 | -48.454 | -1.4322 | -57.3586 | -2.438 | 1702.282 | 2022/11/18 21:00:14
8 |  | WOOD5-C-2 | 0.34 | -43.661 | -1.1272 | -52.5563 | -2.316 | 1700.967 | 2022/11/18 21:09:11
8 |  | WOOD5-C-2 | 0.344 | -42.446 | -1.1229 | -51.3389 | -2.338 | 1706.355 | 2022/11/18 21:18:08
8 |  | WOOD5-C-2 | 0.413 | -42.307 | -1.0483 | -51.1997 | -2.212 | 1700.103 | 2022/11/18 21:27:05
Sample 8 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
8 |  | WOOD5-C-2 | -1.1826 | -53.1134 | -2.326 | 407.4334 |  | 0.15744 | 2.8894 | -19.9618 | -0.25515
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
9 |  | WOOD5-B-2 | 1.605 | -75.346 | 0.24093 | -84.3026 | -8.571 | 2700.869 | 2022/11/18 21:36:53
9 |  | WOOD5-B-2 | 1.324 | -79.685 | -0.062977 | -88.65 | -8.514 | 2731.112 | 2022/11/18 21:45:50
9 |  | WOOD5-B-2 | 1.195 | -80.885 | -0.2025 | -89.8523 | -8.288 | 2748.084 | 2022/11/18 21:54:47
9 |  | WOOD5-B-2 | 1.149 | -81.261 | -0.25225 | -90.229 | -8.843 | 2740.143 | 2022/11/18 22:03:44
Sample 9 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
9 |  | WOOD5-B-2 | -0.069196 | -88.2585 | -8.554 | 1435.0586 | 45.8849 | 0.20502 | 2.7166 | -72.586 | -3.1222
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
10 |  | WOOD5-010-1 | -5.023 | -73.465 | -6.9275 | -82.4179 | 0.048 | 1319.415 | 2022/11/18 22:13:33
10 |  | WOOD5-010-1 | -4.894 | -71.082 | -6.788 | -80.0303 | 0.422 | 1302.388 | 2022/11/18 22:22:29
10 |  | WOOD5-010-1 | -4.849 | -70.574 | -6.7393 | -79.5214 | 0.405 | 1298.882 | 2022/11/18 22:31:26
10 |  | WOOD5-010-1 | -4.884 | -70.58 | -6.7772 | -79.5274 | 0.395 | 1297.644 | 2022/11/18 22:40:24
Sample 10 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
10 |  | WOOD5-010-1 | -6.808 | -80.3743 |  | 9.5889 |  | 0.076151 | 1.3805
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
11 |  | WOOD5-1020-3 | -7.951 | -84.755 | -10.0942 | -93.7298 | 0.508 | 1293.877 | 2022/11/18 22:50:13
11 |  | WOOD5-1020-3 | -8.137 | -86.056 | -10.2954 | -95.0333 | 0.279 | 1294.331 | 2022/11/18 22:59:10
11 |  | WOOD5-1020-3 | -8.194 | -86.542 | -10.357 | -95.5203 | 0.332 | 1292.303 | 2022/11/18 23:08:07
11 |  | WOOD5-1020-3 | -8.151 | -86.406 | -10.3105 | -95.384 | 0.424 | 1295.469 | 2022/11/18 23:17:05
Sample 11 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
11 |  | WOOD5-1020-3 | -10.2643 | -94.9168 |  |  |  | 0.1076 | 0.81593
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
12 |  | WOOD5-D-5 | -0.016 | -90.222 | -1.5122 | -99.2074 | -8.062 | 1859.892 | 2022/11/18 23:26:53
12 |  | WOOD5-D-5 | -0.13 | -92.735 | -1.6355 | -101.7252 | -7.962 | 1854.173 | 2022/11/18 23:35:50
12 |  | WOOD5-D-5 | -0.119 | -93.263 | -1.6236 | -102.2543 | -8.331 | 1845.319 | 2022/11/18 23:44:46
12 |  | WOOD5-D-5 | -0.002 | -92.877 | -1.4971 | -101.8675 | -8.071 | 1874.377 | 2022/11/18 23:53:44
Sample 12 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
12 |  | WOOD5-D-5 | -1.5671 | -101.2636 | -8.1065 | 563.4469 | 49.7806 | 0.067079 | 1.3862 | -41.5102 | 10.9718
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
13 |  | EV | -7.754 | -66.852 | -9.8812 | -75.7922 | 0.086 | 1300.272 | 2022/11/19 00:03:34
13 |  | EV | -7.667 | -62.719 | -9.7871 | -71.6512 | 0.286 | 1292.517 | 2022/11/19 00:12:32
13 |  | EV | -7.63 | -61.621 | -9.7471 | -70.551 | 0.475 | 1291.372 | 2022/11/19 00:21:29
13 |  | EV | -7.695 | -61.319 | -9.8174 | -70.2484 | 0.108 | 1291.428 | 2022/11/19 00:30:26
Sample 13 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
13 |  | EV | -9.8082 | -72.0607 |  |  |  | 0.052284 | 2.5547
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
14 |  | WOOD5-D-1 | -7.643 | -91.026 | -9.7611 | -100.0129 | -0.027 | 1303.147 | 2022/11/19 00:40:16
14 |  | WOOD5-D-1 | -7.619 | -93.792 | -9.7352 | -102.7843 | -0.278 | 1303.562 | 2022/11/19 00:49:13
14 |  | WOOD5-D-1 | -7.651 | -94.497 | -9.7698 | -103.4906 | -0.328 | 1303.913 | 2022/11/19 00:58:09
14 |  | WOOD5-D-1 | -7.643 | -94.812 | -9.7611 | -103.8063 | -0.224 | 1303.406 | 2022/11/19 01:07:06
Sample 14 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
14 |  | WOOD5-D-1 | -9.7568 | -102.5235 |  | 8.5136 |  | 0.013856 | 1.7241
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
15 |  | WOOD5-B-3 | 4.339 | -70.043 | 3.1979 | -78.9893 | -8.864 | 2278.375 | 2022/11/19 01:16:53
15 |  | WOOD5-B-3 | 4.344 | -69.171 | 3.2033 | -78.1156 | -9.411 | 2313.245 | 2022/11/19 01:25:50
15 |  | WOOD5-B-3 | 4.448 | -69.672 | 3.3157 | -78.6176 | -9.461 | 2309.022 | 2022/11/19 01:34:48
15 |  | WOOD5-B-3 | 4.423 | -69.411 | 3.2887 | -78.3561 | -9.293 | 2302.367 | 2022/11/19 01:43:44
Sample 15 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
15 |  | WOOD5-B-3 | 3.2514 | -78.5197 | -9.2572 | 1005.7589 | 54.7435 | 0.05526 | 0.37352 | -60.0472 | -0.48215
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
16 |  | WOOD5-B-4 | 0.102 | -72.471 | -1.3846 | -81.422 | -6.11 | 1862.398 | 2022/11/19 01:53:32
16 |  | WOOD5-B-4 | 0.122 | -72.016 | -1.363 | -80.9661 | -5.89 | 1846.889 | 2022/11/19 02:02:29
16 |  | WOOD5-B-4 | 0.024 | -72.599 | -1.469 | -81.5503 | -5.651 | 1822.203 | 2022/11/19 02:11:26
16 |  | WOOD5-B-4 | 0.073 | -71.315 | -1.416 | -80.2638 | -6.342 | 1901.401 | 2022/11/19 02:20:24
Sample 16 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
16 |  | WOOD5-B-4 | -1.4081 | -81.0506 | -5.9982 | 563.2294 | 16.6108 | 0.042555 | 0.58021 | -38.5312 | -5.3537
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
17 |  | WOOD5-B-1 | -5.845 | -82.982 | -7.8165 | -91.9534 | -1.172 | 1329.383 | 2022/11/19 02:30:12
17 |  | WOOD5-B-1 | -5.865 | -82.614 | -7.8381 | -91.5847 | -1.129 | 1324.849 | 2022/11/19 02:39:09
17 |  | WOOD5-B-1 | -5.83 | -82.666 | -7.8003 | -91.6368 | -1.003 | 1322.708 | 2022/11/19 02:48:07
17 |  | WOOD5-B-1 | -5.87 | -82.616 | -7.8436 | -91.5867 | -0.818 | 1322.574 | 2022/11/19 02:57:04
Sample 17 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
17 |  | WOOD5-B-1 | -7.8246 | -91.6904 | -1.0305 | 29.8851 |  | 0.018484 | 0.17665 |  | 0.7311
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
18 |  | WOOD5-C-1 | -4.535 | -64.918 | -6.3997 | -73.8544 | -1.031 | 1872.703 | 2022/11/19 03:06:53
18 |  | WOOD5-C-1 | -4.437 | -62.862 | -6.2937 | -71.7944 | -0.774 | 1879.365 | 2022/11/19 03:15:50
18 |  | WOOD5-C-1 | -4.37 | -62.185 | -6.2212 | -71.1161 | -0.907 | 1883.973 | 2022/11/19 03:24:47
18 |  | WOOD5-C-1 | -4.347 | -61.894 | -6.1964 | -70.8246 | -1.136 | 1885.354 | 2022/11/19 03:33:45
Sample 18 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
18 |  | WOOD5-C-1 | -6.2778 | -71.8974 | -0.962 | 585.3554 |  | 0.084306 | 1.3638 | -21.9902 | -8.4907
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
19 |  | WOOD5-A-4 | -1.65 | -65.689 | -3.2795 | -74.6269 | -0.819 | 1386.084 | 2022/11/19 03:43:33
19 |  | WOOD5-A-4 | -1.445 | -65.814 | -3.0578 | -74.7521 | -0.893 | 1394.432 | 2022/11/19 03:52:30
19 |  | WOOD5-A-4 | -1.454 | -65.94 | -3.0675 | -74.8784 | -1.034 | 1390.266 | 2022/11/19 04:01:28
19 |  | WOOD5-A-4 | -1.386 | -65.684 | -2.9939 | -74.6219 | -0.883 | 1390.203 | 2022/11/19 04:10:25
Sample 19 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
19 |  | WOOD5-A-4 | -3.0997 | -74.7198 | -0.90725 | 95.2529 |  | 0.11486 | 0.12144 | -5.579 | -3.0447
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
20 |  | WOOD5-1020-2 | -4.465 | -69.957 | -6.324 | -78.9032 | 0.159 | 1290.226 | 2022/11/19 04:20:15
20 |  | WOOD5-1020-2 | -4.535 | -69.473 | -6.3997 | -78.4182 | 0.363 | 1295.955 | 2022/11/19 04:29:11
20 |  | WOOD5-1020-2 | -4.63 | -70.417 | -6.5024 | -79.3641 | 0.265 | 1283.429 | 2022/11/19 04:38:08
20 |  | WOOD5-1020-2 | -4.659 | -70.201 | -6.5338 | -79.1476 | 0.33 | 1284.684 | 2022/11/19 04:47:05
Sample 20 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
20 |  | WOOD5-1020-2 | -6.44 | -78.9583 |  |  |  | 0.088977 | 0.4055
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
21 |  | WOOD5-E-5 | -1 | -51.544 | -2.5765 | -60.4545 | -0.203 | 1360.134 | 2022/11/19 04:56:55
21 |  | WOOD5-E-5 | -0.902 | -49.713 | -2.4705 | -58.62 | 0.19 | 1363.129 | 2022/11/19 05:05:52
21 |  | WOOD5-E-5 | -0.791 | -48.627 | -2.3504 | -57.5319 | -0.425 | 1372.682 | 2022/11/19 05:14:49
21 |  | WOOD5-E-5 | -0.805 | -48.561 | -2.3656 | -57.4658 | 0.07 | 1372.284 | 2022/11/19 05:23:46
Sample 21 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
21 |  | WOOD5-E-5 | -2.4407 | -58.518 |  | 72.0639 |  | 0.097141 | 1.3926 |  | -5.1752
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
22 |  | EV | -7.231 | -59.237 | -9.3155 | -68.1624 | 0.314 | 1289.097 | 2022/11/19 05:33:33
22 |  | EV | -7.499 | -60.137 | -9.6054 | -69.0642 | 0.195 | 1289.904 | 2022/11/19 05:42:29
22 |  | EV | -7.539 | -60.285 | -9.6486 | -69.2124 | 0.56 | 1291.447 | 2022/11/19 05:51:25
22 |  | EV | -7.555 | -60.233 | -9.6659 | -69.1603 | 0.379 | 1289.541 | 2022/11/19 06:00:22
Sample 22 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
22 |  | EV | -9.5589 | -68.8998 |  |  |  | 0.15184 | 0.49448
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
23 |  | WOOD5-2030-2 | -7.197 | -77.284 | -9.2788 | -86.2443 | 0.182 | 1290.817 | 2022/11/19 06:10:09
23 |  | WOOD5-2030-2 | -7.052 | -78.336 | -9.1219 | -87.2984 | 0.399 | 1295.121 | 2022/11/19 06:19:06
23 |  | WOOD5-2030-2 | -7.087 | -78.826 | -9.1598 | -87.7893 | 0.584 | 1293.586 | 2022/11/19 06:28:04
23 |  | WOOD5-2030-2 | -7.062 | -78.845 | -9.1327 | -87.8084 | 0.608 | 1295.919 | 2022/11/19 06:37:01
Sample 23 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
23 |  | WOOD5-2030-2 | -9.1733 | -87.2851 |  |  |  | 0.066646 | 0.73148
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
24 |  | WOOD5-3040-4 | -8.508 | -84.543 | -10.6966 | -93.5174 | 0.262 | 1296.806 | 2022/11/19 06:46:48
24 |  | WOOD5-3040-4 | -8.421 | -84.644 | -10.6026 | -93.6186 | 0.357 | 1300.426 | 2022/11/19 06:55:45
24 |  | WOOD5-3040-4 | -8.531 | -85.075 | -10.7215 | -94.0504 | 0.424 | 1296.838 | 2022/11/19 07:04:42
24 |  | WOOD5-3040-4 | -8.559 | -85.022 | -10.7518 | -93.9973 | 0.312 | 1303.536 | 2022/11/19 07:13:39
Sample 24 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
24 |  | WOOD5-3040-4 | -10.6931 | -93.7959 |  | 4.4081 |  | 0.059601 | 0.26679
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
25 |  | WOOD5-E-1 | -2.249 | -55.355 | -3.9273 | -64.2729 | 0.211 | 1350.921 | 2022/11/19 07:23:27
25 |  | WOOD5-E-1 | -2.064 | -52.545 | -3.7272 | -61.4575 | 0.238 | 1352.895 | 2022/11/19 07:32:24
25 |  | WOOD5-E-1 | -1.895 | -51.654 | -3.5444 | -60.5647 | -0.235 | 1353.458 | 2022/11/19 07:41:21
25 |  | WOOD5-E-1 | -1.837 | -51.546 | -3.4817 | -60.4565 | -0.174 | 1354.082 | 2022/11/19 07:50:18
Sample 25 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
25 |  | WOOD5-E-1 | -3.6702 | -61.6879 |  | 57.8456 |  | 0.18546 | 1.7773 |  | -4.5444
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
26 |  | WOOD5-1020-1 | -5.395 | -71.595 | -7.3298 | -80.5443 | 0.363 | 1329.899 | 2022/11/19 08:00:05
26 |  | WOOD5-1020-1 | -5.575 | -73.689 | -7.5245 | -82.6424 | 0.115 | 1329.475 | 2022/11/19 08:09:02
26 |  | WOOD5-1020-1 | -5.542 | -74.04 | -7.4888 | -82.9941 | 0.037 | 1331.52 | 2022/11/19 08:17:59
26 |  | WOOD5-1020-1 | -5.601 | -74.139 | -7.5526 | -83.0933 | 0.346 | 1330.107 | 2022/11/19 08:26:56
Sample 26 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
26 |  | WOOD5-1020-1 | -7.4739 | -82.3185 |  | 35.2569 |  | 0.092056 | 1.1962 |  | -6.3864
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
27 |  | WOOD5-C-5 | -5.643 | -68.876 | -7.598 | -77.8201 | 0.316 | 1299.138 | 2022/11/19 08:36:45
27 |  | WOOD5-C-5 | -5.601 | -68.076 | -7.5526 | -77.0185 | 0.156 | 1300.904 | 2022/11/19 08:45:41
27 |  | WOOD5-C-5 | -5.669 | -68.237 | -7.6262 | -77.1798 | 0.318 | 1300.852 | 2022/11/19 08:54:38
27 |  | WOOD5-C-5 | -5.565 | -68.015 | -7.5137 | -76.9574 | 0.005 | 1301.945 | 2022/11/19 09:03:35
Sample 27 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
27 |  | WOOD5-C-5 | -7.5726 | -77.244 |  | 5.7164 |  | 0.04588 | 0.39461
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
28 |  | WOOD5-E-3 | -1.709 | -52.085 | -3.3433 | -60.9966 | 0.54 | 1333.791 | 2022/11/19 09:13:24
28 |  | WOOD5-E-3 | -1.544 | -50.204 | -3.1648 | -59.1119 | 0.248 | 1332.143 | 2022/11/19 09:22:21
28 |  | WOOD5-E-3 | -1.523 | -49.887 | -3.1421 | -58.7943 | -0.005 | 1334.766 | 2022/11/19 09:31:18
28 |  | WOOD5-E-3 | -1.313 | -48.851 | -2.915 | -57.7563 | 0.159 | 1338.499 | 2022/11/19 09:40:15
Sample 28 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
28 |  | WOOD5-E-3 | -3.1413 | -59.1648 |  | 39.8064 |  | 0.16241 | 1.3488 |  | -4.7372
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
29 |  | WOOD5-E-2 | -1.7 | -46.99 | -3.3335 | -55.8917 | 0.024 | 1298.079 | 2022/11/19 09:50:00
29 |  | WOOD5-E-2 | -1.411 | -45.341 | -3.021 | -54.2395 | 0.37 | 1305.066 | 2022/11/19 09:58:56
29 |  | WOOD5-E-2 | -1.628 | -45.977 | -3.2557 | -54.8768 | 0.201 | 1303.34 | 2022/11/19 10:07:54
29 |  | WOOD5-E-2 | -1.676 | -46.385 | -3.3076 | -55.2855 | 0.377 | 1301.607 | 2022/11/19 10:16:51
Sample 29 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
29 |  | WOOD5-E-2 | -3.2294 | -55.0734 |  | 7.0296 |  | 0.13194 | 0.69356
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
30 |  | WOOD5-A-5 | -2.976 | -60.542 | -4.7136 | -69.4699 | -0.282 | 1341.861 | 2022/11/19 10:26:36
30 |  | WOOD5-A-5 | -3.133 | -62.524 | -4.8834 | -71.4558 | -0.312 | 1339.109 | 2022/11/19 10:35:33
30 |  | WOOD5-A-5 | -3.057 | -62.317 | -4.8012 | -71.2484 | -0.313 | 1343.908 | 2022/11/19 10:44:30
30 |  | WOOD5-A-5 | -3.112 | -62.653 | -4.8607 | -71.585 | -0.552 | 1340.61 | 2022/11/19 10:53:27
Sample 30 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
30 |  | WOOD5-A-5 | -4.8147 | -70.9398 |  | 46.3786 |  | 0.070088 | 0.98774 |  | -8.6097
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
31 |  | EV | -7.339 | -60.525 | -9.4323 | -69.4529 | 0.147 | 1291.466 | 2022/11/19 11:03:13
31 |  | EV | -7.539 | -60.537 | -9.6486 | -69.4649 | 0.354 | 1287.895 | 2022/11/19 11:12:09
31 |  | EV | -7.604 | -60.762 | -9.7189 | -69.6904 | 0.235 | 1287.718 | 2022/11/19 11:21:06
31 |  | EV | -7.593 | -60.458 | -9.707 | -69.3858 | 0.148 | 1287.991 | 2022/11/19 11:30:03
Sample 31 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
31 |  | EV | -9.6267 | -69.4985 |  |  |  | 0.12315 | 0.13231
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
32 |  | WOOD5-A-3 | -3.893 | -59.288 | -5.7054 | -68.2135 | -0.114 | 1320.243 | 2022/11/19 11:39:50
32 |  | WOOD5-A-3 | -3.694 | -59.079 | -5.4901 | -68.0041 | 0.306 | 1325.908 | 2022/11/19 11:48:47
32 |  | WOOD5-A-3 | -3.648 | -58.873 | -5.4404 | -67.7977 | -0.024 | 1325.019 | 2022/11/19 11:57:44
32 |  | WOOD5-A-3 | -3.677 | -58.914 | -5.4717 | -67.8388 | 0.175 | 1324.582 | 2022/11/19 12:06:40
Sample 32 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
32 |  | WOOD5-A-3 | -5.5269 | -67.9635 |  | 28.9446 |  | 0.11163 | 0.18866 |  | -4.7754
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
33 |  | WOOD5-deep-3 | -11.865 | -99.378 | -14.3274 | -108.3811 | 0.252 | 1255.235 | 2022/11/19 12:16:27
33 |  | WOOD5-deep-3 | -12.25 | -103.851 | -14.7438 | -112.8627 | 0.361 | 1266.306 | 2022/11/19 12:25:24
33 |  | WOOD5-deep-3 | -12.307 | -104.771 | -14.8054 | -113.7845 | 0.189 | 1267.603 | 2022/11/19 12:34:21
33 |  | WOOD5-deep-3 | -12.318 | -104.988 | -14.8173 | -114.0019 | 0.242 | 1272.218 | 2022/11/19 12:43:19
Sample 33 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
33 |  | WOOD5-deep-3 | -14.6735 | -112.2576 |  |  |  | 0.21541 | 2.626
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
34 |  | WOOD5-010-2 | 0.217 | -55.551 | -1.2602 | -64.4693 | 0.272 | 1288.433 | 2022/11/19 12:53:05
34 |  | WOOD5-010-2 | 0.873 | -50.327 | -0.55075 | -59.2352 | 0.448 | 1290.091 | 2022/11/19 13:02:02
34 |  | WOOD5-010-2 | 0.973 | -49.624 | -0.4426 | -58.5308 | 0.508 | 1289.263 | 2022/11/19 13:11:00
34 |  | WOOD5-010-2 | 1.072 | -48.749 | -0.33552 | -57.6541 | 0.155 | 1288.617 | 2022/11/19 13:19:57
Sample 34 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
34 |  | WOOD5-010-2 | -0.64728 | -59.9723 |  |  |  | 0.38647 | 3.061
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
35 |  | WOOD5-E-4 | -1.117 | -49.937 | -2.703 | -58.8444 | -0.601 | 1397.336 | 2022/11/19 13:29:45
35 |  | WOOD5-E-4 | -1.198 | -49.531 | -2.7906 | -58.4376 | -0.378 | 1402.71 | 2022/11/19 13:38:42
35 |  | WOOD5-E-4 | -1.257 | -49.982 | -2.8544 | -58.8895 | -0.637 | 1403.791 | 2022/11/19 13:47:38
35 |  | WOOD5-E-4 | -1.254 | -49.604 | -2.8512 | -58.5108 | -0.249 | 1403.596 | 2022/11/19 13:56:36
Sample 35 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
35 |  | WOOD5-E-4 | -2.7998 | -58.6706 |  | 106.8649 |  | 0.065546 | 0.22901 |  | -3.4427
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
36 |  | WOOD5-A-1 | -2.54 | -53.09 | -4.242 | -62.0035 | -0.053 | 1336.933 | 2022/11/19 14:06:19
36 |  | WOOD5-A-1 | -2.51 | -52.959 | -4.2096 | -61.8723 | 0.168 | 1339.074 | 2022/11/19 14:15:16
36 |  | WOOD5-A-1 | -2.531 | -53.269 | -4.2323 | -62.1829 | 0.225 | 1338.24 | 2022/11/19 14:24:13
36 |  | WOOD5-A-1 | -2.583 | -53.508 | -4.2885 | -62.4223 | -0.125 | 1336.242 | 2022/11/19 14:33:09
Sample 36 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
36 |  | WOOD5-A-1 | -4.2431 | -62.1202 |  | 42.6289 |  | 0.030692 | 0.23779 |  | -2.9924
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
37 |  | WOOD5-deep-2 | -4.924 | -75.113 | -6.8204 | -84.0691 | -0.689 | 1447.826 | 2022/11/19 14:42:54
37 |  | WOOD5-deep-2 | -5.092 | -77.434 | -7.0021 | -86.3946 | -0.853 | 1459.761 | 2022/11/19 14:51:51
37 |  | WOOD5-deep-2 | -5.056 | -77.498 | -6.9632 | -86.4588 | -1.008 | 1465.028 | 2022/11/19 15:00:48
37 |  | WOOD5-deep-2 | -5.07 | -78.085 | -6.9783 | -87.0469 | -0.858 | 1458.281 | 2022/11/19 15:09:44
Sample 37 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
37 |  | WOOD5-deep-2 | -6.941 | -85.9924 | -0.852 | 162.7306 |  | 0.075796 | 1.3128 | -7.2255 | -4.8097
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
38 |  | WOOD5-deep-1 | -14.533 | -114.157 | -17.2129 | -123.1887 | 0.402 | 1290.565 | 2022/11/19 15:19:28
38 |  | WOOD5-deep-1 | -14.933 | -117.82 | -17.6455 | -126.8588 | 0.399 | 1289.615 | 2022/11/19 15:28:25
38 |  | WOOD5-deep-1 | -15.019 | -118.827 | -17.7385 | -127.8677 | 0.533 | 1291.683 | 2022/11/19 15:37:21
38 |  | WOOD5-deep-1 | -15.147 | -119.149 | -17.877 | -128.1903 | 0.388 | 1290.798 | 2022/11/19 15:46:18
Sample 38 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
38 |  | WOOD5-deep-1 | -17.6185 | -126.5264 |  |  |  | 0.26501 | 2.2918
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
39 |  | WOOD5-3040-2 | -5.197 | -78.995 | -7.1157 | -87.9586 | -0.466 | 1393.948 | 2022/11/19 15:56:03
39 |  | WOOD5-3040-2 | -4.674 | -74.782 | -6.55 | -83.7375 | -0.243 | 1396.699 | 2022/11/19 16:04:59
39 |  | WOOD5-3040-2 | -4.61 | -74.184 | -6.4808 | -83.1383 | -0.103 | 1398.926 | 2022/11/19 16:13:55
39 |  | WOOD5-3040-2 | -4.516 | -73.578 | -6.3792 | -82.5312 | -0.255 | 1399.196 | 2022/11/19 16:22:52
Sample 39 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
39 |  | WOOD5-3040-2 | -6.6314 | -84.3414 |  | 102.1989 |  | 0.30547 | 2.4565 |  | -2.7622
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
40 |  | EV | -7.527 | -62.874 | -9.6357 | -71.8065 | 0.427 | 1291.306 | 2022/11/19 16:32:38
40 |  | EV | -7.597 | -61.285 | -9.7114 | -70.2144 | -0.039 | 1288.402 | 2022/11/19 16:41:34
40 |  | EV | -7.641 | -61.009 | -9.759 | -69.9378 | 0.378 | 1287.329 | 2022/11/19 16:50:30
40 |  | EV | -7.619 | -60.934 | -9.7352 | -69.8627 | 0.608 | 1287.696 | 2022/11/19 16:59:27
Sample 40 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
40 |  | EV | -9.7103 | -70.4553 |  |  |  | 0.049383 | 0.91158
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
41 |  | WOOD5-A-2 | -3.543 | -58.795 | -5.3268 | -67.7196 | 0.174 | 1302.566 | 2022/11/19 17:09:12
41 |  | WOOD5-A-2 | -2.889 | -56.487 | -4.6195 | -65.4071 | 0.143 | 1319.928 | 2022/11/19 17:18:10
41 |  | WOOD5-A-2 | -2.997 | -56.922 | -4.7363 | -65.8429 | -0.213 | 1314.079 | 2022/11/19 17:27:06
41 |  | WOOD5-A-2 | -3.071 | -57.549 | -4.8163 | -66.4712 | 0.08 | 1312.247 | 2022/11/19 17:36:02
Sample 41 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
41 |  | WOOD5-A-2 | -4.8747 | -66.3602 |  | 17.2116 |  | 0.28851 | 1.0041
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
42 |  | WOOD5-C-4 | -1.064 | -51.207 | -2.6457 | -60.1169 | 0.024 | 1292.409 | 2022/11/19 17:45:48
42 |  | WOOD5-C-4 | -1.019 | -50.628 | -2.597 | -59.5368 | 0.164 | 1292.496 | 2022/11/19 17:54:45
42 |  | WOOD5-C-4 | -1.033 | -50.643 | -2.6122 | -59.5518 | 0.412 | 1289.644 | 2022/11/19 18:03:42
42 |  | WOOD5-C-4 | -1.035 | -50.541 | -2.6143 | -59.4496 | 0.266 | 1286.58 | 2022/11/19 18:12:38
Sample 42 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
42 |  | WOOD5-C-4 | -2.6173 | -59.6638 |  |  |  | 0.018892 | 0.30483
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
43 |  | WOOD5-2030-3 | -10.578 | -94.53 | -12.9354 | -103.5237 | 0.22 | 1285.622 | 2022/11/19 18:22:23
43 |  | WOOD5-2030-3 | -11.034 | -98.613 | -13.4286 | -107.6146 | 0.444 | 1288.21 | 2022/11/19 18:31:19
43 |  | WOOD5-2030-3 | -11.166 | -100.015 | -13.5714 | -109.0193 | 0.507 | 1287.987 | 2022/11/19 18:40:15
43 |  | WOOD5-2030-3 | -11.22 | -100.453 | -13.6298 | -109.4582 | 0.064 | 1286.854 | 2022/11/19 18:49:12
Sample 43 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
43 |  | WOOD5-2030-3 | -13.3913 | -107.404 |  |  |  | 0.29166 | 2.6985
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
44 |  | WOOD5-3040-3 | -11.302 | -102.242 | -13.7185 | -111.2506 | -0.11 | 1275.508 | 2022/11/19 18:58:57
44 |  | WOOD5-3040-3 | -11.191 | -101.691 | -13.5984 | -110.6986 | 0.332 | 1281.604 | 2022/11/19 19:07:53
44 |  | WOOD5-3040-3 |  |  | -1.4949 | -8.8108 |  |  | 2022/11/19 19:18:37
44 |  | WOOD5-3040-3 |  |  | -1.4949 | -8.8108 |  |  | 2022/11/19 19:28:39
Sample 44 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
44 | * | WOOD5-3040-3 | -7.5767 | -59.8927 |  |  |  | 6.4933 | 58.8708
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
45 |  | WOOD5-010-3 |  |  | -1.4949 | -8.8108 |  |  | 2022/11/19 19:39:30
45 |  | WOOD5-010-3 |  |  | -1.4949 | -8.8108 |  |  | 2022/11/19 19:49:31
45 |  | WOOD5-010-3 |  |  | -1.4949 | -8.8108 |  |  | 2022/11/19 19:59:32
45 |  | WOOD5-010-3 |  |  | -1.4949 | -8.8108 |  |  | 2022/11/19 20:09:32
Sample 45 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
45 | * | WOOD5-010-3 | -1.4949 | -8.8108 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
46 |  | WOOD5-C-5 | -6.541 | -78.666 | -8.5693 | -87.629 | -0.052 | 1222.732 | 2022/11/19 20:20:23
46 |  | WOOD5-C-5 |  |  | -1.4949 | -8.8108 |  |  | 2022/11/19 20:29:20
46 |  | WOOD5-C-5 |  |  | -1.4949 | -8.8108 |  |  | 2022/11/19 20:39:57
Sample 46 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
46 | * | WOOD5-C-5 | -3.853 | -35.0835 |  |  |  | 3.7764 | 45.4178
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
47 |  | SP-21-775 | 3.313 | 6.697 | 2.0882 | -2.1008 | -1.321 | 1305.658 | 2022/11/19 21:08:42
47 |  | SP-21-775 | 3.393 | 7.723 | 2.1747 | -1.0728 | -1.334 | 1306.381 | 2022/11/19 21:17:39
47 |  | SP-21-775 |  |  | -1.4949 | -8.8108 |  |  | 2022/11/19 21:26:36
Sample 47 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
47 | * | SP-21-775 | 0.92266 | -3.9948 | -0.885 |  |  | 1.9363 | 4.1942
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
48 |  | SP-21-787 | 3.091 | 24.543 | 1.8481 | 15.7797 | 0.051 | 1331.667 | 2022/11/19 21:46:23
48 |  | SP-21-787 | 2.228 | 17.608 | 0.91473 | 8.8313 | -0.187 | 1261.034 | 2022/11/19 21:55:21
48 |  | SP-21-787 | 2.748 | 21.605 | 1.4771 | 12.836 | 0.002 | 1294.977 | 2022/11/19 22:04:17
Sample 48 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
48 |  | SP-21-787 | 1.4133 | 12.4823 |  |  |  | 0.43451 | 3.481
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
49 |  | EV |  |  | -1.4949 | -8.8108 |  |  | 2022/11/19 22:14:04
49 |  | EV |  |  | -1.4949 | -8.8108 |  |  | 2022/11/19 22:24:06
49 |  | EV |  |  | -1.4949 | -8.8108 |  |  | 2022/11/19 22:34:06
49 |  | EV |  |  | -1.4949 | -8.8108 |  |  | 2022/11/19 22:44:08
Sample 49 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
49 | * | EV | -1.4949 | -8.8108 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
50 |  | SP-21-625 | 1.552 | -4.699 | 0.18361 | -13.5189 | -0.374 | 1250.482 | 2022/11/19 23:03:54
50 |  | SP-21-625 |  |  | -1.4949 | -8.8108 |  |  | 2022/11/19 23:12:51
50 |  | SP-21-625 | 1.58 | -4.464 | 0.2139 | -13.2834 | -0.479 | 1255.677 | 2022/11/19 23:22:52
Sample 50 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
50 | * | SP-21-625 | -0.36581 | -11.871 |  |  |  | 0.90424 | 2.6477
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
51 |  | SP-21-623 | 2.104 | 0.256 | 0.78062 | -8.5543 | -0.699 | 1310.603 | 2022/11/19 23:41:35
51 |  | SP-21-623 | 1.836 | -2.383 | 0.49077 | -11.1984 | -0.49 | 1280.279 | 2022/11/19 23:50:31
51 |  | SP-21-623 |  |  | -1.4949 | -8.8108 |  |  | 2022/11/19 23:59:29
Sample 51 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
51 | * | SP-21-623 | -0.074513 | -9.5212 |  |  |  | 1.1452 | 1.4554
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
52 |  | SP-21-771 |  |  | -1.4949 | -8.8108 |  |  | 2022/11/20 00:19:16
52 |  | SP-21-771 |  |  | -1.4949 | -8.8108 |  |  | 2022/11/20 00:29:17
52 |  | SP-21-771 | 3.283 | 1.229 | 2.0558 | -7.5794 | -1.18 | 1294.349 | 2022/11/20 00:39:18
Sample 52 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
52 | * | SP-21-771 | -0.31137 | -8.4003 |  |  |  | 1.8954 | 0.70956
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
53 |  | SP-21-789 | 3.322 | 6.576 | 2.0979 | -2.2221 | -1.249 | 1304.709 | 2022/11/20 00:58:02
53 |  | SP-21-789 | 2.958 | 3.764 | 1.7043 | -5.0395 | -0.732 | 1273.919 | 2022/11/20 01:06:59
53 |  | SP-21-789 |  |  | -1.4949 | -8.8108 |  |  | 2022/11/20 01:15:55
Sample 53 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
53 | * | SP-21-789 | 0.76909 | -5.3575 |  |  |  | 1.822 | 3.2995
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
54 |  | SP-21-778 | 3.907 | 8.41 | 2.7306 | -0.38452 | -1.996 | 1366.384 | 2022/11/20 01:35:43
54 |  | SP-21-778 | 3.51 | 5.5 | 2.3013 | -3.3001 | -1.393 | 1325.298 | 2022/11/20 01:44:40
54 |  | SP-21-778 |  |  | -1.4949 | -8.8108 |  |  | 2022/11/20 01:53:36
Sample 54 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
54 | * | SP-21-778 | 1.179 | -4.1652 | -1.1297 |  |  | 2.1503 | 4.271
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
55 |  | SP-21-839 |  |  | -1.4949 | -8.8108 |  |  | 2022/11/20 02:13:24
55 |  | SP-21-839 | 1.855 | 15.833 | 0.51132 | 7.0528 | 0.049 | 1269.813 | 2022/11/20 02:23:25
55 |  | SP-21-839 |  |  | -1.4949 | -8.8108 |  |  | 2022/11/20 02:32:21
Sample 55 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
55 | * | SP-21-839 | -0.82618 | -3.5229 |  |  |  | 1.071 | 9.1412
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
56 |  | SP-21-773 | 4.927 | 20.355 | 3.8338 | 11.5836 | -2.219 | 1541.346 | 2022/11/20 02:52:10
56 |  | SP-21-773 | 4.822 | 18.915 | 3.7202 | 10.1408 | -2.18 | 1512.745 | 2022/11/20 03:01:06
56 |  | SP-21-773 | 4.941 | 19.788 | 3.8489 | 11.0155 | -2.184 | 1524.699 | 2022/11/20 03:10:03
Sample 56 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
56 |  | SP-21-773 | 3.801 | 10.9133 | -2.1943 | 231.27 | 71.9546 | 0.065041 | 0.7254 | -14.277 | 73.7091
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
57 |  | SP-21-772 | 3.328 | 4.845 | 2.1044 | -3.9564 | -1.198 | 1257.073 | 2022/11/20 03:28:46
57 |  | SP-21-772 | 4.203 | 11.478 | 3.0508 | 2.6894 | -1.581 | 1345.472 | 2022/11/20 03:37:42
57 |  | SP-21-772 | 3.363 | 5.294 | 2.1423 | -3.5065 | -0.983 | 1266.989 | 2022/11/20 03:46:39
Sample 57 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
57 |  | SP-21-772 | 2.4325 | -1.5912 | -1.254 |  |  | 0.49539 | 3.7068
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
58 |  | EV |  |  | -1.4949 | -8.8108 |  |  | 2022/11/20 03:56:26
58 |  | EV |  |  | -1.4949 | -8.8108 |  |  | 2022/11/20 04:06:27
58 |  | EV |  |  | -1.4949 | -8.8108 |  |  | 2022/11/20 04:16:28
58 |  | EV |  |  | -1.4949 | -8.8108 |  |  | 2022/11/20 04:26:29
Sample 58 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
58 | * | EV | -1.4949 | -8.8108 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
59 |  | SP-21-626 | 1.636 | -6.559 | 0.27446 | -15.3825 | -0.609 | 1304.828 | 2022/11/20 04:46:18
59 |  | SP-21-626 |  |  | -1.4949 | -8.8108 |  |  | 2022/11/20 04:55:15
59 |  | SP-21-626 | 1.293 | -9.395 | -0.096505 | -18.224 | -0.793 | 1272.28 | 2022/11/20 05:05:16
Sample 59 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
59 | * | SP-21-626 | -0.43899 | -14.1391 |  |  |  | 0.86275 | 4.8189
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
60 |  | SP-21-622 | 2.122 | -3.758 | 0.80009 | -12.5761 | -0.911 | 1315.968 | 2022/11/20 05:23:59
60 |  | SP-21-622 | 1.907 | -5.593 | 0.56756 | -14.4146 | -0.764 | 1295.348 | 2022/11/20 05:32:56
60 |  | SP-21-622 |  |  | -1.4949 | -8.8108 |  |  | 2022/11/20 05:41:52
Sample 60 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
60 | * | SP-21-622 | -0.042428 | -11.9338 |  |  |  | 1.168 | 2.8511
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
61 |  | SP-21-624 |  |  | -1.4949 | -8.8108 |  |  | 2022/11/20 06:01:42
61 |  | SP-21-624 |  |  | -1.4949 | -8.8108 |  |  | 2022/11/20 06:11:42
61 |  | SP-21-624 | 2.352 | 1.324 | 1.0488 | -7.4842 | -0.495 | 1284.366 | 2022/11/20 06:21:43
Sample 61 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
61 | * | SP-21-624 | -0.64701 | -8.3686 |  |  |  | 1.3579 | 0.76441
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
62 |  | SP-21-791 | 2.47 | 4.339 | 1.1765 | -4.4634 | -0.939 | 1314.242 | 2022/11/20 06:40:28
62 |  | SP-21-791 | 2.205 | 1.781 | 0.88986 | -7.0263 | -0.478 | 1284.316 | 2022/11/20 06:49:24
62 |  | SP-21-791 | 1.866 | -0.85 | 0.52322 | -9.6624 | -0.596 | 1254.484 | 2022/11/20 06:58:21
Sample 62 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
62 |  | SP-21-791 | 0.86318 | -7.0507 |  |  |  | 0.30275 | 2.5946
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
63 |  | SP-21-774 | 1.6 | 2.792 | 0.23553 | -6.0134 | -0.477 | 1315.157 | 2022/11/20 07:17:06
63 |  | SP-21-774 | 1.309 | 0.346 | -0.0792 | -8.4641 | -0.529 | 1285.086 | 2022/11/20 07:26:02
63 |  | SP-21-774 | 1.091 | -1.854 | -0.31497 | -10.6684 | -0.421 | 1262.035 | 2022/11/20 07:34:58
Sample 63 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
63 |  | SP-21-774 | -0.052883 | -8.382 |  |  |  | 0.25537 | 2.3241
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
64 |  | EV | -6.829 | -49.326 | -8.8807 | -58.2322 | 0.096 | 1231.873 | 2022/11/20 07:44:45
64 |  | EV |  |  | -1.4949 | -8.8108 |  |  | 2022/11/20 07:53:42
64 |  | EV | -7.765 | -58.228 | -9.8931 | -67.1515 | 0.04 | 1225.951 | 2022/11/20 08:03:43
64 |  | EV |  |  | -1.4949 | -8.8108 |  |  | 2022/11/20 08:12:40
Sample 64 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
64 | * | EV | -5.4409 | -35.7513 |  |  |  | 4.2302 | 31.2601

## Source
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Line | Analysis | Time Code | Port | Inj Nr | d(18_16)Mean | d(D_H)Mean | H2O_Mean | Ignore | Good | Identifier 1 | Identifier 2 | Gas Configuration | Timestamp Mean | d(18_16)_SD | d(D_H)_SD | H2O_SD | d(18_16)_Sl | d(D_H)_Sl | H2O_Sl | baseline_shift | slope_shift | residuals | baseline_curvature | interval | ch4_ppm | h16od_adjust | h16od_shift | n2_flag | DAS Temp | Tray | Sample | Job | Method | Error Code
1 | P-2014 | 2022/11/18 12:44:28 | 2-01 | 1 | -6.107 | -54.516 | 22227 | -1 | 1 | EV | standard | H2O | 1668082758 | 0.372 | 0.773 | 74.769 | 0.002 | 0.002 | 1.727 | 1529.77 | -5.646 | 3.714 | 136.992 | 0.902 | 0.609 | 0 | 0 | 1 | 43.188 | 1 | 1 | 1 | GJB_Rinse | 0
2 | P-2014 | 2022/11/18 12:53:25 | 2-01 | 2 | -6.837 | -57.248 | 18807 | -1 | 1 | EV | standard | H2O | 1668083295 | 0.33 | 0.711 | 71.123 | 0.001 | 0.001 | 1.684 | 1385.05 | -0.911 | 2.982 | 114.069 | 0.883 | 0.33 | 0 | 0 | 1 | 42.688 | 1 | 1 | 1 | GJB_Rinse | 0
3 | P-2014 | 2022/11/18 13:04:11 | 2-01 | 3 | -7.245 | -58.741 | 18340 | -1 | 1 | EV | standard | H2O | 1668083831 | 0.318 | 0.795 | 73.41 | 0 | 0.005 | 1.746 | 1331.37 | 0.313 | 2.92 | 113 | 0.874 | 0.443 | 0 | 0 | 1 | 42.438 | 1 | 1 | 1 | GJB_Rinse | 0
4 | P-2014 | 2022/11/18 13:16:01 | 2-01 | 4 | -7.129 | -58.353 | 18215 | 0 | 1 | EV | standard | H2O | 1668084368 | 0.3 | 0.704 | 73.051 | 0.002 | 0.002 | 1.73 | 1340.422 | 0.77 | 2.879 | 111.762 | 0.873 | 0.41 | 0 | 0 | 1 | 42.312 | 1 | 1 | 1 | GJB_Rinse | 0
5 | P-2014 | 2022/11/18 13:26:08 | 2-01 | 5 | -7.203 | -58.641 | 18454 | 0 | 1 | EV | standard | H2O | 1668084905 | 0.308 | 0.71 | 74.625 | 0.001 | 0.002 | 1.774 | 1331.35 | 1.143 | 2.916 | 113.4 | 0.875 | 0.478 | 0 | 0 | 1 | 42.25 | 1 | 1 | 1 | GJB_Rinse | 0
6 | P-2014 | 2022/11/18 13:35:05 | 2-01 | 6 | -7.273 | -58.81 | 18238 | 0 | 1 | EV | standard | H2O | 1668085442 | 0.295 | 0.636 | 77.308 | 0 | 0.001 | 1.833 | 1307.86 | 1.578 | 2.917 | 112.223 | 0.877 | 0.316 | 0 | 0 | 1 | 42.25 | 1 | 1 | 1 | GJB_Rinse | 0
7 | P-2014 | 2022/11/18 13:44:03 | 2-01 | 7 | -7.395 | -59 | 18226 | 0 | 1 | EV | standard | H2O | 1668085979 | 0.29 | 0.738 | 73.803 | 0.001 | 0.004 | 1.76 | 1301.865 | 1.536 | 2.927 | 112.791 | 0.873 | 0.428 | 0 | 0 | 1 | 42.25 | 1 | 1 | 1 | GJB_Rinse | 0
8 | P-2014 | 2022/11/18 13:53:00 | 2-01 | 8 | -7.349 | -58.974 | 18324 | 0 | 1 | EV | standard | H2O | 1668086515 | 0.326 | 0.669 | 77.66 | 0 | 0.002 | 1.841 | 1300.948 | 1.257 | 2.972 | 114.387 | 0.877 | 0.675 | 0 | 0 | 1 | 42.125 | 1 | 1 | 1 | GJB_Rinse | 0
9 | P-2014 | 2022/11/18 14:01:57 | 2-01 | 9 | -7.42 | -59.346 | 18176 | 0 | 1 | EV | standard | H2O | 1668087052 | 0.302 | 0.699 | 73.854 | 0.001 | 0 | 1.76 | 1298.307 | 1.558 | 2.921 | 112.728 | 0.879 | 0.519 | 0 | 0 | 1 | 42.125 | 1 | 1 | 1 | GJB_Rinse | 0
10 | P-2014 | 2022/11/18 14:10:54 | 2-01 | 10 | -7.411 | -59.315 | 18150 | 0 | 1 | EV | standard | H2O | 1668087590 | 0.294 | 0.694 | 76.153 | 0.001 | 0.002 | 1.8 | 1297.452 | 1.744 | 2.905 | 112.185 | 0.871 | 0.353 | 0 | 0 | 1 | 42.25 | 1 | 1 | 1 | GJB_Rinse | 0
11 | P-2015 | 2022/11/18 14:21:41 | 2-02 | 1 | 4.01 | 28.804 | 18477 | -1 | 1 | PZ | standard | H2O | 1668088177 | 0.289 | 0.751 | 76.323 | -0.001 | -0.002 | 1.81 | 1298.261 | 1.361 | 3.013 | 114.914 | 0.873 | 0.457 | 0 | 0 | 1 | 42.25 | 1 | 2 | 1 | GJB_Rinse | 0
12 | P-2015 | 2022/11/18 14:36:19 | 2-02 | 2 | 4.535 | 36.832 | 18069 | -1 | 1 | PZ | standard | H2O | 1668088715 | 0.299 | 0.771 | 77.392 | 0.001 | -0.005 | 1.852 | 1295.929 | 1.712 | 2.932 | 111.635 | 0.875 | 0.439 | 0 | 0 | 1 | 42.25 | 1 | 2 | 1 | GJB_Rinse | 0
13 | P-2015 | 2022/11/18 14:46:22 | 2-02 | 3 | 4.691 | 38.754 | 18082 | -1 | 1 | PZ | standard | H2O | 1668089252 | 0.291 | 0.752 | 73.247 | 0 | -0.005 | 1.746 | 1295.972 | 1.425 | 2.942 | 113.105 | 0.872 | 0.615 | 0 | 0 | 1 | 42.312 | 1 | 2 | 1 | GJB_Rinse | 0
14 | P-2015 | 2022/11/18 14:55:19 | 2-02 | 4 | 4.763 | 39.807 | 18019 | 0 | 1 | PZ | standard | H2O | 1668089789 | 0.304 | 0.68 | 74.507 | 0 | -0.001 | 1.773 | 1295.894 | 0.77 | 2.958 | 112.029 | 0.879 | 0.487 | 0 | 0 | 1 | 42.375 | 1 | 2 | 1 | GJB_Rinse | 0
15 | P-2015 | 2022/11/18 15:04:15 | 2-02 | 5 | 4.819 | 39.876 | 18002 | 0 | 1 | PZ | standard | H2O | 1668090325 | 0.314 | 0.807 | 72.523 | 0 | -0.004 | 1.739 | 1295.682 | 1.247 | 2.861 | 110.706 | 0.879 | 0.191 | 0 | 0 | 1 | 42.25 | 1 | 2 | 1 | GJB_Rinse | 0
16 | P-2015 | 2022/11/18 15:13:17 | 2-02 | 6 | 4.777 | 40.238 | 18009 | 0 | 1 | PZ | standard | H2O | 1668090862 | 0.304 | 0.734 | 75.412 | 0.001 | -0.002 | 1.779 | 1295.602 | 1.115 | 2.907 | 111.185 | 0.877 | 0.216 | 0 | 0 | 1 | 42.375 | 1 | 2 | 1 | GJB_Rinse | 0
17 | P-2015 | 2022/11/18 15:22:13 | 2-02 | 7 | 4.806 | 40.639 | 17957 | 0 | 1 | PZ | standard | H2O | 1668091398 | 0.32 | 0.755 | 76.879 | 0 | -0.002 | 1.833 | 1294.809 | 1.361 | 2.96 | 111.932 | 0.878 | 0.573 | 0 | 0 | 1 | 42.188 | 1 | 2 | 1 | GJB_Rinse | 0
18 | P-2015 | 2022/11/18 15:31:10 | 2-02 | 8 | 4.84 | 40.948 | 17979 | 0 | 1 | PZ | standard | H2O | 1668091935 | 0.32 | 0.714 | 76.717 | -0.001 | -0.003 | 1.827 | 1295.128 | 1.024 | 2.947 | 111.965 | 0.879 | 0.504 | 0 | 0 | 1 | 42.125 | 1 | 2 | 1 | GJB_Rinse | 0
19 | P-2015 | 2022/11/18 15:40:08 | 2-02 | 9 | 4.859 | 40.953 | 17934 | 0 | 1 | PZ | standard | H2O | 1668092472 | 0.309 | 0.834 | 71.869 | 0 | -0.002 | 1.725 | 1294.939 | 1.253 | 2.931 | 110.375 | 0.878 | 0.444 | 0 | 0 | 1 | 42.188 | 1 | 2 | 1 | GJB_Rinse | 0
20 | P-2015 | 2022/11/18 15:49:32 | 2-02 | 10 | 4.847 | 40.984 | 17909 | 0 | 1 | PZ | standard | H2O | 1668093009 | 0.36 | 0.737 | 76.989 | 0.001 | 0.002 | 1.818 | 1294.768 | 0.979 | 2.882 | 110.705 | 0.875 | 0.214 | 0 | 0 | 1 | 42.125 | 1 | 2 | 1 | GJB_Rinse | 0
21 | P-2015 | 2022/11/18 15:59:34 | 2-02 | 11 | -12.113 | -92.388 | 18531 | -1 | 1 | PZ | standard | H2O | 1668093598 | 0.304 | 0.924 | 76.156 | 0.001 | 0.012 | 1.809 | 1298.378 | 2.904 | 2.923 | 112.527 | 0.879 | 0.384 | 0 | 0 | 1 | 42.125 | 1 | 2 | 1 | GJB_Rinse | 0
22 | P-2016 | 2022/11/18 16:09:23 | 2-03 | 1 | -12.906 | -105.184 | 17941 | -1 | 1 | UT2 | standard | H2O | 1668094135 | 0.324 | 0.86 | 75.918 | 0 | 0.011 | 1.815 | 1294.586 | 2.354 | 2.847 | 109.459 | 0.876 | 0.389 | 0 | 0 | 1 | 42.188 | 1 | 3 | 1 | GJB_Rinse | 0
23 | P-2016 | 2022/11/18 16:18:21 | 2-03 | 2 | -13.147 | -107.874 | 17914 | -1 | 1 | UT2 | standard | H2O | 1668094673 | 0.322 | 0.792 | 76.167 | 0.001 | 0.008 | 1.816 | 1294.372 | 2.171 | 2.839 | 109.774 | 0.879 | 0.415 | 0 | 0 | 1 | 42.188 | 1 | 3 | 1 | GJB_Rinse | 0
24 | P-2016 | 2022/11/18 16:27:18 | 2-03 | 3 | -13.225 | -109.468 | 17914 | 0 | 1 | UT2 | standard | H2O | 1668095210 | 0.305 | 0.802 | 75.49 | 0 | 0.009 | 1.794 | 1294.343 | 1.88 | 2.832 | 110.598 | 0.876 | 0.564 | 0 | 0 | 1 | 42.062 | 1 | 3 | 1 | GJB_Rinse | 0
25 | P-2016 | 2022/11/18 16:36:15 | 2-03 | 4 | -13.24 | -110.097 | 17852 | 0 | 1 | UT2 | standard | H2O | 1668095746 | 0.344 | 0.875 | 76.811 | 0.002 | 0.01 | 1.825 | 1293.844 | 2.634 | 2.808 | 107.877 | 0.876 | 0.27 | 0 | 0 | 1 | 42.062 | 1 | 3 | 1 | GJB_Rinse | 0
26 | P-2016 | 2022/11/18 16:45:12 | 2-03 | 5 | -13.302 | -110.614 | 17915 | 0 | 1 | UT2 | standard | H2O | 1668096283 | 0.322 | 0.755 | 73.05 | 0.001 | 0.005 | 1.751 | 1294.249 | 2.722 | 2.818 | 108.157 | 0.868 | 0.056 | 0 | 0 | 1 | 42.125 | 1 | 3 | 1 | GJB_Rinse | 0
27 | P-2016 | 2022/11/18 16:54:09 | 2-03 | 6 | -13.359 | -111.064 | 17825 | 0 | 1 | UT2 | standard | H2O | 1668096820 | 0.287 | 0.757 | 75.81 | 0 | 0.005 | 1.813 | 1293.664 | 1.883 | 2.824 | 109.897 | 0.871 | 0.37 | 0 | 0 | 1 | 42.125 | 1 | 3 | 1 | GJB_Rinse | 0
28 | P-2016 | 2022/11/18 17:03:06 | 2-03 | 7 | -13.35 | -111.175 | 17755 | 0 | 1 | UT2 | standard | H2O | 1668097357 | 0.34 | 0.787 | 76.493 | 0 | 0.006 | 1.823 | 1293.094 | 1.895 | 2.81 | 109.908 | 0.874 | 0.472 | 0 | 0 | 1 | 42.062 | 1 | 3 | 1 | GJB_Rinse | 0
29 | P-2016 | 2022/11/18 17:12:03 | 2-03 | 8 | -13.305 | -111.049 | 17752 | 0 | 1 | UT2 | standard | H2O | 1668097894 | 0.33 | 0.823 | 74.939 | 0.003 | 0.008 | 1.79 | 1293.282 | 1.962 | 2.776 | 108.05 | 0.876 | 0.281 | 0 | 0 | 1 | 42.188 | 1 | 3 | 1 | GJB_Rinse | 0
30 | P-2016 | 2022/11/18 17:21:00 | 2-03 | 9 | -13.351 | -111.485 | 17868 | 0 | 1 | UT2 | standard | H2O | 1668098432 | 0.342 | 0.7 | 77.125 | 0 | 0.006 | 1.832 | 1293.696 | 2.001 | 2.868 | 110.455 | 0.876 | 0.539 | 0 | 0 | 1 | 42.25 | 1 | 3 | 1 | GJB_Rinse | 0
31 | P-2016 | 2022/11/18 17:29:57 | 2-03 | 10 | -13.451 | -111.754 | 17924 | 0 | 1 | UT2 | standard | H2O | 1668792904 | 0.321 | 0.751 | 74.309 | 0.002 | 0.005 | 1.767 | 1294.369 | 1.845 | 2.831 | 110.474 | 0.87 | 0.345 | 0 | 0 | 1 | 42.188 | 1 | 3 | 1 | GJB_Rinse | 0
32 | P-2017 | 2022/11/18 17:39:48 | 2-04 | 1 | -8.017 | -66.554 | 17994 | -1 | 1 | EV | standard | H2O | 1668793494 | 0.338 | 0.661 | 74.596 | 0.002 | -0.002 | 1.784 | 1294.797 | 2.18 | 2.836 | 109.571 | 0.876 | 0.088 | 0 | 0 | 1 | 42.312 | 1 | 4 | 1 | GJB_Rinse | 0
33 | P-2017 | 2022/11/18 17:48:45 | 2-04 | 2 | -7.77 | -62.362 | 17645 | -1 | 1 | EV | standard | H2O | 1668794031 | 0.311 | 0.722 | 74.623 | 0.001 | 0.001 | 1.778 | 1292.458 | 1.495 | 2.811 | 108.484 | 0.879 | 0.371 | 0 | 0 | 1 | 42.312 | 1 | 4 | 1 | GJB_Rinse | 0
34 | P-2017 | 2022/11/18 17:57:42 | 2-04 | 3 | -7.704 | -61.345 | 17774 | -1 | 1 | EV | standard | H2O | 1668794568 | 0.326 | 0.724 | 77.911 | 0.001 | 0.002 | 1.868 | 1293.273 | 2.262 | 2.835 | 108.38 | 0.879 | 0.343 | 0 | 0 | 1 | 42.312 | 1 | 4 | 1 | GJB_Rinse | 0
35 | P-2017 | 2022/11/18 18:06:39 | 2-04 | 4 | -7.772 | -61.072 | 17612 | 0 | 1 | EV | standard | H2O | 1668795105 | 0.293 | 0.773 | 73.08 | 0.001 | -0.001 | 1.749 | 1292.024 | 1.815 | 2.825 | 108.759 | 0.875 | 0.486 | 0 | 0 | 1 | 42.188 | 1 | 4 | 1 | GJB_Rinse | 0
36 | P-2017 | 2022/11/18 18:15:36 | 2-04 | 5 | -7.641 | -60.439 | 18203 | 0 | 1 | EV | standard | H2O | 1668795643 | 0.303 | 0.795 | 76.173 | 0.001 | 0.001 | 1.802 | 1296.447 | 1.473 | 2.948 | 113.201 | 0.874 | 0.565 | 0 | 0 | 1 | 42.125 | 1 | 4 | 1 | GJB_Rinse | 0
37 | P-2017 | 2022/11/18 18:24:33 | 2-04 | 6 | -7.676 | -60.479 | 17775 | 0 | 1 | EV | standard | H2O | 1668796180 | 0.305 | 0.72 | 76.902 | 0.001 | 0.003 | 1.823 | 1293.239 | 1.934 | 2.806 | 108.806 | 0.878 | 0.217 | 0 | 0 | 1 | 42.125 | 1 | 4 | 1 | GJB_Rinse | 0
38 | P-2017 | 2022/11/18 18:33:31 | 2-04 | 7 | -7.661 | -60.432 | 17742 | 0 | 1 | EV | standard | H2O | 1668796717 | 0.307 | 0.769 | 76.993 | 0.001 | 0 | 1.833 | 1293.214 | 1.632 | 2.826 | 109.183 | 0.879 | 0.326 | 0 | 0 | 1 | 42.188 | 1 | 4 | 1 | GJB_Rinse | 0
39 | P-2017 | 2022/11/18 18:42:29 | 2-04 | 8 | -7.73 | -60.284 | 17778 | 0 | 1 | EV | standard | H2O | 1668797255 | 0.302 | 0.691 | 76.555 | 0.001 | 0.002 | 1.825 | 1293.26 | 1.703 | 2.854 | 109.94 | 0.868 | 0.389 | 0 | 0 | 1 | 42.312 | 1 | 4 | 1 | GJB_Rinse | 0
40 | P-2017 | 2022/11/18 18:51:26 | 2-04 | 9 | -7.625 | -60.196 | 17760 | 0 | 1 | EV | standard | H2O | 1668797793 | 0.323 | 0.713 | 75.231 | 0.001 | 0 | 1.801 | 1292.993 | 2.304 | 2.826 | 109.283 | 0.88 | 0.522 | 0 | 0 | 1 | 42.375 | 1 | 4 | 1 | GJB_Rinse | 0
41 | P-2017 | 2022/11/18 19:00:23 | 2-04 | 10 | -7.672 | -60.306 | 17755 | 0 | 1 | EV | standard | H2O | 1668798330 | 0.302 | 0.688 | 77.37 | 0.001 | 0.003 | 1.85 | 1293.316 | 1.485 | 2.842 | 109.515 | 0.876 | 0.327 | 0 | 0 | 1 | 42.25 | 1 | 4 | 1 | GJB_Rinse | 0
42 | P-2018 | 2022/11/18 19:10:13 | 2-05 | 1 | -6.791 | -82.99 | 17300 | -1 | 1 | WOOD5-D-2 | 22-158 | H2O | 1668798919 | 0.309 | 0.725 | 76.709 | 0.001 | 0.006 | 1.837 | 1326.754 | -1.877 | 2.779 | 109.22 | 0.87 | -0.289 | 0 | 0 | 1 | 42.125 | 1 | 5 | 1 | GJB_Rinse | 0
43 | P-2018 | 2022/11/18 19:19:09 | 2-05 | 2 | -6.874 | -85.569 | 17356 | -1 | 1 | WOOD5-D-2 | 22-158 | H2O | 1668799456 | 0.316 | 0.794 | 75.731 | 0.001 | 0.005 | 1.799 | 1327.776 | -1.522 | 2.731 | 108.02 | 0.876 | -0.478 | 0 | 0 | 1 | 42.188 | 1 | 5 | 1 | GJB_Rinse | 0
44 | P-2018 | 2022/11/18 19:28:06 | 2-05 | 3 | -6.827 | -86.014 | 16906 | -1 | 1 | WOOD5-D-2 | 22-158 | H2O | 1668799993 | 0.323 | 0.835 | 75.455 | 0.001 | 0.005 | 1.799 | 1323.925 | -1.732 | 2.624 | 105.087 | 0.872 | -0.467 | 0 | 0 | 1 | 42.25 | 1 | 5 | 1 | GJB_Rinse | 0
45 | P-2018 | 2022/11/18 19:37:04 | 2-05 | 4 | -6.816 | -86.273 | 16986 | 0 | 1 | WOOD5-D-2 | 22-158 | H2O | 1668800530 | 0.306 | 0.788 | 73.863 | 0 | 0.004 | 1.763 | 1324.796 | -1.822 | 2.676 | 105.977 | 0.877 | -0.498 | 0 | 0 | 1 | 42.312 | 1 | 5 | 1 | GJB_Rinse | 0
46 | P-2019 | 2022/11/18 19:46:54 | 2-06 | 1 | -7.587 | -87.678 | 17778 | -1 | 1 | WOOD5-D-3 | 22-158 | H2O | 1668801120 | 0.305 | 0.736 | 76.058 | 0.001 | 0.004 | 1.806 | 1316.046 | -0.892 | 2.789 | 111.337 | 0.876 | -0.313 | 0 | 0 | 1 | 42.375 | 1 | 6 | 1 | GJB_Rinse | 0
47 | P-2019 | 2022/11/18 19:55:51 | 2-06 | 2 | -7.618 | -87.937 | 17297 | -1 | 1 | WOOD5-D-3 | 22-158 | H2O | 1668801657 | 0.281 | 0.84 | 76.779 | 0.001 | 0.004 | 1.83 | 1312.006 | -1.226 | 2.763 | 108.122 | 0.873 | -0.463 | 0 | 0 | 1 | 42.25 | 1 | 6 | 1 | GJB_Rinse | 0
48 | P-2019 | 2022/11/18 20:04:48 | 2-06 | 3 | -7.582 | -87.918 | 17728 | -1 | 1 | WOOD5-D-3 | 22-158 | H2O | 1668802195 | 0.298 | 0.686 | 75.229 | 0 | 0.004 | 1.781 | 1315.475 | -1.157 | 2.837 | 111.452 | 0.873 | -0.365 | 0 | 0 | 1 | 42.25 | 1 | 6 | 1 | GJB_Rinse | 0
49 | P-2019 | 2022/11/18 20:13:45 | 2-06 | 4 | -7.568 | -87.922 | 17575 | 0 | 1 | WOOD5-D-3 | 22-158 | H2O | 1668802732 | 0.346 | 0.676 | 75.836 | 0.002 | 0.004 | 1.814 | 1314.16 | -1.171 | 2.787 | 110.806 | 0.875 | -0.28 | 0 | 0 | 1 | 42.312 | 1 | 6 | 1 | GJB_Rinse | 0
50 | P-2020 | 2022/11/18 20:23:36 | 2-07 | 1 | -0.831 | -84.193 | 15947 | -1 | 1 | WOOD5-B-5 | 22-158 | H2O | 1668803323 | 0.424 | 1.171 | 72.185 | -0.001 | 0.006 | 1.739 | 1913.458 | -38.559 | 3.496 | 106.571 | 0.878 | -6.537 | 0 | 0 | 1 | 42.312 | 1 | 7 | 1 | GJB_Rinse | 0
51 | P-2020 | 2022/11/18 20:32:33 | 2-07 | 2 | -0.885 | -85.44 | 15871 | -1 | 1 | WOOD5-B-5 | 22-158 | H2O | 1668803860 | 0.36 | 1.245 | 72.55 | 0 | 0.002 | 1.747 | 1920.035 | -38.986 | 3.52 | 106.722 | 0.884 | -6.6 | 0 | 0 | 1 | 42.188 | 1 | 7 | 1 | GJB_Rinse | 0
52 | P-2020 | 2022/11/18 20:41:30 | 2-07 | 3 | -0.809 | -85.029 | 16421 | -1 | 1 | WOOD5-B-5 | 22-158 | H2O | 1668804397 | 0.386 | 1.155 | 74.94 | 0 | 0.004 | 1.796 | 1944.877 | -40.306 | 3.652 | 110.566 | 0.884 | -6.584 | 0 | 0 | 1 | 42.188 | 1 | 7 | 1 | GJB_Rinse | 0
53 | P-2020 | 2022/11/18 20:50:27 | 2-07 | 4 | -0.803 | -85.421 | 16387 | 0 | 1 | WOOD5-B-5 | 22-158 | H2O | 1668804933 | 0.401 | 1.064 | 73.985 | -0.001 | 0.003 | 1.782 | 1947.341 | -40.404 | 3.726 | 111.75 | 0.885 | -6.249 | 0 | 0 | 1 | 42.125 | 1 | 7 | 1 | GJB_Rinse | 0
54 | P-2021 | 2022/11/18 21:00:14 | 2-08 | 1 | 0.058 | -48.454 | 17391 | -1 | 1 | WOOD5-C-2 | 22-158 | H2O | 1668805521 | 0.36 | 0.984 | 77.139 | -0.001 | 0 | 1.84 | 1702.282 | -19.578 | 2.86 | 108.978 | 0.888 | -2.438 | 0 | 0 | 1 | 42.188 | 1 | 8 | 1 | GJB_Rinse | 0
55 | P-2021 | 2022/11/18 21:09:11 | 2-08 | 2 | 0.34 | -43.661 | 17462 | -1 | 1 | WOOD5-C-2 | 22-158 | H2O | 1668806058 | 0.358 | 0.898 | 76.81 | 0.001 | 0.001 | 1.83 | 1700.967 | -20.3 | 2.92 | 110.762 | 0.888 | -2.316 | 0 | 0 | 1 | 42.25 | 1 | 8 | 1 | GJB_Rinse | 0
56 | P-2021 | 2022/11/18 21:18:08 | 2-08 | 3 | 0.344 | -42.446 | 17705 | -1 | 1 | WOOD5-C-2 | 22-158 | H2O | 1668806594 | 0.363 | 0.981 | 75.024 | -0.001 | -0.003 | 1.788 | 1706.355 | -19.719 | 2.946 | 110.478 | 0.881 | -2.338 | 0 | 0 | 1 | 42.375 | 1 | 8 | 1 | GJB_Rinse | 0
57 | P-2021 | 2022/11/18 21:27:05 | 2-08 | 4 | 0.413 | -42.307 | 17473 | 0 | 1 | WOOD5-C-2 | 22-158 | H2O | 1668807131 | 0.349 | 0.946 | 76.668 | 0.001 | -0.002 | 1.833 | 1700.103 | -20.25 | 2.911 | 111.757 | 0.881 | -2.212 | 0 | 0 | 1 | 42.25 | 1 | 8 | 1 | GJB_Rinse | 0
58 | P-2022 | 2022/11/18 21:36:53 | 2-09 | 1 | 1.605 | -75.346 | 16006 | -1 | 1 | WOOD5-B-2 | 22-158 | H2O | 1668807720 | 0.522 | 1.676 | 72.38 | 0.001 | 0.006 | 1.746 | 2700.869 | -70.989 | 4.132 | 105.044 | 0.909 | -8.571 | 0 | 0 | 1 | 42.25 | 1 | 9 | 1 | GJB_Rinse | 0
59 | P-2022 | 2022/11/18 21:45:50 | 2-09 | 2 | 1.324 | -79.685 | 16410 | -1 | 1 | WOOD5-B-2 | 22-158 | H2O | 1668808257 | 0.502 | 1.777 | 71.363 | 0.001 | 0.008 | 1.709 | 2731.112 | -72.603 | 4.175 | 108.596 | 0.913 | -8.514 | 0 | 0 | 1 | 42.375 | 1 | 9 | 1 | GJB_Rinse | 0
60 | P-2022 | 2022/11/18 21:54:47 | 2-09 | 3 | 1.195 | -80.885 | 16439 | -1 | 1 | WOOD5-B-2 | 22-158 | H2O | 1668808793 | 0.521 | 1.866 | 74.637 | -0.001 | -0.003 | 1.783 | 2748.084 | -73.118 | 4.272 | 108.678 | 0.897 | -8.288 | 0 | 0 | 1 | 42.25 | 1 | 9 | 1 | GJB_Rinse | 0
61 | P-2022 | 2022/11/18 22:03:44 | 2-09 | 4 | 1.149 | -81.261 | 16308 | 0 | 1 | WOOD5-B-2 | 22-158 | H2O | 1668809331 | 0.509 | 1.814 | 73.206 | -0.002 | 0.005 | 1.74 | 2740.143 | -73.634 | 4.235 | 108.189 | 0.901 | -8.843 | 0 | 0 | 1 | 42.25 | 1 | 9 | 1 | GJB_Rinse | 0
62 | P-2023 | 2022/11/18 22:13:33 | 2-10 | 1 | -5.023 | -73.465 | 18295 | -1 | 1 | WOOD5-010-1 | 22-158 | H2O | 1668809920 | 0.315 | 0.696 | 76.345 | 0.001 | 0.001 | 1.814 | 1319.415 | 0.422 | 2.906 | 112.211 | 0.875 | 0.048 | 0 | 0 | 1 | 42.125 | 1 | 10 | 1 | GJB_Rinse | 0
63 | P-2023 | 2022/11/18 22:22:29 | 2-10 | 2 | -4.894 | -71.082 | 17689 | -1 | 1 | WOOD5-010-1 | 22-158 | H2O | 1668810456 | 0.324 | 0.78 | 75.107 | 0.002 | 0.004 | 1.795 | 1302.388 | 1.148 | 2.839 | 109.569 | 0.877 | 0.422 | 0 | 0 | 1 | 42.125 | 1 | 10 | 1 | GJB_Rinse | 0
64 | P-2023 | 2022/11/18 22:31:26 | 2-10 | 3 | -4.849 | -70.574 | 17732 | -1 | 1 | WOOD5-010-1 | 22-158 | H2O | 1668810993 | 0.322 | 0.798 | 74.927 | 0 | 0.002 | 1.788 | 1298.882 | 1.335 | 2.799 | 109.748 | 0.873 | 0.405 | 0 | 0 | 1 | 42.125 | 1 | 10 | 1 | GJB_Rinse | 0
65 | P-2023 | 2022/11/18 22:40:24 | 2-10 | 4 | -4.884 | -70.58 | 17753 | 0 | 1 | WOOD5-010-1 | 22-158 | H2O | 1668811530 | 0.315 | 0.735 | 74.482 | 0 | 0.002 | 1.786 | 1297.644 | 1.322 | 2.86 | 109.13 | 0.879 | 0.395 | 0 | 0 | 1 | 42.188 | 1 | 10 | 1 | GJB_Rinse | 0
66 | P-2024 | 2022/11/18 22:50:13 | 2-11 | 1 | -7.951 | -84.755 | 17567 | -1 | 1 | WOOD5-1020-3 | 22-158 | H2O | 1668812120 | 0.278 | 0.707 | 74.585 | 0.001 | 0.003 | 1.779 | 1293.877 | 2.063 | 2.809 | 108.587 | 0.878 | 0.508 | 0 | 0 | 1 | 42.312 | 1 | 11 | 1 | GJB_Rinse | 0
67 | P-2024 | 2022/11/18 22:59:10 | 2-11 | 2 | -8.137 | -86.056 | 17532 | -1 | 1 | WOOD5-1020-3 | 22-158 | H2O | 1668812657 | 0.315 | 0.72 | 73.38 | 0.002 | 0.003 | 1.755 | 1294.331 | 1.656 | 2.763 | 107.797 | 0.871 | 0.279 | 0 | 0 | 1 | 42.25 | 1 | 11 | 1 | GJB_Rinse | 0
68 | P-2024 | 2022/11/18 23:08:07 | 2-11 | 3 | -8.194 | -86.542 | 17307 | -1 | 1 | WOOD5-1020-3 | 22-158 | H2O | 1668813194 | 0.309 | 0.842 | 75.829 | 0 | 0.005 | 1.812 | 1292.303 | 2.053 | 2.766 | 105.503 | 0.872 | 0.332 | 0 | 0 | 1 | 42.062 | 1 | 11 | 1 | GJB_Rinse | 0
69 | P-2024 | 2022/11/18 23:17:05 | 2-11 | 4 | -8.151 | -86.406 | 17795 | 0 | 1 | WOOD5-1020-3 | 22-158 | H2O | 1668813732 | 0.319 | 0.739 | 74.994 | 0 | 0.002 | 1.78 | 1295.469 | 1.044 | 2.835 | 110.307 | 0.877 | 0.424 | 0 | 0 | 1 | 42.125 | 1 | 11 | 1 | GJB_Rinse | 0
70 | P-2025 | 2022/11/18 23:26:53 | 2-12 | 1 | -0.016 | -90.222 | 17606 | -1 | 1 | WOOD5-D-5 | 22-158 | H2O | 1668814320 | 0.403 | 1.092 | 76.224 | -0.001 | 0.001 | 1.82 | 1859.892 | -41.419 | 4.311 | 123.43 | 0.881 | -8.062 | 0 | 0 | 1 | 42.25 | 1 | 12 | 1 | GJB_Rinse | 0
71 | P-2025 | 2022/11/18 23:35:50 | 2-12 | 2 | -0.13 | -92.735 | 17244 | -1 | 1 | WOOD5-D-5 | 22-158 | H2O | 1668814856 | 0.398 | 1.19 | 74.081 | 0.001 | 0.002 | 1.757 | 1854.173 | -41.351 | 4.306 | 121.211 | 0.892 | -7.962 | 0 | 0 | 1 | 42.312 | 1 | 12 | 1 | GJB_Rinse | 0
72 | P-2025 | 2022/11/18 23:44:46 | 2-12 | 3 | -0.119 | -93.263 | 16990 | -1 | 1 | WOOD5-D-5 | 22-158 | H2O | 1668815393 | 0.385 | 1.176 | 72.087 | -0.001 | 0.002 | 1.726 | 1845.319 | -40.61 | 4.191 | 117.588 | 0.884 | -8.331 | 0 | 0 | 1 | 42.188 | 1 | 12 | 1 | GJB_Rinse | 0
73 | P-2025 | 2022/11/18 23:53:44 | 2-12 | 4 | -0.002 | -92.877 | 17677 | 0 | 1 | WOOD5-D-5 | 22-158 | H2O | 1668815931 | 0.399 | 1.175 | 75.349 | -0.001 | 0.002 | 1.787 | 1874.377 | -42.661 | 4.455 | 124.654 | 0.893 | -8.071 | 0 | 0 | 1 | 42.25 | 1 | 12 | 1 | GJB_Rinse | 0
74 | P-2026 | 2022/11/19 00:03:34 | 2-13 | 1 | -7.754 | -66.852 | 17368 | -1 | 1 | EV | standard | H2O | 1668816521 | 0.286 | 0.735 | 77.42 | -0.001 | 0.003 | 1.849 | 1300.272 | 1.027 | 2.732 | 107.095 | 0.878 | 0.086 | 0 | 0 | 1 | 42.312 | 1 | 13 | 1 | GJB_Rinse | 0
75 | P-2026 | 2022/11/19 00:12:32 | 2-13 | 2 | -7.667 | -62.719 | 17071 | -1 | 1 | EV | standard | H2O | 1668817058 | 0.309 | 0.783 | 74.249 | 0.001 | 0.001 | 1.773 | 1292.517 | 1.494 | 2.702 | 104.971 | 0.876 | 0.286 | 0 | 0 | 1 | 42.312 | 1 | 13 | 1 | GJB_Rinse | 0
76 | P-2026 | 2022/11/19 00:21:29 | 2-13 | 3 | -7.63 | -61.621 | 17163 | -1 | 1 | EV | standard | H2O | 1668817595 | 0.317 | 0.797 | 75.479 | 0 | 0.002 | 1.803 | 1291.372 | 2.155 | 2.747 | 105.25 | 0.877 | 0.475 | 0 | 0 | 1 | 42.25 | 1 | 13 | 1 | GJB_Rinse | 0
77 | P-2026 | 2022/11/19 00:30:26 | 2-13 | 4 | -7.695 | -61.319 | 17204 | 0 | 1 | EV | standard | H2O | 1668818132 | 0.311 | 0.74 | 75.615 | 0.001 | 0 | 1.812 | 1291.428 | 1.784 | 2.706 | 104.975 | 0.875 | 0.108 | 0 | 0 | 1 | 42.25 | 1 | 13 | 1 | GJB_Rinse | 0
78 | P-2027 | 2022/11/19 00:40:16 | 2-14 | 1 | -7.643 | -91.026 | 17578 | -1 | 1 | WOOD5-D-1 | 22-158 | H2O | 1668818722 | 0.307 | 0.71 | 76.672 | 0.002 | 0.007 | 1.821 | 1303.147 | -0.251 | 2.811 | 110.601 | 0.876 | -0.027 | 0 | 0 | 1 | 42.188 | 1 | 14 | 1 | GJB_Rinse | 0
79 | P-2027 | 2022/11/19 00:49:13 | 2-14 | 2 | -7.619 | -93.792 | 17557 | -1 | 1 | WOOD5-D-1 | 22-158 | H2O | 1668819259 | 0.297 | 0.804 | 75.711 | 0 | 0.003 | 1.807 | 1303.562 | -0.334 | 2.786 | 108.979 | 0.876 | -0.278 | 0 | 0 | 1 | 42.5 | 1 | 14 | 1 | GJB_Rinse | 0
80 | P-2027 | 2022/11/19 00:58:09 | 2-14 | 3 | -7.651 | -94.497 | 17615 | -1 | 1 | WOOD5-D-1 | 22-158 | H2O | 1668819796 | 0.302 | 0.745 | 76.073 | 0.001 | 0.004 | 1.829 | 1303.913 | -0.03 | 2.781 | 108.812 | 0.877 | -0.328 | 0 | 0 | 1 | 42.438 | 1 | 14 | 1 | GJB_Rinse | 0
81 | P-2027 | 2022/11/19 01:07:06 | 2-14 | 4 | -7.643 | -94.812 | 17602 | 0 | 1 | WOOD5-D-1 | 22-158 | H2O | 1668820333 | 0.333 | 0.737 | 73.242 | 0.002 | 0.001 | 1.747 | 1303.406 | 0.056 | 2.743 | 109.398 | 0.878 | -0.224 | 0 | 0 | 1 | 42.375 | 1 | 14 | 1 | GJB_Rinse | 0
82 | P-2028 | 2022/11/19 01:16:53 | 2-15 | 1 | 4.339 | -70.043 | 15678 | -1 | 1 | WOOD5-B-3 | 22-158 | H2O | 1668820920 | 0.458 | 1.618 | 72.913 | -0.001 | 0.012 | 1.753 | 2278.375 | -59.33 | 4.379 | 110.869 | 0.896 | -8.864 | 0 | 0 | 1 | 42.188 | 1 | 15 | 1 | GJB_Rinse | 0
83 | P-2028 | 2022/11/19 01:25:50 | 2-15 | 2 | 4.344 | -69.171 | 16213 | -1 | 1 | WOOD5-B-3 | 22-158 | H2O | 1668821457 | 0.517 | 1.504 | 71.949 | 0 | 0 | 1.723 | 2313.245 | -60.473 | 4.542 | 112.593 | 0.899 | -9.411 | 0 | 0 | 1 | 42.25 | 1 | 15 | 1 | GJB_Rinse | 0
84 | P-2028 | 2022/11/19 01:34:48 | 2-15 | 3 | 4.448 | -69.672 | 15884 | -1 | 1 | WOOD5-B-3 | 22-158 | H2O | 1668821995 | 0.457 | 1.509 | 70.206 | 0 | 0.007 | 1.68 | 2309.022 | -60.534 | 4.487 | 109.548 | 0.895 | -9.461 | 0 | 0 | 1 | 42.25 | 1 | 15 | 1 | GJB_Rinse | 0
85 | P-2028 | 2022/11/19 01:43:44 | 2-15 | 4 | 4.423 | -69.411 | 15792 | 0 | 1 | WOOD5-B-3 | 22-158 | H2O | 1668822531 | 0.485 | 1.416 | 70.509 | 0 | 0.001 | 1.688 | 2302.367 | -59.852 | 4.427 | 108.057 | 0.9 | -9.293 | 0 | 0 | 1 | 42.375 | 1 | 15 | 1 | GJB_Rinse | 0
86 | P-2029 | 2022/11/19 01:53:32 | 2-16 | 1 | 0.102 | -72.471 | 15598 | -1 | 1 | WOOD5-B-4 | 22-158 | H2O | 1668823118 | 0.386 | 1.197 | 69.697 | -0.001 | 0.001 | 1.686 | 1862.398 | -38.882 | 3.38 | 105.131 | 0.881 | -6.11 | 0 | 0 | 1 | 42.5 | 1 | 16 | 1 | GJB_Rinse | 0
87 | P-2029 | 2022/11/19 02:02:29 | 2-16 | 2 | 0.122 | -72.016 | 15469 | -1 | 1 | WOOD5-B-4 | 22-158 | H2O | 1668823656 | 0.368 | 1.085 | 69.184 | 0.001 | 0.004 | 1.671 | 1846.889 | -38.06 | 3.339 | 105.046 | 0.882 | -5.89 | 0 | 0 | 1 | 42.312 | 1 | 16 | 1 | GJB_Rinse | 0
88 | P-2029 | 2022/11/19 02:11:26 | 2-16 | 3 | 0.024 | -72.599 | 14769 | -1 | 1 | WOOD5-B-4 | 22-158 | H2O | 1668824193 | 0.44 | 1.35 | 69.431 | -0.001 | 0 | 1.671 | 1822.203 | -36.742 | 3.176 | 100.425 | 0.874 | -5.651 | 0 | 0 | 1 | 42.188 | 1 | 16 | 1 | GJB_Rinse | 0
89 | P-2029 | 2022/11/19 02:20:24 | 2-16 | 4 | 0.073 | -71.315 | 16535 | 0 | 1 | WOOD5-B-4 | 22-158 | H2O | 1668824730 | 0.415 | 1.108 | 72.661 | 0.001 | 0.001 | 1.738 | 1901.401 | -40.441 | 3.545 | 110.979 | 0.884 | -6.342 | 0 | 0 | 1 | 42.125 | 1 | 16 | 1 | GJB_Rinse | 0
90 | P-2030 | 2022/11/19 02:30:12 | 2-17 | 1 | -5.845 | -82.982 | 17149 | -1 | 1 | WOOD5-B-1 | 22-158 | H2O | 1668825318 | 0.295 | 0.702 | 73.608 | 0 | 0 | 1.753 | 1329.383 | -3.948 | 2.728 | 108.669 | 0.879 | -1.172 | 0 | 0 | 1 | 42.188 | 1 | 17 | 1 | GJB_Rinse | 0
91 | P-2030 | 2022/11/19 02:39:09 | 2-17 | 2 | -5.865 | -82.614 | 17491 | -1 | 1 | WOOD5-B-1 | 22-158 | H2O | 1668825856 | 0.301 | 0.692 | 73.594 | 0.001 | 0.001 | 1.755 | 1324.849 | -4.152 | 2.788 | 111.851 | 0.873 | -1.129 | 0 | 0 | 1 | 42.188 | 1 | 17 | 1 | GJB_Rinse | 0
92 | P-2030 | 2022/11/19 02:48:07 | 2-17 | 3 | -5.83 | -82.666 | 17421 | -1 | 1 | WOOD5-B-1 | 22-158 | H2O | 1668826394 | 0.329 | 0.804 | 74.666 | 0.002 | 0.002 | 1.786 | 1322.708 | -3.773 | 2.818 | 111.352 | 0.876 | -1.003 | 0 | 0 | 1 | 42.188 | 1 | 17 | 1 | GJB_Rinse | 0
93 | P-2030 | 2022/11/19 02:57:04 | 2-17 | 4 | -5.87 | -82.616 | 17563 | 0 | 1 | WOOD5-B-1 | 22-158 | H2O | 1668826931 | 0.332 | 0.751 | 72.85 | 0.001 | 0.003 | 1.735 | 1322.574 | -3.911 | 2.861 | 114.048 | 0.874 | -0.818 | 0 | 0 | 1 | 42.312 | 1 | 17 | 1 | GJB_Rinse | 0
94 | P-2031 | 2022/11/19 03:06:53 | 2-18 | 1 | -4.535 | -64.918 | 16887 | -1 | 1 | WOOD5-C-1 | 22-158 | H2O | 1668827520 | 0.415 | 1.072 | 72.644 | 0.001 | 0.002 | 1.725 | 1872.703 | -21.733 | 2.647 | 102.069 | 0.885 | -1.031 | 0 | 0 | 1 | 42.25 | 1 | 18 | 1 | GJB_Rinse | 0
95 | P-2031 | 2022/11/19 03:15:50 | 2-18 | 2 | -4.437 | -62.862 | 16887 | -1 | 1 | WOOD5-C-1 | 22-158 | H2O | 1668828057 | 0.395 | 1.014 | 73.25 | 0.001 | 0.001 | 1.762 | 1879.365 | -21.277 | 2.732 | 101.794 | 0.886 | -0.774 | 0 | 0 | 1 | 42.25 | 1 | 18 | 1 | GJB_Rinse | 0
96 | P-2031 | 2022/11/19 03:24:47 | 2-18 | 3 | -4.37 | -62.185 | 16965 | -1 | 1 | WOOD5-C-1 | 22-158 | H2O | 1668828594 | 0.394 | 1.117 | 71.718 | -0.002 | 0 | 1.723 | 1883.973 | -22.181 | 2.702 | 102.265 | 0.887 | -0.907 | 0 | 0 | 1 | 42.312 | 1 | 18 | 1 | GJB_Rinse | 0
97 | P-2031 | 2022/11/19 03:33:45 | 2-18 | 4 | -4.347 | -61.894 | 16969 | 0 | 1 | WOOD5-C-1 | 22-158 | H2O | 1668829132 | 0.347 | 1.054 | 71.276 | 0.001 | 0.004 | 1.712 | 1885.354 | -22.77 | 2.669 | 102.905 | 0.888 | -1.136 | 0 | 0 | 1 | 42.312 | 1 | 18 | 1 | GJB_Rinse | 0
98 | P-2032 | 2022/11/19 03:43:33 | 2-19 | 1 | -1.65 | -65.689 | 16473 | -1 | 1 | WOOD5-A-4 | 22-158 | H2O | 1668829719 | 0.358 | 0.837 | 72.306 | 0.001 | 0.002 | 1.733 | 1386.084 | -5.17 | 2.581 | 103.392 | 0.873 | -0.819 | 0 | 0 | 1 | 42.25 | 1 | 19 | 1 | GJB_Rinse | 0
99 | P-2032 | 2022/11/19 03:52:30 | 2-19 | 2 | -1.445 | -65.814 | 17460 | -1 | 1 | WOOD5-A-4 | 22-158 | H2O | 1668830257 | 0.346 | 0.825 | 72.313 | 0.001 | 0.002 | 1.73 | 1394.432 | -5.805 | 2.771 | 110.152 | 0.88 | -0.893 | 0 | 0 | 1 | 42.25 | 1 | 19 | 1 | GJB_Rinse | 0
100 | P-2032 | 2022/11/19 04:01:28 | 2-19 | 3 | -1.454 | -65.94 | 17239 | -1 | 1 | WOOD5-A-4 | 22-158 | H2O | 1668830794 | 0.351 | 0.776 | 75.498 | 0 | 0.001 | 1.81 | 1390.266 | -5.992 | 2.727 | 108.289 | 0.879 | -1.034 | 0 | 0 | 1 | 42.438 | 1 | 19 | 1 | GJB_Rinse | 0
101 | P-2032 | 2022/11/19 04:10:25 | 2-19 | 4 | -1.386 | -65.684 | 17291 | 0 | 1 | WOOD5-A-4 | 22-158 | H2O | 1668831332 | 0.316 | 0.748 | 76.41 | 0.001 | 0.005 | 1.822 | 1390.203 | -5.349 | 2.706 | 108.984 | 0.873 | -0.883 | 0 | 0 | 1 | 42.25 | 1 | 19 | 1 | GJB_Rinse | 0
102 | P-2033 | 2022/11/19 04:20:15 | 2-20 | 1 | -4.465 | -69.957 | 16817 | -1 | 1 | WOOD5-1020-2 | 22-158 | H2O | 1668831922 | 0.307 | 0.859 | 73.41 | 0.001 | 0.005 | 1.754 | 1290.226 | 0.985 | 2.665 | 103.681 | 0.87 | 0.159 | 0 | 0 | 1 | 42.25 | 1 | 20 | 1 | GJB_Rinse | 0
103 | P-2033 | 2022/11/19 04:29:11 | 2-20 | 2 | -4.535 | -69.473 | 17794 | -1 | 1 | WOOD5-1020-2 | 22-158 | H2O | 1668832458 | 0.319 | 0.735 | 73.435 | 0 | 0.003 | 1.739 | 1295.955 | 1.391 | 2.831 | 109.818 | 0.873 | 0.363 | 0 | 0 | 1 | 42.25 | 1 | 20 | 1 | GJB_Rinse | 0
104 | P-2033 | 2022/11/19 04:38:08 | 2-20 | 3 | -4.63 | -70.417 | 16103 | -1 | 1 | WOOD5-1020-2 | 22-158 | H2O | 1668832995 | 0.308 | 0.813 | 71.013 | 0.001 | 0.002 | 1.702 | 1283.429 | 1.191 | 2.553 | 99.815 | 0.873 | 0.265 | 0 | 0 | 1 | 42.375 | 1 | 20 | 1 | GJB_Rinse | 0
105 | P-2033 | 2022/11/19 04:47:05 | 2-20 | 4 | -4.659 | -70.201 | 16351 | 0 | 1 | WOOD5-1020-2 | 22-158 | H2O | 1668833532 | 0.299 | 0.834 | 71.826 | 0.001 | 0.003 | 1.714 | 1284.684 | 1.642 | 2.602 | 100.886 | 0.872 | 0.33 | 0 | 0 | 1 | 42.375 | 1 | 20 | 1 | GJB_Rinse | 0
106 | P-2034 | 2022/11/19 04:56:55 | 2-21 | 1 | -1 | -51.544 | 16434 | -1 | 1 | WOOD5-E-5 | 22-158 | H2O | 1668834122 | 0.352 | 0.782 | 71.688 | -0.001 | 0 | 1.715 | 1360.134 | -2.694 | 2.578 | 102.175 | 0.875 | -0.203 | 0 | 0 | 1 | 42.25 | 1 | 21 | 1 | GJB_Rinse | 0
107 | P-2034 | 2022/11/19 05:05:52 | 2-21 | 2 | -0.902 | -49.713 | 16618 | -1 | 1 | WOOD5-E-5 | 22-158 | H2O | 1668834659 | 0.305 | 0.863 | 71.125 | 0.001 | -0.001 | 1.699 | 1363.129 | -2.55 | 2.693 | 104.268 | 0.877 | 0.19 | 0 | 0 | 1 | 42.125 | 1 | 21 | 1 | GJB_Rinse | 0
108 | P-2034 | 2022/11/19 05:14:49 | 2-21 | 3 | -0.791 | -48.627 | 17369 | -1 | 1 | WOOD5-E-5 | 22-158 | H2O | 1668835196 | 0.334 | 0.734 | 73.609 | 0.001 | 0 | 1.754 | 1372.682 | -2.612 | 2.732 | 106.381 | 0.878 | -0.425 | 0 | 0 | 1 | 42.125 | 1 | 21 | 1 | GJB_Rinse | 0
109 | P-2034 | 2022/11/19 05:23:46 | 2-21 | 4 | -0.805 | -48.561 | 17358 | 0 | 1 | WOOD5-E-5 | 22-158 | H2O | 1668835733 | 0.357 | 0.789 | 74.173 | 0.001 | 0.002 | 1.769 | 1372.284 | -3.017 | 2.829 | 109.471 | 0.877 | 0.07 | 0 | 0 | 1 | 42.25 | 1 | 21 | 1 | GJB_Rinse | 0
110 | P-2035 | 2022/11/19 05:33:33 | 2-22 | 1 | -7.231 | -59.237 | 16905 | -1 | 1 | EV | standard | H2O | 1668836319 | 0.312 | 0.776 | 76.117 | 0.001 | 0.002 | 1.818 | 1289.097 | 1.991 | 2.667 | 103.532 | 0.876 | 0.314 | 0 | 0 | 1 | 42.188 | 1 | 22 | 1 | GJB_Rinse | 0
111 | P-2035 | 2022/11/19 05:42:29 | 2-22 | 2 | -7.499 | -60.137 | 17136 | -1 | 1 | EV | standard | H2O | 1668836856 | 0.346 | 0.752 | 76.399 | 0.001 | 0.005 | 1.824 | 1289.904 | 1.767 | 2.704 | 104.986 | 0.877 | 0.195 | 0 | 0 | 1 | 42.188 | 1 | 22 | 1 | GJB_Rinse | 0
112 | P-2035 | 2022/11/19 05:51:25 | 2-22 | 3 | -7.539 | -60.285 | 17407 | -1 | 1 | EV | standard | H2O | 1668837391 | 0.326 | 0.802 | 78.706 | 0.002 | 0.007 | 1.88 | 1291.447 | 1.807 | 2.785 | 108.351 | 0.877 | 0.56 | 0 | 0 | 1 | 42.312 | 1 | 22 | 1 | GJB_Rinse | 0
113 | P-2035 | 2022/11/19 06:00:22 | 2-22 | 4 | -7.555 | -60.233 | 17145 | 0 | 1 | EV | standard | H2O | 1668837928 | 0.323 | 0.713 | 72.964 | 0.001 | 0.002 | 1.745 | 1289.541 | 1.836 | 2.722 | 105.477 | 0.875 | 0.379 | 0 | 0 | 1 | 42.312 | 1 | 22 | 1 | GJB_Rinse | 0
114 | P-2036 | 2022/11/19 06:10:09 | 2-23 | 1 | -7.197 | -77.284 | 17237 | -1 | 1 | WOOD5-2030-2 | 22-158 | H2O | 1668838515 | 0.33 | 0.751 | 70.75 | 0.001 | 0.001 | 1.693 | 1290.817 | 1.704 | 2.7 | 105.957 | 0.874 | 0.182 | 0 | 0 | 1 | 42.375 | 1 | 23 | 1 | GJB_Rinse | 0
115 | P-2036 | 2022/11/19 06:19:06 | 2-23 | 2 | -7.052 | -78.336 | 17826 | -1 | 1 | WOOD5-2030-2 | 22-158 | H2O | 1668839053 | 0.316 | 0.721 | 75.384 | 0 | 0.005 | 1.795 | 1295.121 | 1.629 | 2.834 | 110.471 | 0.88 | 0.399 | 0 | 0 | 1 | 42.5 | 1 | 23 | 1 | GJB_Rinse | 0
116 | P-2036 | 2022/11/19 06:28:04 | 2-23 | 3 | -7.087 | -78.826 | 17638 | -1 | 1 | WOOD5-2030-2 | 22-158 | H2O | 1668839591 | 0.303 | 0.692 | 75.256 | 0.001 | 0 | 1.792 | 1293.586 | 1.708 | 2.844 | 109.371 | 0.879 | 0.584 | 0 | 0 | 1 | 42.562 | 1 | 23 | 1 | GJB_Rinse | 0
117 | P-2036 | 2022/11/19 06:37:01 | 2-23 | 4 | -7.062 | -78.845 | 17953 | 0 | 1 | WOOD5-2030-2 | 22-158 | H2O | 1668840128 | 0.342 | 0.732 | 74.544 | 0.001 | 0.002 | 1.765 | 1295.919 | 1.183 | 2.913 | 112.264 | 0.878 | 0.608 | 0 | 0 | 1 | 42.438 | 1 | 23 | 1 | GJB_Rinse | 0
118 | P-2037 | 2022/11/19 06:46:48 | 2-24 | 1 | -8.508 | -84.543 | 16700 | -1 | 1 | WOOD5-3040-4 | 22-158 | H2O | 1668840715 | 0.312 | 0.825 | 74.712 | 0.002 | 0.003 | 1.788 | 1296.806 | 1.414 | 2.627 | 102.84 | 0.875 | 0.262 | 0 | 0 | 1 | 42.25 | 1 | 24 | 1 | GJB_Rinse | 0
119 | P-2037 | 2022/11/19 06:55:45 | 2-24 | 2 | -8.421 | -84.644 | 17145 | -1 | 1 | WOOD5-3040-4 | 22-158 | H2O | 1668841251 | 0.319 | 0.794 | 74.945 | 0.001 | 0.006 | 1.801 | 1300.426 | 1.119 | 2.7 | 106.278 | 0.877 | 0.357 | 0 | 0 | 1 | 42.312 | 1 | 24 | 1 | GJB_Rinse | 0
120 | P-2037 | 2022/11/19 07:04:42 | 2-24 | 3 | -8.531 | -85.075 | 16689 | -1 | 1 | WOOD5-3040-4 | 22-158 | H2O | 1668841789 | 0.275 | 0.833 | 73.158 | 0.001 | 0.006 | 1.753 | 1296.838 | 1.805 | 2.663 | 102.195 | 0.875 | 0.424 | 0 | 0 | 1 | 42.25 | 1 | 24 | 1 | GJB_Rinse | 0
121 | P-2037 | 2022/11/19 07:13:39 | 2-24 | 4 | -8.559 | -85.022 | 17522 | 0 | 1 | WOOD5-3040-4 | 22-158 | H2O | 1668842325 | 0.311 | 0.764 | 75.462 | 0.001 | -0.001 | 1.805 | 1303.536 | 1.187 | 2.811 | 108.181 | 0.876 | 0.312 | 0 | 0 | 1 | 42.312 | 1 | 24 | 1 | GJB_Rinse | 0
122 | P-2038 | 2022/11/19 07:23:27 | 2-25 | 1 | -2.249 | -55.355 | 17170 | -1 | 1 | WOOD5-E-1 | 22-158 | H2O | 1668842914 | 0.321 | 0.829 | 71.051 | 0 | -0.001 | 1.7 | 1350.921 | -1.162 | 2.745 | 106.851 | 0.875 | 0.211 | 0 | 0 | 1 | 42.312 | 1 | 25 | 1 | GJB_Rinse | 0
123 | P-2038 | 2022/11/19 07:32:24 | 2-25 | 2 | -2.064 | -52.545 | 17293 | -1 | 1 | WOOD5-E-1 | 22-158 | H2O | 1668843450 | 0.288 | 0.756 | 73.609 | 0 | 0.001 | 1.765 | 1352.895 | -1.103 | 2.737 | 107.164 | 0.869 | 0.238 | 0 | 0 | 1 | 42.188 | 1 | 25 | 1 | GJB_Rinse | 0
124 | P-2038 | 2022/11/19 07:41:21 | 2-25 | 3 | -1.895 | -51.654 | 17298 | -1 | 1 | WOOD5-E-1 | 22-158 | H2O | 1668843987 | 0.338 | 0.743 | 76.63 | 0 | 0.003 | 1.834 | 1353.458 | -1.273 | 2.721 | 105.439 | 0.875 | -0.235 | 0 | 0 | 1 | 42.125 | 1 | 25 | 1 | GJB_Rinse | 0
125 | P-2038 | 2022/11/19 07:50:18 | 2-25 | 4 | -1.837 | -51.546 | 17335 | 0 | 1 | WOOD5-E-1 | 22-158 | H2O | 1668844524 | 0.326 | 0.827 | 74.082 | 0.001 | 0.004 | 1.779 | 1354.082 | -1.043 | 2.716 | 105.364 | 0.878 | -0.174 | 0 | 0 | 1 | 42.25 | 1 | 25 | 1 | GJB_Rinse | 0
126 | P-2039 | 2022/11/19 08:00:05 | 2-26 | 1 | -5.395 | -71.595 | 16965 | -1 | 1 | WOOD5-1020-1 | 22-158 | H2O | 1668845112 | 0.311 | 0.777 | 72.957 | 0.001 | 0.006 | 1.754 | 1329.899 | 0.082 | 2.714 | 104.605 | 0.876 | 0.363 | 0 | 0 | 1 | 42.125 | 1 | 26 | 1 | GJB_Rinse | 0
127 | P-2039 | 2022/11/19 08:09:02 | 2-26 | 2 | -5.575 | -73.689 | 16938 | -1 | 1 | WOOD5-1020-1 | 22-158 | H2O | 1668845649 | 0.293 | 0.722 | 72.949 | 0 | 0.004 | 1.761 | 1329.475 | 0.256 | 2.635 | 102.898 | 0.878 | 0.115 | 0 | 0 | 1 | 42.25 | 1 | 26 | 1 | GJB_Rinse | 0
128 | P-2039 | 2022/11/19 08:17:59 | 2-26 | 3 | -5.542 | -74.04 | 17144 | -1 | 1 | WOOD5-1020-1 | 22-158 | H2O | 1668846186 | 0.317 | 0.779 | 72.458 | 0.001 | 0.003 | 1.74 | 1331.52 | -0.018 | 2.69 | 104.888 | 0.877 | 0.037 | 0 | 0 | 1 | 42.375 | 1 | 26 | 1 | GJB_Rinse | 0
129 | P-2039 | 2022/11/19 08:26:56 | 2-26 | 4 | -5.601 | -74.139 | 17016 | 0 | 1 | WOOD5-1020-1 | 22-158 | H2O | 1668846722 | 0.329 | 0.791 | 75.89 | 0.001 | 0.003 | 1.814 | 1330.107 | -0.049 | 2.717 | 105.059 | 0.877 | 0.346 | 0 | 0 | 1 | 42.25 | 1 | 26 | 1 | GJB_Rinse | 0
130 | P-2040 | 2022/11/19 08:36:45 | 2-27 | 1 | -5.643 | -68.876 | 17330 | -1 | 1 | WOOD5-C-5 | 22-158 | H2O | 1668847311 | 0.303 | 0.743 | 74.792 | 0 | 0.004 | 1.792 | 1299.138 | 1.236 | 2.768 | 106.997 | 0.879 | 0.316 | 0 | 0 | 1 | 42.312 | 1 | 27 | 1 | GJB_Rinse | 0
131 | P-2040 | 2022/11/19 08:45:41 | 2-27 | 2 | -5.601 | -68.076 | 17601 | -1 | 1 | WOOD5-C-5 | 22-158 | H2O | 1668847847 | 0.302 | 0.759 | 76.088 | 0.001 | 0.004 | 1.806 | 1300.904 | 0.991 | 2.758 | 108.232 | 0.876 | 0.156 | 0 | 0 | 1 | 42.312 | 1 | 27 | 1 | GJB_Rinse | 0
132 | P-2040 | 2022/11/19 08:54:38 | 2-27 | 3 | -5.669 | -68.237 | 17639 | -1 | 1 | WOOD5-C-5 | 22-158 | H2O | 1668848386 | 0.303 | 0.77 | 74.545 | 0 | 0.003 | 1.778 | 1300.852 | 1.042 | 2.811 | 109.727 | 0.878 | 0.318 | 0 | 0 | 1 | 42.25 | 1 | 27 | 1 | GJB_Rinse | 0
133 | P-2040 | 2022/11/19 09:03:35 | 2-27 | 4 | -5.565 | -68.015 | 17744 | 0 | 1 | WOOD5-C-5 | 22-158 | H2O | 1668848922 | 0.334 | 0.809 | 75.701 | 0.001 | 0.003 | 1.802 | 1301.945 | 1.042 | 2.785 | 109.037 | 0.874 | 0.005 | 0 | 0 | 1 | 42.25 | 1 | 27 | 1 | GJB_Rinse | 0
134 | P-2041 | 2022/11/19 09:13:24 | 2-28 | 1 | -1.709 | -52.085 | 17199 | -1 | 1 | WOOD5-E-3 | 22-158 | H2O | 1668849511 | 0.29 | 0.762 | 74.876 | 0 | 0.001 | 1.786 | 1333.791 | -0.686 | 2.759 | 107.183 | 0.878 | 0.54 | 0 | 0 | 1 | 42.25 | 1 | 28 | 1 | GJB_Rinse | 0
135 | P-2041 | 2022/11/19 09:22:21 | 2-28 | 2 | -1.544 | -50.204 | 17012 | -1 | 1 | WOOD5-E-3 | 22-158 | H2O | 1668850047 | 0.337 | 0.706 | 72.639 | 0.001 | 0.002 | 1.735 | 1332.143 | 0.074 | 2.657 | 104.352 | 0.876 | 0.248 | 0 | 0 | 1 | 42.375 | 1 | 28 | 1 | GJB_Rinse | 0
136 | P-2041 | 2022/11/19 09:31:18 | 2-28 | 3 | -1.523 | -49.887 | 17210 | -1 | 1 | WOOD5-E-3 | 22-158 | H2O | 1668850584 | 0.295 | 0.756 | 72.368 | 0.001 | -0.001 | 1.729 | 1334.766 | -0.071 | 2.693 | 104.616 | 0.877 | -0.005 | 0 | 0 | 1 | 42.438 | 1 | 28 | 1 | GJB_Rinse | 0
137 | P-2041 | 2022/11/19 09:40:15 | 2-28 | 4 | -1.313 | -48.851 | 17601 | 0 | 1 | WOOD5-E-3 | 22-158 | H2O | 1668851121 | 0.323 | 0.799 | 73.267 | 0.001 | 0.002 | 1.753 | 1338.499 | -0.427 | 2.787 | 107.896 | 0.88 | 0.159 | 0 | 0 | 1 | 42.5 | 1 | 28 | 1 | GJB_Rinse | 0
138 | P-2042 | 2022/11/19 09:50:00 | 2-29 | 1 | -1.7 | -46.99 | 16538 | -1 | 1 | WOOD5-E-2 | 22-158 | H2O | 1668851706 | 0.332 | 0.848 | 69.614 | 0.001 | -0.001 | 1.665 | 1298.079 | 0.893 | 2.598 | 101.39 | 0.875 | 0.024 | 0 | 0 | 1 | 42.438 | 1 | 29 | 1 | GJB_Rinse | 0
139 | P-2042 | 2022/11/19 09:58:56 | 2-29 | 2 | -1.411 | -45.341 | 17525 | -1 | 1 | WOOD5-E-2 | 22-158 | H2O | 1668852243 | 0.296 | 0.781 | 74.585 | 0.002 | 0.002 | 1.785 | 1305.066 | 0.6 | 2.805 | 109.038 | 0.875 | 0.37 | 0 | 0 | 1 | 42.25 | 1 | 29 | 1 | GJB_Rinse | 0
140 | P-2042 | 2022/11/19 10:07:54 | 2-29 | 3 | -1.628 | -45.977 | 17253 | -1 | 1 | WOOD5-E-2 | 22-158 | H2O | 1668852780 | 0.346 | 0.702 | 73.546 | 0.001 | 0 | 1.754 | 1303.34 | 0.628 | 2.714 | 105.627 | 0.876 | 0.201 | 0 | 0 | 1 | 42.188 | 1 | 29 | 1 | GJB_Rinse | 0
141 | P-2042 | 2022/11/19 10:16:51 | 2-29 | 4 | -1.676 | -46.385 | 17101 | 0 | 1 | WOOD5-E-2 | 22-158 | H2O | 1668853318 | 0.3 | 0.715 | 73.338 | 0.001 | 0 | 1.741 | 1301.607 | 0.975 | 2.719 | 106.135 | 0.871 | 0.377 | 0 | 0 | 1 | 42.25 | 1 | 29 | 1 | GJB_Rinse | 0
142 | P-2043 | 2022/11/19 10:26:36 | 2-30 | 1 | -2.976 | -60.542 | 16543 | -1 | 1 | WOOD5-A-5 | 22-158 | H2O | 1668853903 | 0.342 | 0.783 | 73.993 | 0.001 | 0.003 | 1.771 | 1341.861 | -2.175 | 2.66 | 103.673 | 0.874 | -0.282 | 0 | 0 | 1 | 42.25 | 1 | 30 | 1 | GJB_Rinse | 0
143 | P-2043 | 2022/11/19 10:35:33 | 2-30 | 2 | -3.133 | -62.524 | 16225 | -1 | 1 | WOOD5-A-5 | 22-158 | H2O | 1668854440 | 0.321 | 0.92 | 72.038 | 0.001 | 0.002 | 1.727 | 1339.109 | -1.763 | 2.5 | 100.443 | 0.869 | -0.312 | 0 | 0 | 1 | 42.188 | 1 | 30 | 1 | GJB_Rinse | 0
144 | P-2043 | 2022/11/19 10:44:30 | 2-30 | 3 | -3.057 | -62.317 | 16651 | -1 | 1 | WOOD5-A-5 | 22-158 | H2O | 1668854977 | 0.312 | 0.795 | 73.858 | 0.001 | 0.004 | 1.779 | 1343.908 | -2.541 | 2.626 | 103.761 | 0.875 | -0.313 | 0 | 0 | 1 | 42.125 | 1 | 30 | 1 | GJB_Rinse | 0
145 | P-2043 | 2022/11/19 10:53:27 | 2-30 | 4 | -3.112 | -62.653 | 16341 | 0 | 1 | WOOD5-A-5 | 22-158 | H2O | 1668855514 | 0.327 | 0.722 | 70.642 | 0.001 | 0.004 | 1.69 | 1340.61 | -1.904 | 2.549 | 100.68 | 0.875 | -0.552 | 0 | 0 | 1 | 42.312 | 1 | 30 | 1 | GJB_Rinse | 0
146 | P-2044 | 2022/11/19 11:03:13 | 2-31 | 1 | -7.339 | -60.525 | 17346 | -1 | 1 | EV | standard | H2O | 1668856099 | 0.292 | 0.681 | 74.714 | 0.001 | 0 | 1.79 | 1291.466 | 2.31 | 2.683 | 105.678 | 0.875 | 0.147 | 0 | 0 | 1 | 42.312 | 1 | 31 | 1 | GJB_Rinse | 0
147 | P-2044 | 2022/11/19 11:12:09 | 2-31 | 2 | -7.539 | -60.537 | 16898 | -1 | 1 | EV | standard | H2O | 1668856635 | 0.307 | 0.795 | 74.697 | 0 | 0.004 | 1.786 | 1287.895 | 1.233 | 2.712 | 103.772 | 0.875 | 0.354 | 0 | 0 | 1 | 42.312 | 1 | 31 | 1 | GJB_Rinse | 0
148 | P-2044 | 2022/11/19 11:21:06 | 2-31 | 3 | -7.604 | -60.762 | 16929 | -1 | 1 | EV | standard | H2O | 1668857173 | 0.319 | 0.762 | 70.84 | 0.001 | 0.002 | 1.706 | 1287.718 | 1.94 | 2.662 | 103.366 | 0.874 | 0.235 | 0 | 0 | 1 | 42.125 | 1 | 31 | 1 | GJB_Rinse | 0
149 | P-2044 | 2022/11/19 11:30:03 | 2-31 | 4 | -7.593 | -60.458 | 16969 | 0 | 1 | EV | standard | H2O | 1668857710 | 0.331 | 0.694 | 74.622 | 0.002 | 0.005 | 1.796 | 1287.991 | 1.75 | 2.684 | 103.938 | 0.876 | 0.148 | 0 | 0 | 1 | 42.125 | 1 | 31 | 1 | GJB_Rinse | 0
150 | P-2045 | 2022/11/19 11:39:50 | 2-32 | 1 | -3.893 | -59.288 | 16812 | -1 | 1 | WOOD5-A-3 | 22-158 | H2O | 1668858296 | 0.303 | 0.799 | 71.19 | 0.002 | 0.001 | 1.706 | 1320.243 | 0.164 | 2.6 | 102.324 | 0.874 | -0.114 | 0 | 0 | 1 | 42.25 | 1 | 32 | 1 | GJB_Rinse | 0
151 | P-2045 | 2022/11/19 11:48:47 | 2-32 | 2 | -3.694 | -59.079 | 17421 | -1 | 1 | WOOD5-A-3 | 22-158 | H2O | 1668858833 | 0.304 | 0.746 | 75.731 | 0 | 0.001 | 1.798 | 1325.908 | -0.037 | 2.782 | 108.424 | 0.876 | 0.306 | 0 | 0 | 1 | 42.25 | 1 | 32 | 1 | GJB_Rinse | 0
152 | P-2045 | 2022/11/19 11:57:44 | 2-32 | 3 | -3.648 | -58.873 | 17298 | -1 | 1 | WOOD5-A-3 | 22-158 | H2O | 1668859370 | 0.313 | 0.809 | 76.114 | 0.001 | 0.004 | 1.814 | 1325.019 | -0.022 | 2.745 | 106.662 | 0.874 | -0.024 | 0 | 0 | 1 | 42.125 | 1 | 32 | 1 | GJB_Rinse | 0
153 | P-2045 | 2022/11/19 12:06:40 | 2-32 | 4 | -3.677 | -58.914 | 17232 | 0 | 1 | WOOD5-A-3 | 22-158 | H2O | 1668859907 | 0.302 | 0.751 | 74.508 | 0.001 | 0.001 | 1.777 | 1324.582 | -0.362 | 2.779 | 106.484 | 0.877 | 0.175 | 0 | 0 | 1 | 42.125 | 1 | 32 | 1 | GJB_Rinse | 0
154 | P-2046 | 2022/11/19 12:16:27 | 2-33 | 1 | -11.865 | -99.378 | 12201 | -1 | 1 | WOOD5-deep-3 | 22-158 | H2O | 1668860494 | 0.324 | 0.838 | 62.691 | 0.002 | 0.005 | 1.55 | 1255.235 | 1.764 | 1.866 | 75.614 | 0.863 | 0.252 | 0 | 0 | 1 | 42.188 | 1 | 33 | 1 | GJB_Rinse | 0
155 | P-2046 | 2022/11/19 12:25:24 | 2-33 | 2 | -12.25 | -103.851 | 13877 | -1 | 1 | WOOD5-deep-3 | 22-158 | H2O | 1668861031 | 0.307 | 0.851 | 66.09 | 0.001 | 0.004 | 1.613 | 1266.306 | 2.15 | 2.147 | 85.599 | 0.868 | 0.361 | 0 | 0 | 1 | 42.5 | 1 | 33 | 1 | GJB_Rinse | 0
156 | P-2046 | 2022/11/19 12:34:21 | 2-33 | 3 | -12.307 | -104.771 | 14047 | -1 | 1 | WOOD5-deep-3 | 22-158 | H2O | 1668861568 | 0.293 | 0.834 | 64.953 | 0 | 0.004 | 1.577 | 1267.603 | 1.906 | 2.151 | 85.734 | 0.868 | 0.189 | 0 | 0 | 1 | 42.5 | 1 | 33 | 1 | GJB_Rinse | 0
157 | P-2046 | 2022/11/19 12:43:19 | 2-33 | 4 | -12.318 | -104.988 | 14719 | 0 | 1 | WOOD5-deep-3 | 22-158 | H2O | 1668862105 | 0.298 | 0.916 | 66.458 | 0.001 | 0.008 | 1.611 | 1272.218 | 1.68 | 2.298 | 90.537 | 0.869 | 0.242 | 0 | 0 | 1 | 42.625 | 1 | 33 | 1 | GJB_Rinse | 0
158 | P-2047 | 2022/11/19 12:53:05 | 2-34 | 1 | 0.217 | -55.551 | 16990 | -1 | 1 | WOOD5-010-2 | 22-158 | H2O | 1668862691 | 0.302 | 0.811 | 70.998 | 0.001 | 0 | 1.703 | 1288.433 | 0.556 | 2.727 | 105.948 | 0.878 | 0.272 | 0 | 0 | 1 | 42.625 | 1 | 34 | 1 | GJB_Rinse | 0
159 | P-2047 | 2022/11/19 13:02:02 | 2-34 | 2 | 0.873 | -50.327 | 17269 | -1 | 1 | WOOD5-010-2 | 22-158 | H2O | 1668863229 | 0.354 | 0.732 | 72.859 | 0 | -0.001 | 1.738 | 1290.091 | 1.315 | 2.76 | 107.564 | 0.867 | 0.448 | 0 | 0 | 1 | 42.375 | 1 | 34 | 1 | GJB_Rinse | 0
160 | P-2047 | 2022/11/19 13:11:00 | 2-34 | 3 | 0.973 | -49.624 | 17171 | -1 | 1 | WOOD5-010-2 | 22-158 | H2O | 1668863767 | 0.301 | 0.799 | 74.473 | 0.001 | 0 | 1.773 | 1289.263 | 1.521 | 2.736 | 106.806 | 0.869 | 0.508 | 0 | 0 | 1 | 42.25 | 1 | 34 | 1 | GJB_Rinse | 0
161 | P-2047 | 2022/11/19 13:19:57 | 2-34 | 4 | 1.072 | -48.749 | 17063 | 0 | 1 | WOOD5-010-2 | 22-158 | H2O | 1668864303 | 0.325 | 0.685 | 71.239 | 0 | -0.003 | 1.699 | 1288.617 | 2.062 | 2.661 | 103.671 | 0.878 | 0.155 | 0 | 0 | 1 | 42.188 | 1 | 34 | 1 | GJB_Rinse | 0
162 | P-2048 | 2022/11/19 13:29:45 | 2-35 | 1 | -1.117 | -49.937 | 17032 | -1 | 1 | WOOD5-E-4 | 22-158 | H2O | 1668864891 | 0.314 | 0.728 | 74.431 | 0 | 0.002 | 1.779 | 1397.336 | -4.466 | 2.675 | 105.641 | 0.875 | -0.601 | 0 | 0 | 1 | 42.375 | 1 | 35 | 1 | GJB_Rinse | 0
163 | P-2048 | 2022/11/19 13:38:42 | 2-35 | 2 | -1.198 | -49.531 | 17326 | -1 | 1 | WOOD5-E-4 | 22-158 | H2O | 1668865429 | 0.322 | 0.757 | 74.31 | 0 | -0.001 | 1.77 | 1402.71 | -4.271 | 2.786 | 107.853 | 0.88 | -0.378 | 0 | 0 | 1 | 42.312 | 1 | 35 | 1 | GJB_Rinse | 0
164 | P-2048 | 2022/11/19 13:47:38 | 2-35 | 3 | -1.257 | -49.982 | 17368 | -1 | 1 | WOOD5-E-4 | 22-158 | H2O | 1668865965 | 0.344 | 0.73 | 72.488 | 0.001 | 0.001 | 1.732 | 1403.791 | -4.516 | 2.762 | 106.75 | 0.876 | -0.637 | 0 | 0 | 1 | 42.125 | 1 | 35 | 1 | GJB_Rinse | 0
165 | P-2048 | 2022/11/19 13:56:36 | 2-35 | 4 | -1.254 | -49.604 | 17357 | 0 | 1 | WOOD5-E-4 | 22-158 | H2O | 1668866502 | 0.321 | 0.803 | 73.051 | 0 | 0.003 | 1.745 | 1403.596 | -4.749 | 2.796 | 108.981 | 0.871 | -0.249 | 0 | 0 | 1 | 42.188 | 1 | 35 | 1 | GJB_Rinse | 0
166 | P-2049 | 2022/11/19 14:06:19 | 2-36 | 1 | -2.54 | -53.09 | 17314 | -1 | 1 | WOOD5-A-1 | 22-158 | H2O | 1668867086 | 0.31 | 0.766 | 75.14 | 0.001 | 0.003 | 1.781 | 1336.933 | -1.047 | 2.732 | 106.568 | 0.867 | -0.053 | 0 | 0 | 1 | 42.125 | 1 | 36 | 1 | GJB_Rinse | 0
167 | P-2049 | 2022/11/19 14:15:16 | 2-36 | 2 | -2.51 | -52.959 | 17633 | -1 | 1 | WOOD5-A-1 | 22-158 | H2O | 1668867622 | 0.339 | 0.792 | 76.967 | 0.002 | 0.005 | 1.838 | 1339.074 | -0.745 | 2.813 | 109.093 | 0.877 | 0.168 | 0 | 0 | 1 | 42.062 | 1 | 36 | 1 | GJB_Rinse | 0
168 | P-2049 | 2022/11/19 14:24:13 | 2-36 | 3 | -2.531 | -53.269 | 17566 | -1 | 1 | WOOD5-A-1 | 22-158 | H2O | 1668868159 | 0.329 | 0.701 | 75.092 | 0.002 | 0.003 | 1.79 | 1338.24 | -0.825 | 2.858 | 109.471 | 0.878 | 0.225 | 0 | 0 | 1 | 42.125 | 1 | 36 | 1 | GJB_Rinse | 0
169 | P-2049 | 2022/11/19 14:33:09 | 2-36 | 4 | -2.583 | -53.508 | 17368 | 0 | 1 | WOOD5-A-1 | 22-158 | H2O | 1668868696 | 0.305 | 0.78 | 73.357 | 0.002 | 0.002 | 1.746 | 1336.242 | -0.144 | 2.707 | 105.894 | 0.874 | -0.125 | 0 | 0 | 1 | 42.188 | 1 | 36 | 1 | GJB_Rinse | 0
170 | P-2050 | 2022/11/19 14:42:54 | 2-37 | 1 | -4.924 | -75.113 | 16492 | -1 | 1 | WOOD5-deep-2 | 22-158 | H2O | 1668869280 | 0.338 | 0.845 | 72.327 | 0.001 | 0.003 | 1.739 | 1447.826 | -7.124 | 2.63 | 103.617 | 0.876 | -0.689 | 0 | 0 | 1 | 42.25 | 1 | 37 | 1 | GJB_Rinse | 0
171 | P-2050 | 2022/11/19 14:51:51 | 2-37 | 2 | -5.092 | -77.434 | 17090 | -1 | 1 | WOOD5-deep-2 | 22-158 | H2O | 1668869817 | 0.343 | 0.778 | 73.208 | 0.001 | 0.005 | 1.756 | 1459.761 | -7.508 | 2.748 | 107.13 | 0.873 | -0.853 | 0 | 0 | 1 | 42.25 | 1 | 37 | 1 | GJB_Rinse | 0
172 | P-2050 | 2022/11/19 15:00:48 | 2-37 | 3 | -5.056 | -77.498 | 17432 | -1 | 1 | WOOD5-deep-2 | 22-158 | H2O | 1668870354 | 0.332 | 0.772 | 74.02 | 0.001 | 0.005 | 1.77 | 1465.028 | -7.186 | 2.697 | 107.144 | 0.88 | -1.008 | 0 | 0 | 1 | 42.312 | 1 | 37 | 1 | GJB_Rinse | 0
173 | P-2050 | 2022/11/19 15:09:44 | 2-37 | 4 | -5.07 | -78.085 | 17011 | 0 | 1 | WOOD5-deep-2 | 22-158 | H2O | 1668870890 | 0.311 | 0.816 | 74.75 | 0.001 | 0.005 | 1.792 | 1458.281 | -7.084 | 2.682 | 105.866 | 0.879 | -0.858 | 0 | 0 | 1 | 42.312 | 1 | 37 | 1 | GJB_Rinse | 0
174 | P-2051 | 2022/11/19 15:19:28 | 2-38 | 1 | -14.533 | -114.157 | 16927 | -1 | 1 | WOOD5-deep-1 | 22-158 | H2O | 1668871475 | 0.315 | 0.808 | 73.734 | 0.002 | 0.01 | 1.761 | 1290.565 | 2.423 | 2.646 | 103.612 | 0.872 | 0.402 | 0 | 0 | 1 | 42.188 | 1 | 38 | 1 | GJB_Rinse | 0
175 | P-2051 | 2022/11/19 15:28:25 | 2-38 | 2 | -14.933 | -117.82 | 17067 | -1 | 1 | WOOD5-deep-1 | 22-158 | H2O | 1668872011 | 0.319 | 0.776 | 73.995 | 0.001 | 0.007 | 1.77 | 1289.615 | 2.061 | 2.691 | 105.126 | 0.875 | 0.399 | 0 | 0 | 1 | 42.188 | 1 | 38 | 1 | GJB_Rinse | 0
176 | P-2051 | 2022/11/19 15:37:21 | 2-38 | 3 | -15.019 | -118.827 | 17440 | -1 | 1 | WOOD5-deep-1 | 22-158 | H2O | 1668872548 | 0.348 | 0.757 | 75.786 | 0.002 | 0.006 | 1.802 | 1291.683 | 2.807 | 2.759 | 107.144 | 0.875 | 0.533 | 0 | 0 | 1 | 42.188 | 1 | 38 | 1 | GJB_Rinse | 0
177 | P-2051 | 2022/11/19 15:46:18 | 2-38 | 4 | -15.147 | -119.149 | 17336 | 0 | 1 | WOOD5-deep-1 | 22-158 | H2O | 1668873085 | 0.303 | 0.693 | 74.415 | 0.001 | 0.002 | 1.767 | 1290.798 | 2.449 | 2.724 | 106.067 | 0.871 | 0.388 | 0 | 0 | 1 | 42.188 | 1 | 38 | 1 | GJB_Rinse | 0
178 | P-2052 | 2022/11/19 15:56:03 | 2-39 | 1 | -5.197 | -78.995 | 17226 | -1 | 1 | WOOD5-3040-2 | 22-158 | H2O | 1668873670 | 0.316 | 0.811 | 74.245 | 0.001 | 0.001 | 1.771 | 1393.948 | -3.614 | 2.705 | 106.385 | 0.873 | -0.466 | 0 | 0 | 1 | 42.375 | 1 | 39 | 1 | GJB_Rinse | 0
179 | P-2052 | 2022/11/19 16:04:59 | 2-39 | 2 | -4.674 | -74.782 | 17332 | -1 | 1 | WOOD5-3040-2 | 22-158 | H2O | 1668874205 | 0.331 | 0.766 | 75.386 | 0 | 0.002 | 1.802 | 1396.699 | -3.783 | 2.745 | 108.08 | 0.876 | -0.243 | 0 | 0 | 1 | 42.25 | 1 | 39 | 1 | GJB_Rinse | 0
180 | P-2052 | 2022/11/19 16:13:55 | 2-39 | 3 | -4.61 | -74.184 | 17477 | -1 | 1 | WOOD5-3040-2 | 22-158 | H2O | 1668874741 | 0.295 | 0.731 | 74.204 | 0.001 | 0.002 | 1.767 | 1398.926 | -3.559 | 2.783 | 108.661 | 0.878 | -0.103 | 0 | 0 | 1 | 42.188 | 1 | 39 | 1 | GJB_Rinse | 0
181 | P-2052 | 2022/11/19 16:22:52 | 2-39 | 4 | -4.516 | -73.578 | 17493 | 0 | 1 | WOOD5-3040-2 | 22-158 | H2O | 1668875278 | 0.332 | 0.679 | 75.396 | 0.001 | 0.001 | 1.791 | 1399.196 | -3.649 | 2.761 | 108.821 | 0.876 | -0.255 | 0 | 0 | 1 | 42.188 | 1 | 39 | 1 | GJB_Rinse | 0
182 | P-2053 | 2022/11/19 16:32:38 | 2-40 | 1 | -7.527 | -62.874 | 17178 | -1 | 1 | EV | standard | H2O | 1668875864 | 0.315 | 0.754 | 73.494 | 0.001 | -0.001 | 1.764 | 1291.306 | 1.592 | 2.708 | 106.026 | 0.872 | 0.427 | 0 | 0 | 1 | 42.188 | 1 | 40 | 1 | GJB_Rinse | 0
183 | P-2053 | 2022/11/19 16:41:34 | 2-40 | 2 | -7.597 | -61.285 | 16910 | -1 | 1 | EV | standard | H2O | 1668876400 | 0.311 | 0.738 | 73.611 | 0.001 | 0 | 1.772 | 1288.402 | 1.231 | 2.642 | 103.138 | 0.876 | -0.039 | 0 | 0 | 1 | 42.25 | 1 | 40 | 1 | GJB_Rinse | 0
184 | P-2053 | 2022/11/19 16:50:30 | 2-40 | 3 | -7.641 | -61.009 | 16881 | -1 | 1 | EV | standard | H2O | 1668876937 | 0.319 | 0.778 | 74.816 | 0 | 0.001 | 1.789 | 1287.329 | 1.97 | 2.665 | 103.55 | 0.875 | 0.378 | 0 | 0 | 1 | 42.188 | 1 | 40 | 1 | GJB_Rinse | 0
185 | P-2053 | 2022/11/19 16:59:27 | 2-40 | 4 | -7.619 | -60.934 | 16969 | 0 | 1 | EV | standard | H2O | 1668877473 | 0.307 | 0.781 | 73.288 | 0.001 | 0.002 | 1.751 | 1287.696 | 2.013 | 2.733 | 105.848 | 0.871 | 0.608 | 0 | 0 | 1 | 42.312 | 1 | 40 | 1 | GJB_Rinse | 0
186 | P-2054 | 2022/11/19 17:09:12 | 2-41 | 1 | -3.543 | -58.795 | 16026 | -1 | 1 | WOOD5-A-2 | 22-158 | H2O | 1668878059 | 0.304 | 0.752 | 69.169 | 0.001 | 0.002 | 1.668 | 1302.566 | 0.3 | 2.527 | 99.926 | 0.868 | 0.174 | 0 | 0 | 1 | 42.312 | 1 | 41 | 1 | GJB_Rinse | 0
187 | P-2054 | 2022/11/19 17:18:10 | 2-41 | 2 | -2.889 | -56.487 | 18033 | -1 | 1 | WOOD5-A-2 | 22-158 | H2O | 1668878596 | 0.324 | 0.677 | 74.688 | 0.002 | 0.003 | 1.774 | 1319.928 | -0.295 | 2.903 | 112.501 | 0.871 | 0.143 | 0 | 0 | 1 | 42.312 | 1 | 41 | 1 | GJB_Rinse | 0
188 | P-2054 | 2022/11/19 17:27:06 | 2-41 | 3 | -2.997 | -56.922 | 17308 | -1 | 1 | WOOD5-A-2 | 22-158 | H2O | 1668879132 | 0.299 | 0.808 | 72.624 | 0 | 0.003 | 1.744 | 1314.079 | -0.191 | 2.727 | 106.943 | 0.878 | -0.213 | 0 | 0 | 1 | 42.25 | 1 | 41 | 1 | GJB_Rinse | 0
189 | P-2054 | 2022/11/19 17:36:02 | 2-41 | 4 | -3.071 | -57.549 | 17096 | 0 | 1 | WOOD5-A-2 | 22-158 | H2O | 1668879668 | 0.281 | 0.729 | 73.502 | 0 | 0 | 1.755 | 1312.247 | -0.24 | 2.724 | 106.206 | 0.877 | 0.08 | 0 | 0 | 1 | 42.188 | 1 | 41 | 1 | GJB_Rinse | 0
190 | P-2055 | 2022/11/19 17:45:48 | 2-42 | 1 | -1.064 | -51.207 | 16230 | -1 | 1 | WOOD5-C-4 | 22-158 | H2O | 1668880254 | 0.294 | 0.794 | 74.485 | 0.001 | 0.005 | 1.796 | 1292.409 | 0.599 | 2.529 | 100.351 | 0.869 | 0.024 | 0 | 0 | 1 | 42.188 | 1 | 42 | 1 | GJB_Rinse | 0
191 | P-2055 | 2022/11/19 17:54:45 | 2-42 | 2 | -1.019 | -50.628 | 16276 | -1 | 1 | WOOD5-C-4 | 22-158 | H2O | 1668880791 | 0.306 | 0.855 | 72.096 | 0 | 0.003 | 1.728 | 1292.496 | 0.542 | 2.587 | 101.718 | 0.874 | 0.164 | 0 | 0 | 1 | 42.125 | 1 | 42 | 1 | GJB_Rinse | 0
192 | P-2055 | 2022/11/19 18:03:42 | 2-42 | 3 | -1.033 | -50.643 | 15945 | -1 | 1 | WOOD5-C-4 | 22-158 | H2O | 1668881328 | 0.321 | 0.811 | 71.234 | 0.001 | 0.004 | 1.705 | 1289.644 | 0.899 | 2.534 | 99.902 | 0.87 | 0.412 | 0 | 0 | 1 | 42.188 | 1 | 42 | 1 | GJB_Rinse | 0
193 | P-2055 | 2022/11/19 18:12:38 | 2-42 | 4 | -1.035 | -50.541 | 15529 | 0 | 1 | WOOD5-C-4 | 22-158 | H2O | 1668881865 | 0.318 | 0.875 | 69.635 | 0.001 | 0.003 | 1.685 | 1286.58 | 0.605 | 2.488 | 97.646 | 0.871 | 0.266 | 0 | 0 | 1 | 42.25 | 1 | 42 | 1 | GJB_Rinse | 0
194 | P-2056 | 2022/11/19 18:22:23 | 2-43 | 1 | -10.578 | -94.53 | 16621 | -1 | 1 | WOOD5-2030-3 | 22-158 | H2O | 1668882449 | 0.299 | 0.716 | 69.417 | 0.001 | 0.006 | 1.671 | 1285.622 | 1.982 | 2.623 | 101.626 | 0.869 | 0.22 | 0 | 0 | 1 | 42.188 | 1 | 43 | 1 | GJB_Rinse | 0
195 | P-2056 | 2022/11/19 18:31:19 | 2-43 | 2 | -11.034 | -98.613 | 17004 | -1 | 1 | WOOD5-2030-3 | 22-158 | H2O | 1668882986 | 0.318 | 0.816 | 74.64 | 0.001 | 0.006 | 1.787 | 1288.21 | 2.186 | 2.655 | 103.954 | 0.874 | 0.444 | 0 | 0 | 1 | 42.25 | 1 | 43 | 1 | GJB_Rinse | 0
196 | P-2056 | 2022/11/19 18:40:15 | 2-43 | 3 | -11.166 | -100.015 | 16995 | -1 | 1 | WOOD5-2030-3 | 22-158 | H2O | 1668883522 | 0.326 | 0.844 | 73.192 | 0.001 | 0.006 | 1.755 | 1287.987 | 1.955 | 2.736 | 105.051 | 0.874 | 0.507 | 0 | 0 | 1 | 42.25 | 1 | 43 | 1 | GJB_Rinse | 0
197 | P-2056 | 2022/11/19 18:49:12 | 2-43 | 4 | -11.22 | -100.453 | 16819 | 0 | 1 | WOOD5-2030-3 | 22-158 | H2O | 1668884058 | 0.332 | 0.768 | 70.524 | 0.002 | 0.005 | 1.691 | 1286.854 | 2.494 | 2.631 | 101.792 | 0.873 | 0.064 | 0 | 0 | 1 | 42.062 | 1 | 43 | 1 | GJB_Rinse | 0
198 | P-2057 | 2022/11/19 18:58:57 | 2-44 | 1 | -11.302 | -102.242 | 15205 | -1 | 1 | WOOD5-3040-3 | 22-158 | H2O | 1668884644 | 0.315 | 0.813 | 67.938 | 0.002 | 0.003 | 1.643 | 1275.508 | 2.313 | 2.317 | 91.162 | 0.866 | -0.11 | 0 | 0 | 1 | 42.25 | 1 | 44 | 1 | GJB_Rinse | 0
199 | P-2057 | 2022/11/19 19:07:53 | 2-44 | 2 | -11.191 | -101.691 | 16128 | -1 | 1 | WOOD5-3040-3 | 22-158 | H2O | 1668885180 | 0.307 | 0.714 | 71.318 | 0 | 0.003 | 1.719 | 1281.604 | 2.136 | 2.561 | 99.322 | 0.87 | 0.332 | 0 | 0 | 1 | 42.25 | 1 | 44 | 1 | GJB_Rinse | 0
200 | P-2057 | 2022/11/19 19:18:37 | 2-44 | 3 | [5 empty] | WOOD5-3040-3 | 22-158 | H2O | [17 empty] | 1 | 44 | 1 | GJB_Rinse | 0
201 | P-2057 | 2022/11/19 19:28:39 | 2-44 | 4 | [5 empty] | WOOD5-3040-3 | 22-158 | H2O | [17 empty] | 1 | 44 | 1 | GJB_Rinse | 0
202 | P-2058 | 2022/11/19 19:39:30 | 2-45 | 1 | [5 empty] | WOOD5-010-3 | 22-158 | H2O | [17 empty] | 1 | 45 | 1 | GJB_Rinse | 0
203 | P-2058 | 2022/11/19 19:49:31 | 2-45 | 2 | [5 empty] | WOOD5-010-3 | 22-158 | H2O | [17 empty] | 1 | 45 | 1 | GJB_Rinse | 0
204 | P-2058 | 2022/11/19 19:59:32 | 2-45 | 3 | [5 empty] | WOOD5-010-3 | 22-158 | H2O | [17 empty] | 1 | 45 | 1 | GJB_Rinse | 0
205 | P-2058 | 2022/11/19 20:09:32 | 2-45 | 4 | [5 empty] | WOOD5-010-3 | 22-158 | H2O | [17 empty] | 1 | 45 | 1 | GJB_Rinse | 0
206 | P-2059 | 2022/11/19 20:20:23 | 2-46 | 1 | -6.541 | -78.666 | 6657 | -1 | 1 | WOOD5-C-5 | 22-158 | H2O | 1668889567 | 0.324 | 1.428 | 19.421 | -0.001 | -0.005 | 1.116 | 1222.732 | 0.158 | 1.025 | 42.461 | 0.857 | -0.052 | 0 | 0 | 1 | 42 | 1 | 46 | 1 | GJB_Rinse | 0
207 | P-2059 | 2022/11/19 20:29:20 | 2-46 | 2 | [5 empty] | WOOD5-C-5 | 22-158 | H2O | [17 empty] | 1 | 46 | 1 | GJB_Rinse | 0
208 | P-2059 | 2022/11/19 20:39:57 | 2-46 | 3 | [5 empty] | WOOD5-C-5 | 22-158 | H2O | [17 empty] | 1 | 46 | 1 | GJB_Rinse | 0
211 | P-2060 | 2022/11/19 21:08:42 | 2-47 | 2 | 3.313 | 6.697 | 12280 | -1 | 1 | SP-21-775 | 22-258 | H2O | 1668892429 | 0.304 | 0.896 | 86.781 | -0.001 | -0.006 | 2.156 | 1305.658 | -4.976 | 1.949 | 79.329 | 0.866 | -1.321 | 0 | 0 | 1 | 43.938 | 1 | 47 | 1 | GJB_Rinse | 0
212 | P-2060 | 2022/11/19 21:17:39 | 2-47 | 3 | 3.393 | 7.723 | 12635 | -1 | 1 | SP-21-775 | 22-258 | H2O | 1668892966 | 0.285 | 0.8 | 80.024 | 0.001 | -0.002 | 1.972 | 1306.381 | -3.923 | 1.957 | 79.979 | 0.866 | -1.334 | 0 | 0 | 1 | 43.438 | 1 | 47 | 1 | GJB_Rinse | 0
213 | P-2060 | 2022/11/19 21:26:36 | 2-47 | 4 | [5 empty] | SP-21-775 | 22-258 | H2O | [17 empty] | 1 | 47 | 1 | GJB_Rinse | 0
215 | P-2061 | 2022/11/19 21:46:23 | 2-48 | 2 | 3.091 | 24.543 | 21487 | -1 | 1 | SP-21-787 | 22-258 | H2O | 1668894690 | 0.324 | 0.701 | 117.441 | 0 | 0.001 | 2.759 | 1331.667 | -1.867 | 3.611 | 136.101 | 0.887 | 0.051 | 0 | 0 | 1 | 41.312 | 1 | 48 | 1 | GJB_Rinse | 0
216 | P-2061 | 2022/11/19 21:55:21 | 2-48 | 3 | 2.228 | 17.608 | 12315 | -1 | 1 | SP-21-787 | 22-258 | H2O | 1668895228 | 0.314 | 0.88 | 85.926 | 0 | -0.003 | 2.135 | 1261.034 | 0.723 | 1.892 | 76.109 | 0.865 | -0.187 | 0 | 0 | 1 | 41.75 | 1 | 48 | 1 | GJB_Rinse | 0
217 | P-2061 | 2022/11/19 22:04:17 | 2-48 | 4 | 2.748 | 21.605 | 16901 | 0 | 1 | SP-21-787 | 22-258 | H2O | 1668895764 | 0.281 | 0.804 | 89.783 | 0.001 | 0.001 | 2.147 | 1294.977 | 0.468 | 2.709 | 104.687 | 0.876 | 0.002 | 0 | 0 | 1 | 41.875 | 1 | 48 | 1 | GJB_Rinse | 0
218 | P-2062 | 2022/11/19 22:14:04 | 2-49 | 1 | [5 empty] | EV | standard | H2O | [17 empty] | 1 | 49 | 1 | GJB_Rinse | 0
219 | P-2062 | 2022/11/19 22:24:06 | 2-49 | 2 | [5 empty] | EV | standard | H2O | [17 empty] | 1 | 49 | 1 | GJB_Rinse | 0
220 | P-2062 | 2022/11/19 22:34:06 | 2-49 | 3 | [5 empty] | EV | standard | H2O | [17 empty] | 1 | 49 | 1 | GJB_Rinse | 0
221 | P-2062 | 2022/11/19 22:44:08 | 2-49 | 4 | [5 empty] | EV | standard | H2O | [17 empty] | 1 | 49 | 1 | GJB_Rinse | 0
223 | P-2063 | 2022/11/19 23:03:54 | 2-50 | 2 | 1.552 | -4.699 | 9721 | -1 | 1 | SP-21-625 | 22-258 | H2O | 1668899343 | 0.371 | 0.911 | 75.08 | 0 | 0 | 1.934 | 1250.482 | -0.986 | 1.493 | 61.919 | 0.861 | -0.374 | 0 | 0 | 1 | 41.5 | 1 | 50 | 1 | GJB_Rinse | 0
224 | P-2063 | 2022/11/19 23:12:51 | 2-50 | 3 | [5 empty] | SP-21-625 | 22-258 | H2O | [17 empty] | 1 | 50 | 1 | GJB_Rinse | 0
225 | P-2063 | 2022/11/19 23:22:52 | 2-50 | 4 | 1.58 | -4.464 | 10370 | 0 | 1 | SP-21-625 | 22-258 | H2O | 1668900480 | 0.342 | 1.111 | 62.614 | 0 | -0.001 | 1.591 | 1255.677 | -0.929 | 1.599 | 65.686 | 0.862 | -0.479 | 0 | 0 | 1 | 42.062 | 1 | 50 | 1 | GJB_Rinse | 0
227 | P-2064 | 2022/11/19 23:41:35 | 2-51 | 2 | 2.104 | 0.256 | 16730 | -1 | 1 | SP-21-623 | 22-258 | H2O | 1668901602 | 0.312 | 0.892 | 103.131 | 0.001 | -0.002 | 2.459 | 1310.603 | -2.492 | 2.617 | 104.878 | 0.873 | -0.699 | 0 | 0 | 1 | 41.312 | 1 | 51 | 1 | GJB_Rinse | 0
228 | P-2064 | 2022/11/19 23:50:31 | 2-51 | 3 | 1.836 | -2.383 | 13244 | -1 | 1 | SP-21-623 | 22-258 | H2O | 1668902138 | 0.318 | 0.918 | 83.983 | 0 | 0.001 | 2.066 | 1280.279 | -1.399 | 2.051 | 83.26 | 0.868 | -0.49 | 0 | 0 | 1 | 41.75 | 1 | 51 | 1 | GJB_Rinse | 0
229 | P-2064 | 2022/11/19 23:59:29 | 2-51 | 4 | [5 empty] | SP-21-623 | 22-258 | H2O | [17 empty] | 1 | 51 | 1 | GJB_Rinse | 0
231 | P-2065 | 2022/11/20 00:19:16 | 2-52 | 2 | [5 empty] | SP-21-771 | 22-258 | H2O | [17 empty] | 1 | 52 | 1 | GJB_Rinse | 0
232 | P-2065 | 2022/11/20 00:29:17 | 2-52 | 3 | [5 empty] | SP-21-771 | 22-258 | H2O | [17 empty] | 1 | 52 | 1 | GJB_Rinse | 0
233 | P-2065 | 2022/11/20 00:39:18 | 2-52 | 4 | 3.283 | 1.229 | 12062 | 0 | 1 | SP-21-771 | 22-258 | H2O | 1668905066 | 0.29 | 0.843 | 73.461 | 0.001 | 0.001 | 1.828 | 1294.349 | -4.358 | 1.937 | 78.91 | 0.864 | -1.18 | 0 | 0 | 1 | 42 | 1 | 52 | 1 | GJB_Rinse | 0
235 | P-2066 | 2022/11/20 00:58:02 | 2-53 | 2 | 3.322 | 6.576 | 14083 | -1 | 1 | SP-21-789 | 22-258 | H2O | 1668906189 | 0.28 | 0.813 | 93.871 | 0.001 | 0.001 | 2.294 | 1304.709 | -4.211 | 2.197 | 90.126 | 0.869 | -1.249 | 0 | 0 | 1 | 41.375 | 1 | 53 | 1 | GJB_Rinse | 0
236 | P-2066 | 2022/11/20 01:06:59 | 2-53 | 3 | 2.958 | 3.764 | 10955 | -1 | 1 | SP-21-789 | 22-258 | H2O | 1668906727 | 0.325 | 0.857 | 78.018 | 0.001 | 0.001 | 1.956 | 1273.919 | -2.444 | 1.752 | 70.516 | 0.863 | -0.732 | 0 | 0 | 1 | 42.188 | 1 | 53 | 1 | GJB_Rinse | 0
237 | P-2066 | 2022/11/20 01:15:55 | 2-53 | 4 | [5 empty] | SP-21-789 | 22-258 | H2O | [17 empty] | 1 | 53 | 1 | GJB_Rinse | 0
239 | P-2067 | 2022/11/20 01:35:43 | 2-54 | 2 | 3.907 | 8.41 | 15350 | -1 | 1 | SP-21-778 | 22-258 | H2O | 1668908450 | 0.292 | 0.888 | 99.753 | 0 | 0.003 | 2.412 | 1366.384 | -7.016 | 2.428 | 97.032 | 0.872 | -1.996 | 0 | 0 | 1 | 41.25 | 1 | 54 | 1 | GJB_Rinse | 0
240 | P-2067 | 2022/11/20 01:44:40 | 2-54 | 3 | 3.51 | 5.5 | 12351 | -1 | 1 | SP-21-778 | 22-258 | H2O | 1668908987 | 0.33 | 0.886 | 82.273 | 0.001 | 0 | 2.032 | 1325.298 | -5.495 | 1.951 | 78.97 | 0.864 | -1.393 | 0 | 0 | 1 | 41.812 | 1 | 54 | 1 | GJB_Rinse | 0
241 | P-2067 | 2022/11/20 01:53:36 | 2-54 | 4 | [5 empty] | SP-21-778 | 22-258 | H2O | [17 empty] | 1 | 54 | 1 | GJB_Rinse | 0
243 | P-2068 | 2022/11/20 02:13:24 | 2-55 | 2 | [5 empty] | SP-21-839 | 22-258 | H2O | [17 empty] | 1 | 55 | 1 | GJB_Rinse | 0
244 | P-2068 | 2022/11/20 02:23:25 | 2-55 | 3 | 1.855 | 15.833 | 14242 | -1 | 1 | SP-21-839 | 22-258 | H2O | 1668911312 | 0.309 | 0.821 | 89.346 | 0 | -0.001 | 2.183 | 1269.813 | 0.843 | 2.199 | 88.183 | 0.869 | 0.049 | 0 | 0 | 1 | 41.875 | 1 | 55 | 1 | GJB_Rinse | 0
245 | P-2068 | 2022/11/20 02:32:21 | 2-55 | 4 | [5 empty] | SP-21-839 | 22-258 | H2O | [17 empty] | 1 | 55 | 1 | GJB_Rinse | 0
247 | P-2069 | 2022/11/20 02:52:10 | 2-56 | 2 | 4.927 | 20.355 | 29609 | -1 | 0 | SP-21-773 | 22-258 | H2O | 1668913030 | 0.356 | 0.756 | 117.347 | 0.001 | 0 | 2.954 | 1541.346 | -15.083 | 5.18 | 191.81 | 0.947 | -2.219 | 0 | 0 | 1 | 41.375 | 1 | 56 | 1 | GJB_Rinse | 0
248 | P-2069 | 2022/11/20 03:01:06 | 2-56 | 3 | 4.822 | 18.915 | 27804 | -1 | 0 | SP-21-773 | 22-258 | H2O | 1668913573 | 0.35 | 0.708 | 120.334 | 0.001 | 0.001 | 2.768 | 1512.745 | -13.144 | 4.746 | 176.689 | 0.937 | -2.18 | 0 | 0 | 1 | 41.938 | 1 | 56 | 1 | GJB_Rinse | 0
249 | P-2069 | 2022/11/20 03:10:03 | 2-56 | 4 | 4.941 | 19.788 | 28618 | -1 | 0 | SP-21-773 | 22-258 | H2O | 1668914110 | 0.344 | 0.787 | 118.068 | -0.001 | -0.001 | 2.705 | 1524.699 | -14.604 | 4.938 | 184.875 | 0.932 | -2.184 | 0 | 0 | 1 | 42.125 | 1 | 56 | 1 | GJB_Rinse | 0
251 | P-2070 | 2022/11/20 03:28:46 | 2-57 | 2 | 3.328 | 4.845 | 8476 | -1 | 1 | SP-21-772 | 22-258 | H2O | 1668915236 | 0.349 | 1.171 | 93.167 | 0 | 0.002 | 2.5 | 1257.073 | -3.186 | 1.307 | 54.627 | 0.86 | -1.198 | 0 | 0 | 1 | 41.688 | 1 | 57 | 1 | GJB_Rinse | 0
252 | P-2070 | 2022/11/20 03:37:42 | 2-57 | 3 | 4.203 | 11.478 | 17802 | -1 | 1 | SP-21-772 | 22-258 | H2O | 1668915769 | 0.327 | 0.759 | 102.929 | 0 | 0.001 | 2.443 | 1345.472 | -6.057 | 2.919 | 114.485 | 0.877 | -1.581 | 0 | 0 | 1 | 42 | 1 | 57 | 1 | GJB_Rinse | 0
253 | P-2070 | 2022/11/20 03:46:39 | 2-57 | 4 | 3.363 | 5.294 | 9855 | 0 | 1 | SP-21-772 | 22-258 | H2O | 1668916307 | 0.288 | 1.006 | 77.87 | 0 | 0.004 | 2.006 | 1266.989 | -2.987 | 1.545 | 63.612 | 0.862 | -0.983 | 0 | 0 | 1 | 42.062 | 1 | 57 | 1 | GJB_Rinse | 0
254 | P-2071 | 2022/11/20 03:56:26 | 2-58 | 1 | [5 empty] | EV | standard | H2O | [17 empty] | 1 | 58 | 1 | GJB_Rinse | 0
255 | P-2071 | 2022/11/20 04:06:27 | 2-58 | 2 | [5 empty] | EV | standard | H2O | [17 empty] | 1 | 58 | 1 | GJB_Rinse | 0
256 | P-2071 | 2022/11/20 04:16:28 | 2-58 | 3 | [5 empty] | EV | standard | H2O | [17 empty] | 1 | 58 | 1 | GJB_Rinse | 0
257 | P-2071 | 2022/11/20 04:26:29 | 2-58 | 4 | [5 empty] | EV | standard | H2O | [17 empty] | 1 | 58 | 1 | GJB_Rinse | 0
259 | P-2072 | 2022/11/20 04:46:18 | 2-59 | 2 | 1.636 | -6.559 | 16415 | -1 | 1 | SP-21-626 | 22-258 | H2O | 1668919885 | 0.331 | 0.719 | 95.726 | 0.001 | 0.001 | 2.296 | 1304.828 | -2.403 | 2.594 | 103.867 | 0.874 | -0.609 | 0 | 0 | 1 | 41.625 | 1 | 59 | 1 | GJB_Rinse | 0
260 | P-2072 | 2022/11/20 04:55:15 | 2-59 | 3 | [5 empty] | SP-21-626 | 22-258 | H2O | [17 empty] | 1 | 59 | 1 | GJB_Rinse | 0
261 | P-2072 | 2022/11/20 05:05:16 | 2-59 | 4 | 1.293 | -9.395 | 12429 | 0 | 1 | SP-21-626 | 22-258 | H2O | 1668921024 | 0.349 | 0.943 | 74.716 | 0.001 | 0.004 | 1.857 | 1272.28 | -1.266 | 1.887 | 77.548 | 0.864 | -0.793 | 0 | 0 | 1 | 42.125 | 1 | 59 | 1 | GJB_Rinse | 0
263 | P-2073 | 2022/11/20 05:23:59 | 2-60 | 2 | 2.122 | -3.758 | 17214 | -1 | 1 | SP-21-622 | 22-258 | H2O | 1668922146 | 0.323 | 0.818 | 104.355 | 0.002 | 0.003 | 2.501 | 1315.968 | -3.362 | 2.724 | 109.682 | 0.87 | -0.911 | 0 | 0 | 1 | 41.312 | 1 | 60 | 1 | GJB_Rinse | 0
264 | P-2073 | 2022/11/20 05:32:56 | 2-60 | 3 | 1.907 | -5.593 | 14886 | -1 | 1 | SP-21-622 | 22-258 | H2O | 1668922683 | 0.315 | 0.925 | 89.462 | 0.002 | 0.003 | 2.177 | 1295.348 | -2.55 | 2.356 | 94.507 | 0.872 | -0.764 | 0 | 0 | 1 | 41.938 | 1 | 60 | 1 | GJB_Rinse | 0
265 | P-2073 | 2022/11/20 05:41:52 | 2-60 | 4 | [5 empty] | SP-21-622 | 22-258 | H2O | [17 empty] | 1 | 60 | 1 | GJB_Rinse | 0
267 | P-2074 | 2022/11/20 06:01:42 | 2-61 | 2 | [5 empty] | SP-21-624 | 22-258 | H2O | [17 empty] | 1 | 61 | 1 | GJB_Rinse | 0
268 | P-2074 | 2022/11/20 06:11:42 | 2-61 | 3 | [5 empty] | SP-21-624 | 22-258 | H2O | [17 empty] | 1 | 61 | 1 | GJB_Rinse | 0
269 | P-2074 | 2022/11/20 06:21:43 | 2-61 | 4 | 2.352 | 1.324 | 13680 | 0 | 1 | SP-21-624 | 22-258 | H2O | 1668925610 | 0.341 | 0.821 | 74.043 | 0.001 | 0.001 | 1.81 | 1284.366 | -1.579 | 2.125 | 86.289 | 0.867 | -0.495 | 0 | 0 | 1 | 42.062 | 1 | 61 | 1 | GJB_Rinse | 0
271 | P-2075 | 2022/11/20 06:40:28 | 2-62 | 2 | 2.47 | 4.339 | 16083 | -1 | 1 | SP-21-791 | 22-258 | H2O | 1668926735 | 0.318 | 0.781 | 98.487 | 0.001 | 0.004 | 2.377 | 1314.242 | -3.665 | 2.55 | 102.23 | 0.869 | -0.939 | 0 | 0 | 1 | 41.25 | 1 | 62 | 1 | GJB_Rinse | 0
272 | P-2075 | 2022/11/20 06:49:24 | 2-62 | 3 | 2.205 | 1.781 | 12869 | -1 | 1 | SP-21-791 | 22-258 | H2O | 1668927271 | 0.321 | 0.976 | 83.751 | 0.001 | 0.002 | 2.053 | 1284.316 | -1.88 | 1.989 | 81.422 | 0.865 | -0.478 | 0 | 0 | 1 | 41.875 | 1 | 62 | 1 | GJB_Rinse | 0
273 | P-2075 | 2022/11/20 06:58:21 | 2-62 | 4 | 1.866 | -0.85 | 9421 | 0 | 1 | SP-21-791 | 22-258 | H2O | 1668927810 | 0.326 | 1.036 | 70.156 | 0 | 0 | 1.834 | 1254.484 | -1.645 | 1.439 | 60.606 | 0.86 | -0.596 | 0 | 0 | 1 | 42.125 | 1 | 62 | 1 | GJB_Rinse | 0
275 | P-2076 | 2022/11/20 07:17:06 | 2-63 | 2 | 1.6 | 2.792 | 16373 | -1 | 1 | SP-21-774 | 22-258 | H2O | 1668928933 | 0.302 | 0.739 | 102.684 | 0.001 | 0.003 | 2.477 | 1315.157 | -2.006 | 2.561 | 102.072 | 0.874 | -0.477 | 0 | 0 | 1 | 41.438 | 1 | 63 | 1 | GJB_Rinse | 0
276 | P-2076 | 2022/11/20 07:26:02 | 2-63 | 3 | 1.309 | 0.346 | 13080 | -1 | 1 | SP-21-774 | 22-258 | H2O | 1668929470 | 0.346 | 0.906 | 87.278 | 0.002 | 0.003 | 2.14 | 1285.086 | -1.418 | 1.991 | 82.226 | 0.865 | -0.529 | 0 | 0 | 1 | 41.875 | 1 | 63 | 1 | GJB_Rinse | 0
277 | P-2076 | 2022/11/20 07:34:58 | 2-63 | 4 | 1.091 | -1.854 | 10432 | 0 | 1 | SP-21-774 | 22-258 | H2O | 1668930007 | 0.31 | 1.001 | 75.103 | 0.001 | 0.002 | 1.924 | 1262.035 | -0.837 | 1.583 | 65.854 | 0.861 | -0.421 | 0 | 0 | 1 | 42.062 | 1 | 63 | 1 | GJB_Rinse | 0
278 | P-2077 | 2022/11/20 07:44:45 | 2-64 | 1 | -6.829 | -49.326 | 8629 | -1 | 1 | EV | standard | H2O | 1668930596 | 0.328 | 1.223 | 64.325 | 0.001 | 0.012 | 1.717 | 1231.873 | 0.782 | 1.305 | 54.175 | 0.859 | 0.096 | 0 | 0 | 1 | 42.25 | 1 | 64 | 1 | GJB_Rinse | 0
279 | P-2077 | 2022/11/20 07:53:42 | 2-64 | 2 | [5 empty] | EV | standard | H2O | [17 empty] | 1 | 64 | 1 | GJB_Rinse | 0
280 | P-2077 | 2022/11/20 08:03:43 | 2-64 | 3 | -7.765 | -58.228 | 7804 | -1 | 1 | EV | standard | H2O | 1668931736 | 0.349 | 1.154 | 53.426 | 0.002 | 0.005 | 1.478 | 1225.951 | 0.922 | 1.157 | 48.667 | 0.857 | 0.04 | 0 | 0 | 1 | 42.375 | 1 | 64 | 1 | GJB_Rinse | 0
281 | P-2077 | 2022/11/20 08:12:40 | 2-64 | 4 | [5 empty] | EV | standard | H2O | [17 empty] | 1 | 64 | 1 | GJB_Rinse | 0

## Instructions
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Variable | Instruction | Comment
StandardsFile | standards file.csv
OutputBase | ChemCorrect
CorrectionType | M
timecode | Time Code
identifier | Identifier 1
sample | Sample
injection | Inj Nr | Inj Nr
line | Line
job | Job
Variable | Instruction | Comment
baseline_shift_MEAN | mean_standards(baseline_shift) | The mean of the organic baseline of the standards.
baseline_shift_STD | std_standards(baseline_shift) | The standard deviation of the organic baseline of the standards.
ch4_ppm_MEAN | mean_standards(ch4_ppm) | The mean methane concentration of the standards.
ch4_ppm_STD | std_standards(ch4_ppm) | The standard deviation of the methane concentrations of the standards.
residuals_MEAN | mean_standards(residuals) | The mean difference between the measured and modeled spectra of the standards.
residuals_STD | std_standards(residuals) | The standard deviation of the difference between measured and modeled spectra of the standards.
slope_shift_MEAN | mean_standards(slope_shift) | The mean organic slope of the standards.
slope_shift_STD | std_standards(slope_shift) | The standard deviation of the organic slope of the standards.
baseline_curvature_MEAN | mean_standards(baseline_curvature) | The mean of the baseline curvature of the standards.
baseline_curvature_STD | std_standards(baseline_curvature) | The standard deviation of the baseline curvature of the standards.
residuals_MEAN_ABS | abs(residuals_MEAN) | The absolute value of the mean difference between the measured and modeled spectra of the standards.
baseline_shift_MEAN_ABS | abs(baseline_shift_MEAN) | The absolute value of the mean of the organic baseline of the standards.
ch4_ppm_MEAN_ABS | abs(ch4_ppm_MEAN) | The absolute value of the mean methane concentration of the standards.
slope_shift_MEAN_ABS | abs(slope_shift_MEAN) | The absolute value of the mean organic slope of the standards.
residuals_15SIGMA | evaluation(residuals_STD*1.5) | One and half times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_15SIGMA | evaluation(residuals_MEAN+residuals_15SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_2SIGMA | evaluation(residuals_STD*2) | Twice the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_2SIGMA | evaluation(residuals_MEAN+residuals_2SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_3SIGMA | evaluation(residuals_STD*3) | Three times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_3SIGMA | evaluation(residuals_MEAN+residuals_3SIGMA) | The mean difference between measured and modeled spectra of the standards plus three standard deviation.
residuals_10PCT | evaluation(residuals_MEAN*0.1) | 30% of the difference between measured and modeled spectra of the standards.
ch4_ppm_3SIGMA | evaluation(ch4_ppm_STD*3) | Three times the standard deviation of the methanol amplitudes of the standards.
ch4_ppm_PLUS_3SIGMA | evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA) | The mean methanol amplitude of the standards plus three standard deviation.
HALF | float(0.5) | A fixed value of 0.5 or half.
baseline_shift_HALFSIGMA | evaluation(baseline_shift_STD*HALF) | Half the standard deviation of the organic baseline of the standards.
baseline_shift_PLUS_HALFSIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA) | The absolute value of the mean organic baseline of the standards plus one-half standard deviation.
baseline_shift_PLUS_1SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD) | The absolute value of the mean organic baseline of the standards plus one standard deviation.
baseline_shift_3SIGMA | evaluation(baseline_shift_STD*3) | Three times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_3SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA) | The absolute value of the mean organic baseline of the standards plus three standard deviation.
baseline_shift_18SIGMA | evaluation(baseline_shift_STD*18) | Eighteen times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_18SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA) | The absolute value of the mean organic baseline of the standards plus eighteen standard deviation.
slope_shift_5SIGMA | evaluation(slope_shift_STD*5) | Five times the standard deviation of the organic slope of the standards.
slope_shift_PLUS_5SIGMA | evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA) | The absolute value of the mean organic slope of the standards plus five standard deviation.
baseline_curvature_3SIGMA | evaluation(baseline_curvature_STD*3) | Three times the standard deviation of the baseline curvature of the standards.
baseline_curvature_PLUS_3SIGMA | evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
TRUE | evaluation(1>0)
FALSE | evaluation(1<0)
PROC_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the uncalibrated d18O value in order to calibrate it.
PROC_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the uncalibrated dD value in order to calibrate it.
Variable | Instruction | Comment
IS_STANDARD | is_standard_sample() | Standard recognition.
IS_NOT_STANDARD | bool_xor(IS_STANDARD,TRUE)
SAMP_residuals_MEAN | mean_all(residuals) | The mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_residuals_STD | std_all(residuals) | One standard deviation of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN | mean_all(ch4_ppm) | The mean methane concentration of a sample.
SAMP_ch4_ppm_STD | std_all(ch4_ppm) | The standard deviation of the methane concentrations of a sample.
SAMP_baseline_shift_MEAN | mean_all(baseline_shift) | The mean of the organic baseline of a sample.
SAMP_baseline_shift_STD | std_all(baseline_shift) | The standard deviation of the organic baseline of a sample.
SAMP_slope_shift_MEAN | mean_all(slope_shift) | The mean organic slope of a sample.
SAMP_slope_shift_STD | std_all(slope_shift) | The standard deviation of the organic slope of a sample.
SAMP_d18O_STD | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
SAMP_dD_STD | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_baseline_curvature_MEAN | mean_all(baseline_curvature) | The mean of the baseline curvature of the ssamples
SAMP_baseline_curvature_STD | std_all(baseline_curvature) | The standard deviation of the baseline curvature of the samples
SAMP_residuals_MEAN_ABS | abs(SAMP_residuals_MEAN) | Absolute value of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN_ABS | abs(SAMP_ch4_ppm_MEAN) | Absolute value of the mean methane concentration of a sample.
SAMP_baseline_shift_MEAN_ABS | abs(SAMP_baseline_shift_MEAN) | Absolute value of the mean organic baseline of a sample.
SAMP_slope_shift_MEAN_ABS | abs(SAMP_slope_shift_MEAN) | Absolute value of the mean organic slope of a sample.
SAMP_residuals_DIFFERENTIAL | evaluation(SAMP_residuals_MEAN-residuals_MEAN)
SAMP_residuals_DIFFERENTIAL_ABS | abs(SAMP_residuals_DIFFERENTIAL) | Absolute value of the difference between the mean of the sample residual and the mean of all standards residuals
SAMP_residuals_DIFFERENTIAL_PCT | evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
SAMP_baseline_curvature_DIFFERENTIAL | evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
SAMP_STD_DELTA_18_16O_CURR_MEAN | current_standards_mean(d(18_16)Mean) | The d18O mean of the last run standard.
SAMP_STD_DELTA_DH_CURR_MEAN | current_standards_mean(d(D_H)Mean) | The dD mean of the last run standard.
known_dD | known_standard(d(D_H)Mean) | The dD value of the known Standard.
known_d18O | known_standard(d(18_16)Mean) | The d18O value of the known Standard.
d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
measured_d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
measured_dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
measured_d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
measured_dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the mean uncalibrated d18O value in order to calibrate it.
SAMP_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the mean uncalibrated dD value in order to calibrate it.
SAMP_TEST_VALUE | evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
residuals_GT_3SIGMA | evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA) | Threshold of red trigger. If the mean difference between the measured spectra of the sample and the modeled spectra of the standard is greater than that of the mean difference of the standard plus three standard deviations, this will trigger.
residuals_GE_2SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GE_15SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GT_10PCT | evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
ch4_ppm_GE_3SIGMA | evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA) | True if the mean methane concentration of the sample is greater than or equal to that of the standard plus three standard deviations, otherwise false.
baseline_shift_GT_1SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus one standard deviation, otherwise false.
baseline_shift_GT_3SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus three standard deviation, otherwise false. Sets yellow trigger for "other", trigger threshold is approximately sample to sample dD precision specification
baseline_shift_GT_18SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus eighteen standard deviation, otherwise false. Sets red trigger for "other", trigger threshold is approximately 24 hour dD drift specification
slope_shift_GT_5SIGMA | evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA) | True if the mean organic slope of the sample is greater than that of the standard plus five standard deviations, otherwise false.
baseline_curvature_GT_3SIGMA | evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
RED | truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA)) | Red trigger. If the sample is not a standard and the threshold is triggered (measured-modeled > standard+3stdev's), then RED.
NOT_RED | bool_xor(RED,TRUE) | If the sample is determined to be "red" than it can not be "not red."
YELLOW | truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA))) | Truth evaluation used to determine whether or not a given sample is yellow.
NOT_YELLOW | bool_xor(YELLOW,TRUE) | If sample is determined to be "yellow" than it can not be "not yellow."
GREEN | truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW)) | This is the pre-requisite for a sample to be colored green. It must not be a standard, must not be yellow, and must not be red.
NOT_GREEN | bool_xor(GREEN,TRUE) | Pre-requisite for being "not-green."
CYAN | truth_evaluation((IS_STANDARD)) | This is the pre-requisite for a sample to be a standard and therefore cyan colored.
NOT_CYAN | bool_xor(CYAN,TRUE) | Pre-requisite for being not a standard.
METHANE | truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA)) | This is the threshold for amounts of methane being included in the main ChemCorrect spreadsheet. Methane will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured methane concentration must be greater than or equal to that of the standard plus three standard deviations.
OTHER1 | truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA)) | This is the threshold for amounts of "OTHER1" being included in the main ChemCorrect spreadsheet. This column (OTHER1) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic baseline must be greater than that of the standard plus one standard deviation.
OTHER2 | truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA)) | This is the threshold for amounts of "OTHER2" being included in the main ChemCorrect spreadsheet. This column (OTHER2) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic slope must be greater than that of the standard plus three standard deviations.
OTHER3 | truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | truth_evaluation(RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN)) | Organic baseline differential showing the difference between the measured baseline and the standard baseline.
Variable | Instruction | Comment
DTL_TEST | math_evaluation((1 + 2)) | Test evaluation of 1+2 being equal to 3.
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample | FALSE | FALSE | FALSE | FALSE | Testing Sample
B | identifier | TRUE | Name | RED | YELLOW | GREEN | CYAN | Testing Name
C | SAMP_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d<sup>18</sup>O Mean | RED | YELLOW | GREEN | CYAN | Testing18 16 Mean
D | SAMP_DELTA_DH_CALIBRATE | TRUE | Calibrated d<sup>2</sup>H Mean | RED | YELLOW | GREEN | CYAN | This is a test to see that I can use tool tips
E | SAMP_ch4_ppm_MEAN | METHANE | CH<sub>4</sub> | RED | YELLOW | FALSE | FALSE
F | baseline_shift_DIFF | OTHER1 | C<sub>2+</sub> alcohols | RED | YELLOW | FALSE | FALSE
G | SAMP_residuals_DIFFERENTIAL_PCT | OTHER3 | relative %deviation | RED | YELLOW | FALSE | FALSE
H | SAMP_d18O_STD | TRUE | Uncalibrated d<sup>18</sup>O precision | RED | YELLOW | GREEN | CYAN
I | SAMP_dD_STD | TRUE | Uncalibrated d<sup>2</sup>H precision | RED | YELLOW | GREEN | CYAN
J | SAMP_slope_shift_MEAN | OTHER2 | slope | RED | YELLOW | FALSE | FALSE
K | SAMP_baseline_curvature_DIFFERENTIAL | OTHER4 | curvature | RED | YELLOW | FALSE | FALSE
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample
B | identifier | TRUE | Name
C | col(d(18_16)Mean) | TRUE | Measured d(18_16) Mean
D | col(d(D_H)Mean) | TRUE | Measured d(D_H) Mean
E | DTL_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d(18_16) Mean
F | DTL_DELTA_DH_CALIBRATE | TRUE | Calibrated d(D_H) Mean
G | col(ch4_ppm) | METHANE | CH<sub>4</sub>
H | col(baseline_shift) | OTHER1 | Other
I | col(Time Code) | TRUE | Time Code
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | d18O_mean | TRUE | d(18_16) Mean | RED | YELLOW | GREEN | CYAN
B | dD_mean | TRUE | d(D_H) Mean | RED | YELLOW | GREEN | CYAN
C | XY_X(d18O_mean) XY_Y(dD_mean) | TRUE | d18O x dD | gmwl | best_fit

## Process Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Process Variables
Variable | Value | Operations
baseline_shift_MEAN | 1294.99338095
varray: [1298.261, 1295.929, 1295.972, 1295.894, 1295.682, 1295.602, 1294.809, 1295.128, 1294.939, 1294.768, 1298.378, 1294.586, 1294.372, 1294.343, 1293.844, 1294.249, 1293.664, 1293.094, 1293.282, 1293.696, 1294.369]
inst: mean_standards
operation: mean_standards(baseline_shift)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0, 12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: baseline_shift
baseline_shift_STD | 1.39561246326
varray: [1298.261, 1295.929, 1295.972, 1295.894, 1295.682, 1295.602, 1294.809, 1295.128, 1294.939, 1294.768, 1298.378, 1294.586, 1294.372, 1294.343, 1293.844, 1294.249, 1293.664, 1293.094, 1293.282, 1293.696, 1294.369]
inst: std_standards
operation: std_standards(baseline_shift)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0, 12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: baseline_shift
ch4_ppm_MEAN | 0.391666666667
varray: [0.457, 0.439, 0.615, 0.487, 0.191, 0.216, 0.573, 0.504, 0.444, 0.214, 0.384, 0.389, 0.415, 0.564, 0.27, 0.056, 0.37, 0.472, 0.281, 0.539, 0.345]
inst: mean_standards
operation: mean_standards(ch4_ppm)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0, 12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: ch4_ppm
ch4_ppm_STD | 0.144437645139
varray: [0.457, 0.439, 0.615, 0.487, 0.191, 0.216, 0.573, 0.504, 0.444, 0.214, 0.384, 0.389, 0.415, 0.564, 0.27, 0.056, 0.37, 0.472, 0.281, 0.539, 0.345]
inst: std_standards
operation: std_standards(ch4_ppm)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0, 12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: ch4_ppm
residuals_MEAN | 2.88138095238
varray: [3.013, 2.932, 2.942, 2.958, 2.861, 2.907, 2.96, 2.947, 2.931, 2.882, 2.923, 2.847, 2.839, 2.832, 2.808, 2.818, 2.824, 2.81, 2.776, 2.868, 2.831]
inst: mean_standards
operation: mean_standards(residuals)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0, 12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: residuals
residuals_STD | 0.0641143323996
varray: [3.013, 2.932, 2.942, 2.958, 2.861, 2.907, 2.96, 2.947, 2.931, 2.882, 2.923, 2.847, 2.839, 2.832, 2.808, 2.818, 2.824, 2.81, 2.776, 2.868, 2.831]
inst: std_standards
operation: std_standards(residuals)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0, 12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: residuals
slope_shift_MEAN | 1.738
varray: [1.361, 1.712, 1.425, 0.77, 1.247, 1.115, 1.361, 1.024, 1.253, 0.979, 2.904, 2.354, 2.171, 1.88, 2.634, 2.722, 1.883, 1.895, 1.962, 2.001, 1.845]
inst: mean_standards
operation: mean_standards(slope_shift)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0, 12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: slope_shift
slope_shift_STD | 0.598516666435
varray: [1.361, 1.712, 1.425, 0.77, 1.247, 1.115, 1.361, 1.024, 1.253, 0.979, 2.904, 2.354, 2.171, 1.88, 2.634, 2.722, 1.883, 1.895, 1.962, 2.001, 1.845]
inst: std_standards
operation: std_standards(slope_shift)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0, 12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: slope_shift
baseline_curvature_MEAN | 110.748904762
varray: [114.914, 111.635, 113.105, 112.029, 110.706, 111.185, 111.932, 111.965, 110.375, 110.705, 112.527, 109.459, 109.774, 110.598, 107.877, 108.157, 109.897, 109.908, 108.05, 110.455, 110.474]
inst: mean_standards
operation: mean_standards(baseline_curvature)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0, 12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: baseline_curvature
baseline_curvature_STD | 1.70556096064
varray: [114.914, 111.635, 113.105, 112.029, 110.706, 111.185, 111.932, 111.965, 110.375, 110.705, 112.527, 109.459, 109.774, 110.598, 107.877, 108.157, 109.897, 109.908, 108.05, 110.455, 110.474]
inst: std_standards
operation: std_standards(baseline_curvature)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0, 12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: baseline_curvature
residuals_MEAN_ABS | 2.88138095238
operation: abs(residuals_MEAN)
inst: abs
source_col: residuals_MEAN
baseline_shift_MEAN_ABS | 1294.99338095
operation: abs(baseline_shift_MEAN)
inst: abs
source_col: baseline_shift_MEAN
ch4_ppm_MEAN_ABS | 0.391666666667
operation: abs(ch4_ppm_MEAN)
inst: abs
source_col: ch4_ppm_MEAN
slope_shift_MEAN_ABS | 1.738
operation: abs(slope_shift_MEAN)
inst: abs
source_col: slope_shift_MEAN
residuals_15SIGMA | 0.0961714985994
operation: evaluation(residuals_STD*1.5)
inst: evaluation
factors: residuals_STD*1.5
residuals_PLUS_15SIGMA | 2.97755245098
operation: evaluation(residuals_MEAN+residuals_15SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_15SIGMA
residuals_2SIGMA | 0.128228664799
operation: evaluation(residuals_STD*2)
inst: evaluation
factors: residuals_STD*2
residuals_PLUS_2SIGMA | 3.00960961718
operation: evaluation(residuals_MEAN+residuals_2SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_2SIGMA
residuals_3SIGMA | 0.192342997199
operation: evaluation(residuals_STD*3)
inst: evaluation
factors: residuals_STD*3
residuals_PLUS_3SIGMA | 3.07372394958
operation: evaluation(residuals_MEAN+residuals_3SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_3SIGMA
residuals_10PCT | 0.288138095238
operation: evaluation(residuals_MEAN*0.1)
inst: evaluation
factors: residuals_MEAN*0.1
ch4_ppm_3SIGMA | 0.433312935417
operation: evaluation(ch4_ppm_STD*3)
inst: evaluation
factors: ch4_ppm_STD*3
ch4_ppm_PLUS_3SIGMA | 0.824979602084
operation: evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA)
inst: evaluation
factors: ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA
HALF | 0.5
operation: float(0.5)
inst: float
baseline_shift_HALFSIGMA | 0.697806231632
operation: evaluation(baseline_shift_STD*HALF)
inst: evaluation
factors: baseline_shift_STD*HALF
baseline_shift_PLUS_HALFSIGMA | 1295.69118718
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA
baseline_shift_PLUS_1SIGMA | 1296.38899342
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_STD
baseline_shift_3SIGMA | 4.18683738979
operation: evaluation(baseline_shift_STD*3)
inst: evaluation
factors: baseline_shift_STD*3
baseline_shift_PLUS_3SIGMA | 1299.18021834
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_3SIGMA
baseline_shift_18SIGMA | 25.1210243388
operation: evaluation(baseline_shift_STD*18)
inst: evaluation
factors: baseline_shift_STD*18
baseline_shift_PLUS_18SIGMA | 1320.11440529
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_18SIGMA
slope_shift_5SIGMA | 2.99258333217
operation: evaluation(slope_shift_STD*5)
inst: evaluation
factors: slope_shift_STD*5
slope_shift_PLUS_5SIGMA | 4.73058333217
operation: evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA)
inst: evaluation
factors: slope_shift_MEAN_ABS+slope_shift_5SIGMA
baseline_curvature_3SIGMA | 5.11668288193
operation: evaluation(baseline_curvature_STD*3)
inst: evaluation
factors: baseline_curvature_STD*3
baseline_curvature_PLUS_3SIGMA | 115.865587644
operation: evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
inst: evaluation
factors: baseline_curvature_3SIGMA+baseline_curvature_MEAN
TRUE | True
operation: evaluation(1>0)
inst: evaluation
factors: 1>0
FALSE | False
operation: evaluation(1<0)
inst: evaluation
factors: 1<0
PROC_DELTA_18_16O_CALIBRATE | -5.11974409871
actual_array: [ 1.93 -15.84]
varray: [-6.107, -6.837, -7.245, -7.129, -7.203, -7.273, -7.395, -7.349, -7.42, -7.411, 4.01, 4.535, 4.691, 4.763, 4.819, 4.777, 4.806, 4.84, 4.859, 4.847, -12.113, -12.906, -13.147, -13.225, -13.24, -13.302, -13.359, -13.35, -13.305, -13.351, -13.451, -8.017, -7.77, -7.704, -7.772, -7.641, -7.676, -7.661, -7.73, -7.625, -7.672, -6.791, -6.874, -6.827, -6.816, -7.587, -7.618, -7.582, -7.568, -0.831, -0.885, -0.809, -0.803, 0.058, 0.34, 0.344, 0.413, 1.605, 1.324, 1.195, 1.149, -5.023, -4.894, -4.849, -4.884, -7.951, -8.137, -8.194, -8.151, -0.016, -0.13, -0.119, -0.002, -7.754, -7.667, -7.63, -7.695, -7.643, -7.619, -7.651, -7.643, 4.339, 4.344, 4.448, 4.423, 0.102, 0.122, 0.024, 0.073, -5.845, -5.865, -5.83, -5.87, -4.535, -4.437, -4.37, -4.347, -1.65, -1.445, -1.454, -1.386, -4.465, -4.535, -4.63, -4.659, -1.0, -0.902, -0.791, -0.805, -7.231, -7.499, -7.539, -7.555, -7.197, -7.052, -7.087, -7.062, -8.508, -8.421, -8.531, -8.559, -2.249, -2.064, -1.895, -1.837, -5.395, -5.575, -5.542, -5.601, -5.643, -5.601, -5.669, -5.565, -1.709, -1.544, -1.523, -1.313, -1.7, -1.411, -1.628, -1.676, -2.976, -3.133, -3.057, -3.112, -7.339, -7.539, -7.604, -7.593, -3.893, -3.694, -3.648, -3.677, -11.865, -12.25, -12.307, -12.318, 0.217, 0.873, 0.973, 1.072, -1.117, -1.198, -1.257, -1.254, -2.54, -2.51, -2.531, -2.583, -4.924, -5.092, -5.056, -5.07, -14.533, -14.933, -15.019, -15.147, -5.197, -4.674, -4.61, -4.516, -7.527, -7.597, -7.641, -7.619, -3.543, -2.889, -2.997, -3.071, -1.064, -1.019, -1.033, -1.035, -10.578, -11.034, -11.166, -11.22, -11.302, -11.191, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, -6.541, 0.0, 0.0, 3.313, 3.393, 0.0, 3.091, 2.228, 2.748, 0.0, 0.0, 0.0, 0.0, 1.552, 0.0, 1.58, 2.104, 1.836, 0.0, 0.0, 0.0, 3.283, 3.322, 2.958, 0.0, 3.907, 3.51, 0.0, 0.0, 1.855, 0.0, 4.927, 4.822, 4.941, 3.328, 4.203, 3.363, 0.0, 0.0, 0.0, 0.0, 1.636, 0.0, 1.293, 2.122, 1.907, 0.0, 0.0, 0.0, 2.352, 2.47, 2.205, 1.866, 1.6, 1.309, 1.091, -6.829, 0.0, -7.765, 0.0]
fit_rsq: 1.0
fit_intercept: 1.38222915025
except: True
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -3.35153962264
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [1.0, 538.0, 1184.0, 1894.0, 2501.0, 3038.0, 3576.0, 4113.0, 4650.0, 5187.0, 5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0, 12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0, 17721.0, 18258.0, 18795.0, 19332.0, 19869.0, 20406.0, 20944.0, 21482.0, 22019.0, 22556.0, 23146.0, 23682.0, 24219.0, 24757.0, 25347.0, 25884.0, 26421.0, 26958.0, 27549.0, 28086.0, 28623.0, 29160.0, 29747.0, 30284.0, 30821.0, 31358.0, 31946.0, 32483.0, 33020.0, 33557.0, 34146.0, 34682.0, 35219.0, 35757.0, 36346.0, 36883.0, 37420.0, 37958.0, 38546.0, 39083.0, 39619.0, 40157.0, 40747.0, 41285.0, 41822.0, 42359.0, 42949.0, 43486.0, 44022.0, 44559.0, 45146.0, 45683.0, 46221.0, 46757.0, 47345.0, 47882.0, 48419.0, 48957.0, 49545.0, 50082.0, 50620.0, 51157.0, 51746.0, 52283.0, 52820.0, 53358.0, 53946.0, 54483.0, 55021.0, 55558.0, 56148.0, 56684.0, 57221.0, 57758.0, 58348.0, 58885.0, 59422.0, 59959.0, 60546.0, 61082.0, 61618.0, 62155.0, 62742.0, 63279.0, 63817.0, 64354.0, 64941.0, 65478.0, 66015.0, 66552.0, 67140.0, 67677.0, 68214.0, 68751.0, 69338.0, 69875.0, 70412.0, 70949.0, 71538.0, 72074.0, 72611.0, 73148.0, 73737.0, 74274.0, 74811.0, 75348.0, 75933.0, 76469.0, 77007.0, 77544.0, 78129.0, 78666.0, 79203.0, 79740.0, 80326.0, 80862.0, 81399.0, 81936.0, 82523.0, 83060.0, 83597.0, 84133.0, 84720.0, 85257.0, 85794.0, 86332.0, 86918.0, 87455.0, 87993.0, 88530.0, 89118.0, 89655.0, 90191.0, 90729.0, 91312.0, 91849.0, 92386.0, 92922.0, 93507.0, 94044.0, 94581.0, 95117.0, 95701.0, 96238.0, 96774.0, 97311.0, 97896.0, 98432.0, 98968.0, 99505.0, 100091.0, 100627.0, 101163.0, 101700.0, 102285.0, 102823.0, 103359.0, 103895.0, 104481.0, 105018.0, 105555.0, 106091.0, 106676.0, 107212.0, 107748.0, 108285.0, 108870.0, 109406.0, 110050.0, 110652.0, 111303.0, 111904.0, 112505.0, 113105.0, 113756.0, 114293.0, 114930.0, 116655.0, 117192.0, 117729.0, 118916.0, 119454.0, 119990.0, 120577.0, 121179.0, 121779.0, 122381.0, 123567.0, 124104.0, 124705.0, 125828.0, 126364.0, 126902.0, 128089.0, 128690.0, 129291.0, 130415.0, 130952.0, 131488.0, 132676.0, 133213.0, 133749.0, 134937.0, 135538.0, 136074.0, 137263.0, 137799.0, 138336.0, 139459.0, 139995.0, 140532.0, 141119.0, 141720.0, 142321.0, 142922.0, 144111.0, 144648.0, 145249.0, 146372.0, 146909.0, 147445.0, 148635.0, 149235.0, 149836.0, 150961.0, 151497.0, 152034.0, 153159.0, 153695.0, 154231.0, 154818.0, 155355.0, 155956.0, 156493.0]
source_col: d(18_16)Mean
PROC_DELTA_DH_CALIBRATE | -58.8159247278
actual_array: [ 18.1 -119. ]
varray: [-54.516, -57.248, -58.741, -58.353, -58.641, -58.81, -59.0, -58.974, -59.346, -59.315, 28.804, 36.832, 38.754, 39.807, 39.876, 40.238, 40.639, 40.948, 40.953, 40.984, -92.388, -105.184, -107.874, -109.468, -110.097, -110.614, -111.064, -111.175, -111.049, -111.485, -111.754, -66.554, -62.362, -61.345, -61.072, -60.439, -60.479, -60.432, -60.284, -60.196, -60.306, -82.99, -85.569, -86.014, -86.273, -87.678, -87.937, -87.918, -87.922, -84.193, -85.44, -85.029, -85.421, -48.454, -43.661, -42.446, -42.307, -75.346, -79.685, -80.885, -81.261, -73.465, -71.082, -70.574, -70.58, -84.755, -86.056, -86.542, -86.406, -90.222, -92.735, -93.263, -92.877, -66.852, -62.719, -61.621, -61.319, -91.026, -93.792, -94.497, -94.812, -70.043, -69.171, -69.672, -69.411, -72.471, -72.016, -72.599, -71.315, -82.982, -82.614, -82.666, -82.616, -64.918, -62.862, -62.185, -61.894, -65.689, -65.814, -65.94, -65.684, -69.957, -69.473, -70.417, -70.201, -51.544, -49.713, -48.627, -48.561, -59.237, -60.137, -60.285, -60.233, -77.284, -78.336, -78.826, -78.845, -84.543, -84.644, -85.075, -85.022, -55.355, -52.545, -51.654, -51.546, -71.595, -73.689, -74.04, -74.139, -68.876, -68.076, -68.237, -68.015, -52.085, -50.204, -49.887, -48.851, -46.99, -45.341, -45.977, -46.385, -60.542, -62.524, -62.317, -62.653, -60.525, -60.537, -60.762, -60.458, -59.288, -59.079, -58.873, -58.914, -99.378, -103.851, -104.771, -104.988, -55.551, -50.327, -49.624, -48.749, -49.937, -49.531, -49.982, -49.604, -53.09, -52.959, -53.269, -53.508, -75.113, -77.434, -77.498, -78.085, -114.157, -117.82, -118.827, -119.149, -78.995, -74.782, -74.184, -73.578, -62.874, -61.285, -61.009, -60.934, -58.795, -56.487, -56.922, -57.549, -51.207, -50.628, -50.643, -50.541, -94.53, -98.613, -100.015, -100.453, -102.242, -101.691, 0.0, 0.0, 0.0, 0.0, 0.0, 0.0, -78.666, 0.0, 0.0, 6.697, 7.723, 0.0, 24.543, 17.608, 21.605, 0.0, 0.0, 0.0, 0.0, -4.699, 0.0, -4.464, 0.256, -2.383, 0.0, 0.0, 0.0, 1.229, 6.576, 3.764, 0.0, 8.41, 5.5, 0.0, 0.0, 15.833, 0.0, 20.355, 18.915, 19.788, 4.845, 11.478, 5.294, 0.0, 0.0, 0.0, 0.0, -6.559, 0.0, -9.395, -3.758, -5.593, 0.0, 0.0, 0.0, 1.324, 4.339, 1.781, -0.85, 2.792, 0.346, -1.854, -49.326, 0.0, -58.228, 0.0]
fit_rsq: 1.0
fit_intercept: 8.79377478947
except: True
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -49.9085584906
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [1.0, 538.0, 1184.0, 1894.0, 2501.0, 3038.0, 3576.0, 4113.0, 4650.0, 5187.0, 5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0, 12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0, 17721.0, 18258.0, 18795.0, 19332.0, 19869.0, 20406.0, 20944.0, 21482.0, 22019.0, 22556.0, 23146.0, 23682.0, 24219.0, 24757.0, 25347.0, 25884.0, 26421.0, 26958.0, 27549.0, 28086.0, 28623.0, 29160.0, 29747.0, 30284.0, 30821.0, 31358.0, 31946.0, 32483.0, 33020.0, 33557.0, 34146.0, 34682.0, 35219.0, 35757.0, 36346.0, 36883.0, 37420.0, 37958.0, 38546.0, 39083.0, 39619.0, 40157.0, 40747.0, 41285.0, 41822.0, 42359.0, 42949.0, 43486.0, 44022.0, 44559.0, 45146.0, 45683.0, 46221.0, 46757.0, 47345.0, 47882.0, 48419.0, 48957.0, 49545.0, 50082.0, 50620.0, 51157.0, 51746.0, 52283.0, 52820.0, 53358.0, 53946.0, 54483.0, 55021.0, 55558.0, 56148.0, 56684.0, 57221.0, 57758.0, 58348.0, 58885.0, 59422.0, 59959.0, 60546.0, 61082.0, 61618.0, 62155.0, 62742.0, 63279.0, 63817.0, 64354.0, 64941.0, 65478.0, 66015.0, 66552.0, 67140.0, 67677.0, 68214.0, 68751.0, 69338.0, 69875.0, 70412.0, 70949.0, 71538.0, 72074.0, 72611.0, 73148.0, 73737.0, 74274.0, 74811.0, 75348.0, 75933.0, 76469.0, 77007.0, 77544.0, 78129.0, 78666.0, 79203.0, 79740.0, 80326.0, 80862.0, 81399.0, 81936.0, 82523.0, 83060.0, 83597.0, 84133.0, 84720.0, 85257.0, 85794.0, 86332.0, 86918.0, 87455.0, 87993.0, 88530.0, 89118.0, 89655.0, 90191.0, 90729.0, 91312.0, 91849.0, 92386.0, 92922.0, 93507.0, 94044.0, 94581.0, 95117.0, 95701.0, 96238.0, 96774.0, 97311.0, 97896.0, 98432.0, 98968.0, 99505.0, 100091.0, 100627.0, 101163.0, 101700.0, 102285.0, 102823.0, 103359.0, 103895.0, 104481.0, 105018.0, 105555.0, 106091.0, 106676.0, 107212.0, 107748.0, 108285.0, 108870.0, 109406.0, 110050.0, 110652.0, 111303.0, 111904.0, 112505.0, 113105.0, 113756.0, 114293.0, 114930.0, 116655.0, 117192.0, 117729.0, 118916.0, 119454.0, 119990.0, 120577.0, 121179.0, 121779.0, 122381.0, 123567.0, 124104.0, 124705.0, 125828.0, 126364.0, 126902.0, 128089.0, 128690.0, 129291.0, 130415.0, 130952.0, 131488.0, 132676.0, 133213.0, 133749.0, 134937.0, 135538.0, 136074.0, 137263.0, 137799.0, 138336.0, 139459.0, 139995.0, 140532.0, 141119.0, 141720.0, 142321.0, 142922.0, 144111.0, 144648.0, 145249.0, 146372.0, 146909.0, 147445.0, 148635.0, 149235.0, 149836.0, 150961.0, 151497.0, 152034.0, 153159.0, 153695.0, 154231.0, 154818.0, 155355.0, 155956.0, 156493.0]
source_col: d(D_H)Mean

## 1 - Sample 1 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 1
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.0053
varray: [3.714, 2.982, 2.92, 2.879, 2.916, 2.917, 2.927, 2.972, 2.921, 2.905]
inst: mean_all
operation: mean_all(residuals)
tarray: [1.0, 538.0, 1184.0, 1894.0, 2501.0, 3038.0, 3576.0, 4113.0, 4650.0, 5187.0]
source_col: residuals
SAMP_residuals_STD | 0.25082178268
varray: [3.714, 2.982, 2.92, 2.879, 2.916, 2.917, 2.927, 2.972, 2.921, 2.905]
inst: std_all
operation: std_all(residuals)
tarray: [1.0, 538.0, 1184.0, 1894.0, 2501.0, 3038.0, 3576.0, 4113.0, 4650.0, 5187.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4561
varray: [0.609, 0.33, 0.443, 0.41, 0.478, 0.316, 0.428, 0.675, 0.519, 0.353]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [1.0, 538.0, 1184.0, 1894.0, 2501.0, 3038.0, 3576.0, 4113.0, 4650.0, 5187.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.11784024025
varray: [0.609, 0.33, 0.443, 0.41, 0.478, 0.316, 0.428, 0.675, 0.519, 0.353]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [1.0, 538.0, 1184.0, 1894.0, 2501.0, 3038.0, 3576.0, 4113.0, 4650.0, 5187.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1342.4394
varray: [1529.77, 1385.05, 1331.37, 1340.422, 1331.35, 1307.86, 1301.865, 1300.948, 1298.307, 1297.452]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [1.0, 538.0, 1184.0, 1894.0, 2501.0, 3038.0, 3576.0, 4113.0, 4650.0, 5187.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 71.2531334534
varray: [1529.77, 1385.05, 1331.37, 1340.422, 1331.35, 1307.86, 1301.865, 1300.948, 1298.307, 1297.452]
inst: std_all
operation: std_all(baseline_shift)
tarray: [1.0, 538.0, 1184.0, 1894.0, 2501.0, 3038.0, 3576.0, 4113.0, 4650.0, 5187.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.3342
varray: [-5.646, -0.911, 0.313, 0.77, 1.143, 1.578, 1.536, 1.257, 1.558, 1.744]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [1.0, 538.0, 1184.0, 1894.0, 2501.0, 3038.0, 3576.0, 4113.0, 4650.0, 5187.0]
source_col: slope_shift
SAMP_slope_shift_STD | 2.24811001115
varray: [-5.646, -0.911, 0.313, 0.77, 1.143, 1.578, 1.536, 1.257, 1.558, 1.744]
inst: std_all
operation: std_all(slope_shift)
tarray: [1.0, 538.0, 1184.0, 1894.0, 2501.0, 3038.0, 3576.0, 4113.0, 4650.0, 5187.0]
source_col: slope_shift
SAMP_d18O_STD | 0.401684356457
varray: [-6.107, -6.837, -7.245, -7.129, -7.203, -7.273, -7.395, -7.349, -7.42, -7.411]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1184.0, 1894.0, 2501.0, 3038.0, 3576.0, 4113.0, 4650.0, 5187.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.45553863264
varray: [-54.516, -57.248, -58.741, -58.353, -58.641, -58.81, -59.0, -58.974, -59.346, -59.315]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1184.0, 1894.0, 2501.0, 3038.0, 3576.0, 4113.0, 4650.0, 5187.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 115.3537
varray: [136.992, 114.069, 113.0, 111.762, 113.4, 112.223, 112.791, 114.387, 112.728, 112.185]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [1.0, 538.0, 1184.0, 1894.0, 2501.0, 3038.0, 3576.0, 4113.0, 4650.0, 5187.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 7.64752552645
varray: [136.992, 114.069, 113.0, 111.762, 113.4, 112.223, 112.791, 114.387, 112.728, 112.185]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [1.0, 538.0, 1184.0, 1894.0, 2501.0, 3038.0, 3576.0, 4113.0, 4650.0, 5187.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.0053
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4561
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1342.4394
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.3342
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.123919047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.123919047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 4.30068254309
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 4.6047952381
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 0
varray: [-6.107, -6.837, -7.245, -7.129, -7.203, -7.273, -7.395, -7.349, -7.42, -7.411]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [1.0, 538.0, 1184.0, 1894.0, 2501.0, 3038.0, 3576.0, 4113.0, 4650.0, 5187.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 0
varray: [-54.516, -57.248, -58.741, -58.353, -58.641, -58.81, -59.0, -58.974, -59.346, -59.315]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [1.0, 538.0, 1184.0, 1894.0, 2501.0, 3038.0, 3576.0, 4113.0, 4650.0, 5187.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.1369
varray: [-6.107, -6.837, -7.245, -7.129, -7.203, -7.273, -7.395, -7.349, -7.42, -7.411]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1184.0, 1894.0, 2501.0, 3038.0, 3576.0, 4113.0, 4650.0, 5187.0]
source_col: d(18_16)Mean
dD_mean | -58.2944
varray: [-54.516, -57.248, -58.741, -58.353, -58.641, -58.81, -59.0, -58.974, -59.346, -59.315]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1184.0, 1894.0, 2501.0, 3038.0, 3576.0, 4113.0, 4650.0, 5187.0]
source_col: d(D_H)Mean
d18O_stdev | 0.401684356457
varray: [-6.107, -6.837, -7.245, -7.129, -7.203, -7.273, -7.395, -7.349, -7.42, -7.411]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1184.0, 1894.0, 2501.0, 3038.0, 3576.0, 4113.0, 4650.0, 5187.0]
source_col: d(18_16)Mean
dD_stdev | 1.45553863264
varray: [-54.516, -57.248, -58.741, -58.353, -58.641, -58.81, -59.0, -58.974, -59.346, -59.315]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1184.0, 1894.0, 2501.0, 3038.0, 3576.0, 4113.0, 4650.0, 5187.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.1369
varray: [-6.107, -6.837, -7.245, -7.129, -7.203, -7.273, -7.395, -7.349, -7.42, -7.411]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1184.0, 1894.0, 2501.0, 3038.0, 3576.0, 4113.0, 4650.0, 5187.0]
source_col: d(18_16)Mean
measured_dD_mean | -58.2944
varray: [-54.516, -57.248, -58.741, -58.353, -58.641, -58.81, -59.0, -58.974, -59.346, -59.315]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1184.0, 1894.0, 2501.0, 3038.0, 3576.0, 4113.0, 4650.0, 5187.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.401684356457
varray: [-6.107, -6.837, -7.245, -7.129, -7.203, -7.273, -7.395, -7.349, -7.42, -7.411]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1184.0, 1894.0, 2501.0, 3038.0, 3576.0, 4113.0, 4650.0, 5187.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.45553863264
varray: [-54.516, -57.248, -58.741, -58.353, -58.641, -58.81, -59.0, -58.974, -59.346, -59.315]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1184.0, 1894.0, 2501.0, 3038.0, 3576.0, 4113.0, 4650.0, 5187.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.21374982294
actual_array: [ 1.93 -15.84]
varray: [-6.107, -6.837, -7.245, -7.129, -7.203, -7.273, -7.395, -7.349, -7.42, -7.411]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.1369
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [1.0, 538.0, 1184.0, 1894.0, 2501.0, 3038.0, 3576.0, 4113.0, 4650.0, 5187.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -67.2179931881
actual_array: [ 18.1 -119. ]
varray: [-54.516, -57.248, -58.741, -58.353, -58.641, -58.81, -59.0, -58.974, -59.346, -59.315]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -58.2944
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [1.0, 538.0, 1184.0, 1894.0, 2501.0, 3038.0, 3576.0, 4113.0, 4650.0, 5187.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1291.98808095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 47.4460190476
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 2 - Sample 2 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 2
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.93236363636
varray: [3.013, 2.932, 2.942, 2.958, 2.861, 2.907, 2.96, 2.947, 2.931, 2.882, 2.923]
inst: mean_all
operation: mean_all(residuals)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0]
source_col: residuals
SAMP_residuals_STD | 0.0407339483165
varray: [3.013, 2.932, 2.942, 2.958, 2.861, 2.907, 2.96, 2.947, 2.931, 2.882, 2.923]
inst: std_all
operation: std_all(residuals)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.411272727273
varray: [0.457, 0.439, 0.615, 0.487, 0.191, 0.216, 0.573, 0.504, 0.444, 0.214, 0.384]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.145683280378
varray: [0.457, 0.439, 0.615, 0.487, 0.191, 0.216, 0.573, 0.504, 0.444, 0.214, 0.384]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1295.942
varray: [1298.261, 1295.929, 1295.972, 1295.894, 1295.682, 1295.602, 1294.809, 1295.128, 1294.939, 1294.768, 1298.378]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.25813353822
varray: [1298.261, 1295.929, 1295.972, 1295.894, 1295.682, 1295.602, 1294.809, 1295.128, 1294.939, 1294.768, 1298.378]
inst: std_all
operation: std_all(baseline_shift)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.37736363636
varray: [1.361, 1.712, 1.425, 0.77, 1.247, 1.115, 1.361, 1.024, 1.253, 0.979, 2.904]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.565610338082
varray: [1.361, 1.712, 1.425, 0.77, 1.247, 1.115, 1.361, 1.024, 1.253, 0.979, 2.904]
inst: std_all
operation: std_all(slope_shift)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0]
source_col: slope_shift
SAMP_d18O_STD | 5.0736769919
varray: [4.01, 4.535, 4.691, 4.763, 4.819, 4.777, 4.806, 4.84, 4.859, 4.847, -12.113]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 39.7079553788
varray: [28.804, 36.832, 38.754, 39.807, 39.876, 40.238, 40.639, 40.948, 40.953, 40.984, -92.388]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.916181818
varray: [114.914, 111.635, 113.105, 112.029, 110.706, 111.185, 111.932, 111.965, 110.375, 110.705, 112.527]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.29263667116
varray: [114.914, 111.635, 113.105, 112.029, 110.706, 111.185, 111.932, 111.965, 110.375, 110.705, 112.527]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.93236363636
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.411272727273
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1295.942
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.37736363636
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0509826839827
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0509826839827
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.76938366794
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.16727705628
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 3.16672727273
varray: [4.01, 4.535, 4.691, 4.763, 4.819, 4.777, 4.806, 4.84, 4.859, 4.847, -12.113]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 26.8588181818
varray: [28.804, 36.832, 38.754, 39.807, 39.876, 40.238, 40.639, 40.948, 40.953, 40.984, -92.388]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0]
source_col: d(D_H)Mean
known_dD | 18.1
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 1.93
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 3.16672727273
varray: [4.01, 4.535, 4.691, 4.763, 4.819, 4.777, 4.806, 4.84, 4.859, 4.847, -12.113]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0]
source_col: d(18_16)Mean
dD_mean | 26.8588181818
varray: [28.804, 36.832, 38.754, 39.807, 39.876, 40.238, 40.639, 40.948, 40.953, 40.984, -92.388]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0]
source_col: d(D_H)Mean
d18O_stdev | 5.0736769919
varray: [4.01, 4.535, 4.691, 4.763, 4.819, 4.777, 4.806, 4.84, 4.859, 4.847, -12.113]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0]
source_col: d(18_16)Mean
dD_stdev | 39.7079553788
varray: [28.804, 36.832, 38.754, 39.807, 39.876, 40.238, 40.639, 40.948, 40.953, 40.984, -92.388]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0]
source_col: d(D_H)Mean
measured_d18O_mean | 3.16672727273
varray: [4.01, 4.535, 4.691, 4.763, 4.819, 4.777, 4.806, 4.84, 4.859, 4.847, -12.113]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0]
source_col: d(18_16)Mean
measured_dD_mean | 26.8588181818
varray: [28.804, 36.832, 38.754, 39.807, 39.876, 40.238, 40.639, 40.948, 40.953, 40.984, -92.388]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 5.0736769919
varray: [4.01, 4.535, 4.691, 4.763, 4.819, 4.777, 4.806, 4.84, 4.859, 4.847, -12.113]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0]
source_col: d(18_16)Mean
measured_dD_stdev | 39.7079553788
varray: [28.804, 36.832, 38.754, 39.807, 39.876, 40.238, 40.639, 40.948, 40.953, 40.984, -92.388]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.93
actual_array: [ 1.93 -15.84]
varray: [4.01, 4.535, 4.691, 4.763, 4.819, 4.777, 4.806, 4.84, 4.859, 4.847, -12.113]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 3.16672727273
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 18.1
actual_array: [ 18.1 -119. ]
varray: [28.804, 36.832, 38.754, 39.807, 39.876, 40.238, 40.639, 40.948, 40.953, 40.984, -92.388]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 26.8588181818
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [5834.0, 6712.0, 7315.0, 7852.0, 8388.0, 8930.0, 9466.0, 10003.0, 10541.0, 11105.0, 11707.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.06101732
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.948619047619
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 3 - Sample 3 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 3
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8253
varray: [2.847, 2.839, 2.832, 2.808, 2.818, 2.824, 2.81, 2.776, 2.868, 2.831]
inst: mean_all
operation: mean_all(residuals)
tarray: [12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: residuals
SAMP_residuals_STD | 0.0248955596219
varray: [2.847, 2.839, 2.832, 2.808, 2.818, 2.824, 2.81, 2.776, 2.868, 2.831]
inst: std_all
operation: std_all(residuals)
tarray: [12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3701
varray: [0.389, 0.415, 0.564, 0.27, 0.056, 0.37, 0.472, 0.281, 0.539, 0.345]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.147621475402
varray: [0.389, 0.415, 0.564, 0.27, 0.056, 0.37, 0.472, 0.281, 0.539, 0.345]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1293.9499
varray: [1294.586, 1294.372, 1294.343, 1293.844, 1294.249, 1293.664, 1293.094, 1293.282, 1293.696, 1294.369]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.50973837516
varray: [1294.586, 1294.372, 1294.343, 1293.844, 1294.249, 1293.664, 1293.094, 1293.282, 1293.696, 1294.369]
inst: std_all
operation: std_all(baseline_shift)
tarray: [12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.1347
varray: [2.354, 2.171, 1.88, 2.634, 2.722, 1.883, 1.895, 1.962, 2.001, 1.845]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.326758507498
varray: [2.354, 2.171, 1.88, 2.634, 2.722, 1.883, 1.895, 1.962, 2.001, 1.845]
inst: std_all
operation: std_all(slope_shift)
tarray: [12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: slope_shift
SAMP_d18O_STD | 0.151302640059
varray: [-12.906, -13.147, -13.225, -13.24, -13.302, -13.359, -13.35, -13.305, -13.351, -13.451]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.03503902228
varray: [-105.184, -107.874, -109.468, -110.097, -110.614, -111.064, -111.175, -111.049, -111.485, -111.754]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.4649
varray: [109.459, 109.774, 110.598, 107.877, 108.157, 109.897, 109.908, 108.05, 110.455, 110.474]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.05383389478
varray: [109.459, 109.774, 110.598, 107.877, 108.157, 109.897, 109.908, 108.05, 110.455, 110.474]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8253
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3701
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1293.9499
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.1347
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.056080952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.056080952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.94632203474
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.2840047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-12.906, -13.147, -13.225, -13.24, -13.302, -13.359, -13.35, -13.305, -13.351, -13.451]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-105.184, -107.874, -109.468, -110.097, -110.614, -111.064, -111.175, -111.049, -111.485, -111.754]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: d(D_H)Mean
known_dD | -119.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | -15.84
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -13.2636
varray: [-12.906, -13.147, -13.225, -13.24, -13.302, -13.359, -13.35, -13.305, -13.351, -13.451]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: d(18_16)Mean
dD_mean | -109.9764
varray: [-105.184, -107.874, -109.468, -110.097, -110.614, -111.064, -111.175, -111.049, -111.485, -111.754]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: d(D_H)Mean
d18O_stdev | 0.151302640059
varray: [-12.906, -13.147, -13.225, -13.24, -13.302, -13.359, -13.35, -13.305, -13.351, -13.451]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: d(18_16)Mean
dD_stdev | 2.03503902228
varray: [-105.184, -107.874, -109.468, -110.097, -110.614, -111.064, -111.175, -111.049, -111.485, -111.754]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: d(D_H)Mean
measured_d18O_mean | -13.2636
varray: [-12.906, -13.147, -13.225, -13.24, -13.302, -13.359, -13.35, -13.305, -13.351, -13.451]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: d(18_16)Mean
measured_dD_mean | -109.9764
varray: [-105.184, -107.874, -109.468, -110.097, -110.614, -111.064, -111.175, -111.049, -111.485, -111.754]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.151302640059
varray: [-12.906, -13.147, -13.225, -13.24, -13.302, -13.359, -13.35, -13.305, -13.351, -13.451]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.03503902228
varray: [-105.184, -107.874, -109.468, -110.097, -110.614, -111.064, -111.175, -111.049, -111.485, -111.754]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -15.84
actual_array: [ 1.93 -15.84]
varray: [-12.906, -13.147, -13.225, -13.24, -13.302, -13.359, -13.35, -13.305, -13.351, -13.451]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -13.2636
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -119.0
actual_array: [ 18.1 -119. ]
varray: [-105.184, -107.874, -109.468, -110.097, -110.614, -111.064, -111.175, -111.049, -111.485, -111.754]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -109.9764
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [12296.0, 12834.0, 13371.0, 13908.0, 14445.0, 14982.0, 15519.0, 16056.0, 16593.0, 17130.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.16808095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.04348095238
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 4 - Sample 4 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 4
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8409
varray: [2.836, 2.811, 2.835, 2.825, 2.948, 2.806, 2.826, 2.854, 2.826, 2.842]
inst: mean_all
operation: mean_all(residuals)
tarray: [17721.0, 18258.0, 18795.0, 19332.0, 19869.0, 20406.0, 20944.0, 21482.0, 22019.0, 22556.0]
source_col: residuals
SAMP_residuals_STD | 0.0401537323573
varray: [2.836, 2.811, 2.835, 2.825, 2.948, 2.806, 2.826, 2.854, 2.826, 2.842]
inst: std_all
operation: std_all(residuals)
tarray: [17721.0, 18258.0, 18795.0, 19332.0, 19869.0, 20406.0, 20944.0, 21482.0, 22019.0, 22556.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3634
varray: [0.088, 0.371, 0.343, 0.486, 0.565, 0.217, 0.326, 0.389, 0.522, 0.327]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [17721.0, 18258.0, 18795.0, 19332.0, 19869.0, 20406.0, 20944.0, 21482.0, 22019.0, 22556.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.142188607139
varray: [0.088, 0.371, 0.343, 0.486, 0.565, 0.217, 0.326, 0.389, 0.522, 0.327]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [17721.0, 18258.0, 18795.0, 19332.0, 19869.0, 20406.0, 20944.0, 21482.0, 22019.0, 22556.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1293.5021
varray: [1294.797, 1292.458, 1293.273, 1292.024, 1296.447, 1293.239, 1293.214, 1293.26, 1292.993, 1293.316]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [17721.0, 18258.0, 18795.0, 19332.0, 19869.0, 20406.0, 20944.0, 21482.0, 22019.0, 22556.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.25526734204
varray: [1294.797, 1292.458, 1293.273, 1292.024, 1296.447, 1293.239, 1293.214, 1293.26, 1292.993, 1293.316]
inst: std_all
operation: std_all(baseline_shift)
tarray: [17721.0, 18258.0, 18795.0, 19332.0, 19869.0, 20406.0, 20944.0, 21482.0, 22019.0, 22556.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.8283
varray: [2.18, 1.495, 2.262, 1.815, 1.473, 1.934, 1.632, 1.703, 2.304, 1.485]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [17721.0, 18258.0, 18795.0, 19332.0, 19869.0, 20406.0, 20944.0, 21482.0, 22019.0, 22556.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.326401950566
varray: [2.18, 1.495, 2.262, 1.815, 1.473, 1.934, 1.632, 1.703, 2.304, 1.485]
inst: std_all
operation: std_all(slope_shift)
tarray: [17721.0, 18258.0, 18795.0, 19332.0, 19869.0, 20406.0, 20944.0, 21482.0, 22019.0, 22556.0]
source_col: slope_shift
SAMP_d18O_STD | 0.113672824858
varray: [-8.017, -7.77, -7.704, -7.772, -7.641, -7.676, -7.661, -7.73, -7.625, -7.672]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [17721.0, 18258.0, 18795.0, 19332.0, 19869.0, 20406.0, 20944.0, 21482.0, 22019.0, 22556.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.94856833199
varray: [-66.554, -62.362, -61.345, -61.072, -60.439, -60.479, -60.432, -60.284, -60.196, -60.306]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [17721.0, 18258.0, 18795.0, 19332.0, 19869.0, 20406.0, 20944.0, 21482.0, 22019.0, 22556.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.5122
varray: [109.571, 108.484, 108.38, 108.759, 113.201, 108.806, 109.183, 109.94, 109.283, 109.515]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [17721.0, 18258.0, 18795.0, 19332.0, 19869.0, 20406.0, 20944.0, 21482.0, 22019.0, 22556.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.38883362735
varray: [109.571, 108.484, 108.38, 108.759, 113.201, 108.806, 109.183, 109.94, 109.283, 109.515]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [17721.0, 18258.0, 18795.0, 19332.0, 19869.0, 20406.0, 20944.0, 21482.0, 22019.0, 22556.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8409
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3634
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1293.5021
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.8283
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.040480952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.040480952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.40491497133
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.2367047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-8.017, -7.77, -7.704, -7.772, -7.641, -7.676, -7.661, -7.73, -7.625, -7.672]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [17721.0, 18258.0, 18795.0, 19332.0, 19869.0, 20406.0, 20944.0, 21482.0, 22019.0, 22556.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-66.554, -62.362, -61.345, -61.072, -60.439, -60.479, -60.432, -60.284, -60.196, -60.306]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [17721.0, 18258.0, 18795.0, 19332.0, 19869.0, 20406.0, 20944.0, 21482.0, 22019.0, 22556.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.7268
varray: [-8.017, -7.77, -7.704, -7.772, -7.641, -7.676, -7.661, -7.73, -7.625, -7.672]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [17721.0, 18258.0, 18795.0, 19332.0, 19869.0, 20406.0, 20944.0, 21482.0, 22019.0, 22556.0]
source_col: d(18_16)Mean
dD_mean | -61.3469
varray: [-66.554, -62.362, -61.345, -61.072, -60.439, -60.479, -60.432, -60.284, -60.196, -60.306]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [17721.0, 18258.0, 18795.0, 19332.0, 19869.0, 20406.0, 20944.0, 21482.0, 22019.0, 22556.0]
source_col: d(D_H)Mean
d18O_stdev | 0.113672824858
varray: [-8.017, -7.77, -7.704, -7.772, -7.641, -7.676, -7.661, -7.73, -7.625, -7.672]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [17721.0, 18258.0, 18795.0, 19332.0, 19869.0, 20406.0, 20944.0, 21482.0, 22019.0, 22556.0]
source_col: d(18_16)Mean
dD_stdev | 1.94856833199
varray: [-66.554, -62.362, -61.345, -61.072, -60.439, -60.479, -60.432, -60.284, -60.196, -60.306]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [17721.0, 18258.0, 18795.0, 19332.0, 19869.0, 20406.0, 20944.0, 21482.0, 22019.0, 22556.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.7268
varray: [-8.017, -7.77, -7.704, -7.772, -7.641, -7.676, -7.661, -7.73, -7.625, -7.672]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [17721.0, 18258.0, 18795.0, 19332.0, 19869.0, 20406.0, 20944.0, 21482.0, 22019.0, 22556.0]
source_col: d(18_16)Mean
measured_dD_mean | -61.3469
varray: [-66.554, -62.362, -61.345, -61.072, -60.439, -60.479, -60.432, -60.284, -60.196, -60.306]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [17721.0, 18258.0, 18795.0, 19332.0, 19869.0, 20406.0, 20944.0, 21482.0, 22019.0, 22556.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.113672824858
varray: [-8.017, -7.77, -7.704, -7.772, -7.641, -7.676, -7.661, -7.73, -7.625, -7.672]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [17721.0, 18258.0, 18795.0, 19332.0, 19869.0, 20406.0, 20944.0, 21482.0, 22019.0, 22556.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.94856833199
varray: [-66.554, -62.362, -61.345, -61.072, -60.439, -60.479, -60.432, -60.284, -60.196, -60.306]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [17721.0, 18258.0, 18795.0, 19332.0, 19869.0, 20406.0, 20944.0, 21482.0, 22019.0, 22556.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.85174825268
actual_array: [ 1.93 -15.84]
varray: [-8.017, -7.77, -7.704, -7.772, -7.641, -7.676, -7.661, -7.73, -7.625, -7.672]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.7268
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [17721.0, 18258.0, 18795.0, 19332.0, 19869.0, 20406.0, 20944.0, 21482.0, 22019.0, 22556.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -70.2763999021
actual_array: [ 18.1 -119. ]
varray: [-66.554, -62.362, -61.345, -61.072, -60.439, -60.479, -60.432, -60.284, -60.196, -60.306]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -61.3469
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [17721.0, 18258.0, 18795.0, 19332.0, 19869.0, 20406.0, 20944.0, 21482.0, 22019.0, 22556.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.15248095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.49128095238
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 5 - Sample 5 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 5
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.7025
varray: [2.779, 2.731, 2.624, 2.676]
inst: mean_all
operation: mean_all(residuals)
tarray: [23146.0, 23682.0, 24219.0, 24757.0]
source_col: residuals
SAMP_residuals_STD | 0.0671540517517
varray: [2.779, 2.731, 2.624, 2.676]
inst: std_all
operation: std_all(residuals)
tarray: [23146.0, 23682.0, 24219.0, 24757.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.433
varray: [-0.289, -0.478, -0.467, -0.498]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [23146.0, 23682.0, 24219.0, 24757.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0968538417755
varray: [-0.289, -0.478, -0.467, -0.498]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [23146.0, 23682.0, 24219.0, 24757.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1325.81275
varray: [1326.754, 1327.776, 1323.925, 1324.796]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [23146.0, 23682.0, 24219.0, 24757.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.76424514075
varray: [1326.754, 1327.776, 1323.925, 1324.796]
inst: std_all
operation: std_all(baseline_shift)
tarray: [23146.0, 23682.0, 24219.0, 24757.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -1.73825
varray: [-1.877, -1.522, -1.732, -1.822]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [23146.0, 23682.0, 24219.0, 24757.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.15606489035
varray: [-1.877, -1.522, -1.732, -1.822]
inst: std_all
operation: std_all(slope_shift)
tarray: [23146.0, 23682.0, 24219.0, 24757.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0347658836601
varray: [-6.791, -6.874, -6.827, -6.816]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23146.0, 23682.0, 24219.0, 24757.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.50926659894
varray: [-82.99, -85.569, -86.014, -86.273]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23146.0, 23682.0, 24219.0, 24757.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 107.076
varray: [109.22, 108.02, 105.087, 105.977]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [23146.0, 23682.0, 24219.0, 24757.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.88430305418
varray: [109.22, 108.02, 105.087, 105.977]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [23146.0, 23682.0, 24219.0, 24757.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.7025
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.433
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1325.81275
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.73825
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.178880952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.178880952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.20816738006
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.6729047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-6.791, -6.874, -6.827, -6.816]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [23146.0, 23682.0, 24219.0, 24757.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-82.99, -85.569, -86.014, -86.273]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [23146.0, 23682.0, 24219.0, 24757.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.827
varray: [-6.791, -6.874, -6.827, -6.816]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23146.0, 23682.0, 24219.0, 24757.0]
source_col: d(18_16)Mean
dD_mean | -85.2115
varray: [-82.99, -85.569, -86.014, -86.273]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23146.0, 23682.0, 24219.0, 24757.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0347658836601
varray: [-6.791, -6.874, -6.827, -6.816]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23146.0, 23682.0, 24219.0, 24757.0]
source_col: d(18_16)Mean
dD_stdev | 1.50926659894
varray: [-82.99, -85.569, -86.014, -86.273]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23146.0, 23682.0, 24219.0, 24757.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.827
varray: [-6.791, -6.874, -6.827, -6.816]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23146.0, 23682.0, 24219.0, 24757.0]
source_col: d(18_16)Mean
measured_dD_mean | -85.2115
varray: [-82.99, -85.569, -86.014, -86.273]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23146.0, 23682.0, 24219.0, 24757.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0347658836601
varray: [-6.791, -6.874, -6.827, -6.816]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23146.0, 23682.0, 24219.0, 24757.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.50926659894
varray: [-82.99, -85.569, -86.014, -86.273]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23146.0, 23682.0, 24219.0, 24757.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.87858163618
actual_array: [ 1.93 -15.84]
varray: [-6.791, -6.874, -6.827, -6.816]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -6.827
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [23146.0, 23682.0, 24219.0, 24757.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -94.1871788921
actual_array: [ 18.1 -119. ]
varray: [-82.99, -85.569, -86.014, -86.273]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -85.2115
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [23146.0, 23682.0, 24219.0, 24757.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.29088095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 30.8193690476
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 6 - Sample 6 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 6
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.794
varray: [2.789, 2.763, 2.837, 2.787]
inst: mean_all
operation: mean_all(residuals)
tarray: [25347.0, 25884.0, 26421.0, 26958.0]
source_col: residuals
SAMP_residuals_STD | 0.031005375878
varray: [2.789, 2.763, 2.837, 2.787]
inst: std_all
operation: std_all(residuals)
tarray: [25347.0, 25884.0, 26421.0, 26958.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.35525
varray: [-0.313, -0.463, -0.365, -0.28]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [25347.0, 25884.0, 26421.0, 26958.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0799015018632
varray: [-0.313, -0.463, -0.365, -0.28]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [25347.0, 25884.0, 26421.0, 26958.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1314.42175
varray: [1316.046, 1312.006, 1315.475, 1314.16]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [25347.0, 25884.0, 26421.0, 26958.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.79368194412
varray: [1316.046, 1312.006, 1315.475, 1314.16]
inst: std_all
operation: std_all(baseline_shift)
tarray: [25347.0, 25884.0, 26421.0, 26958.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -1.1115
varray: [-0.892, -1.226, -1.157, -1.171]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [25347.0, 25884.0, 26421.0, 26958.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.149332961309
varray: [-0.892, -1.226, -1.157, -1.171]
inst: std_all
operation: std_all(slope_shift)
tarray: [25347.0, 25884.0, 26421.0, 26958.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0210930478278
varray: [-7.587, -7.618, -7.582, -7.568]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [25347.0, 25884.0, 26421.0, 26958.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.124103115728
varray: [-87.678, -87.937, -87.918, -87.922]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [25347.0, 25884.0, 26421.0, 26958.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.42925
varray: [111.337, 108.122, 111.452, 110.806]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [25347.0, 25884.0, 26421.0, 26958.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.5636890942
varray: [111.337, 108.122, 111.452, 110.806]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [25347.0, 25884.0, 26421.0, 26958.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.794
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.35525
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1314.42175
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.1115
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.087380952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.087380952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.03260671966
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.319654761905
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-7.587, -7.618, -7.582, -7.568]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [25347.0, 25884.0, 26421.0, 26958.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-87.678, -87.937, -87.918, -87.922]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [25347.0, 25884.0, 26421.0, 26958.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.58875
varray: [-7.587, -7.618, -7.582, -7.568]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [25347.0, 25884.0, 26421.0, 26958.0]
source_col: d(18_16)Mean
dD_mean | -87.86375
varray: [-87.678, -87.937, -87.918, -87.922]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [25347.0, 25884.0, 26421.0, 26958.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0210930478278
varray: [-7.587, -7.618, -7.582, -7.568]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [25347.0, 25884.0, 26421.0, 26958.0]
source_col: d(18_16)Mean
dD_stdev | 0.124103115728
varray: [-87.678, -87.937, -87.918, -87.922]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [25347.0, 25884.0, 26421.0, 26958.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.58875
varray: [-7.587, -7.618, -7.582, -7.568]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [25347.0, 25884.0, 26421.0, 26958.0]
source_col: d(18_16)Mean
measured_dD_mean | -87.86375
varray: [-87.678, -87.937, -87.918, -87.922]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [25347.0, 25884.0, 26421.0, 26958.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0210930478278
varray: [-7.587, -7.618, -7.582, -7.568]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [25347.0, 25884.0, 26421.0, 26958.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.124103115728
varray: [-87.678, -87.937, -87.918, -87.922]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [25347.0, 25884.0, 26421.0, 26958.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.70244212753
actual_array: [ 1.93 -15.84]
varray: [-7.587, -7.618, -7.582, -7.568]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.58875
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [25347.0, 25884.0, 26421.0, 26958.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -96.8445611058
actual_array: [ 18.1 -119. ]
varray: [-87.678, -87.937, -87.918, -87.922]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -87.86375
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [25347.0, 25884.0, 26421.0, 26958.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.19938095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 19.4283690476
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 7 - Sample 7 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 7
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.5985
varray: [3.496, 3.52, 3.652, 3.726]
inst: mean_all
operation: mean_all(residuals)
tarray: [27549.0, 28086.0, 28623.0, 29160.0]
source_col: residuals
SAMP_residuals_STD | 0.109219961545
varray: [3.496, 3.52, 3.652, 3.726]
inst: std_all
operation: std_all(residuals)
tarray: [27549.0, 28086.0, 28623.0, 29160.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -6.4925
varray: [-6.537, -6.6, -6.584, -6.249]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [27549.0, 28086.0, 28623.0, 29160.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.164520515438
varray: [-6.537, -6.6, -6.584, -6.249]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [27549.0, 28086.0, 28623.0, 29160.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1931.42775
varray: [1913.458, 1920.035, 1944.877, 1947.341]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [27549.0, 28086.0, 28623.0, 29160.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 17.1932207449
varray: [1913.458, 1920.035, 1944.877, 1947.341]
inst: std_all
operation: std_all(baseline_shift)
tarray: [27549.0, 28086.0, 28623.0, 29160.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -39.56375
varray: [-38.559, -38.986, -40.306, -40.404]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [27549.0, 28086.0, 28623.0, 29160.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.930998165054
varray: [-38.559, -38.986, -40.306, -40.404]
inst: std_all
operation: std_all(slope_shift)
tarray: [27549.0, 28086.0, 28623.0, 29160.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0373273804778
varray: [-0.831, -0.885, -0.809, -0.803]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [27549.0, 28086.0, 28623.0, 29160.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.583440585378
varray: [-84.193, -85.44, -85.029, -85.421]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [27549.0, 28086.0, 28623.0, 29160.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.90225
varray: [106.571, 106.722, 110.566, 111.75]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [27549.0, 28086.0, 28623.0, 29160.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.64990319006
varray: [106.571, 106.722, 110.566, 111.75]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [27549.0, 28086.0, 28623.0, 29160.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.5985
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 6.4925
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1931.42775
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 39.56375
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.717119047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.717119047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 24.8880331851
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.8466547619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-0.831, -0.885, -0.809, -0.803]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [27549.0, 28086.0, 28623.0, 29160.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-84.193, -85.44, -85.029, -85.421]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [27549.0, 28086.0, 28623.0, 29160.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.832
varray: [-0.831, -0.885, -0.809, -0.803]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [27549.0, 28086.0, 28623.0, 29160.0]
source_col: d(18_16)Mean
dD_mean | -85.02075
varray: [-84.193, -85.44, -85.029, -85.421]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [27549.0, 28086.0, 28623.0, 29160.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0373273804778
varray: [-0.831, -0.885, -0.809, -0.803]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [27549.0, 28086.0, 28623.0, 29160.0]
source_col: d(18_16)Mean
dD_stdev | 0.583440585378
varray: [-84.193, -85.44, -85.029, -85.421]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [27549.0, 28086.0, 28623.0, 29160.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.832
varray: [-0.831, -0.885, -0.809, -0.803]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [27549.0, 28086.0, 28623.0, 29160.0]
source_col: d(18_16)Mean
measured_dD_mean | -85.02075
varray: [-84.193, -85.44, -85.029, -85.421]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [27549.0, 28086.0, 28623.0, 29160.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0373273804778
varray: [-0.831, -0.885, -0.809, -0.803]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [27549.0, 28086.0, 28623.0, 29160.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.583440585378
varray: [-84.193, -85.44, -85.029, -85.421]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [27549.0, 28086.0, 28623.0, 29160.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.39476982697
actual_array: [ 1.93 -15.84]
varray: [-0.831, -0.885, -0.809, -0.803]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -0.832
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [27549.0, 28086.0, 28623.0, 29160.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -93.996059783
actual_array: [ 18.1 -119. ]
varray: [-84.193, -85.44, -85.029, -85.421]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -85.02075
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [27549.0, 28086.0, 28623.0, 29160.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1291.39488095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 636.434369048
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 8 - Sample 8 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 8
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.90925
varray: [2.86, 2.92, 2.946, 2.911]
inst: mean_all
operation: mean_all(residuals)
tarray: [29747.0, 30284.0, 30821.0, 31358.0]
source_col: residuals
SAMP_residuals_STD | 0.0360312364484
varray: [2.86, 2.92, 2.946, 2.911]
inst: std_all
operation: std_all(residuals)
tarray: [29747.0, 30284.0, 30821.0, 31358.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.326
varray: [-2.438, -2.316, -2.338, -2.212]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [29747.0, 30284.0, 30821.0, 31358.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0927074250892
varray: [-2.438, -2.316, -2.338, -2.212]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [29747.0, 30284.0, 30821.0, 31358.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1702.42675
varray: [1702.282, 1700.967, 1706.355, 1700.103]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [29747.0, 30284.0, 30821.0, 31358.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.76783819072
varray: [1702.282, 1700.967, 1706.355, 1700.103]
inst: std_all
operation: std_all(baseline_shift)
tarray: [29747.0, 30284.0, 30821.0, 31358.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -19.96175
varray: [-19.578, -20.3, -19.719, -20.25]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [29747.0, 30284.0, 30821.0, 31358.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.366830001499
varray: [-19.578, -20.3, -19.719, -20.25]
inst: std_all
operation: std_all(slope_shift)
tarray: [29747.0, 30284.0, 30821.0, 31358.0]
source_col: slope_shift
SAMP_d18O_STD | 0.157440729588
varray: [0.058, 0.34, 0.344, 0.413]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [29747.0, 30284.0, 30821.0, 31358.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.8893970536
varray: [-48.454, -43.661, -42.446, -42.307]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [29747.0, 30284.0, 30821.0, 31358.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.49375
varray: [108.978, 110.762, 110.478, 111.757]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [29747.0, 30284.0, 30821.0, 31358.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.14971079697
varray: [108.978, 110.762, 110.478, 111.757]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [29747.0, 30284.0, 30821.0, 31358.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.90925
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.326
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1702.42675
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 19.96175
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.027869047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.027869047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.9672114892
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.255154761905
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [0.058, 0.34, 0.344, 0.413]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [29747.0, 30284.0, 30821.0, 31358.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-48.454, -43.661, -42.446, -42.307]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [29747.0, 30284.0, 30821.0, 31358.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.28875
varray: [0.058, 0.34, 0.344, 0.413]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [29747.0, 30284.0, 30821.0, 31358.0]
source_col: d(18_16)Mean
dD_mean | -44.217
varray: [-48.454, -43.661, -42.446, -42.307]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [29747.0, 30284.0, 30821.0, 31358.0]
source_col: d(D_H)Mean
d18O_stdev | 0.157440729588
varray: [0.058, 0.34, 0.344, 0.413]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [29747.0, 30284.0, 30821.0, 31358.0]
source_col: d(18_16)Mean
dD_stdev | 2.8893970536
varray: [-48.454, -43.661, -42.446, -42.307]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [29747.0, 30284.0, 30821.0, 31358.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.28875
varray: [0.058, 0.34, 0.344, 0.413]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [29747.0, 30284.0, 30821.0, 31358.0]
source_col: d(18_16)Mean
measured_dD_mean | -44.217
varray: [-48.454, -43.661, -42.446, -42.307]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [29747.0, 30284.0, 30821.0, 31358.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.157440729588
varray: [0.058, 0.34, 0.344, 0.413]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [29747.0, 30284.0, 30821.0, 31358.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.8893970536
varray: [-48.454, -43.661, -42.446, -42.307]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [29747.0, 30284.0, 30821.0, 31358.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.18263770267
actual_array: [ 1.93 -15.84]
varray: [0.058, 0.34, 0.344, 0.413]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.28875
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [29747.0, 30284.0, 30821.0, 31358.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -53.1133528357
actual_array: [ 18.1 -119. ]
varray: [-48.454, -43.661, -42.446, -42.307]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -44.217
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [29747.0, 30284.0, 30821.0, 31358.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.08413095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 407.433369048
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 9 - Sample 9 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 9
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 4.2035
varray: [4.132, 4.175, 4.272, 4.235]
inst: mean_all
operation: mean_all(residuals)
tarray: [31946.0, 32483.0, 33020.0, 33557.0]
source_col: residuals
SAMP_residuals_STD | 0.0622066448755
varray: [4.132, 4.175, 4.272, 4.235]
inst: std_all
operation: std_all(residuals)
tarray: [31946.0, 32483.0, 33020.0, 33557.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -8.554
varray: [-8.571, -8.514, -8.288, -8.843]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [31946.0, 32483.0, 33020.0, 33557.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.22815637912
varray: [-8.571, -8.514, -8.288, -8.843]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [31946.0, 32483.0, 33020.0, 33557.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 2730.052
varray: [2700.869, 2731.112, 2748.084, 2740.143]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [31946.0, 32483.0, 33020.0, 33557.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 20.6539132208
varray: [2700.869, 2731.112, 2748.084, 2740.143]
inst: std_all
operation: std_all(baseline_shift)
tarray: [31946.0, 32483.0, 33020.0, 33557.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -72.586
varray: [-70.989, -72.603, -73.118, -73.634]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [31946.0, 32483.0, 33020.0, 33557.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.14484729695
varray: [-70.989, -72.603, -73.118, -73.634]
inst: std_all
operation: std_all(slope_shift)
tarray: [31946.0, 32483.0, 33020.0, 33557.0]
source_col: slope_shift
SAMP_d18O_STD | 0.205016056282
varray: [1.605, 1.324, 1.195, 1.149]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [31946.0, 32483.0, 33020.0, 33557.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.71661276531
varray: [-75.346, -79.685, -80.885, -81.261]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [31946.0, 32483.0, 33020.0, 33557.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 107.62675
varray: [105.044, 108.596, 108.678, 108.189]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [31946.0, 32483.0, 33020.0, 33557.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.73505953308
varray: [105.044, 108.596, 108.678, 108.189]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [31946.0, 32483.0, 33020.0, 33557.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 4.2035
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 8.554
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 2730.052
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 72.586
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 1.32211904762
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.32211904762
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 45.8849096829
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.1221547619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [1.605, 1.324, 1.195, 1.149]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [31946.0, 32483.0, 33020.0, 33557.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-75.346, -79.685, -80.885, -81.261]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [31946.0, 32483.0, 33020.0, 33557.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.31825
varray: [1.605, 1.324, 1.195, 1.149]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [31946.0, 32483.0, 33020.0, 33557.0]
source_col: d(18_16)Mean
dD_mean | -79.29425
varray: [-75.346, -79.685, -80.885, -81.261]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [31946.0, 32483.0, 33020.0, 33557.0]
source_col: d(D_H)Mean
d18O_stdev | 0.205016056282
varray: [1.605, 1.324, 1.195, 1.149]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [31946.0, 32483.0, 33020.0, 33557.0]
source_col: d(18_16)Mean
dD_stdev | 2.71661276531
varray: [-75.346, -79.685, -80.885, -81.261]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [31946.0, 32483.0, 33020.0, 33557.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.31825
varray: [1.605, 1.324, 1.195, 1.149]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [31946.0, 32483.0, 33020.0, 33557.0]
source_col: d(18_16)Mean
measured_dD_mean | -79.29425
varray: [-75.346, -79.685, -80.885, -81.261]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [31946.0, 32483.0, 33020.0, 33557.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.205016056282
varray: [1.605, 1.324, 1.195, 1.149]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [31946.0, 32483.0, 33020.0, 33557.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.71661276531
varray: [-75.346, -79.685, -80.885, -81.261]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [31946.0, 32483.0, 33020.0, 33557.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.0691957914854
actual_array: [ 1.93 -15.84]
varray: [1.605, 1.324, 1.195, 1.149]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 1.31825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [31946.0, 32483.0, 33020.0, 33557.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -88.2584787682
actual_array: [ 18.1 -119. ]
varray: [-75.346, -79.685, -80.885, -81.261]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -79.29425
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [31946.0, 32483.0, 33020.0, 33557.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1290.78988095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1435.05861905
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 10 - Sample 10 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 10
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.851
varray: [2.906, 2.839, 2.799, 2.86]
inst: mean_all
operation: mean_all(residuals)
tarray: [34146.0, 34682.0, 35219.0, 35757.0]
source_col: residuals
SAMP_residuals_STD | 0.0445495978283
varray: [2.906, 2.839, 2.799, 2.86]
inst: std_all
operation: std_all(residuals)
tarray: [34146.0, 34682.0, 35219.0, 35757.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3175
varray: [0.048, 0.422, 0.405, 0.395]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [34146.0, 34682.0, 35219.0, 35757.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.180012036635
varray: [0.048, 0.422, 0.405, 0.395]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [34146.0, 34682.0, 35219.0, 35757.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1304.58225
varray: [1319.415, 1302.388, 1298.882, 1297.644]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [34146.0, 34682.0, 35219.0, 35757.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 10.090546215
varray: [1319.415, 1302.388, 1298.882, 1297.644]
inst: std_all
operation: std_all(baseline_shift)
tarray: [34146.0, 34682.0, 35219.0, 35757.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.05675
varray: [0.422, 1.148, 1.335, 1.322]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [34146.0, 34682.0, 35219.0, 35757.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.431669144137
varray: [0.422, 1.148, 1.335, 1.322]
inst: std_all
operation: std_all(slope_shift)
tarray: [34146.0, 34682.0, 35219.0, 35757.0]
source_col: slope_shift
SAMP_d18O_STD | 0.076151165454
varray: [-5.023, -4.894, -4.849, -4.884]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34146.0, 34682.0, 35219.0, 35757.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.38051617762
varray: [-73.465, -71.082, -70.574, -70.58]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34146.0, 34682.0, 35219.0, 35757.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.1645
varray: [112.211, 109.569, 109.748, 109.13]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [34146.0, 34682.0, 35219.0, 35757.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.38881784263
varray: [112.211, 109.569, 109.748, 109.13]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [34146.0, 34682.0, 35219.0, 35757.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.851
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3175
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1304.58225
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.05675
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.030380952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.030380952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.05438860335
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.584404761905
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-5.023, -4.894, -4.849, -4.884]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [34146.0, 34682.0, 35219.0, 35757.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-73.465, -71.082, -70.574, -70.58]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [34146.0, 34682.0, 35219.0, 35757.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.9125
varray: [-5.023, -4.894, -4.849, -4.884]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [34146.0, 34682.0, 35219.0, 35757.0]
source_col: d(18_16)Mean
dD_mean | -71.42525
varray: [-73.465, -71.082, -70.574, -70.58]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [34146.0, 34682.0, 35219.0, 35757.0]
source_col: d(D_H)Mean
d18O_stdev | 0.076151165454
varray: [-5.023, -4.894, -4.849, -4.884]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34146.0, 34682.0, 35219.0, 35757.0]
source_col: d(18_16)Mean
dD_stdev | 1.38051617762
varray: [-73.465, -71.082, -70.574, -70.58]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34146.0, 34682.0, 35219.0, 35757.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.9125
varray: [-5.023, -4.894, -4.849, -4.884]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [34146.0, 34682.0, 35219.0, 35757.0]
source_col: d(18_16)Mean
measured_dD_mean | -71.42525
varray: [-73.465, -71.082, -70.574, -70.58]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [34146.0, 34682.0, 35219.0, 35757.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.076151165454
varray: [-5.023, -4.894, -4.849, -4.884]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34146.0, 34682.0, 35219.0, 35757.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.38051617762
varray: [-73.465, -71.082, -70.574, -70.58]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34146.0, 34682.0, 35219.0, 35757.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.8079798499
actual_array: [ 1.93 -15.84]
varray: [-5.023, -4.894, -4.849, -4.884]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -4.9125
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [34146.0, 34682.0, 35219.0, 35757.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -80.374251927
actual_array: [ 18.1 -119. ]
varray: [-73.465, -71.082, -70.574, -70.58]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -71.42525
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [34146.0, 34682.0, 35219.0, 35757.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.14238095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 9.58886904762
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 11 - Sample 11 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 11
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.79325
varray: [2.809, 2.763, 2.766, 2.835]
inst: mean_all
operation: mean_all(residuals)
tarray: [36346.0, 36883.0, 37420.0, 37958.0]
source_col: residuals
SAMP_residuals_STD | 0.0348747759849
varray: [2.809, 2.763, 2.766, 2.835]
inst: std_all
operation: std_all(residuals)
tarray: [36346.0, 36883.0, 37420.0, 37958.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.38575
varray: [0.508, 0.279, 0.332, 0.424]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [36346.0, 36883.0, 37420.0, 37958.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.101147993884
varray: [0.508, 0.279, 0.332, 0.424]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [36346.0, 36883.0, 37420.0, 37958.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1293.995
varray: [1293.877, 1294.331, 1292.303, 1295.469]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [36346.0, 36883.0, 37420.0, 37958.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.31178758443
varray: [1293.877, 1294.331, 1292.303, 1295.469]
inst: std_all
operation: std_all(baseline_shift)
tarray: [36346.0, 36883.0, 37420.0, 37958.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.704
varray: [2.063, 1.656, 2.053, 1.044]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [36346.0, 36883.0, 37420.0, 37958.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.479091501351
varray: [2.063, 1.656, 2.053, 1.044]
inst: std_all
operation: std_all(slope_shift)
tarray: [36346.0, 36883.0, 37420.0, 37958.0]
source_col: slope_shift
SAMP_d18O_STD | 0.107602276928
varray: [-7.951, -8.137, -8.194, -8.151]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [36346.0, 36883.0, 37420.0, 37958.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.815933156984
varray: [-84.755, -86.056, -86.542, -86.406]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [36346.0, 36883.0, 37420.0, 37958.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.0485
varray: [108.587, 107.797, 105.503, 110.307]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [36346.0, 36883.0, 37420.0, 37958.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.99446124722
varray: [108.587, 107.797, 105.503, 110.307]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [36346.0, 36883.0, 37420.0, 37958.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.79325
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.38575
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1293.995
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.704
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.088130952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.088130952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.0586359054
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.7004047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-7.951, -8.137, -8.194, -8.151]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [36346.0, 36883.0, 37420.0, 37958.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-84.755, -86.056, -86.542, -86.406]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [36346.0, 36883.0, 37420.0, 37958.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.10825
varray: [-7.951, -8.137, -8.194, -8.151]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [36346.0, 36883.0, 37420.0, 37958.0]
source_col: d(18_16)Mean
dD_mean | -85.93975
varray: [-84.755, -86.056, -86.542, -86.406]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [36346.0, 36883.0, 37420.0, 37958.0]
source_col: d(D_H)Mean
d18O_stdev | 0.107602276928
varray: [-7.951, -8.137, -8.194, -8.151]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [36346.0, 36883.0, 37420.0, 37958.0]
source_col: d(18_16)Mean
dD_stdev | 0.815933156984
varray: [-84.755, -86.056, -86.542, -86.406]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [36346.0, 36883.0, 37420.0, 37958.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.10825
varray: [-7.951, -8.137, -8.194, -8.151]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [36346.0, 36883.0, 37420.0, 37958.0]
source_col: d(18_16)Mean
measured_dD_mean | -85.93975
varray: [-84.755, -86.056, -86.542, -86.406]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [36346.0, 36883.0, 37420.0, 37958.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.107602276928
varray: [-7.951, -8.137, -8.194, -8.151]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [36346.0, 36883.0, 37420.0, 37958.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.815933156984
varray: [-84.755, -86.056, -86.542, -86.406]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [36346.0, 36883.0, 37420.0, 37958.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.2643003819
actual_array: [ 1.93 -15.84]
varray: [-7.951, -8.137, -8.194, -8.151]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -8.10825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [36346.0, 36883.0, 37420.0, 37958.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -94.9168380861
actual_array: [ 18.1 -119. ]
varray: [-84.755, -86.056, -86.542, -86.406]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -85.93975
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [36346.0, 36883.0, 37420.0, 37958.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.20013095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.998380952381
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 12 - Sample 12 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 12
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 4.31575
varray: [4.311, 4.306, 4.191, 4.455]
inst: mean_all
operation: mean_all(residuals)
tarray: [38546.0, 39083.0, 39619.0, 40157.0]
source_col: residuals
SAMP_residuals_STD | 0.10812145948
varray: [4.311, 4.306, 4.191, 4.455]
inst: std_all
operation: std_all(residuals)
tarray: [38546.0, 39083.0, 39619.0, 40157.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -8.1065
varray: [-8.062, -7.962, -8.331, -8.071]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [38546.0, 39083.0, 39619.0, 40157.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.157608163917
varray: [-8.062, -7.962, -8.331, -8.071]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [38546.0, 39083.0, 39619.0, 40157.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1858.44025
varray: [1859.892, 1854.173, 1845.319, 1874.377]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [38546.0, 39083.0, 39619.0, 40157.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 12.1992377187
varray: [1859.892, 1854.173, 1845.319, 1874.377]
inst: std_all
operation: std_all(baseline_shift)
tarray: [38546.0, 39083.0, 39619.0, 40157.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -41.51025
varray: [-41.419, -41.351, -40.61, -42.661]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [38546.0, 39083.0, 39619.0, 40157.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.850169149836
varray: [-41.419, -41.351, -40.61, -42.661]
inst: std_all
operation: std_all(slope_shift)
tarray: [38546.0, 39083.0, 39619.0, 40157.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0670789336031
varray: [-0.016, -0.13, -0.119, -0.002]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [38546.0, 39083.0, 39619.0, 40157.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.38623648175
varray: [-90.222, -92.735, -93.263, -92.877]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [38546.0, 39083.0, 39619.0, 40157.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 121.72075
varray: [123.43, 121.211, 117.588, 124.654]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [38546.0, 39083.0, 39619.0, 40157.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.10187893327
varray: [123.43, 121.211, 117.588, 124.654]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [38546.0, 39083.0, 39619.0, 40157.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 4.31575
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 8.1065
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1858.44025
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 41.51025
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 1.43436904762
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.43436904762
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 49.7806111488
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 10.9718452381
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-0.016, -0.13, -0.119, -0.002]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [38546.0, 39083.0, 39619.0, 40157.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-90.222, -92.735, -93.263, -92.877]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [38546.0, 39083.0, 39619.0, 40157.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.06675
varray: [-0.016, -0.13, -0.119, -0.002]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [38546.0, 39083.0, 39619.0, 40157.0]
source_col: d(18_16)Mean
dD_mean | -92.27425
varray: [-90.222, -92.735, -93.263, -92.877]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [38546.0, 39083.0, 39619.0, 40157.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0670789336031
varray: [-0.016, -0.13, -0.119, -0.002]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [38546.0, 39083.0, 39619.0, 40157.0]
source_col: d(18_16)Mean
dD_stdev | 1.38623648175
varray: [-90.222, -92.735, -93.263, -92.877]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [38546.0, 39083.0, 39619.0, 40157.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.06675
varray: [-0.016, -0.13, -0.119, -0.002]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [38546.0, 39083.0, 39619.0, 40157.0]
source_col: d(18_16)Mean
measured_dD_mean | -92.27425
varray: [-90.222, -92.735, -93.263, -92.877]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [38546.0, 39083.0, 39619.0, 40157.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0670789336031
varray: [-0.016, -0.13, -0.119, -0.002]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [38546.0, 39083.0, 39619.0, 40157.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.38623648175
varray: [-90.222, -92.735, -93.263, -92.877]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [38546.0, 39083.0, 39619.0, 40157.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.56712395758
actual_array: [ 1.93 -15.84]
varray: [-0.016, -0.13, -0.119, -0.002]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -0.06675
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [38546.0, 39083.0, 39619.0, 40157.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -101.263595606
actual_array: [ 18.1 -119. ]
varray: [-90.222, -92.735, -93.263, -92.877]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -92.27425
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [38546.0, 39083.0, 39619.0, 40157.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1290.67763095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 563.446869048
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 13 - Sample 13 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 13
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.72175
varray: [2.732, 2.702, 2.747, 2.706]
inst: mean_all
operation: mean_all(residuals)
tarray: [40747.0, 41285.0, 41822.0, 42359.0]
source_col: residuals
SAMP_residuals_STD | 0.0214534379529
varray: [2.732, 2.702, 2.747, 2.706]
inst: std_all
operation: std_all(residuals)
tarray: [40747.0, 41285.0, 41822.0, 42359.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.23875
varray: [0.086, 0.286, 0.475, 0.108]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [40747.0, 41285.0, 41822.0, 42359.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.181176479342
varray: [0.086, 0.286, 0.475, 0.108]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [40747.0, 41285.0, 41822.0, 42359.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1293.89725
varray: [1300.272, 1292.517, 1291.372, 1291.428]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [40747.0, 41285.0, 41822.0, 42359.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 4.28239071664
varray: [1300.272, 1292.517, 1291.372, 1291.428]
inst: std_all
operation: std_all(baseline_shift)
tarray: [40747.0, 41285.0, 41822.0, 42359.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.615
varray: [1.027, 1.494, 2.155, 1.784]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [40747.0, 41285.0, 41822.0, 42359.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.476286328448
varray: [1.027, 1.494, 2.155, 1.784]
inst: std_all
operation: std_all(slope_shift)
tarray: [40747.0, 41285.0, 41822.0, 42359.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0522844782576
varray: [-7.754, -7.667, -7.63, -7.695]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [40747.0, 41285.0, 41822.0, 42359.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.55466806389
varray: [-66.852, -62.719, -61.621, -61.319]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [40747.0, 41285.0, 41822.0, 42359.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 105.57275
varray: [107.095, 104.971, 105.25, 104.975]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [40747.0, 41285.0, 41822.0, 42359.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.02320098221
varray: [107.095, 104.971, 105.25, 104.975]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [40747.0, 41285.0, 41822.0, 42359.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.72175
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.23875
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1293.89725
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.615
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.159630952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.159630952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.54008494604
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.1761547619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-7.754, -7.667, -7.63, -7.695]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [40747.0, 41285.0, 41822.0, 42359.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-66.852, -62.719, -61.621, -61.319]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [40747.0, 41285.0, 41822.0, 42359.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.6865
varray: [-7.754, -7.667, -7.63, -7.695]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [40747.0, 41285.0, 41822.0, 42359.0]
source_col: d(18_16)Mean
dD_mean | -63.12775
varray: [-66.852, -62.719, -61.621, -61.319]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [40747.0, 41285.0, 41822.0, 42359.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0522844782576
varray: [-7.754, -7.667, -7.63, -7.695]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [40747.0, 41285.0, 41822.0, 42359.0]
source_col: d(18_16)Mean
dD_stdev | 2.55466806389
varray: [-66.852, -62.719, -61.621, -61.319]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [40747.0, 41285.0, 41822.0, 42359.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.6865
varray: [-7.754, -7.667, -7.63, -7.695]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [40747.0, 41285.0, 41822.0, 42359.0]
source_col: d(18_16)Mean
measured_dD_mean | -63.12775
varray: [-66.852, -62.719, -61.621, -61.319]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [40747.0, 41285.0, 41822.0, 42359.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0522844782576
varray: [-7.754, -7.667, -7.63, -7.695]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [40747.0, 41285.0, 41822.0, 42359.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.55466806389
varray: [-66.852, -62.719, -61.621, -61.319]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [40747.0, 41285.0, 41822.0, 42359.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.80816232842
actual_array: [ 1.93 -15.84]
varray: [-7.754, -7.667, -7.63, -7.695]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.6865
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [40747.0, 41285.0, 41822.0, 42359.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -72.0606959207
actual_array: [ 18.1 -119. ]
varray: [-66.852, -62.719, -61.621, -61.319]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -63.12775
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [40747.0, 41285.0, 41822.0, 42359.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.27163095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.09613095238
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 14 - Sample 14 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 14
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.78025
varray: [2.811, 2.786, 2.781, 2.743]
inst: mean_all
operation: mean_all(residuals)
tarray: [42949.0, 43486.0, 44022.0, 44559.0]
source_col: residuals
SAMP_residuals_STD | 0.0280876603986
varray: [2.811, 2.786, 2.781, 2.743]
inst: std_all
operation: std_all(residuals)
tarray: [42949.0, 43486.0, 44022.0, 44559.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.21425
varray: [-0.027, -0.278, -0.328, -0.224]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [42949.0, 43486.0, 44022.0, 44559.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.13185945801
varray: [-0.027, -0.278, -0.328, -0.224]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [42949.0, 43486.0, 44022.0, 44559.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1303.507
varray: [1303.147, 1303.562, 1303.913, 1303.406]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [42949.0, 43486.0, 44022.0, 44559.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.320240534599
varray: [1303.147, 1303.562, 1303.913, 1303.406]
inst: std_all
operation: std_all(baseline_shift)
tarray: [42949.0, 43486.0, 44022.0, 44559.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -0.13975
varray: [-0.251, -0.334, -0.03, 0.056]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [42949.0, 43486.0, 44022.0, 44559.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.183005236719
varray: [-0.251, -0.334, -0.03, 0.056]
inst: std_all
operation: std_all(slope_shift)
tarray: [42949.0, 43486.0, 44022.0, 44559.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0138564064606
varray: [-7.643, -7.619, -7.651, -7.643]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [42949.0, 43486.0, 44022.0, 44559.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.72407083671
varray: [-91.026, -93.792, -94.497, -94.812]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [42949.0, 43486.0, 44022.0, 44559.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.4475
varray: [110.601, 108.979, 108.812, 109.398]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [42949.0, 43486.0, 44022.0, 44559.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.80754050466
varray: [110.601, 108.979, 108.812, 109.398]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [42949.0, 43486.0, 44022.0, 44559.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.78025
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.21425
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1303.507
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.13975
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.101130952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.101130952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.50980845825
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.3014047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-7.643, -7.619, -7.651, -7.643]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [42949.0, 43486.0, 44022.0, 44559.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-91.026, -93.792, -94.497, -94.812]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [42949.0, 43486.0, 44022.0, 44559.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.639
varray: [-7.643, -7.619, -7.651, -7.643]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [42949.0, 43486.0, 44022.0, 44559.0]
source_col: d(18_16)Mean
dD_mean | -93.53175
varray: [-91.026, -93.792, -94.497, -94.812]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [42949.0, 43486.0, 44022.0, 44559.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0138564064606
varray: [-7.643, -7.619, -7.651, -7.643]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [42949.0, 43486.0, 44022.0, 44559.0]
source_col: d(18_16)Mean
dD_stdev | 1.72407083671
varray: [-91.026, -93.792, -94.497, -94.812]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [42949.0, 43486.0, 44022.0, 44559.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.639
varray: [-7.643, -7.619, -7.651, -7.643]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [42949.0, 43486.0, 44022.0, 44559.0]
source_col: d(18_16)Mean
measured_dD_mean | -93.53175
varray: [-91.026, -93.792, -94.497, -94.812]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [42949.0, 43486.0, 44022.0, 44559.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0138564064606
varray: [-7.643, -7.619, -7.651, -7.643]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [42949.0, 43486.0, 44022.0, 44559.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.72407083671
varray: [-91.026, -93.792, -94.497, -94.812]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [42949.0, 43486.0, 44022.0, 44559.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.75678934078
actual_array: [ 1.93 -15.84]
varray: [-7.643, -7.619, -7.651, -7.643]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.639
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [42949.0, 43486.0, 44022.0, 44559.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -102.523528921
actual_array: [ 18.1 -119. ]
varray: [-91.026, -93.792, -94.497, -94.812]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -93.53175
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [42949.0, 43486.0, 44022.0, 44559.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.21313095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 8.51361904762
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 15 - Sample 15 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 15
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 4.45875
varray: [4.379, 4.542, 4.487, 4.427]
inst: mean_all
operation: mean_all(residuals)
tarray: [45146.0, 45683.0, 46221.0, 46757.0]
source_col: residuals
SAMP_residuals_STD | 0.0709383535191
varray: [4.379, 4.542, 4.487, 4.427]
inst: std_all
operation: std_all(residuals)
tarray: [45146.0, 45683.0, 46221.0, 46757.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -9.25725
varray: [-8.864, -9.411, -9.461, -9.293]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [45146.0, 45683.0, 46221.0, 46757.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.271463165089
varray: [-8.864, -9.411, -9.461, -9.293]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [45146.0, 45683.0, 46221.0, 46757.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 2300.75225
varray: [2278.375, 2313.245, 2309.022, 2302.367]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [45146.0, 45683.0, 46221.0, 46757.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 15.5756894203
varray: [2278.375, 2313.245, 2309.022, 2302.367]
inst: std_all
operation: std_all(baseline_shift)
tarray: [45146.0, 45683.0, 46221.0, 46757.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -60.04725
varray: [-59.33, -60.473, -60.534, -59.852]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [45146.0, 45683.0, 46221.0, 46757.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.568846420398
varray: [-59.33, -60.473, -60.534, -59.852]
inst: std_all
operation: std_all(slope_shift)
tarray: [45146.0, 45683.0, 46221.0, 46757.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0552599915551
varray: [4.339, 4.344, 4.448, 4.423]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45146.0, 45683.0, 46221.0, 46757.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.373516063912
varray: [-70.043, -69.171, -69.672, -69.411]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45146.0, 45683.0, 46221.0, 46757.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.26675
varray: [110.869, 112.593, 109.548, 108.057]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [45146.0, 45683.0, 46221.0, 46757.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.92991716143
varray: [110.869, 112.593, 109.548, 108.057]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [45146.0, 45683.0, 46221.0, 46757.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 4.45875
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 9.25725
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 2300.75225
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 60.04725
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 1.57736904762
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.57736904762
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 54.74350923
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.482154761905
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [4.339, 4.344, 4.448, 4.423]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [45146.0, 45683.0, 46221.0, 46757.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-70.043, -69.171, -69.672, -69.411]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [45146.0, 45683.0, 46221.0, 46757.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 4.3885
varray: [4.339, 4.344, 4.448, 4.423]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45146.0, 45683.0, 46221.0, 46757.0]
source_col: d(18_16)Mean
dD_mean | -69.57425
varray: [-70.043, -69.171, -69.672, -69.411]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45146.0, 45683.0, 46221.0, 46757.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0552599915551
varray: [4.339, 4.344, 4.448, 4.423]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45146.0, 45683.0, 46221.0, 46757.0]
source_col: d(18_16)Mean
dD_stdev | 0.373516063912
varray: [-70.043, -69.171, -69.672, -69.411]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45146.0, 45683.0, 46221.0, 46757.0]
source_col: d(D_H)Mean
measured_d18O_mean | 4.3885
varray: [4.339, 4.344, 4.448, 4.423]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45146.0, 45683.0, 46221.0, 46757.0]
source_col: d(18_16)Mean
measured_dD_mean | -69.57425
varray: [-70.043, -69.171, -69.672, -69.411]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45146.0, 45683.0, 46221.0, 46757.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0552599915551
varray: [4.339, 4.344, 4.448, 4.423]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45146.0, 45683.0, 46221.0, 46757.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.373516063912
varray: [-70.043, -69.171, -69.672, -69.411]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45146.0, 45683.0, 46221.0, 46757.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 3.25139189946
actual_array: [ 1.93 -15.84]
varray: [4.339, 4.344, 4.448, 4.423]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 4.3885
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [45146.0, 45683.0, 46221.0, 46757.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -78.5196701653
actual_array: [ 18.1 -119. ]
varray: [-70.043, -69.171, -69.672, -69.411]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -69.57425
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [45146.0, 45683.0, 46221.0, 46757.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1290.53463095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1005.75886905
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 16 - Sample 16 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 16
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.36
varray: [3.38, 3.339, 3.176, 3.545]
inst: mean_all
operation: mean_all(residuals)
tarray: [47345.0, 47882.0, 48419.0, 48957.0]
source_col: residuals
SAMP_residuals_STD | 0.151571765181
varray: [3.38, 3.339, 3.176, 3.545]
inst: std_all
operation: std_all(residuals)
tarray: [47345.0, 47882.0, 48419.0, 48957.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -5.99825
varray: [-6.11, -5.89, -5.651, -6.342]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [47345.0, 47882.0, 48419.0, 48957.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.296058974981
varray: [-6.11, -5.89, -5.651, -6.342]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [47345.0, 47882.0, 48419.0, 48957.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1858.22275
varray: [1862.398, 1846.889, 1822.203, 1901.401]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [47345.0, 47882.0, 48419.0, 48957.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 33.2047700627
varray: [1862.398, 1846.889, 1822.203, 1901.401]
inst: std_all
operation: std_all(baseline_shift)
tarray: [47345.0, 47882.0, 48419.0, 48957.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -38.53125
varray: [-38.882, -38.06, -36.742, -40.441]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [47345.0, 47882.0, 48419.0, 48957.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.54851140885
varray: [-38.882, -38.06, -36.742, -40.441]
inst: std_all
operation: std_all(slope_shift)
tarray: [47345.0, 47882.0, 48419.0, 48957.0]
source_col: slope_shift
SAMP_d18O_STD | 0.042554866545
varray: [0.102, 0.122, 0.024, 0.073]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [47345.0, 47882.0, 48419.0, 48957.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.580207649611
varray: [-72.471, -72.016, -72.599, -71.315]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [47345.0, 47882.0, 48419.0, 48957.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 105.39525
varray: [105.131, 105.046, 100.425, 110.979]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [47345.0, 47882.0, 48419.0, 48957.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 4.32332637175
varray: [105.131, 105.046, 100.425, 110.979]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [47345.0, 47882.0, 48419.0, 48957.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.36
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 5.99825
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1858.22275
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 38.53125
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.478619047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.478619047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 16.6107521195
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.3536547619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [0.102, 0.122, 0.024, 0.073]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [47345.0, 47882.0, 48419.0, 48957.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-72.471, -72.016, -72.599, -71.315]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [47345.0, 47882.0, 48419.0, 48957.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.08025
varray: [0.102, 0.122, 0.024, 0.073]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [47345.0, 47882.0, 48419.0, 48957.0]
source_col: d(18_16)Mean
dD_mean | -72.10025
varray: [-72.471, -72.016, -72.599, -71.315]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [47345.0, 47882.0, 48419.0, 48957.0]
source_col: d(D_H)Mean
d18O_stdev | 0.042554866545
varray: [0.102, 0.122, 0.024, 0.073]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [47345.0, 47882.0, 48419.0, 48957.0]
source_col: d(18_16)Mean
dD_stdev | 0.580207649611
varray: [-72.471, -72.016, -72.599, -71.315]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [47345.0, 47882.0, 48419.0, 48957.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.08025
varray: [0.102, 0.122, 0.024, 0.073]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [47345.0, 47882.0, 48419.0, 48957.0]
source_col: d(18_16)Mean
measured_dD_mean | -72.10025
varray: [-72.471, -72.016, -72.599, -71.315]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [47345.0, 47882.0, 48419.0, 48957.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.042554866545
varray: [0.102, 0.122, 0.024, 0.073]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [47345.0, 47882.0, 48419.0, 48957.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.580207649611
varray: [-72.471, -72.016, -72.599, -71.315]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [47345.0, 47882.0, 48419.0, 48957.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.40813808002
actual_array: [ 1.93 -15.84]
varray: [0.102, 0.122, 0.024, 0.073]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.08025
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [47345.0, 47882.0, 48419.0, 48957.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -81.05055808
actual_array: [ 18.1 -119. ]
varray: [-72.471, -72.016, -72.599, -71.315]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -72.10025
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [47345.0, 47882.0, 48419.0, 48957.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1291.63338095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 563.229369048
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 17 - Sample 17 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 17
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.79875
varray: [2.728, 2.788, 2.818, 2.861]
inst: mean_all
operation: mean_all(residuals)
tarray: [49545.0, 50082.0, 50620.0, 51157.0]
source_col: residuals
SAMP_residuals_STD | 0.0558770972761
varray: [2.728, 2.788, 2.818, 2.861]
inst: std_all
operation: std_all(residuals)
tarray: [49545.0, 50082.0, 50620.0, 51157.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0305
varray: [-1.172, -1.129, -1.003, -0.818]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [49545.0, 50082.0, 50620.0, 51157.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.158783920261
varray: [-1.172, -1.129, -1.003, -0.818]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [49545.0, 50082.0, 50620.0, 51157.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1324.8785
varray: [1329.383, 1324.849, 1322.708, 1322.574]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [49545.0, 50082.0, 50620.0, 51157.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.17874089329
varray: [1329.383, 1324.849, 1322.708, 1322.574]
inst: std_all
operation: std_all(baseline_shift)
tarray: [49545.0, 50082.0, 50620.0, 51157.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -3.946
varray: [-3.948, -4.152, -3.773, -3.911]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [49545.0, 50082.0, 50620.0, 51157.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.156624817106
varray: [-3.948, -4.152, -3.773, -3.911]
inst: std_all
operation: std_all(slope_shift)
tarray: [49545.0, 50082.0, 50620.0, 51157.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0184842275107
varray: [-5.845, -5.865, -5.83, -5.87]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [49545.0, 50082.0, 50620.0, 51157.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.176645596228
varray: [-82.982, -82.614, -82.666, -82.616]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [49545.0, 50082.0, 50620.0, 51157.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.48
varray: [108.669, 111.851, 111.352, 114.048]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [49545.0, 50082.0, 50620.0, 51157.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.20985444468
varray: [108.669, 111.851, 111.352, 114.048]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [49545.0, 50082.0, 50620.0, 51157.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.79875
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0305
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1324.8785
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 3.946
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.082630952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.082630952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.86775520997
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.731095238095
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-5.845, -5.865, -5.83, -5.87]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [49545.0, 50082.0, 50620.0, 51157.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-82.982, -82.614, -82.666, -82.616]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [49545.0, 50082.0, 50620.0, 51157.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.8525
varray: [-5.845, -5.865, -5.83, -5.87]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [49545.0, 50082.0, 50620.0, 51157.0]
source_col: d(18_16)Mean
dD_mean | -82.7195
varray: [-82.982, -82.614, -82.666, -82.616]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [49545.0, 50082.0, 50620.0, 51157.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0184842275107
varray: [-5.845, -5.865, -5.83, -5.87]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [49545.0, 50082.0, 50620.0, 51157.0]
source_col: d(18_16)Mean
dD_stdev | 0.176645596228
varray: [-82.982, -82.614, -82.666, -82.616]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [49545.0, 50082.0, 50620.0, 51157.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.8525
varray: [-5.845, -5.865, -5.83, -5.87]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [49545.0, 50082.0, 50620.0, 51157.0]
source_col: d(18_16)Mean
measured_dD_mean | -82.7195
varray: [-82.982, -82.614, -82.666, -82.616]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [49545.0, 50082.0, 50620.0, 51157.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0184842275107
varray: [-5.845, -5.865, -5.83, -5.87]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [49545.0, 50082.0, 50620.0, 51157.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.176645596228
varray: [-82.982, -82.614, -82.666, -82.616]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [49545.0, 50082.0, 50620.0, 51157.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.824624237
actual_array: [ 1.93 -15.84]
varray: [-5.845, -5.865, -5.83, -5.87]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -5.8525
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [49545.0, 50082.0, 50620.0, 51157.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -91.6903567689
actual_array: [ 18.1 -119. ]
varray: [-82.982, -82.614, -82.666, -82.616]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -82.7195
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [49545.0, 50082.0, 50620.0, 51157.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.19463095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 29.8851190476
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 18 - Sample 18 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 18
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.6875
varray: [2.647, 2.732, 2.702, 2.669]
inst: mean_all
operation: mean_all(residuals)
tarray: [51746.0, 52283.0, 52820.0, 53358.0]
source_col: residuals
SAMP_residuals_STD | 0.0372961123979
varray: [2.647, 2.732, 2.702, 2.669]
inst: std_all
operation: std_all(residuals)
tarray: [51746.0, 52283.0, 52820.0, 53358.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.962
varray: [-1.031, -0.774, -0.907, -1.136]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [51746.0, 52283.0, 52820.0, 53358.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.156424635741
varray: [-1.031, -0.774, -0.907, -1.136]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [51746.0, 52283.0, 52820.0, 53358.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1880.34875
varray: [1872.703, 1879.365, 1883.973, 1885.354]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [51746.0, 52283.0, 52820.0, 53358.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 5.70417895786
varray: [1872.703, 1879.365, 1883.973, 1885.354]
inst: std_all
operation: std_all(baseline_shift)
tarray: [51746.0, 52283.0, 52820.0, 53358.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -21.99025
varray: [-21.733, -21.277, -22.181, -22.77]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [51746.0, 52283.0, 52820.0, 53358.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.63752091469
varray: [-21.733, -21.277, -22.181, -22.77]
inst: std_all
operation: std_all(slope_shift)
tarray: [51746.0, 52283.0, 52820.0, 53358.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0843064845272
varray: [-4.535, -4.437, -4.37, -4.347]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [51746.0, 52283.0, 52820.0, 53358.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.36384979012
varray: [-64.918, -62.862, -62.185, -61.894]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [51746.0, 52283.0, 52820.0, 53358.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 102.25825
varray: [102.069, 101.794, 102.265, 102.905]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [51746.0, 52283.0, 52820.0, 53358.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.47246684187
varray: [102.069, 101.794, 102.265, 102.905]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [51746.0, 52283.0, 52820.0, 53358.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.6875
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.962
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1880.34875
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 21.99025
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.193880952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.193880952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.72875109488
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -8.4906547619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-4.535, -4.437, -4.37, -4.347]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [51746.0, 52283.0, 52820.0, 53358.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-64.918, -62.862, -62.185, -61.894]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [51746.0, 52283.0, 52820.0, 53358.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.42225
varray: [-4.535, -4.437, -4.37, -4.347]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [51746.0, 52283.0, 52820.0, 53358.0]
source_col: d(18_16)Mean
dD_mean | -62.96475
varray: [-64.918, -62.862, -62.185, -61.894]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [51746.0, 52283.0, 52820.0, 53358.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0843064845272
varray: [-4.535, -4.437, -4.37, -4.347]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [51746.0, 52283.0, 52820.0, 53358.0]
source_col: d(18_16)Mean
dD_stdev | 1.36384979012
varray: [-64.918, -62.862, -62.185, -61.894]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [51746.0, 52283.0, 52820.0, 53358.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.42225
varray: [-4.535, -4.437, -4.37, -4.347]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [51746.0, 52283.0, 52820.0, 53358.0]
source_col: d(18_16)Mean
measured_dD_mean | -62.96475
varray: [-64.918, -62.862, -62.185, -61.894]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [51746.0, 52283.0, 52820.0, 53358.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0843064845272
varray: [-4.535, -4.437, -4.37, -4.347]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [51746.0, 52283.0, 52820.0, 53358.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.36384979012
varray: [-64.918, -62.862, -62.185, -61.894]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [51746.0, 52283.0, 52820.0, 53358.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.27775654057
actual_array: [ 1.93 -15.84]
varray: [-4.535, -4.437, -4.37, -4.347]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -4.42225
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [51746.0, 52283.0, 52820.0, 53358.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -71.8973805089
actual_array: [ 18.1 -119. ]
varray: [-64.918, -62.862, -62.185, -61.894]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -62.96475
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [51746.0, 52283.0, 52820.0, 53358.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.30588095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 585.355369048
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 19 - Sample 19 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 19
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.69625
varray: [2.581, 2.771, 2.727, 2.706]
inst: mean_all
operation: mean_all(residuals)
tarray: [53946.0, 54483.0, 55021.0, 55558.0]
source_col: residuals
SAMP_residuals_STD | 0.0814672735929
varray: [2.581, 2.771, 2.727, 2.706]
inst: std_all
operation: std_all(residuals)
tarray: [53946.0, 54483.0, 55021.0, 55558.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.90725
varray: [-0.819, -0.893, -1.034, -0.883]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [53946.0, 54483.0, 55021.0, 55558.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0906361774716
varray: [-0.819, -0.893, -1.034, -0.883]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [53946.0, 54483.0, 55021.0, 55558.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1390.24625
varray: [1386.084, 1394.432, 1390.266, 1390.203]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [53946.0, 54483.0, 55021.0, 55558.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.40818078306
varray: [1386.084, 1394.432, 1390.266, 1390.203]
inst: std_all
operation: std_all(baseline_shift)
tarray: [53946.0, 54483.0, 55021.0, 55558.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -5.579
varray: [-5.17, -5.805, -5.992, -5.349]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [53946.0, 54483.0, 55021.0, 55558.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.383764684844
varray: [-5.17, -5.805, -5.992, -5.349]
inst: std_all
operation: std_all(slope_shift)
tarray: [53946.0, 54483.0, 55021.0, 55558.0]
source_col: slope_shift
SAMP_d18O_STD | 0.114863324579
varray: [-1.65, -1.445, -1.454, -1.386]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [53946.0, 54483.0, 55021.0, 55558.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.121436883469
varray: [-65.689, -65.814, -65.94, -65.684]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [53946.0, 54483.0, 55021.0, 55558.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 107.70425
varray: [103.392, 110.152, 108.289, 108.984]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [53946.0, 54483.0, 55021.0, 55558.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.97582889013
varray: [103.392, 110.152, 108.289, 108.984]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [53946.0, 54483.0, 55021.0, 55558.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.69625
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.90725
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1390.24625
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 5.579
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.185130952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.185130952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.42507726123
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.0446547619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-1.65, -1.445, -1.454, -1.386]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [53946.0, 54483.0, 55021.0, 55558.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-65.689, -65.814, -65.94, -65.684]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [53946.0, 54483.0, 55021.0, 55558.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.48375
varray: [-1.65, -1.445, -1.454, -1.386]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [53946.0, 54483.0, 55021.0, 55558.0]
source_col: d(18_16)Mean
dD_mean | -65.78175
varray: [-65.689, -65.814, -65.94, -65.684]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [53946.0, 54483.0, 55021.0, 55558.0]
source_col: d(D_H)Mean
d18O_stdev | 0.114863324579
varray: [-1.65, -1.445, -1.454, -1.386]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [53946.0, 54483.0, 55021.0, 55558.0]
source_col: d(18_16)Mean
dD_stdev | 0.121436883469
varray: [-65.689, -65.814, -65.94, -65.684]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [53946.0, 54483.0, 55021.0, 55558.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.48375
varray: [-1.65, -1.445, -1.454, -1.386]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [53946.0, 54483.0, 55021.0, 55558.0]
source_col: d(18_16)Mean
measured_dD_mean | -65.78175
varray: [-65.689, -65.814, -65.94, -65.684]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [53946.0, 54483.0, 55021.0, 55558.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.114863324579
varray: [-1.65, -1.445, -1.454, -1.386]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [53946.0, 54483.0, 55021.0, 55558.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.121436883469
varray: [-65.689, -65.814, -65.94, -65.684]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [53946.0, 54483.0, 55021.0, 55558.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.0996612943
actual_array: [ 1.93 -15.84]
varray: [-1.65, -1.445, -1.454, -1.386]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -1.48375
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [53946.0, 54483.0, 55021.0, 55558.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -74.7198315206
actual_array: [ 18.1 -119. ]
varray: [-65.689, -65.814, -65.94, -65.684]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -65.78175
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [53946.0, 54483.0, 55021.0, 55558.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.29713095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 95.2528690476
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 20 - Sample 20 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 20
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.66275
varray: [2.665, 2.831, 2.553, 2.602]
inst: mean_all
operation: mean_all(residuals)
tarray: [56148.0, 56684.0, 57221.0, 57758.0]
source_col: residuals
SAMP_residuals_STD | 0.12117308557
varray: [2.665, 2.831, 2.553, 2.602]
inst: std_all
operation: std_all(residuals)
tarray: [56148.0, 56684.0, 57221.0, 57758.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.27925
varray: [0.159, 0.363, 0.265, 0.33]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [56148.0, 56684.0, 57221.0, 57758.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0899124574239
varray: [0.159, 0.363, 0.265, 0.33]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [56148.0, 56684.0, 57221.0, 57758.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1288.5735
varray: [1290.226, 1295.955, 1283.429, 1284.684]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [56148.0, 56684.0, 57221.0, 57758.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 5.73908787759
varray: [1290.226, 1295.955, 1283.429, 1284.684]
inst: std_all
operation: std_all(baseline_shift)
tarray: [56148.0, 56684.0, 57221.0, 57758.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.30225
varray: [0.985, 1.391, 1.191, 1.642]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [56148.0, 56684.0, 57221.0, 57758.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.280672258456
varray: [0.985, 1.391, 1.191, 1.642]
inst: std_all
operation: std_all(slope_shift)
tarray: [56148.0, 56684.0, 57221.0, 57758.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0889770569679
varray: [-4.465, -4.535, -4.63, -4.659]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56148.0, 56684.0, 57221.0, 57758.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.405500513769
varray: [-69.957, -69.473, -70.417, -70.201]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56148.0, 56684.0, 57221.0, 57758.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 103.55
varray: [103.681, 109.818, 99.815, 100.886]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [56148.0, 56684.0, 57221.0, 57758.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 4.48523897542
varray: [103.681, 109.818, 99.815, 100.886]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [56148.0, 56684.0, 57221.0, 57758.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.66275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.27925
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1288.5735
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.30225
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.218630952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.218630952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -7.58771422433
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -7.1989047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-4.465, -4.535, -4.63, -4.659]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [56148.0, 56684.0, 57221.0, 57758.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-69.957, -69.473, -70.417, -70.201]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [56148.0, 56684.0, 57221.0, 57758.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.57225
varray: [-4.465, -4.535, -4.63, -4.659]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56148.0, 56684.0, 57221.0, 57758.0]
source_col: d(18_16)Mean
dD_mean | -70.012
varray: [-69.957, -69.473, -70.417, -70.201]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56148.0, 56684.0, 57221.0, 57758.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0889770569679
varray: [-4.465, -4.535, -4.63, -4.659]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56148.0, 56684.0, 57221.0, 57758.0]
source_col: d(18_16)Mean
dD_stdev | 0.405500513769
varray: [-69.957, -69.473, -70.417, -70.201]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56148.0, 56684.0, 57221.0, 57758.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.57225
varray: [-4.465, -4.535, -4.63, -4.659]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56148.0, 56684.0, 57221.0, 57758.0]
source_col: d(18_16)Mean
measured_dD_mean | -70.012
varray: [-69.957, -69.473, -70.417, -70.201]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56148.0, 56684.0, 57221.0, 57758.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0889770569679
varray: [-4.465, -4.535, -4.63, -4.659]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56148.0, 56684.0, 57221.0, 57758.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.405500513769
varray: [-69.957, -69.473, -70.417, -70.201]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56148.0, 56684.0, 57221.0, 57758.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.43998702787
actual_array: [ 1.93 -15.84]
varray: [-4.465, -4.535, -4.63, -4.659]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -4.57225
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [56148.0, 56684.0, 57221.0, 57758.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -78.9582672297
actual_array: [ 18.1 -119. ]
varray: [-69.957, -69.473, -70.417, -70.201]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -70.012
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [56148.0, 56684.0, 57221.0, 57758.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.33063095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -6.41988095238
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 21 - Sample 21 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 21
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.708
varray: [2.578, 2.693, 2.732, 2.829]
inst: mean_all
operation: mean_all(residuals)
tarray: [58348.0, 58885.0, 59422.0, 59959.0]
source_col: residuals
SAMP_residuals_STD | 0.103829989245
varray: [2.578, 2.693, 2.732, 2.829]
inst: std_all
operation: std_all(residuals)
tarray: [58348.0, 58885.0, 59422.0, 59959.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.092
varray: [-0.203, 0.19, -0.425, 0.07]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [58348.0, 58885.0, 59422.0, 59959.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.27627160549
varray: [-0.203, 0.19, -0.425, 0.07]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [58348.0, 58885.0, 59422.0, 59959.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1367.05725
varray: [1360.134, 1363.129, 1372.682, 1372.284]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [58348.0, 58885.0, 59422.0, 59959.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 6.38538087483
varray: [1360.134, 1363.129, 1372.682, 1372.284]
inst: std_all
operation: std_all(baseline_shift)
tarray: [58348.0, 58885.0, 59422.0, 59959.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -2.71825
varray: [-2.694, -2.55, -2.612, -3.017]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [58348.0, 58885.0, 59422.0, 59959.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.207715149504
varray: [-2.694, -2.55, -2.612, -3.017]
inst: std_all
operation: std_all(slope_shift)
tarray: [58348.0, 58885.0, 59422.0, 59959.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0971407912946
varray: [-1.0, -0.902, -0.791, -0.805]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [58348.0, 58885.0, 59422.0, 59959.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.39255744944
varray: [-51.544, -49.713, -48.627, -48.561]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [58348.0, 58885.0, 59422.0, 59959.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 105.57375
varray: [102.175, 104.268, 106.381, 109.471]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [58348.0, 58885.0, 59422.0, 59959.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.11430542015
varray: [102.175, 104.268, 106.381, 109.471]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [58348.0, 58885.0, 59422.0, 59959.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.708
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.092
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1367.05725
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.71825
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.173380952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.173380952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.01728668463
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.1751547619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-1.0, -0.902, -0.791, -0.805]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [58348.0, 58885.0, 59422.0, 59959.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-51.544, -49.713, -48.627, -48.561]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [58348.0, 58885.0, 59422.0, 59959.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.8745
varray: [-1.0, -0.902, -0.791, -0.805]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [58348.0, 58885.0, 59422.0, 59959.0]
source_col: d(18_16)Mean
dD_mean | -49.61125
varray: [-51.544, -49.713, -48.627, -48.561]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [58348.0, 58885.0, 59422.0, 59959.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0971407912946
varray: [-1.0, -0.902, -0.791, -0.805]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [58348.0, 58885.0, 59422.0, 59959.0]
source_col: d(18_16)Mean
dD_stdev | 1.39255744944
varray: [-51.544, -49.713, -48.627, -48.561]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [58348.0, 58885.0, 59422.0, 59959.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.8745
varray: [-1.0, -0.902, -0.791, -0.805]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [58348.0, 58885.0, 59422.0, 59959.0]
source_col: d(18_16)Mean
measured_dD_mean | -49.61125
varray: [-51.544, -49.713, -48.627, -48.561]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [58348.0, 58885.0, 59422.0, 59959.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0971407912946
varray: [-1.0, -0.902, -0.791, -0.805]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [58348.0, 58885.0, 59422.0, 59959.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.39255744944
varray: [-51.544, -49.713, -48.627, -48.561]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [58348.0, 58885.0, 59422.0, 59959.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.44073513171
actual_array: [ 1.93 -15.84]
varray: [-1.0, -0.902, -0.791, -0.805]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -0.8745
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [58348.0, 58885.0, 59422.0, 59959.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -58.518040933
actual_array: [ 18.1 -119. ]
varray: [-51.544, -49.713, -48.627, -48.561]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -49.61125
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [58348.0, 58885.0, 59422.0, 59959.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.28538095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 72.0638690476
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 22 - Sample 22 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 22
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.7195
varray: [2.667, 2.704, 2.785, 2.722]
inst: mean_all
operation: mean_all(residuals)
tarray: [60546.0, 61082.0, 61618.0, 62155.0]
source_col: residuals
SAMP_residuals_STD | 0.0493051721425
varray: [2.667, 2.704, 2.785, 2.722]
inst: std_all
operation: std_all(residuals)
tarray: [60546.0, 61082.0, 61618.0, 62155.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.362
varray: [0.314, 0.195, 0.56, 0.379]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [60546.0, 61082.0, 61618.0, 62155.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.15240953601
varray: [0.314, 0.195, 0.56, 0.379]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [60546.0, 61082.0, 61618.0, 62155.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1289.99725
varray: [1289.097, 1289.904, 1291.447, 1289.541]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [60546.0, 61082.0, 61618.0, 62155.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.02128754521
varray: [1289.097, 1289.904, 1291.447, 1289.541]
inst: std_all
operation: std_all(baseline_shift)
tarray: [60546.0, 61082.0, 61618.0, 62155.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.85025
varray: [1.991, 1.767, 1.807, 1.836]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [60546.0, 61082.0, 61618.0, 62155.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0980046767592
varray: [1.991, 1.767, 1.807, 1.836]
inst: std_all
operation: std_all(slope_shift)
tarray: [60546.0, 61082.0, 61618.0, 62155.0]
source_col: slope_shift
SAMP_d18O_STD | 0.151837632577
varray: [-7.231, -7.499, -7.539, -7.555]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [60546.0, 61082.0, 61618.0, 62155.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.494481546673
varray: [-59.237, -60.137, -60.285, -60.233]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [60546.0, 61082.0, 61618.0, 62155.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 105.5865
varray: [103.532, 104.986, 108.351, 105.477]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [60546.0, 61082.0, 61618.0, 62155.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.01957264456
varray: [103.532, 104.986, 108.351, 105.477]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [60546.0, 61082.0, 61618.0, 62155.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.7195
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.362
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1289.99725
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.85025
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.161880952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.161880952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.61817250326
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.1624047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-7.231, -7.499, -7.539, -7.555]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [60546.0, 61082.0, 61618.0, 62155.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-59.237, -60.137, -60.285, -60.233]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [60546.0, 61082.0, 61618.0, 62155.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.456
varray: [-7.231, -7.499, -7.539, -7.555]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [60546.0, 61082.0, 61618.0, 62155.0]
source_col: d(18_16)Mean
dD_mean | -59.973
varray: [-59.237, -60.137, -60.285, -60.233]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [60546.0, 61082.0, 61618.0, 62155.0]
source_col: d(D_H)Mean
d18O_stdev | 0.151837632577
varray: [-7.231, -7.499, -7.539, -7.555]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [60546.0, 61082.0, 61618.0, 62155.0]
source_col: d(18_16)Mean
dD_stdev | 0.494481546673
varray: [-59.237, -60.137, -60.285, -60.233]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [60546.0, 61082.0, 61618.0, 62155.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.456
varray: [-7.231, -7.499, -7.539, -7.555]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [60546.0, 61082.0, 61618.0, 62155.0]
source_col: d(18_16)Mean
measured_dD_mean | -59.973
varray: [-59.237, -60.137, -60.285, -60.233]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [60546.0, 61082.0, 61618.0, 62155.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.151837632577
varray: [-7.231, -7.499, -7.539, -7.555]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [60546.0, 61082.0, 61618.0, 62155.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.494481546673
varray: [-59.237, -60.137, -60.285, -60.233]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [60546.0, 61082.0, 61618.0, 62155.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.55886814627
actual_array: [ 1.93 -15.84]
varray: [-7.231, -7.499, -7.539, -7.555]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.456
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [60546.0, 61082.0, 61618.0, 62155.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -68.8998413487
actual_array: [ 18.1 -119. ]
varray: [-59.237, -60.137, -60.285, -60.233]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -59.973
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [60546.0, 61082.0, 61618.0, 62155.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.27388095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -4.99613095238
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 23 - Sample 23 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 23
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.82275
varray: [2.7, 2.834, 2.844, 2.913]
inst: mean_all
operation: mean_all(residuals)
tarray: [62742.0, 63279.0, 63817.0, 64354.0]
source_col: residuals
SAMP_residuals_STD | 0.0890519511297
varray: [2.7, 2.834, 2.844, 2.913]
inst: std_all
operation: std_all(residuals)
tarray: [62742.0, 63279.0, 63817.0, 64354.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.44325
varray: [0.182, 0.399, 0.584, 0.608]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [62742.0, 63279.0, 63817.0, 64354.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.197621481626
varray: [0.182, 0.399, 0.584, 0.608]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [62742.0, 63279.0, 63817.0, 64354.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1293.86075
varray: [1290.817, 1295.121, 1293.586, 1295.919]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [62742.0, 63279.0, 63817.0, 64354.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.24829748551
varray: [1290.817, 1295.121, 1293.586, 1295.919]
inst: std_all
operation: std_all(baseline_shift)
tarray: [62742.0, 63279.0, 63817.0, 64354.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.556
varray: [1.704, 1.629, 1.708, 1.183]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [62742.0, 63279.0, 63817.0, 64354.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.251307248867
varray: [1.704, 1.629, 1.708, 1.183]
inst: std_all
operation: std_all(slope_shift)
tarray: [62742.0, 63279.0, 63817.0, 64354.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0666458300771
varray: [-7.197, -7.052, -7.087, -7.062]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [62742.0, 63279.0, 63817.0, 64354.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.731478582507
varray: [-77.284, -78.336, -78.826, -78.845]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [62742.0, 63279.0, 63817.0, 64354.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.51575
varray: [105.957, 110.471, 109.371, 112.264]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [62742.0, 63279.0, 63817.0, 64354.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.65524856024
varray: [105.957, 110.471, 109.371, 112.264]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [62742.0, 63279.0, 63817.0, 64354.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.82275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.44325
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1293.86075
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.556
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.058630952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.058630952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.03482126626
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.2331547619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-7.197, -7.052, -7.087, -7.062]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [62742.0, 63279.0, 63817.0, 64354.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-77.284, -78.336, -78.826, -78.845]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [62742.0, 63279.0, 63817.0, 64354.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.0995
varray: [-7.197, -7.052, -7.087, -7.062]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [62742.0, 63279.0, 63817.0, 64354.0]
source_col: d(18_16)Mean
dD_mean | -78.32275
varray: [-77.284, -78.336, -78.826, -78.845]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [62742.0, 63279.0, 63817.0, 64354.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0666458300771
varray: [-7.197, -7.052, -7.087, -7.062]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [62742.0, 63279.0, 63817.0, 64354.0]
source_col: d(18_16)Mean
dD_stdev | 0.731478582507
varray: [-77.284, -78.336, -78.826, -78.845]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [62742.0, 63279.0, 63817.0, 64354.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.0995
varray: [-7.197, -7.052, -7.087, -7.062]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [62742.0, 63279.0, 63817.0, 64354.0]
source_col: d(18_16)Mean
measured_dD_mean | -78.32275
varray: [-77.284, -78.336, -78.826, -78.845]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [62742.0, 63279.0, 63817.0, 64354.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0666458300771
varray: [-7.197, -7.052, -7.087, -7.062]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [62742.0, 63279.0, 63817.0, 64354.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.731478582507
varray: [-77.284, -78.336, -78.826, -78.845]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [62742.0, 63279.0, 63817.0, 64354.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.17330035478
actual_array: [ 1.93 -15.84]
varray: [-7.197, -7.052, -7.087, -7.062]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.0995
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [62742.0, 63279.0, 63817.0, 64354.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -87.2850988754
actual_array: [ 18.1 -119. ]
varray: [-77.284, -78.336, -78.826, -78.845]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -78.32275
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [62742.0, 63279.0, 63817.0, 64354.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.17063095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.13263095238
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 24 - Sample 24 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 24
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.70025
varray: [2.627, 2.7, 2.663, 2.811]
inst: mean_all
operation: mean_all(residuals)
tarray: [64941.0, 65478.0, 66015.0, 66552.0]
source_col: residuals
SAMP_residuals_STD | 0.0796215004464
varray: [2.627, 2.7, 2.663, 2.811]
inst: std_all
operation: std_all(residuals)
tarray: [64941.0, 65478.0, 66015.0, 66552.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.33875
varray: [0.262, 0.357, 0.424, 0.312]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [64941.0, 65478.0, 66015.0, 66552.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0688155747875
varray: [0.262, 0.357, 0.424, 0.312]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [64941.0, 65478.0, 66015.0, 66552.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1299.4015
varray: [1296.806, 1300.426, 1296.838, 1303.536]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [64941.0, 65478.0, 66015.0, 66552.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.23789247299
varray: [1296.806, 1300.426, 1296.838, 1303.536]
inst: std_all
operation: std_all(baseline_shift)
tarray: [64941.0, 65478.0, 66015.0, 66552.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.38125
varray: [1.414, 1.119, 1.805, 1.187]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [64941.0, 65478.0, 66015.0, 66552.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.309378274393
varray: [1.414, 1.119, 1.805, 1.187]
inst: std_all
operation: std_all(slope_shift)
tarray: [64941.0, 65478.0, 66015.0, 66552.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0596007550288
varray: [-8.508, -8.421, -8.531, -8.559]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [64941.0, 65478.0, 66015.0, 66552.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.266789555018
varray: [-84.543, -84.644, -85.075, -85.022]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [64941.0, 65478.0, 66015.0, 66552.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 104.8735
varray: [102.84, 106.278, 102.195, 108.181]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [64941.0, 65478.0, 66015.0, 66552.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.84145860431
varray: [102.84, 106.278, 102.195, 108.181]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [64941.0, 65478.0, 66015.0, 66552.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.70025
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.33875
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1299.4015
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.38125
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.181130952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.181130952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.28625493728
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.8754047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-8.508, -8.421, -8.531, -8.559]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [64941.0, 65478.0, 66015.0, 66552.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-84.543, -84.644, -85.075, -85.022]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [64941.0, 65478.0, 66015.0, 66552.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.50475
varray: [-8.508, -8.421, -8.531, -8.559]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [64941.0, 65478.0, 66015.0, 66552.0]
source_col: d(18_16)Mean
dD_mean | -84.821
varray: [-84.543, -84.644, -85.075, -85.022]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [64941.0, 65478.0, 66015.0, 66552.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0596007550288
varray: [-8.508, -8.421, -8.531, -8.559]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [64941.0, 65478.0, 66015.0, 66552.0]
source_col: d(18_16)Mean
dD_stdev | 0.266789555018
varray: [-84.543, -84.644, -85.075, -85.022]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [64941.0, 65478.0, 66015.0, 66552.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.50475
varray: [-8.508, -8.421, -8.531, -8.559]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [64941.0, 65478.0, 66015.0, 66552.0]
source_col: d(18_16)Mean
measured_dD_mean | -84.821
varray: [-84.543, -84.644, -85.075, -85.022]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [64941.0, 65478.0, 66015.0, 66552.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0596007550288
varray: [-8.508, -8.421, -8.531, -8.559]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [64941.0, 65478.0, 66015.0, 66552.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.266789555018
varray: [-84.543, -84.644, -85.075, -85.022]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [64941.0, 65478.0, 66015.0, 66552.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.6931296367
actual_array: [ 1.93 -15.84]
varray: [-8.508, -8.421, -8.531, -8.559]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -8.50475
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [64941.0, 65478.0, 66015.0, 66552.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -93.7959232585
actual_array: [ 18.1 -119. ]
varray: [-84.543, -84.644, -85.075, -85.022]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -84.821
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [64941.0, 65478.0, 66015.0, 66552.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.29313095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 4.40811904762
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 25 - Sample 25 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 25
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.72975
varray: [2.745, 2.737, 2.721, 2.716]
inst: mean_all
operation: mean_all(residuals)
tarray: [67140.0, 67677.0, 68214.0, 68751.0]
source_col: residuals
SAMP_residuals_STD | 0.0135492927245
varray: [2.745, 2.737, 2.721, 2.716]
inst: std_all
operation: std_all(residuals)
tarray: [67140.0, 67677.0, 68214.0, 68751.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.01
varray: [0.211, 0.238, -0.235, -0.174]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [67140.0, 67677.0, 68214.0, 68751.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.2491759753
varray: [0.211, 0.238, -0.235, -0.174]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [67140.0, 67677.0, 68214.0, 68751.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1352.839
varray: [1350.921, 1352.895, 1353.458, 1354.082]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [67140.0, 67677.0, 68214.0, 68751.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.3674879646
varray: [1350.921, 1352.895, 1353.458, 1354.082]
inst: std_all
operation: std_all(baseline_shift)
tarray: [67140.0, 67677.0, 68214.0, 68751.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -1.14525
varray: [-1.162, -1.103, -1.273, -1.043]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [67140.0, 67677.0, 68214.0, 68751.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0980488823666
varray: [-1.162, -1.103, -1.273, -1.043]
inst: std_all
operation: std_all(slope_shift)
tarray: [67140.0, 67677.0, 68214.0, 68751.0]
source_col: slope_shift
SAMP_d18O_STD | 0.185458665655
varray: [-2.249, -2.064, -1.895, -1.837]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67140.0, 67677.0, 68214.0, 68751.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.77729963709
varray: [-55.355, -52.545, -51.654, -51.546]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67140.0, 67677.0, 68214.0, 68751.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 106.2045
varray: [106.851, 107.164, 105.439, 105.364]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [67140.0, 67677.0, 68214.0, 68751.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.936488654496
varray: [106.851, 107.164, 105.439, 105.364]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [67140.0, 67677.0, 68214.0, 68751.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.72975
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.01
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1352.839
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.14525
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.151630952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.151630952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.26244029814
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.5444047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-2.249, -2.064, -1.895, -1.837]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [67140.0, 67677.0, 68214.0, 68751.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-55.355, -52.545, -51.654, -51.546]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [67140.0, 67677.0, 68214.0, 68751.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.01125
varray: [-2.249, -2.064, -1.895, -1.837]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [67140.0, 67677.0, 68214.0, 68751.0]
source_col: d(18_16)Mean
dD_mean | -52.775
varray: [-55.355, -52.545, -51.654, -51.546]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [67140.0, 67677.0, 68214.0, 68751.0]
source_col: d(D_H)Mean
d18O_stdev | 0.185458665655
varray: [-2.249, -2.064, -1.895, -1.837]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67140.0, 67677.0, 68214.0, 68751.0]
source_col: d(18_16)Mean
dD_stdev | 1.77729963709
varray: [-55.355, -52.545, -51.654, -51.546]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67140.0, 67677.0, 68214.0, 68751.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.01125
varray: [-2.249, -2.064, -1.895, -1.837]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [67140.0, 67677.0, 68214.0, 68751.0]
source_col: d(18_16)Mean
measured_dD_mean | -52.775
varray: [-55.355, -52.545, -51.654, -51.546]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [67140.0, 67677.0, 68214.0, 68751.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.185458665655
varray: [-2.249, -2.064, -1.895, -1.837]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67140.0, 67677.0, 68214.0, 68751.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.77729963709
varray: [-55.355, -52.545, -51.654, -51.546]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67140.0, 67677.0, 68214.0, 68751.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.67017184132
actual_array: [ 1.93 -15.84]
varray: [-2.249, -2.064, -1.895, -1.837]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -2.01125
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [67140.0, 67677.0, 68214.0, 68751.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -61.6879129203
actual_array: [ 18.1 -119. ]
varray: [-55.355, -52.545, -51.654, -51.546]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -52.775
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [67140.0, 67677.0, 68214.0, 68751.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.26363095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 57.8456190476
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 26 - Sample 26 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 26
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.689
varray: [2.714, 2.635, 2.69, 2.717]
inst: mean_all
operation: mean_all(residuals)
tarray: [69338.0, 69875.0, 70412.0, 70949.0]
source_col: residuals
SAMP_residuals_STD | 0.0379736750921
varray: [2.714, 2.635, 2.69, 2.717]
inst: std_all
operation: std_all(residuals)
tarray: [69338.0, 69875.0, 70412.0, 70949.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.21525
varray: [0.363, 0.115, 0.037, 0.346]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [69338.0, 69875.0, 70412.0, 70949.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.164061726189
varray: [0.363, 0.115, 0.037, 0.346]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [69338.0, 69875.0, 70412.0, 70949.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1330.25025
varray: [1329.899, 1329.475, 1331.52, 1330.107]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [69338.0, 69875.0, 70412.0, 70949.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.886411257074
varray: [1329.899, 1329.475, 1331.52, 1330.107]
inst: std_all
operation: std_all(baseline_shift)
tarray: [69338.0, 69875.0, 70412.0, 70949.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.06775
varray: [0.082, 0.256, -0.018, -0.049]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [69338.0, 69875.0, 70412.0, 70949.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.137386013359
varray: [0.082, 0.256, -0.018, -0.049]
inst: std_all
operation: std_all(slope_shift)
tarray: [69338.0, 69875.0, 70412.0, 70949.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0920556896666
varray: [-5.395, -5.575, -5.542, -5.601]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [69338.0, 69875.0, 70412.0, 70949.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.19618487284
varray: [-71.595, -73.689, -74.04, -74.139]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [69338.0, 69875.0, 70412.0, 70949.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 104.3625
varray: [104.605, 102.898, 104.888, 105.059]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [69338.0, 69875.0, 70412.0, 70949.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.994120884668
varray: [104.605, 102.898, 104.888, 105.059]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [69338.0, 69875.0, 70412.0, 70949.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.689
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.21525
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1330.25025
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.06775
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.192380952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.192380952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.6766927234
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -6.3864047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-5.395, -5.575, -5.542, -5.601]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [69338.0, 69875.0, 70412.0, 70949.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-71.595, -73.689, -74.04, -74.139]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [69338.0, 69875.0, 70412.0, 70949.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.52825
varray: [-5.395, -5.575, -5.542, -5.601]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [69338.0, 69875.0, 70412.0, 70949.0]
source_col: d(18_16)Mean
dD_mean | -73.36575
varray: [-71.595, -73.689, -74.04, -74.139]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [69338.0, 69875.0, 70412.0, 70949.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0920556896666
varray: [-5.395, -5.575, -5.542, -5.601]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [69338.0, 69875.0, 70412.0, 70949.0]
source_col: d(18_16)Mean
dD_stdev | 1.19618487284
varray: [-71.595, -73.689, -74.04, -74.139]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [69338.0, 69875.0, 70412.0, 70949.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.52825
varray: [-5.395, -5.575, -5.542, -5.601]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [69338.0, 69875.0, 70412.0, 70949.0]
source_col: d(18_16)Mean
measured_dD_mean | -73.36575
varray: [-71.595, -73.689, -74.04, -74.139]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [69338.0, 69875.0, 70412.0, 70949.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0920556896666
varray: [-5.395, -5.575, -5.542, -5.601]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [69338.0, 69875.0, 70412.0, 70949.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.19618487284
varray: [-71.595, -73.689, -74.04, -74.139]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [69338.0, 69875.0, 70412.0, 70949.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.47393600028
actual_array: [ 1.93 -15.84]
varray: [-5.395, -5.575, -5.542, -5.601]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -5.52825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [69338.0, 69875.0, 70412.0, 70949.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -82.318506875
actual_array: [ 18.1 -119. ]
varray: [-71.595, -73.689, -74.04, -74.139]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -73.36575
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [69338.0, 69875.0, 70412.0, 70949.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.30438095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 35.2568690476
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 27 - Sample 27 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 27
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.7805
varray: [2.768, 2.758, 2.811, 2.785]
inst: mean_all
operation: mean_all(residuals)
tarray: [71538.0, 72074.0, 72611.0, 73148.0]
source_col: residuals
SAMP_residuals_STD | 0.0231876403859
varray: [2.768, 2.758, 2.811, 2.785]
inst: std_all
operation: std_all(residuals)
tarray: [71538.0, 72074.0, 72611.0, 73148.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.19875
varray: [0.316, 0.156, 0.318, 0.005]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [71538.0, 72074.0, 72611.0, 73148.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.149816276374
varray: [0.316, 0.156, 0.318, 0.005]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [71538.0, 72074.0, 72611.0, 73148.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1300.70975
varray: [1299.138, 1300.904, 1300.852, 1301.945]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [71538.0, 72074.0, 72611.0, 73148.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.16249856631
varray: [1299.138, 1300.904, 1300.852, 1301.945]
inst: std_all
operation: std_all(baseline_shift)
tarray: [71538.0, 72074.0, 72611.0, 73148.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.07775
varray: [1.236, 0.991, 1.042, 1.042]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [71538.0, 72074.0, 72611.0, 73148.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.108204667182
varray: [1.236, 0.991, 1.042, 1.042]
inst: std_all
operation: std_all(slope_shift)
tarray: [71538.0, 72074.0, 72611.0, 73148.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0458802789878
varray: [-5.643, -5.601, -5.669, -5.565]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [71538.0, 72074.0, 72611.0, 73148.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.394606132745
varray: [-68.876, -68.076, -68.237, -68.015]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [71538.0, 72074.0, 72611.0, 73148.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.49825
varray: [106.997, 108.232, 109.727, 109.037]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [71538.0, 72074.0, 72611.0, 73148.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.17256396414
varray: [106.997, 108.232, 109.727, 109.037]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [71538.0, 72074.0, 72611.0, 73148.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.7805
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.19875
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1300.70975
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.07775
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.100880952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.100880952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.501132063
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.2506547619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-5.643, -5.601, -5.669, -5.565]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [71538.0, 72074.0, 72611.0, 73148.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-68.876, -68.076, -68.237, -68.015]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [71538.0, 72074.0, 72611.0, 73148.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.6195
varray: [-5.643, -5.601, -5.669, -5.565]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [71538.0, 72074.0, 72611.0, 73148.0]
source_col: d(18_16)Mean
dD_mean | -68.301
varray: [-68.876, -68.076, -68.237, -68.015]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [71538.0, 72074.0, 72611.0, 73148.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0458802789878
varray: [-5.643, -5.601, -5.669, -5.565]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [71538.0, 72074.0, 72611.0, 73148.0]
source_col: d(18_16)Mean
dD_stdev | 0.394606132745
varray: [-68.876, -68.076, -68.237, -68.015]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [71538.0, 72074.0, 72611.0, 73148.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.6195
varray: [-5.643, -5.601, -5.669, -5.565]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [71538.0, 72074.0, 72611.0, 73148.0]
source_col: d(18_16)Mean
measured_dD_mean | -68.301
varray: [-68.876, -68.076, -68.237, -68.015]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [71538.0, 72074.0, 72611.0, 73148.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0458802789878
varray: [-5.643, -5.601, -5.669, -5.565]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [71538.0, 72074.0, 72611.0, 73148.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.394606132745
varray: [-68.876, -68.076, -68.237, -68.015]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [71538.0, 72074.0, 72611.0, 73148.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.57262621339
actual_array: [ 1.93 -15.84]
varray: [-5.643, -5.601, -5.669, -5.565]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -5.6195
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [71538.0, 72074.0, 72611.0, 73148.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -77.2439563738
actual_array: [ 18.1 -119. ]
varray: [-68.876, -68.076, -68.237, -68.015]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -68.301
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [71538.0, 72074.0, 72611.0, 73148.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.21288095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 5.71636904762
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 28 - Sample 28 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 28
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.724
varray: [2.759, 2.657, 2.693, 2.787]
inst: mean_all
operation: mean_all(residuals)
tarray: [73737.0, 74274.0, 74811.0, 75348.0]
source_col: residuals
SAMP_residuals_STD | 0.0595650904473
varray: [2.759, 2.657, 2.693, 2.787]
inst: std_all
operation: std_all(residuals)
tarray: [73737.0, 74274.0, 74811.0, 75348.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.2355
varray: [0.54, 0.248, -0.005, 0.159]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [73737.0, 74274.0, 74811.0, 75348.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.228450578171
varray: [0.54, 0.248, -0.005, 0.159]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [73737.0, 74274.0, 74811.0, 75348.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1334.79975
varray: [1333.791, 1332.143, 1334.766, 1338.499]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [73737.0, 74274.0, 74811.0, 75348.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.69329332169
varray: [1333.791, 1332.143, 1334.766, 1338.499]
inst: std_all
operation: std_all(baseline_shift)
tarray: [73737.0, 74274.0, 74811.0, 75348.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -0.2775
varray: [-0.686, 0.074, -0.071, -0.427]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [73737.0, 74274.0, 74811.0, 75348.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.344198101874
varray: [-0.686, 0.074, -0.071, -0.427]
inst: std_all
operation: std_all(slope_shift)
tarray: [73737.0, 74274.0, 74811.0, 75348.0]
source_col: slope_shift
SAMP_d18O_STD | 0.162413823303
varray: [-1.709, -1.544, -1.523, -1.313]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [73737.0, 74274.0, 74811.0, 75348.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.34884255938
varray: [-52.085, -50.204, -49.887, -48.851]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [73737.0, 74274.0, 74811.0, 75348.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 106.01175
varray: [107.183, 104.352, 104.616, 107.896]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [73737.0, 74274.0, 74811.0, 75348.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.7911925962
varray: [107.183, 104.352, 104.616, 107.896]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [73737.0, 74274.0, 74811.0, 75348.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.724
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.2355
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1334.79975
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.2775
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.157380952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.157380952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.46199738882
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.7371547619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-1.709, -1.544, -1.523, -1.313]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [73737.0, 74274.0, 74811.0, 75348.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-52.085, -50.204, -49.887, -48.851]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [73737.0, 74274.0, 74811.0, 75348.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.52225
varray: [-1.709, -1.544, -1.523, -1.313]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [73737.0, 74274.0, 74811.0, 75348.0]
source_col: d(18_16)Mean
dD_mean | -50.25675
varray: [-52.085, -50.204, -49.887, -48.851]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [73737.0, 74274.0, 74811.0, 75348.0]
source_col: d(D_H)Mean
d18O_stdev | 0.162413823303
varray: [-1.709, -1.544, -1.523, -1.313]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [73737.0, 74274.0, 74811.0, 75348.0]
source_col: d(18_16)Mean
dD_stdev | 1.34884255938
varray: [-52.085, -50.204, -49.887, -48.851]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [73737.0, 74274.0, 74811.0, 75348.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.52225
varray: [-1.709, -1.544, -1.523, -1.313]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [73737.0, 74274.0, 74811.0, 75348.0]
source_col: d(18_16)Mean
measured_dD_mean | -50.25675
varray: [-52.085, -50.204, -49.887, -48.851]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [73737.0, 74274.0, 74811.0, 75348.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.162413823303
varray: [-1.709, -1.544, -1.523, -1.313]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [73737.0, 74274.0, 74811.0, 75348.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.34884255938
varray: [-52.085, -50.204, -49.887, -48.851]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [73737.0, 74274.0, 74811.0, 75348.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.14130045271
actual_array: [ 1.93 -15.84]
varray: [-1.709, -1.544, -1.523, -1.313]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -1.52225
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [73737.0, 74274.0, 74811.0, 75348.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -59.1647900022
actual_array: [ 18.1 -119. ]
varray: [-52.085, -50.204, -49.887, -48.851]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -50.25675
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [73737.0, 74274.0, 74811.0, 75348.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.26938095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 39.8063690476
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 29 - Sample 29 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 29
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.709
varray: [2.598, 2.805, 2.714, 2.719]
inst: mean_all
operation: mean_all(residuals)
tarray: [75933.0, 76469.0, 77007.0, 77544.0]
source_col: residuals
SAMP_residuals_STD | 0.0849745059807
varray: [2.598, 2.805, 2.714, 2.719]
inst: std_all
operation: std_all(residuals)
tarray: [75933.0, 76469.0, 77007.0, 77544.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.243
varray: [0.024, 0.37, 0.201, 0.377]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [75933.0, 76469.0, 77007.0, 77544.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.167142653643
varray: [0.024, 0.37, 0.201, 0.377]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [75933.0, 76469.0, 77007.0, 77544.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1302.023
varray: [1298.079, 1305.066, 1303.34, 1301.607]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [75933.0, 76469.0, 77007.0, 77544.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.98454519148
varray: [1298.079, 1305.066, 1303.34, 1301.607]
inst: std_all
operation: std_all(baseline_shift)
tarray: [75933.0, 76469.0, 77007.0, 77544.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.774
varray: [0.893, 0.6, 0.628, 0.975]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [75933.0, 76469.0, 77007.0, 77544.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.188108124935
varray: [0.893, 0.6, 0.628, 0.975]
inst: std_all
operation: std_all(slope_shift)
tarray: [75933.0, 76469.0, 77007.0, 77544.0]
source_col: slope_shift
SAMP_d18O_STD | 0.131940327421
varray: [-1.7, -1.411, -1.628, -1.676]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [75933.0, 76469.0, 77007.0, 77544.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.693559117884
varray: [-46.99, -45.341, -45.977, -46.385]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [75933.0, 76469.0, 77007.0, 77544.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 105.5475
varray: [101.39, 109.038, 105.627, 106.135]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [75933.0, 76469.0, 77007.0, 77544.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.15276984042
varray: [101.39, 109.038, 105.627, 106.135]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [75933.0, 76469.0, 77007.0, 77544.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.709
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.243
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1302.023
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.774
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.172380952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.172380952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.98258110364
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.2014047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-1.7, -1.411, -1.628, -1.676]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [75933.0, 76469.0, 77007.0, 77544.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-46.99, -45.341, -45.977, -46.385]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [75933.0, 76469.0, 77007.0, 77544.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.60375
varray: [-1.7, -1.411, -1.628, -1.676]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [75933.0, 76469.0, 77007.0, 77544.0]
source_col: d(18_16)Mean
dD_mean | -46.17325
varray: [-46.99, -45.341, -45.977, -46.385]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [75933.0, 76469.0, 77007.0, 77544.0]
source_col: d(D_H)Mean
d18O_stdev | 0.131940327421
varray: [-1.7, -1.411, -1.628, -1.676]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [75933.0, 76469.0, 77007.0, 77544.0]
source_col: d(18_16)Mean
dD_stdev | 0.693559117884
varray: [-46.99, -45.341, -45.977, -46.385]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [75933.0, 76469.0, 77007.0, 77544.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.60375
varray: [-1.7, -1.411, -1.628, -1.676]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [75933.0, 76469.0, 77007.0, 77544.0]
source_col: d(18_16)Mean
measured_dD_mean | -46.17325
varray: [-46.99, -45.341, -45.977, -46.385]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [75933.0, 76469.0, 77007.0, 77544.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.131940327421
varray: [-1.7, -1.411, -1.628, -1.676]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [75933.0, 76469.0, 77007.0, 77544.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.693559117884
varray: [-46.99, -45.341, -45.977, -46.385]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [75933.0, 76469.0, 77007.0, 77544.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.22944568415
actual_array: [ 1.93 -15.84]
varray: [-1.7, -1.411, -1.628, -1.676]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -1.60375
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [75933.0, 76469.0, 77007.0, 77544.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -55.0733882605
actual_array: [ 18.1 -119. ]
varray: [-46.99, -45.341, -45.977, -46.385]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -46.17325
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [75933.0, 76469.0, 77007.0, 77544.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.28438095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 7.02961904762
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 30 - Sample 30 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 30
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.58375
varray: [2.66, 2.5, 2.626, 2.549]
inst: mean_all
operation: mean_all(residuals)
tarray: [78129.0, 78666.0, 79203.0, 79740.0]
source_col: residuals
SAMP_residuals_STD | 0.0726194418412
varray: [2.66, 2.5, 2.626, 2.549]
inst: std_all
operation: std_all(residuals)
tarray: [78129.0, 78666.0, 79203.0, 79740.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.36475
varray: [-0.282, -0.312, -0.313, -0.552]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [78129.0, 78666.0, 79203.0, 79740.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.125659261497
varray: [-0.282, -0.312, -0.313, -0.552]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [78129.0, 78666.0, 79203.0, 79740.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1341.372
varray: [1341.861, 1339.109, 1343.908, 1340.61]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [78129.0, 78666.0, 79203.0, 79740.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.03078227948
varray: [1341.861, 1339.109, 1343.908, 1340.61]
inst: std_all
operation: std_all(baseline_shift)
tarray: [78129.0, 78666.0, 79203.0, 79740.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -2.09575
varray: [-2.175, -1.763, -2.541, -1.904]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [78129.0, 78666.0, 79203.0, 79740.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.342548658344
varray: [-2.175, -1.763, -2.541, -1.904]
inst: std_all
operation: std_all(slope_shift)
tarray: [78129.0, 78666.0, 79203.0, 79740.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0700880398737
varray: [-2.976, -3.133, -3.057, -3.112]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78129.0, 78666.0, 79203.0, 79740.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.987743893932
varray: [-60.542, -62.524, -62.317, -62.653]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78129.0, 78666.0, 79203.0, 79740.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 102.13925
varray: [103.673, 100.443, 103.761, 100.68]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [78129.0, 78666.0, 79203.0, 79740.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.82474991437
varray: [103.673, 100.443, 103.761, 100.68]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [78129.0, 78666.0, 79203.0, 79740.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.58375
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.36475
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1341.372
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.09575
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.297630952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.297630952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -10.3294551224
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -8.6096547619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-2.976, -3.133, -3.057, -3.112]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [78129.0, 78666.0, 79203.0, 79740.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-60.542, -62.524, -62.317, -62.653]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [78129.0, 78666.0, 79203.0, 79740.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.0695
varray: [-2.976, -3.133, -3.057, -3.112]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [78129.0, 78666.0, 79203.0, 79740.0]
source_col: d(18_16)Mean
dD_mean | -62.009
varray: [-60.542, -62.524, -62.317, -62.653]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [78129.0, 78666.0, 79203.0, 79740.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0700880398737
varray: [-2.976, -3.133, -3.057, -3.112]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78129.0, 78666.0, 79203.0, 79740.0]
source_col: d(18_16)Mean
dD_stdev | 0.987743893932
varray: [-60.542, -62.524, -62.317, -62.653]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78129.0, 78666.0, 79203.0, 79740.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.0695
varray: [-2.976, -3.133, -3.057, -3.112]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [78129.0, 78666.0, 79203.0, 79740.0]
source_col: d(18_16)Mean
measured_dD_mean | -62.009
varray: [-60.542, -62.524, -62.317, -62.653]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [78129.0, 78666.0, 79203.0, 79740.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0700880398737
varray: [-2.976, -3.133, -3.057, -3.112]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78129.0, 78666.0, 79203.0, 79740.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.987743893932
varray: [-60.542, -62.524, -62.317, -62.653]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78129.0, 78666.0, 79203.0, 79740.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.81470792924
actual_array: [ 1.93 -15.84]
varray: [-2.976, -3.133, -3.057, -3.112]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -3.0695
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [78129.0, 78666.0, 79203.0, 79740.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -70.939781093
actual_array: [ 18.1 -119. ]
varray: [-60.542, -62.524, -62.317, -62.653]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -62.009
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [78129.0, 78666.0, 79203.0, 79740.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.40963095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 46.3786190476
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 31 - Sample 31 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 31
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.68525
varray: [2.683, 2.712, 2.662, 2.684]
inst: mean_all
operation: mean_all(residuals)
tarray: [80326.0, 80862.0, 81399.0, 81936.0]
source_col: residuals
SAMP_residuals_STD | 0.0205162537191
varray: [2.683, 2.712, 2.662, 2.684]
inst: std_all
operation: std_all(residuals)
tarray: [80326.0, 80862.0, 81399.0, 81936.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.221
varray: [0.147, 0.354, 0.235, 0.148]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [80326.0, 80862.0, 81399.0, 81936.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0977922969018
varray: [0.147, 0.354, 0.235, 0.148]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [80326.0, 80862.0, 81399.0, 81936.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1288.7675
varray: [1291.466, 1287.895, 1287.718, 1287.991]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [80326.0, 80862.0, 81399.0, 81936.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.802550138
varray: [1291.466, 1287.895, 1287.718, 1287.991]
inst: std_all
operation: std_all(baseline_shift)
tarray: [80326.0, 80862.0, 81399.0, 81936.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.80825
varray: [2.31, 1.233, 1.94, 1.75]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [80326.0, 80862.0, 81399.0, 81936.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.448485135391
varray: [2.31, 1.233, 1.94, 1.75]
inst: std_all
operation: std_all(slope_shift)
tarray: [80326.0, 80862.0, 81399.0, 81936.0]
source_col: slope_shift
SAMP_d18O_STD | 0.123154036339
varray: [-7.339, -7.539, -7.604, -7.593]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [80326.0, 80862.0, 81399.0, 81936.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.132314020421
varray: [-60.525, -60.537, -60.762, -60.458]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [80326.0, 80862.0, 81399.0, 81936.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 104.1885
varray: [105.678, 103.772, 103.366, 103.938]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [80326.0, 80862.0, 81399.0, 81936.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.02165535611
varray: [105.678, 103.772, 103.366, 103.938]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [80326.0, 80862.0, 81399.0, 81936.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.68525
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.221
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1288.7675
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.80825
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.196130952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.196130952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.8068386521
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -6.5604047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-7.339, -7.539, -7.604, -7.593]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [80326.0, 80862.0, 81399.0, 81936.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-60.525, -60.537, -60.762, -60.458]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [80326.0, 80862.0, 81399.0, 81936.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.51875
varray: [-7.339, -7.539, -7.604, -7.593]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [80326.0, 80862.0, 81399.0, 81936.0]
source_col: d(18_16)Mean
dD_mean | -60.5705
varray: [-60.525, -60.537, -60.762, -60.458]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [80326.0, 80862.0, 81399.0, 81936.0]
source_col: d(D_H)Mean
d18O_stdev | 0.123154036339
varray: [-7.339, -7.539, -7.604, -7.593]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [80326.0, 80862.0, 81399.0, 81936.0]
source_col: d(18_16)Mean
dD_stdev | 0.132314020421
varray: [-60.525, -60.537, -60.762, -60.458]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [80326.0, 80862.0, 81399.0, 81936.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.51875
varray: [-7.339, -7.539, -7.604, -7.593]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [80326.0, 80862.0, 81399.0, 81936.0]
source_col: d(18_16)Mean
measured_dD_mean | -60.5705
varray: [-60.525, -60.537, -60.762, -60.458]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [80326.0, 80862.0, 81399.0, 81936.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.123154036339
varray: [-7.339, -7.539, -7.604, -7.593]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [80326.0, 80862.0, 81399.0, 81936.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.132314020421
varray: [-60.525, -60.537, -60.762, -60.458]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [80326.0, 80862.0, 81399.0, 81936.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.62673456679
actual_array: [ 1.93 -15.84]
varray: [-7.339, -7.539, -7.604, -7.593]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.51875
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [80326.0, 80862.0, 81399.0, 81936.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -69.4984975359
actual_array: [ 18.1 -119. ]
varray: [-60.525, -60.537, -60.762, -60.458]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -60.5705
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [80326.0, 80862.0, 81399.0, 81936.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.30813095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -6.22588095238
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 32 - Sample 32 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 32
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.7265
varray: [2.6, 2.782, 2.745, 2.779]
inst: mean_all
operation: mean_all(residuals)
tarray: [82523.0, 83060.0, 83597.0, 84133.0]
source_col: residuals
SAMP_residuals_STD | 0.0859864330384
varray: [2.6, 2.782, 2.745, 2.779]
inst: std_all
operation: std_all(residuals)
tarray: [82523.0, 83060.0, 83597.0, 84133.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.08575
varray: [-0.114, 0.306, -0.024, 0.175]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [82523.0, 83060.0, 83597.0, 84133.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.190105891545
varray: [-0.114, 0.306, -0.024, 0.175]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [82523.0, 83060.0, 83597.0, 84133.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1323.938
varray: [1320.243, 1325.908, 1325.019, 1324.582]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [82523.0, 83060.0, 83597.0, 84133.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.52436275787
varray: [1320.243, 1325.908, 1325.019, 1324.582]
inst: std_all
operation: std_all(baseline_shift)
tarray: [82523.0, 83060.0, 83597.0, 84133.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -0.06425
varray: [0.164, -0.037, -0.022, -0.362]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [82523.0, 83060.0, 83597.0, 84133.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.21854118605
varray: [0.164, -0.037, -0.022, -0.362]
inst: std_all
operation: std_all(slope_shift)
tarray: [82523.0, 83060.0, 83597.0, 84133.0]
source_col: slope_shift
SAMP_d18O_STD | 0.111627356265
varray: [-3.893, -3.694, -3.648, -3.677]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [82523.0, 83060.0, 83597.0, 84133.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.188662838595
varray: [-59.288, -59.079, -58.873, -58.914]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [82523.0, 83060.0, 83597.0, 84133.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 105.9735
varray: [102.324, 108.424, 106.662, 106.484]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [82523.0, 83060.0, 83597.0, 84133.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.58575862498
varray: [102.324, 108.424, 106.662, 106.484]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [82523.0, 83060.0, 83597.0, 84133.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.7265
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.08575
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1323.938
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.06425
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.154880952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.154880952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.37523343635
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.7754047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-3.893, -3.694, -3.648, -3.677]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [82523.0, 83060.0, 83597.0, 84133.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-59.288, -59.079, -58.873, -58.914]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [82523.0, 83060.0, 83597.0, 84133.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.728
varray: [-3.893, -3.694, -3.648, -3.677]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [82523.0, 83060.0, 83597.0, 84133.0]
source_col: d(18_16)Mean
dD_mean | -59.0385
varray: [-59.288, -59.079, -58.873, -58.914]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [82523.0, 83060.0, 83597.0, 84133.0]
source_col: d(D_H)Mean
d18O_stdev | 0.111627356265
varray: [-3.893, -3.694, -3.648, -3.677]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [82523.0, 83060.0, 83597.0, 84133.0]
source_col: d(18_16)Mean
dD_stdev | 0.188662838595
varray: [-59.288, -59.079, -58.873, -58.914]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [82523.0, 83060.0, 83597.0, 84133.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.728
varray: [-3.893, -3.694, -3.648, -3.677]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [82523.0, 83060.0, 83597.0, 84133.0]
source_col: d(18_16)Mean
measured_dD_mean | -59.0385
varray: [-59.288, -59.079, -58.873, -58.914]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [82523.0, 83060.0, 83597.0, 84133.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.111627356265
varray: [-3.893, -3.694, -3.648, -3.677]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [82523.0, 83060.0, 83597.0, 84133.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.188662838595
varray: [-59.288, -59.079, -58.873, -58.914]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [82523.0, 83060.0, 83597.0, 84133.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.5268997685
actual_array: [ 1.93 -15.84]
varray: [-3.893, -3.694, -3.648, -3.677]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -3.728
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [82523.0, 83060.0, 83597.0, 84133.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -67.9635330524
actual_array: [ 18.1 -119. ]
varray: [-59.288, -59.079, -58.873, -58.914]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -59.0385
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [82523.0, 83060.0, 83597.0, 84133.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.26688095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 28.9446190476
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 33 - Sample 33 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 33
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1155
varray: [1.866, 2.147, 2.151, 2.298]
inst: mean_all
operation: mean_all(residuals)
tarray: [84720.0, 85257.0, 85794.0, 86332.0]
source_col: residuals
SAMP_residuals_STD | 0.180563008393
varray: [1.866, 2.147, 2.151, 2.298]
inst: std_all
operation: std_all(residuals)
tarray: [84720.0, 85257.0, 85794.0, 86332.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.261
varray: [0.252, 0.361, 0.189, 0.242]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [84720.0, 85257.0, 85794.0, 86332.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0721710930128
varray: [0.252, 0.361, 0.189, 0.242]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [84720.0, 85257.0, 85794.0, 86332.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1265.3405
varray: [1255.235, 1266.306, 1267.603, 1272.218]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [84720.0, 85257.0, 85794.0, 86332.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 7.19889419749
varray: [1255.235, 1266.306, 1267.603, 1272.218]
inst: std_all
operation: std_all(baseline_shift)
tarray: [84720.0, 85257.0, 85794.0, 86332.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.875
varray: [1.764, 2.15, 1.906, 1.68]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [84720.0, 85257.0, 85794.0, 86332.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.205695567932
varray: [1.764, 2.15, 1.906, 1.68]
inst: std_all
operation: std_all(slope_shift)
tarray: [84720.0, 85257.0, 85794.0, 86332.0]
source_col: slope_shift
SAMP_d18O_STD | 0.215405044819
varray: [-11.865, -12.25, -12.307, -12.318]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [84720.0, 85257.0, 85794.0, 86332.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.62600038081
varray: [-99.378, -103.851, -104.771, -104.988]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [84720.0, 85257.0, 85794.0, 86332.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 84.371
varray: [75.614, 85.599, 85.734, 90.537]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [84720.0, 85257.0, 85794.0, 86332.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 6.27349870487
varray: [75.614, 85.599, 85.734, 90.537]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [84720.0, 85257.0, 85794.0, 86332.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1155
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.261
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1265.3405
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.875
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.765880952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.765880952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -26.58034342
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -26.3779047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-11.865, -12.25, -12.307, -12.318]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [84720.0, 85257.0, 85794.0, 86332.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-99.378, -103.851, -104.771, -104.988]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [84720.0, 85257.0, 85794.0, 86332.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -12.185
varray: [-11.865, -12.25, -12.307, -12.318]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [84720.0, 85257.0, 85794.0, 86332.0]
source_col: d(18_16)Mean
dD_mean | -103.247
varray: [-99.378, -103.851, -104.771, -104.988]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [84720.0, 85257.0, 85794.0, 86332.0]
source_col: d(D_H)Mean
d18O_stdev | 0.215405044819
varray: [-11.865, -12.25, -12.307, -12.318]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [84720.0, 85257.0, 85794.0, 86332.0]
source_col: d(18_16)Mean
dD_stdev | 2.62600038081
varray: [-99.378, -103.851, -104.771, -104.988]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [84720.0, 85257.0, 85794.0, 86332.0]
source_col: d(D_H)Mean
measured_d18O_mean | -12.185
varray: [-11.865, -12.25, -12.307, -12.318]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [84720.0, 85257.0, 85794.0, 86332.0]
source_col: d(18_16)Mean
measured_dD_mean | -103.247
varray: [-99.378, -103.851, -104.771, -104.988]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [84720.0, 85257.0, 85794.0, 86332.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.215405044819
varray: [-11.865, -12.25, -12.307, -12.318]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [84720.0, 85257.0, 85794.0, 86332.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.62600038081
varray: [-99.378, -103.851, -104.771, -104.988]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [84720.0, 85257.0, 85794.0, 86332.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -14.6734546426
actual_array: [ 1.93 -15.84]
varray: [-11.865, -12.25, -12.307, -12.318]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -12.185
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [84720.0, 85257.0, 85794.0, 86332.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -112.257578332
actual_array: [ 18.1 -119. ]
varray: [-99.378, -103.851, -104.771, -104.988]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -103.247
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [84720.0, 85257.0, 85794.0, 86332.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.87788095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -29.6528809524
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 34 - Sample 34 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 34
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.721
varray: [2.727, 2.76, 2.736, 2.661]
inst: mean_all
operation: mean_all(residuals)
tarray: [86918.0, 87455.0, 87993.0, 88530.0]
source_col: residuals
SAMP_residuals_STD | 0.0423556371691
varray: [2.727, 2.76, 2.736, 2.661]
inst: std_all
operation: std_all(residuals)
tarray: [86918.0, 87455.0, 87993.0, 88530.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.34575
varray: [0.272, 0.448, 0.508, 0.155]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [86918.0, 87455.0, 87993.0, 88530.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.161869031421
varray: [0.272, 0.448, 0.508, 0.155]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [86918.0, 87455.0, 87993.0, 88530.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1289.101
varray: [1288.433, 1290.091, 1289.263, 1288.617]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [86918.0, 87455.0, 87993.0, 88530.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.74984976273
varray: [1288.433, 1290.091, 1289.263, 1288.617]
inst: std_all
operation: std_all(baseline_shift)
tarray: [86918.0, 87455.0, 87993.0, 88530.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.3635
varray: [0.556, 1.315, 1.521, 2.062]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [86918.0, 87455.0, 87993.0, 88530.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.623729909496
varray: [0.556, 1.315, 1.521, 2.062]
inst: std_all
operation: std_all(slope_shift)
tarray: [86918.0, 87455.0, 87993.0, 88530.0]
source_col: slope_shift
SAMP_d18O_STD | 0.386468950887
varray: [0.217, 0.873, 0.973, 1.072]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [86918.0, 87455.0, 87993.0, 88530.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.0609996597
varray: [-55.551, -50.327, -49.624, -48.749]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [86918.0, 87455.0, 87993.0, 88530.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 105.99725
varray: [105.948, 107.564, 106.806, 103.671]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [86918.0, 87455.0, 87993.0, 88530.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.68549169384
varray: [105.948, 107.564, 106.806, 103.671]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [86918.0, 87455.0, 87993.0, 88530.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.721
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.34575
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1289.101
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.3635
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.160380952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.160380952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.56611413178
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.7516547619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [0.217, 0.873, 0.973, 1.072]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [86918.0, 87455.0, 87993.0, 88530.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-55.551, -50.327, -49.624, -48.749]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [86918.0, 87455.0, 87993.0, 88530.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.78375
varray: [0.217, 0.873, 0.973, 1.072]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [86918.0, 87455.0, 87993.0, 88530.0]
source_col: d(18_16)Mean
dD_mean | -51.06275
varray: [-55.551, -50.327, -49.624, -48.749]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [86918.0, 87455.0, 87993.0, 88530.0]
source_col: d(D_H)Mean
d18O_stdev | 0.386468950887
varray: [0.217, 0.873, 0.973, 1.072]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [86918.0, 87455.0, 87993.0, 88530.0]
source_col: d(18_16)Mean
dD_stdev | 3.0609996597
varray: [-55.551, -50.327, -49.624, -48.749]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [86918.0, 87455.0, 87993.0, 88530.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.78375
varray: [0.217, 0.873, 0.973, 1.072]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [86918.0, 87455.0, 87993.0, 88530.0]
source_col: d(18_16)Mean
measured_dD_mean | -51.06275
varray: [-55.551, -50.327, -49.624, -48.749]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [86918.0, 87455.0, 87993.0, 88530.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.386468950887
varray: [0.217, 0.873, 0.973, 1.072]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [86918.0, 87455.0, 87993.0, 88530.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.0609996597
varray: [-55.551, -50.327, -49.624, -48.749]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [86918.0, 87455.0, 87993.0, 88530.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.647277094575
actual_array: [ 1.93 -15.84]
varray: [0.217, 0.873, 0.973, 1.072]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.78375
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [86918.0, 87455.0, 87993.0, 88530.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -59.9723496456
actual_array: [ 18.1 -119. ]
varray: [-55.551, -50.327, -49.624, -48.749]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -51.06275
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [86918.0, 87455.0, 87993.0, 88530.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.27238095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -5.89238095238
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 35 - Sample 35 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 35
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.75475
varray: [2.675, 2.786, 2.762, 2.796]
inst: mean_all
operation: mean_all(residuals)
tarray: [89118.0, 89655.0, 90191.0, 90729.0]
source_col: residuals
SAMP_residuals_STD | 0.0550477065826
varray: [2.675, 2.786, 2.762, 2.796]
inst: std_all
operation: std_all(residuals)
tarray: [89118.0, 89655.0, 90191.0, 90729.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.46625
varray: [-0.601, -0.378, -0.637, -0.249]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [89118.0, 89655.0, 90191.0, 90729.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.184660724934
varray: [-0.601, -0.378, -0.637, -0.249]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [89118.0, 89655.0, 90191.0, 90729.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1401.85825
varray: [1397.336, 1402.71, 1403.791, 1403.596]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [89118.0, 89655.0, 90191.0, 90729.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.05131232696
varray: [1397.336, 1402.71, 1403.791, 1403.596]
inst: std_all
operation: std_all(baseline_shift)
tarray: [89118.0, 89655.0, 90191.0, 90729.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -4.5005
varray: [-4.466, -4.271, -4.516, -4.749]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [89118.0, 89655.0, 90191.0, 90729.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.196513782383
varray: [-4.466, -4.271, -4.516, -4.749]
inst: std_all
operation: std_all(slope_shift)
tarray: [89118.0, 89655.0, 90191.0, 90729.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0655464212092
varray: [-1.117, -1.198, -1.257, -1.254]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89118.0, 89655.0, 90191.0, 90729.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.229013100062
varray: [-49.937, -49.531, -49.982, -49.604]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89118.0, 89655.0, 90191.0, 90729.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 107.30625
varray: [105.641, 107.853, 106.75, 108.981]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [89118.0, 89655.0, 90191.0, 90729.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.43598917707
varray: [105.641, 107.853, 106.75, 108.981]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [89118.0, 89655.0, 90191.0, 90729.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.75475
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.46625
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1401.85825
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 4.5005
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.126630952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.126630952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.39480077344
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.4426547619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-1.117, -1.198, -1.257, -1.254]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [89118.0, 89655.0, 90191.0, 90729.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-49.937, -49.531, -49.982, -49.604]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [89118.0, 89655.0, 90191.0, 90729.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.2065
varray: [-1.117, -1.198, -1.257, -1.254]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [89118.0, 89655.0, 90191.0, 90729.0]
source_col: d(18_16)Mean
dD_mean | -49.7635
varray: [-49.937, -49.531, -49.982, -49.604]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [89118.0, 89655.0, 90191.0, 90729.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0655464212092
varray: [-1.117, -1.198, -1.257, -1.254]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89118.0, 89655.0, 90191.0, 90729.0]
source_col: d(18_16)Mean
dD_stdev | 0.229013100062
varray: [-49.937, -49.531, -49.982, -49.604]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89118.0, 89655.0, 90191.0, 90729.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.2065
varray: [-1.117, -1.198, -1.257, -1.254]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [89118.0, 89655.0, 90191.0, 90729.0]
source_col: d(18_16)Mean
measured_dD_mean | -49.7635
varray: [-49.937, -49.531, -49.982, -49.604]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [89118.0, 89655.0, 90191.0, 90729.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0655464212092
varray: [-1.117, -1.198, -1.257, -1.254]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89118.0, 89655.0, 90191.0, 90729.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.229013100062
varray: [-49.937, -49.531, -49.982, -49.604]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89118.0, 89655.0, 90191.0, 90729.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.79980527694
actual_array: [ 1.93 -15.84]
varray: [-1.117, -1.198, -1.257, -1.254]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -1.2065
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [89118.0, 89655.0, 90191.0, 90729.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -58.670585543
actual_array: [ 18.1 -119. ]
varray: [-49.937, -49.531, -49.982, -49.604]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -49.7635
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [89118.0, 89655.0, 90191.0, 90729.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.23863095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 106.864869048
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 36 - Sample 36 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 36
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.7775
varray: [2.732, 2.813, 2.858, 2.707]
inst: mean_all
operation: mean_all(residuals)
tarray: [91312.0, 91849.0, 92386.0, 92922.0]
source_col: residuals
SAMP_residuals_STD | 0.0701925922017
varray: [2.732, 2.813, 2.858, 2.707]
inst: std_all
operation: std_all(residuals)
tarray: [91312.0, 91849.0, 92386.0, 92922.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.05375
varray: [-0.053, 0.168, 0.225, -0.125]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [91312.0, 91849.0, 92386.0, 92922.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.16904314045
varray: [-0.053, 0.168, 0.225, -0.125]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [91312.0, 91849.0, 92386.0, 92922.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1337.62225
varray: [1336.933, 1339.074, 1338.24, 1336.242]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [91312.0, 91849.0, 92386.0, 92922.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.27401579661
varray: [1336.933, 1339.074, 1338.24, 1336.242]
inst: std_all
operation: std_all(baseline_shift)
tarray: [91312.0, 91849.0, 92386.0, 92922.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -0.69025
varray: [-1.047, -0.745, -0.825, -0.144]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [91312.0, 91849.0, 92386.0, 92922.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.385925187051
varray: [-1.047, -0.745, -0.825, -0.144]
inst: std_all
operation: std_all(slope_shift)
tarray: [91312.0, 91849.0, 92386.0, 92922.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0306920185064
varray: [-2.54, -2.51, -2.531, -2.583]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [91312.0, 91849.0, 92386.0, 92922.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.237793327633
varray: [-53.09, -52.959, -53.269, -53.508]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [91312.0, 91849.0, 92386.0, 92922.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 107.7565
varray: [106.568, 109.093, 109.471, 105.894]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [91312.0, 91849.0, 92386.0, 92922.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.7895233071
varray: [106.568, 109.093, 109.471, 105.894]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [91312.0, 91849.0, 92386.0, 92922.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.7775
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.05375
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1337.62225
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.69025
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.103880952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.103880952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.60524880596
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.9924047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-2.54, -2.51, -2.531, -2.583]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [91312.0, 91849.0, 92386.0, 92922.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-53.09, -52.959, -53.269, -53.508]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [91312.0, 91849.0, 92386.0, 92922.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.541
varray: [-2.54, -2.51, -2.531, -2.583]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [91312.0, 91849.0, 92386.0, 92922.0]
source_col: d(18_16)Mean
dD_mean | -53.2065
varray: [-53.09, -52.959, -53.269, -53.508]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [91312.0, 91849.0, 92386.0, 92922.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0306920185064
varray: [-2.54, -2.51, -2.531, -2.583]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [91312.0, 91849.0, 92386.0, 92922.0]
source_col: d(18_16)Mean
dD_stdev | 0.237793327633
varray: [-53.09, -52.959, -53.269, -53.508]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [91312.0, 91849.0, 92386.0, 92922.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.541
varray: [-2.54, -2.51, -2.531, -2.583]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [91312.0, 91849.0, 92386.0, 92922.0]
source_col: d(18_16)Mean
measured_dD_mean | -53.2065
varray: [-53.09, -52.959, -53.269, -53.508]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [91312.0, 91849.0, 92386.0, 92922.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0306920185064
varray: [-2.54, -2.51, -2.531, -2.583]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [91312.0, 91849.0, 92386.0, 92922.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.237793327633
varray: [-53.09, -52.959, -53.269, -53.508]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [91312.0, 91849.0, 92386.0, 92922.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.24311584564
actual_array: [ 1.93 -15.84]
varray: [-2.54, -2.51, -2.531, -2.583]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -2.541
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [91312.0, 91849.0, 92386.0, 92922.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -62.1202478907
actual_array: [ 18.1 -119. ]
varray: [-53.09, -52.959, -53.269, -53.508]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -53.2065
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [91312.0, 91849.0, 92386.0, 92922.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.21588095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 42.6288690476
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 37 - Sample 37 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 37
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.68925
varray: [2.63, 2.748, 2.697, 2.682]
inst: mean_all
operation: mean_all(residuals)
tarray: [93507.0, 94044.0, 94581.0, 95117.0]
source_col: residuals
SAMP_residuals_STD | 0.0485618162758
varray: [2.63, 2.748, 2.697, 2.682]
inst: std_all
operation: std_all(residuals)
tarray: [93507.0, 94044.0, 94581.0, 95117.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.852
varray: [-0.689, -0.853, -1.008, -0.858]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [93507.0, 94044.0, 94581.0, 95117.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.130309887064
varray: [-0.689, -0.853, -1.008, -0.858]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [93507.0, 94044.0, 94581.0, 95117.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1457.724
varray: [1447.826, 1459.761, 1465.028, 1458.281]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [93507.0, 94044.0, 94581.0, 95117.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 7.20597987785
varray: [1447.826, 1459.761, 1465.028, 1458.281]
inst: std_all
operation: std_all(baseline_shift)
tarray: [93507.0, 94044.0, 94581.0, 95117.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -7.2255
varray: [-7.124, -7.508, -7.186, -7.084]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [93507.0, 94044.0, 94581.0, 95117.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.1929516347
varray: [-7.124, -7.508, -7.186, -7.084]
inst: std_all
operation: std_all(slope_shift)
tarray: [93507.0, 94044.0, 94581.0, 95117.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0757957782465
varray: [-4.924, -5.092, -5.056, -5.07]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [93507.0, 94044.0, 94581.0, 95117.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.3127742888
varray: [-75.113, -77.434, -77.498, -78.085]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [93507.0, 94044.0, 94581.0, 95117.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 105.93925
varray: [103.617, 107.13, 107.144, 105.866]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [93507.0, 94044.0, 94581.0, 95117.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.66007216209
varray: [103.617, 107.13, 107.144, 105.866]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [93507.0, 94044.0, 94581.0, 95117.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.68925
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.852
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1457.724
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 7.2255
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.192130952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.192130952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.66801632815
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.8096547619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-4.924, -5.092, -5.056, -5.07]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [93507.0, 94044.0, 94581.0, 95117.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-75.113, -77.434, -77.498, -78.085]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [93507.0, 94044.0, 94581.0, 95117.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.0355
varray: [-4.924, -5.092, -5.056, -5.07]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [93507.0, 94044.0, 94581.0, 95117.0]
source_col: d(18_16)Mean
dD_mean | -77.0325
varray: [-75.113, -77.434, -77.498, -78.085]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [93507.0, 94044.0, 94581.0, 95117.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0757957782465
varray: [-4.924, -5.092, -5.056, -5.07]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [93507.0, 94044.0, 94581.0, 95117.0]
source_col: d(18_16)Mean
dD_stdev | 1.3127742888
varray: [-75.113, -77.434, -77.498, -78.085]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [93507.0, 94044.0, 94581.0, 95117.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.0355
varray: [-4.924, -5.092, -5.056, -5.07]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [93507.0, 94044.0, 94581.0, 95117.0]
source_col: d(18_16)Mean
measured_dD_mean | -77.0325
varray: [-75.113, -77.434, -77.498, -78.085]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [93507.0, 94044.0, 94581.0, 95117.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0757957782465
varray: [-4.924, -5.092, -5.056, -5.07]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [93507.0, 94044.0, 94581.0, 95117.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.3127742888
varray: [-75.113, -77.434, -77.498, -78.085]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [93507.0, 94044.0, 94581.0, 95117.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.94100884949
actual_array: [ 1.93 -15.84]
varray: [-4.924, -5.092, -5.056, -5.07]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -5.0355
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [93507.0, 94044.0, 94581.0, 95117.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -85.9923521882
actual_array: [ 18.1 -119. ]
varray: [-75.113, -77.434, -77.498, -78.085]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -77.0325
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [93507.0, 94044.0, 94581.0, 95117.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.30413095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 162.730619048
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 38 - Sample 38 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 38
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.705
varray: [2.646, 2.691, 2.759, 2.724]
inst: mean_all
operation: mean_all(residuals)
tarray: [95701.0, 96238.0, 96774.0, 97311.0]
source_col: residuals
SAMP_residuals_STD | 0.0481456124688
varray: [2.646, 2.691, 2.759, 2.724]
inst: std_all
operation: std_all(residuals)
tarray: [95701.0, 96238.0, 96774.0, 97311.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4305
varray: [0.402, 0.399, 0.533, 0.388]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [95701.0, 96238.0, 96774.0, 97311.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0685978619686
varray: [0.402, 0.399, 0.533, 0.388]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [95701.0, 96238.0, 96774.0, 97311.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1290.66525
varray: [1290.565, 1289.615, 1291.683, 1290.798]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [95701.0, 96238.0, 96774.0, 97311.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.849806399521
varray: [1290.565, 1289.615, 1291.683, 1290.798]
inst: std_all
operation: std_all(baseline_shift)
tarray: [95701.0, 96238.0, 96774.0, 97311.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.435
varray: [2.423, 2.061, 2.807, 2.449]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [95701.0, 96238.0, 96774.0, 97311.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.304740326617
varray: [2.423, 2.061, 2.807, 2.449]
inst: std_all
operation: std_all(slope_shift)
tarray: [95701.0, 96238.0, 96774.0, 97311.0]
source_col: slope_shift
SAMP_d18O_STD | 0.265010691608
varray: [-14.533, -14.933, -15.019, -15.147]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [95701.0, 96238.0, 96774.0, 97311.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.29184254477
varray: [-114.157, -117.82, -118.827, -119.149]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [95701.0, 96238.0, 96774.0, 97311.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 105.48725
varray: [103.612, 105.126, 107.144, 106.067]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [95701.0, 96238.0, 96774.0, 97311.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.49755297625
varray: [103.612, 105.126, 107.144, 106.067]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [95701.0, 96238.0, 96774.0, 97311.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.705
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4305
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1290.66525
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.435
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.176380952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.176380952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.12140342759
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.2616547619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-14.533, -14.933, -15.019, -15.147]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [95701.0, 96238.0, 96774.0, 97311.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-114.157, -117.82, -118.827, -119.149]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [95701.0, 96238.0, 96774.0, 97311.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -14.908
varray: [-14.533, -14.933, -15.019, -15.147]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [95701.0, 96238.0, 96774.0, 97311.0]
source_col: d(18_16)Mean
dD_mean | -117.48825
varray: [-114.157, -117.82, -118.827, -119.149]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [95701.0, 96238.0, 96774.0, 97311.0]
source_col: d(D_H)Mean
d18O_stdev | 0.265010691608
varray: [-14.533, -14.933, -15.019, -15.147]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [95701.0, 96238.0, 96774.0, 97311.0]
source_col: d(18_16)Mean
dD_stdev | 2.29184254477
varray: [-114.157, -117.82, -118.827, -119.149]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [95701.0, 96238.0, 96774.0, 97311.0]
source_col: d(D_H)Mean
measured_d18O_mean | -14.908
varray: [-14.533, -14.933, -15.019, -15.147]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [95701.0, 96238.0, 96774.0, 97311.0]
source_col: d(18_16)Mean
measured_dD_mean | -117.48825
varray: [-114.157, -117.82, -118.827, -119.149]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [95701.0, 96238.0, 96774.0, 97311.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.265010691608
varray: [-14.533, -14.933, -15.019, -15.147]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [95701.0, 96238.0, 96774.0, 97311.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.29184254477
varray: [-114.157, -117.82, -118.827, -119.149]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [95701.0, 96238.0, 96774.0, 97311.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -17.6184787555
actual_array: [ 1.93 -15.84]
varray: [-14.533, -14.933, -15.019, -15.147]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -14.908
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [95701.0, 96238.0, 96774.0, 97311.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -126.526385741
actual_array: [ 18.1 -119. ]
varray: [-114.157, -117.82, -118.827, -119.149]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -117.48825
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [95701.0, 96238.0, 96774.0, 97311.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.28838095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -4.32813095238
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 39 - Sample 39 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 39
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.7485
varray: [2.705, 2.745, 2.783, 2.761]
inst: mean_all
operation: mean_all(residuals)
tarray: [97896.0, 98432.0, 98968.0, 99505.0]
source_col: residuals
SAMP_residuals_STD | 0.0329190927376
varray: [2.705, 2.745, 2.783, 2.761]
inst: std_all
operation: std_all(residuals)
tarray: [97896.0, 98432.0, 98968.0, 99505.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.26675
varray: [-0.466, -0.243, -0.103, -0.255]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [97896.0, 98432.0, 98968.0, 99505.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.149684946916
varray: [-0.466, -0.243, -0.103, -0.255]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [97896.0, 98432.0, 98968.0, 99505.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1397.19225
varray: [1393.948, 1396.699, 1398.926, 1399.196]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [97896.0, 98432.0, 98968.0, 99505.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.43511510679
varray: [1393.948, 1396.699, 1398.926, 1399.196]
inst: std_all
operation: std_all(baseline_shift)
tarray: [97896.0, 98432.0, 98968.0, 99505.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -3.65125
varray: [-3.614, -3.783, -3.559, -3.649]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [97896.0, 98432.0, 98968.0, 99505.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0953253201761
varray: [-3.614, -3.783, -3.559, -3.649]
inst: std_all
operation: std_all(slope_shift)
tarray: [97896.0, 98432.0, 98968.0, 99505.0]
source_col: slope_shift
SAMP_d18O_STD | 0.305471629888
varray: [-5.197, -4.674, -4.61, -4.516]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [97896.0, 98432.0, 98968.0, 99505.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.45651234816
varray: [-78.995, -74.782, -74.184, -73.578]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [97896.0, 98432.0, 98968.0, 99505.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 107.98675
varray: [106.385, 108.08, 108.661, 108.821]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [97896.0, 98432.0, 98968.0, 99505.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.11428373855
varray: [106.385, 108.08, 108.661, 108.821]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [97896.0, 98432.0, 98968.0, 99505.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.7485
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.26675
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1397.19225
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 3.65125
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.132880952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.132880952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.61171065461
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.7621547619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-5.197, -4.674, -4.61, -4.516]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [97896.0, 98432.0, 98968.0, 99505.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-78.995, -74.782, -74.184, -73.578]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [97896.0, 98432.0, 98968.0, 99505.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.74925
varray: [-5.197, -4.674, -4.61, -4.516]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [97896.0, 98432.0, 98968.0, 99505.0]
source_col: d(18_16)Mean
dD_mean | -75.38475
varray: [-78.995, -74.782, -74.184, -73.578]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [97896.0, 98432.0, 98968.0, 99505.0]
source_col: d(D_H)Mean
d18O_stdev | 0.305471629888
varray: [-5.197, -4.674, -4.61, -4.516]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [97896.0, 98432.0, 98968.0, 99505.0]
source_col: d(18_16)Mean
dD_stdev | 2.45651234816
varray: [-78.995, -74.782, -74.184, -73.578]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [97896.0, 98432.0, 98968.0, 99505.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.74925
varray: [-5.197, -4.674, -4.61, -4.516]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [97896.0, 98432.0, 98968.0, 99505.0]
source_col: d(18_16)Mean
measured_dD_mean | -75.38475
varray: [-78.995, -74.782, -74.184, -73.578]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [97896.0, 98432.0, 98968.0, 99505.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.305471629888
varray: [-5.197, -4.674, -4.61, -4.516]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [97896.0, 98432.0, 98968.0, 99505.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.45651234816
varray: [-78.995, -74.782, -74.184, -73.578]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [97896.0, 98432.0, 98968.0, 99505.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.63141900289
actual_array: [ 1.93 -15.84]
varray: [-5.197, -4.674, -4.61, -4.516]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -4.74925
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [97896.0, 98432.0, 98968.0, 99505.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -84.3414137236
actual_array: [ 18.1 -119. ]
varray: [-78.995, -74.782, -74.184, -73.578]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -75.38475
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [97896.0, 98432.0, 98968.0, 99505.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.24488095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 102.198869048
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 40 - Sample 40 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 40
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.687
varray: [2.708, 2.642, 2.665, 2.733]
inst: mean_all
operation: mean_all(residuals)
tarray: [100091.0, 100627.0, 101163.0, 101700.0]
source_col: residuals
SAMP_residuals_STD | 0.0410933895738
varray: [2.708, 2.642, 2.665, 2.733]
inst: std_all
operation: std_all(residuals)
tarray: [100091.0, 100627.0, 101163.0, 101700.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3435
varray: [0.427, -0.039, 0.378, 0.608]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [100091.0, 100627.0, 101163.0, 101700.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.27351355847
varray: [0.427, -0.039, 0.378, 0.608]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [100091.0, 100627.0, 101163.0, 101700.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1288.68325
varray: [1291.306, 1288.402, 1287.329, 1287.696]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [100091.0, 100627.0, 101163.0, 101700.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.80430732323
varray: [1291.306, 1288.402, 1287.329, 1287.696]
inst: std_all
operation: std_all(baseline_shift)
tarray: [100091.0, 100627.0, 101163.0, 101700.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.7015
varray: [1.592, 1.231, 1.97, 2.013]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [100091.0, 100627.0, 101163.0, 101700.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.366280857631
varray: [1.592, 1.231, 1.97, 2.013]
inst: std_all
operation: std_all(slope_shift)
tarray: [100091.0, 100627.0, 101163.0, 101700.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0493828580245
varray: [-7.527, -7.597, -7.641, -7.619]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100091.0, 100627.0, 101163.0, 101700.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.911580495623
varray: [-62.874, -61.285, -61.009, -60.934]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100091.0, 100627.0, 101163.0, 101700.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 104.6405
varray: [106.026, 103.138, 103.55, 105.848]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [100091.0, 100627.0, 101163.0, 101700.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.50823992344
varray: [106.026, 103.138, 103.55, 105.848]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [100091.0, 100627.0, 101163.0, 101700.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.687
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3435
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1288.68325
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.7015
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.194380952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.194380952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.74610388537
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -6.1084047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-7.527, -7.597, -7.641, -7.619]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [100091.0, 100627.0, 101163.0, 101700.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-62.874, -61.285, -61.009, -60.934]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [100091.0, 100627.0, 101163.0, 101700.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.596
varray: [-7.527, -7.597, -7.641, -7.619]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [100091.0, 100627.0, 101163.0, 101700.0]
source_col: d(18_16)Mean
dD_mean | -61.5255
varray: [-62.874, -61.285, -61.009, -60.934]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [100091.0, 100627.0, 101163.0, 101700.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0493828580245
varray: [-7.527, -7.597, -7.641, -7.619]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100091.0, 100627.0, 101163.0, 101700.0]
source_col: d(18_16)Mean
dD_stdev | 0.911580495623
varray: [-62.874, -61.285, -61.009, -60.934]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100091.0, 100627.0, 101163.0, 101700.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.596
varray: [-7.527, -7.597, -7.641, -7.619]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [100091.0, 100627.0, 101163.0, 101700.0]
source_col: d(18_16)Mean
measured_dD_mean | -61.5255
varray: [-62.874, -61.285, -61.009, -60.934]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [100091.0, 100627.0, 101163.0, 101700.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0493828580245
varray: [-7.527, -7.597, -7.641, -7.619]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100091.0, 100627.0, 101163.0, 101700.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.911580495623
varray: [-62.874, -61.285, -61.009, -60.934]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100091.0, 100627.0, 101163.0, 101700.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.71028326775
actual_array: [ 1.93 -15.84]
varray: [-7.527, -7.597, -7.641, -7.619]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.596
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [100091.0, 100627.0, 101163.0, 101700.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -70.4553455005
actual_array: [ 18.1 -119. ]
varray: [-62.874, -61.285, -61.009, -60.934]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -61.5255
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [100091.0, 100627.0, 101163.0, 101700.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.30638095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -6.31013095238
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 41 - Sample 41 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 41
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.72025
varray: [2.527, 2.903, 2.727, 2.724]
inst: mean_all
operation: mean_all(residuals)
tarray: [102285.0, 102823.0, 103359.0, 103895.0]
source_col: residuals
SAMP_residuals_STD | 0.153625898424
varray: [2.527, 2.903, 2.727, 2.724]
inst: std_all
operation: std_all(residuals)
tarray: [102285.0, 102823.0, 103359.0, 103895.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.046
varray: [0.174, 0.143, -0.213, 0.08]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [102285.0, 102823.0, 103359.0, 103895.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.177040485012
varray: [0.174, 0.143, -0.213, 0.08]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [102285.0, 102823.0, 103359.0, 103895.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1312.205
varray: [1302.566, 1319.928, 1314.079, 1312.247]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [102285.0, 102823.0, 103359.0, 103895.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 7.21268997439
varray: [1302.566, 1319.928, 1314.079, 1312.247]
inst: std_all
operation: std_all(baseline_shift)
tarray: [102285.0, 102823.0, 103359.0, 103895.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -0.1065
varray: [0.3, -0.295, -0.191, -0.24]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [102285.0, 102823.0, 103359.0, 103895.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.274309435978
varray: [0.3, -0.295, -0.191, -0.24]
inst: std_all
operation: std_all(slope_shift)
tarray: [102285.0, 102823.0, 103359.0, 103895.0]
source_col: slope_shift
SAMP_d18O_STD | 0.28851343123
varray: [-3.543, -2.889, -2.997, -3.071]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [102285.0, 102823.0, 103359.0, 103895.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.00406287154
varray: [-58.795, -56.487, -56.922, -57.549]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [102285.0, 102823.0, 103359.0, 103895.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 106.394
varray: [99.926, 112.501, 106.943, 106.206]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [102285.0, 102823.0, 103359.0, 103895.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 5.14675360462
varray: [99.926, 112.501, 106.943, 106.206]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [102285.0, 102823.0, 103359.0, 103895.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.72025
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.046
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1312.205
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.1065
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.161130952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.161130952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.59214331752
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.3549047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-3.543, -2.889, -2.997, -3.071]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [102285.0, 102823.0, 103359.0, 103895.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-58.795, -56.487, -56.922, -57.549]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [102285.0, 102823.0, 103359.0, 103895.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.125
varray: [-3.543, -2.889, -2.997, -3.071]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [102285.0, 102823.0, 103359.0, 103895.0]
source_col: d(18_16)Mean
dD_mean | -57.43825
varray: [-58.795, -56.487, -56.922, -57.549]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [102285.0, 102823.0, 103359.0, 103895.0]
source_col: d(D_H)Mean
d18O_stdev | 0.28851343123
varray: [-3.543, -2.889, -2.997, -3.071]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [102285.0, 102823.0, 103359.0, 103895.0]
source_col: d(18_16)Mean
dD_stdev | 1.00406287154
varray: [-58.795, -56.487, -56.922, -57.549]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [102285.0, 102823.0, 103359.0, 103895.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.125
varray: [-3.543, -2.889, -2.997, -3.071]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [102285.0, 102823.0, 103359.0, 103895.0]
source_col: d(18_16)Mean
measured_dD_mean | -57.43825
varray: [-58.795, -56.487, -56.922, -57.549]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [102285.0, 102823.0, 103359.0, 103895.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.28851343123
varray: [-3.543, -2.889, -2.997, -3.071]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [102285.0, 102823.0, 103359.0, 103895.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.00406287154
varray: [-58.795, -56.487, -56.922, -57.549]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [102285.0, 102823.0, 103359.0, 103895.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.87473320954
actual_array: [ 1.93 -15.84]
varray: [-3.543, -2.889, -2.997, -3.071]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -3.125
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [102285.0, 102823.0, 103359.0, 103895.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -66.3601865024
actual_array: [ 18.1 -119. ]
varray: [-58.795, -56.487, -56.922, -57.549]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -57.43825
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [102285.0, 102823.0, 103359.0, 103895.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.27313095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 17.2116190476
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 42 - Sample 42 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 42
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.5345
varray: [2.529, 2.587, 2.534, 2.488]
inst: mean_all
operation: mean_all(residuals)
tarray: [104481.0, 105018.0, 105555.0, 106091.0]
source_col: residuals
SAMP_residuals_STD | 0.0406160887662
varray: [2.529, 2.587, 2.534, 2.488]
inst: std_all
operation: std_all(residuals)
tarray: [104481.0, 105018.0, 105555.0, 106091.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.2165
varray: [0.024, 0.164, 0.412, 0.266]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [104481.0, 105018.0, 105555.0, 106091.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.163791534173
varray: [0.024, 0.164, 0.412, 0.266]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [104481.0, 105018.0, 105555.0, 106091.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1290.28225
varray: [1292.409, 1292.496, 1289.644, 1286.58]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [104481.0, 105018.0, 105555.0, 106091.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.80105770201
varray: [1292.409, 1292.496, 1289.644, 1286.58]
inst: std_all
operation: std_all(baseline_shift)
tarray: [104481.0, 105018.0, 105555.0, 106091.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.66125
varray: [0.599, 0.542, 0.899, 0.605]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [104481.0, 105018.0, 105555.0, 106091.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.161022513954
varray: [0.599, 0.542, 0.899, 0.605]
inst: std_all
operation: std_all(slope_shift)
tarray: [104481.0, 105018.0, 105555.0, 106091.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0188922382651
varray: [-1.064, -1.019, -1.033, -1.035]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [104481.0, 105018.0, 105555.0, 106091.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.304834791321
varray: [-51.207, -50.628, -50.643, -50.541]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [104481.0, 105018.0, 105555.0, 106091.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 99.90425
varray: [100.351, 101.718, 99.902, 97.646]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [104481.0, 105018.0, 105555.0, 106091.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.69203750451
varray: [100.351, 101.718, 99.902, 97.646]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [104481.0, 105018.0, 105555.0, 106091.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.5345
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.2165
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1290.28225
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.66125
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.346880952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.346880952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -12.038704986
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -10.8446547619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-1.064, -1.019, -1.033, -1.035]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [104481.0, 105018.0, 105555.0, 106091.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-51.207, -50.628, -50.643, -50.541]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [104481.0, 105018.0, 105555.0, 106091.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.03775
varray: [-1.064, -1.019, -1.033, -1.035]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [104481.0, 105018.0, 105555.0, 106091.0]
source_col: d(18_16)Mean
dD_mean | -50.75475
varray: [-51.207, -50.628, -50.643, -50.541]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [104481.0, 105018.0, 105555.0, 106091.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0188922382651
varray: [-1.064, -1.019, -1.033, -1.035]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [104481.0, 105018.0, 105555.0, 106091.0]
source_col: d(18_16)Mean
dD_stdev | 0.304834791321
varray: [-51.207, -50.628, -50.643, -50.541]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [104481.0, 105018.0, 105555.0, 106091.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.03775
varray: [-1.064, -1.019, -1.033, -1.035]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [104481.0, 105018.0, 105555.0, 106091.0]
source_col: d(18_16)Mean
measured_dD_mean | -50.75475
varray: [-51.207, -50.628, -50.643, -50.541]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [104481.0, 105018.0, 105555.0, 106091.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0188922382651
varray: [-1.064, -1.019, -1.033, -1.035]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [104481.0, 105018.0, 105555.0, 106091.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.304834791321
varray: [-51.207, -50.628, -50.643, -50.541]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [104481.0, 105018.0, 105555.0, 106091.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.61729597872
actual_array: [ 1.93 -15.84]
varray: [-1.064, -1.019, -1.033, -1.035]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -1.03775
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [104481.0, 105018.0, 105555.0, 106091.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -59.6637536529
actual_array: [ 18.1 -119. ]
varray: [-51.207, -50.628, -50.643, -50.541]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -50.75475
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [104481.0, 105018.0, 105555.0, 106091.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.45888095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -4.71113095238
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 43 - Sample 43 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 43
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.66125
varray: [2.623, 2.655, 2.736, 2.631]
inst: mean_all
operation: mean_all(residuals)
tarray: [106676.0, 107212.0, 107748.0, 108285.0]
source_col: residuals
SAMP_residuals_STD | 0.0516551062336
varray: [2.623, 2.655, 2.736, 2.631]
inst: std_all
operation: std_all(residuals)
tarray: [106676.0, 107212.0, 107748.0, 108285.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.30875
varray: [0.22, 0.444, 0.507, 0.064]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [106676.0, 107212.0, 107748.0, 108285.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.204429898335
varray: [0.22, 0.444, 0.507, 0.064]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [106676.0, 107212.0, 107748.0, 108285.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1287.16825
varray: [1285.622, 1288.21, 1287.987, 1286.854]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [106676.0, 107212.0, 107748.0, 108285.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.18957089602
varray: [1285.622, 1288.21, 1287.987, 1286.854]
inst: std_all
operation: std_all(baseline_shift)
tarray: [106676.0, 107212.0, 107748.0, 108285.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.15425
varray: [1.982, 2.186, 1.955, 2.494]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [106676.0, 107212.0, 107748.0, 108285.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.248869945956
varray: [1.982, 2.186, 1.955, 2.494]
inst: std_all
operation: std_all(slope_shift)
tarray: [106676.0, 107212.0, 107748.0, 108285.0]
source_col: slope_shift
SAMP_d18O_STD | 0.29165904752
varray: [-10.578, -11.034, -11.166, -11.22]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [106676.0, 107212.0, 107748.0, 108285.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.69847319485
varray: [-94.53, -98.613, -100.015, -100.453]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [106676.0, 107212.0, 107748.0, 108285.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 103.10575
varray: [101.626, 103.954, 105.051, 101.792]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [106676.0, 107212.0, 107748.0, 108285.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.67522384076
varray: [101.626, 103.954, 105.051, 101.792]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [106676.0, 107212.0, 107748.0, 108285.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.66125
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.30875
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1287.16825
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.15425
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.220130952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.220130952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -7.63977259581
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -7.6431547619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-10.578, -11.034, -11.166, -11.22]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [106676.0, 107212.0, 107748.0, 108285.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-94.53, -98.613, -100.015, -100.453]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [106676.0, 107212.0, 107748.0, 108285.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -10.9995
varray: [-10.578, -11.034, -11.166, -11.22]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [106676.0, 107212.0, 107748.0, 108285.0]
source_col: d(18_16)Mean
dD_mean | -98.40275
varray: [-94.53, -98.613, -100.015, -100.453]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [106676.0, 107212.0, 107748.0, 108285.0]
source_col: d(D_H)Mean
d18O_stdev | 0.29165904752
varray: [-10.578, -11.034, -11.166, -11.22]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [106676.0, 107212.0, 107748.0, 108285.0]
source_col: d(18_16)Mean
dD_stdev | 2.69847319485
varray: [-94.53, -98.613, -100.015, -100.453]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [106676.0, 107212.0, 107748.0, 108285.0]
source_col: d(D_H)Mean
measured_d18O_mean | -10.9995
varray: [-10.578, -11.034, -11.166, -11.22]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [106676.0, 107212.0, 107748.0, 108285.0]
source_col: d(18_16)Mean
measured_dD_mean | -98.40275
varray: [-94.53, -98.613, -100.015, -100.453]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [106676.0, 107212.0, 107748.0, 108285.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.29165904752
varray: [-10.578, -11.034, -11.166, -11.22]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [106676.0, 107212.0, 107748.0, 108285.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.69847319485
varray: [-94.53, -98.613, -100.015, -100.453]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [106676.0, 107212.0, 107748.0, 108285.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -13.3912930247
actual_array: [ 1.93 -15.84]
varray: [-10.578, -11.034, -11.166, -11.22]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -10.9995
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [106676.0, 107212.0, 107748.0, 108285.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -107.403954508
actual_array: [ 18.1 -119. ]
varray: [-94.53, -98.613, -100.015, -100.453]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -98.40275
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [106676.0, 107212.0, 107748.0, 108285.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.33213095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -7.82513095238
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 44 - Sample 44 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 44
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.2195
varray: [2.317, 2.561, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [108870.0, 109406.0, 110050.0, 110652.0]
source_col: residuals
SAMP_residuals_STD | 1.41167619044
varray: [2.317, 2.561, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [108870.0, 109406.0, 110050.0, 110652.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0555
varray: [-0.11, 0.332, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [108870.0, 109406.0, 110050.0, 110652.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.191488032698
varray: [-0.11, 0.332, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [108870.0, 109406.0, 110050.0, 110652.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 639.278
varray: [1275.508, 1281.604, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [108870.0, 109406.0, 110050.0, 110652.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 738.178845932
varray: [1275.508, 1281.604, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [108870.0, 109406.0, 110050.0, 110652.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.11225
varray: [2.313, 2.136, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [108870.0, 109406.0, 110050.0, 110652.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.28634686224
varray: [2.313, 2.136, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [108870.0, 109406.0, 110050.0, 110652.0]
source_col: slope_shift
SAMP_d18O_STD | 6.49332792821
varray: [-11.302, -11.191, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [108870.0, 109406.0, 110050.0, 110652.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 58.8708159808
varray: [-102.242, -101.691, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [108870.0, 109406.0, 110050.0, 110652.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 47.621
varray: [91.162, 99.322, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [108870.0, 109406.0, 110050.0, 110652.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 55.0888112173
varray: [91.162, 99.322, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [108870.0, 109406.0, 110050.0, 110652.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.2195
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0555
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 639.278
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.11225
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.66188095238
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.66188095238
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -57.6765439852
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -63.1279047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-11.302, -11.191, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [108870.0, 109406.0, 110050.0, 110652.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-102.242, -101.691, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [108870.0, 109406.0, 110050.0, 110652.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.62325
varray: [-11.302, -11.191, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [108870.0, 109406.0, 110050.0, 110652.0]
source_col: d(18_16)Mean
dD_mean | -50.98325
varray: [-102.242, -101.691, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [108870.0, 109406.0, 110050.0, 110652.0]
source_col: d(D_H)Mean
d18O_stdev | 6.49332792821
varray: [-11.302, -11.191, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [108870.0, 109406.0, 110050.0, 110652.0]
source_col: d(18_16)Mean
dD_stdev | 58.8708159808
varray: [-102.242, -101.691, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [108870.0, 109406.0, 110050.0, 110652.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.62325
varray: [-11.302, -11.191, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [108870.0, 109406.0, 110050.0, 110652.0]
source_col: d(18_16)Mean
measured_dD_mean | -50.98325
varray: [-102.242, -101.691, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [108870.0, 109406.0, 110050.0, 110652.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 6.49332792821
varray: [-11.302, -11.191, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [108870.0, 109406.0, 110050.0, 110652.0]
source_col: d(18_16)Mean
measured_dD_stdev | 58.8708159808
varray: [-102.242, -101.691, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [108870.0, 109406.0, 110050.0, 110652.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.57668197557
actual_array: [ 1.93 -15.84]
varray: [-11.302, -11.191, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 1.38222915025
except: True
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -5.62325
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [108870.0, 109406.0, 110050.0, 110652.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -59.8926958098
actual_array: [ 18.1 -119. ]
varray: [-102.242, -101.691, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 8.79377478947
except: True
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -50.98325
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [108870.0, 109406.0, 110050.0, 110652.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1293.77388095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -655.715380952
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 45 - Sample 45 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 45
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [111303.0, 111904.0, 112505.0, 113105.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [111303.0, 111904.0, 112505.0, 113105.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [111303.0, 111904.0, 112505.0, 113105.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [111303.0, 111904.0, 112505.0, 113105.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [111303.0, 111904.0, 112505.0, 113105.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [111303.0, 111904.0, 112505.0, 113105.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [111303.0, 111904.0, 112505.0, 113105.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [111303.0, 111904.0, 112505.0, 113105.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [111303.0, 111904.0, 112505.0, 113105.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [111303.0, 111904.0, 112505.0, 113105.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [111303.0, 111904.0, 112505.0, 113105.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [111303.0, 111904.0, 112505.0, 113105.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.88138095238
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.88138095238
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -110.748904762
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [111303.0, 111904.0, 112505.0, 113105.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [111303.0, 111904.0, 112505.0, 113105.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [111303.0, 111904.0, 112505.0, 113105.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [111303.0, 111904.0, 112505.0, 113105.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [111303.0, 111904.0, 112505.0, 113105.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [111303.0, 111904.0, 112505.0, 113105.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [111303.0, 111904.0, 112505.0, 113105.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [111303.0, 111904.0, 112505.0, 113105.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [111303.0, 111904.0, 112505.0, 113105.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [111303.0, 111904.0, 112505.0, 113105.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 1.38222915025
except: True
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [111303.0, 111904.0, 112505.0, 113105.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 8.79377478947
except: True
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [111303.0, 111904.0, 112505.0, 113105.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1294.99338095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1294.99338095
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 46 - Sample 46 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 46
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.341666666667
varray: [1.025, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [113756.0, 114293.0, 114930.0]
source_col: residuals
SAMP_residuals_STD | 0.591784025919
varray: [1.025, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [113756.0, 114293.0, 114930.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.0173333333333
varray: [-0.052, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [113756.0, 114293.0, 114930.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0300222139979
varray: [-0.052, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [113756.0, 114293.0, 114930.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 407.577333333
varray: [1222.732, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [113756.0, 114293.0, 114930.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 705.944649347
varray: [1222.732, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [113756.0, 114293.0, 114930.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0526666666667
varray: [0.158, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [113756.0, 114293.0, 114930.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.091221342532
varray: [0.158, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [113756.0, 114293.0, 114930.0]
source_col: slope_shift
SAMP_d18O_STD | 3.77644811077
varray: [-6.541, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [113756.0, 114293.0, 114930.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 45.4178362761
varray: [-78.666, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [113756.0, 114293.0, 114930.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 14.1536666667
varray: [42.461, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [113756.0, 114293.0, 114930.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 24.5148697801
varray: [42.461, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [113756.0, 114293.0, 114930.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.341666666667
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0173333333333
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 407.577333333
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0526666666667
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.53971428571
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.53971428571
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -88.1422598291
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -96.5952380952
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-6.541, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [113756.0, 114293.0, 114930.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-78.666, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [113756.0, 114293.0, 114930.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.18033333333
varray: [-6.541, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [113756.0, 114293.0, 114930.0]
source_col: d(18_16)Mean
dD_mean | -26.222
varray: [-78.666, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [113756.0, 114293.0, 114930.0]
source_col: d(D_H)Mean
d18O_stdev | 3.77644811077
varray: [-6.541, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [113756.0, 114293.0, 114930.0]
source_col: d(18_16)Mean
dD_stdev | 45.4178362761
varray: [-78.666, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [113756.0, 114293.0, 114930.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.18033333333
varray: [-6.541, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [113756.0, 114293.0, 114930.0]
source_col: d(18_16)Mean
measured_dD_mean | -26.222
varray: [-78.666, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [113756.0, 114293.0, 114930.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 3.77644811077
varray: [-6.541, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [113756.0, 114293.0, 114930.0]
source_col: d(18_16)Mean
measured_dD_stdev | 45.4178362761
varray: [-78.666, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [113756.0, 114293.0, 114930.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.85304165173
actual_array: [ 1.93 -15.84]
varray: [-6.541, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 1.38222915025
except: True
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -2.18033333333
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [113756.0, 114293.0, 114930.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -35.0835317649
actual_array: [ 18.1 -119. ]
varray: [-78.666, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 8.79377478947
except: True
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -26.222
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [113756.0, 114293.0, 114930.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1294.65171429
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -887.416047619
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 47 - Sample 47 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 47
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.302
varray: [1.949, 1.957, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [116655.0, 117192.0, 117729.0]
source_col: residuals
SAMP_residuals_STD | 1.12757217064
varray: [1.949, 1.957, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [116655.0, 117192.0, 117729.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.885
varray: [-1.321, -1.334, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [116655.0, 117192.0, 117729.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.766460044621
varray: [-1.321, -1.334, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [116655.0, 117192.0, 117729.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 870.679666667
varray: [1305.658, 1306.381, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [116655.0, 117192.0, 117729.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 754.030796548
varray: [1305.658, 1306.381, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [116655.0, 117192.0, 117729.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -2.96633333333
varray: [-4.976, -3.923, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [116655.0, 117192.0, 117729.0]
source_col: slope_shift
SAMP_slope_shift_STD | 2.62231812207
varray: [-4.976, -3.923, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [116655.0, 117192.0, 117729.0]
source_col: slope_shift
SAMP_d18O_STD | 1.93626866249
varray: [3.313, 3.393, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [116655.0, 117192.0, 117729.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 4.19418673086
varray: [6.697, 7.723, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [116655.0, 117192.0, 117729.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 53.1026666667
varray: [79.329, 79.979, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [116655.0, 117192.0, 117729.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 45.9894067186
varray: [79.329, 79.979, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [116655.0, 117192.0, 117729.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.302
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.885
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 870.679666667
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.96633333333
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.57938095238
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.57938095238
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -54.8133335537
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -57.6462380952
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [3.313, 3.393, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [116655.0, 117192.0, 117729.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [6.697, 7.723, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [116655.0, 117192.0, 117729.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.23533333333
varray: [3.313, 3.393, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [116655.0, 117192.0, 117729.0]
source_col: d(18_16)Mean
dD_mean | 4.80666666667
varray: [6.697, 7.723, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [116655.0, 117192.0, 117729.0]
source_col: d(D_H)Mean
d18O_stdev | 1.93626866249
varray: [3.313, 3.393, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [116655.0, 117192.0, 117729.0]
source_col: d(18_16)Mean
dD_stdev | 4.19418673086
varray: [6.697, 7.723, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [116655.0, 117192.0, 117729.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.23533333333
varray: [3.313, 3.393, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [116655.0, 117192.0, 117729.0]
source_col: d(18_16)Mean
measured_dD_mean | 4.80666666667
varray: [6.697, 7.723, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [116655.0, 117192.0, 117729.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 1.93626866249
varray: [3.313, 3.393, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [116655.0, 117192.0, 117729.0]
source_col: d(18_16)Mean
measured_dD_stdev | 4.19418673086
varray: [6.697, 7.723, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [116655.0, 117192.0, 117729.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 0.922663382275
actual_array: [ 1.93 -15.84]
varray: [3.313, 3.393, 0.0]
fit_rsq: 1.0
fit_intercept: 1.38222915025
except: True
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 2.23533333333
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [116655.0, 117192.0, 117729.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -3.99482334226
actual_array: [ 18.1 -119. ]
varray: [6.697, 7.723, 0.0]
fit_rsq: 1.0
fit_intercept: 8.79377478947
except: True
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 4.80666666667
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [116655.0, 117192.0, 117729.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1293.69138095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -424.313714286
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 48 - Sample 48 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 48
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.73733333333
varray: [3.611, 1.892, 2.709]
inst: mean_all
operation: mean_all(residuals)
tarray: [118916.0, 119454.0, 119990.0]
source_col: residuals
SAMP_residuals_STD | 0.859850180749
varray: [3.611, 1.892, 2.709]
inst: std_all
operation: std_all(residuals)
tarray: [118916.0, 119454.0, 119990.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.0446666666667
varray: [0.051, -0.187, 0.002]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [118916.0, 119454.0, 119990.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.125675508089
varray: [0.051, -0.187, 0.002]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [118916.0, 119454.0, 119990.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1295.89266667
varray: [1331.667, 1261.034, 1294.977]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [118916.0, 119454.0, 119990.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 35.3254017151
varray: [1331.667, 1261.034, 1294.977]
inst: std_all
operation: std_all(baseline_shift)
tarray: [118916.0, 119454.0, 119990.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -0.225333333333
varray: [-1.867, 0.723, 0.468]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [118916.0, 119454.0, 119990.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.42743067549
varray: [-1.867, 0.723, 0.468]
inst: std_all
operation: std_all(slope_shift)
tarray: [118916.0, 119454.0, 119990.0]
source_col: slope_shift
SAMP_d18O_STD | 0.434514671789
varray: [3.091, 2.228, 2.748]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [118916.0, 119454.0, 119990.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.480950014
varray: [24.543, 17.608, 21.605]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [118916.0, 119454.0, 119990.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 105.632333333
varray: [136.101, 76.109, 104.687]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [118916.0, 119454.0, 119990.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 30.0071700987
varray: [136.101, 76.109, 104.687]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [118916.0, 119454.0, 119990.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.73733333333
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0446666666667
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1295.89266667
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.225333333333
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.144047619048
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.144047619048
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.99925630898
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.11657142857
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [3.091, 2.228, 2.748]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [118916.0, 119454.0, 119990.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [24.543, 17.608, 21.605]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [118916.0, 119454.0, 119990.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.689
varray: [3.091, 2.228, 2.748]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [118916.0, 119454.0, 119990.0]
source_col: d(18_16)Mean
dD_mean | 21.252
varray: [24.543, 17.608, 21.605]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [118916.0, 119454.0, 119990.0]
source_col: d(D_H)Mean
d18O_stdev | 0.434514671789
varray: [3.091, 2.228, 2.748]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [118916.0, 119454.0, 119990.0]
source_col: d(18_16)Mean
dD_stdev | 3.480950014
varray: [24.543, 17.608, 21.605]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [118916.0, 119454.0, 119990.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.689
varray: [3.091, 2.228, 2.748]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [118916.0, 119454.0, 119990.0]
source_col: d(18_16)Mean
measured_dD_mean | 21.252
varray: [24.543, 17.608, 21.605]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [118916.0, 119454.0, 119990.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.434514671789
varray: [3.091, 2.228, 2.748]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [118916.0, 119454.0, 119990.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.480950014
varray: [24.543, 17.608, 21.605]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [118916.0, 119454.0, 119990.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.41332047832
actual_array: [ 1.93 -15.84]
varray: [3.091, 2.228, 2.748]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 2.689
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [118916.0, 119454.0, 119990.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 12.4823323926
actual_array: [ 18.1 -119. ]
varray: [24.543, 17.608, 21.605]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 21.252
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [118916.0, 119454.0, 119990.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.25604762
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.899285714286
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 49 - Sample 49 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 49
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [120577.0, 121179.0, 121779.0, 122381.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [120577.0, 121179.0, 121779.0, 122381.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [120577.0, 121179.0, 121779.0, 122381.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [120577.0, 121179.0, 121779.0, 122381.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [120577.0, 121179.0, 121779.0, 122381.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [120577.0, 121179.0, 121779.0, 122381.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [120577.0, 121179.0, 121779.0, 122381.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [120577.0, 121179.0, 121779.0, 122381.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [120577.0, 121179.0, 121779.0, 122381.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [120577.0, 121179.0, 121779.0, 122381.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [120577.0, 121179.0, 121779.0, 122381.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [120577.0, 121179.0, 121779.0, 122381.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.88138095238
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.88138095238
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -110.748904762
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [120577.0, 121179.0, 121779.0, 122381.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [120577.0, 121179.0, 121779.0, 122381.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [120577.0, 121179.0, 121779.0, 122381.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [120577.0, 121179.0, 121779.0, 122381.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [120577.0, 121179.0, 121779.0, 122381.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [120577.0, 121179.0, 121779.0, 122381.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [120577.0, 121179.0, 121779.0, 122381.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [120577.0, 121179.0, 121779.0, 122381.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [120577.0, 121179.0, 121779.0, 122381.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [120577.0, 121179.0, 121779.0, 122381.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 1.38222915025
except: True
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [120577.0, 121179.0, 121779.0, 122381.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 8.79377478947
except: True
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [120577.0, 121179.0, 121779.0, 122381.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1294.99338095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1294.99338095
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 50 - Sample 50 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 50
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.03066666667
varray: [1.493, 0.0, 1.599]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [123567.0, 124104.0, 124705.0]
source_col: residuals
SAMP_residuals_STD | 0.894155653862
varray: [1.493, 0.0, 1.599]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [123567.0, 124104.0, 124705.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.284333333333
varray: [-0.374, 0.0, -0.479]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [123567.0, 124104.0, 124705.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.25177436989
varray: [-0.374, 0.0, -0.479]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [123567.0, 124104.0, 124705.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 835.386333333
varray: [1250.482, 0.0, 1255.677]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [123567.0, 124104.0, 124705.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 723.470449601
varray: [1250.482, 0.0, 1255.677]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [123567.0, 124104.0, 124705.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -0.638333333333
varray: [-0.986, 0.0, -0.929]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [123567.0, 124104.0, 124705.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.553547047082
varray: [-0.986, 0.0, -0.929]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [123567.0, 124104.0, 124705.0]
source_col: slope_shift
SAMP_d18O_STD | 0.904238906484
varray: [1.552, 0.0, 1.58]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [123567.0, 124104.0, 124705.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.64773872075
varray: [-4.699, 0.0, -4.464]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [123567.0, 124104.0, 124705.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 42.535
varray: [61.919, 0.0, 65.686]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [123567.0, 124104.0, 124705.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 36.8845122104
varray: [61.919, 0.0, 65.686]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [123567.0, 124104.0, 124705.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.03066666667
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.284333333333
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 835.386333333
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.638333333333
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.85071428571
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.85071428571
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -64.2301145284
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -68.2139047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [1.552, 0.0, 1.58]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [123567.0, 124104.0, 124705.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-4.699, 0.0, -4.464]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [123567.0, 124104.0, 124705.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.044
varray: [1.552, 0.0, 1.58]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [123567.0, 124104.0, 124705.0]
source_col: d(18_16)Mean
dD_mean | -3.05433333333
varray: [-4.699, 0.0, -4.464]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [123567.0, 124104.0, 124705.0]
source_col: d(D_H)Mean
d18O_stdev | 0.904238906484
varray: [1.552, 0.0, 1.58]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [123567.0, 124104.0, 124705.0]
source_col: d(18_16)Mean
dD_stdev | 2.64773872075
varray: [-4.699, 0.0, -4.464]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [123567.0, 124104.0, 124705.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.044
varray: [1.552, 0.0, 1.58]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [123567.0, 124104.0, 124705.0]
source_col: d(18_16)Mean
measured_dD_mean | -3.05433333333
varray: [-4.699, 0.0, -4.464]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [123567.0, 124104.0, 124705.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.904238906484
varray: [1.552, 0.0, 1.58]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [123567.0, 124104.0, 124705.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.64773872075
varray: [-4.699, 0.0, -4.464]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [123567.0, 124104.0, 124705.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.365807199104
actual_array: [ 1.93 -15.84]
varray: [1.552, 0.0, 1.58]
fit_rsq: 1.0
fit_intercept: 1.38222915025
except: True
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 1.044
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [123567.0, 124104.0, 124705.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -11.8710347031
actual_array: [ 18.1 -119. ]
varray: [-4.699, 0.0, -4.464]
fit_rsq: 1.0
fit_intercept: 8.79377478947
except: True
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -3.05433333333
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [123567.0, 124104.0, 124705.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1293.96271429
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -459.607047619
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 51 - Sample 51 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 51
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.556
varray: [2.617, 2.051, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [125828.0, 126364.0, 126902.0]
source_col: residuals
SAMP_residuals_STD | 1.37693173397
varray: [2.617, 2.051, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [125828.0, 126364.0, 126902.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.396333333333
varray: [-0.699, -0.49, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [125828.0, 126364.0, 126902.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.358790096482
varray: [-0.699, -0.49, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [125828.0, 126364.0, 126902.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 863.627333333
varray: [1310.603, 1280.279, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [125828.0, 126364.0, 126902.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 748.076877336
varray: [1310.603, 1280.279, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [125828.0, 126364.0, 126902.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -1.297
varray: [-2.492, -1.399, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [125828.0, 126364.0, 126902.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.24912729535
varray: [-2.492, -1.399, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [125828.0, 126364.0, 126902.0]
source_col: slope_shift
SAMP_d18O_STD | 1.14524640726
varray: [2.104, 1.836, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [125828.0, 126364.0, 126902.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.45536627692
varray: [0.256, -2.383, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [125828.0, 126364.0, 126902.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.7126666667
varray: [104.878, 83.26, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [125828.0, 126364.0, 126902.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 55.3759279952
varray: [104.878, 83.26, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [125828.0, 126364.0, 126902.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.556
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.396333333333
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 863.627333333
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.297
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.32538095238
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.32538095238
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -45.9981159827
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -48.0362380952
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [2.104, 1.836, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [125828.0, 126364.0, 126902.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [0.256, -2.383, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [125828.0, 126364.0, 126902.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.31333333333
varray: [2.104, 1.836, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [125828.0, 126364.0, 126902.0]
source_col: d(18_16)Mean
dD_mean | -0.709
varray: [0.256, -2.383, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [125828.0, 126364.0, 126902.0]
source_col: d(D_H)Mean
d18O_stdev | 1.14524640726
varray: [2.104, 1.836, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [125828.0, 126364.0, 126902.0]
source_col: d(18_16)Mean
dD_stdev | 1.45536627692
varray: [0.256, -2.383, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [125828.0, 126364.0, 126902.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.31333333333
varray: [2.104, 1.836, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [125828.0, 126364.0, 126902.0]
source_col: d(18_16)Mean
measured_dD_mean | -0.709
varray: [0.256, -2.383, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [125828.0, 126364.0, 126902.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 1.14524640726
varray: [2.104, 1.836, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [125828.0, 126364.0, 126902.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.45536627692
varray: [0.256, -2.383, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [125828.0, 126364.0, 126902.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.074513346347
actual_array: [ 1.93 -15.84]
varray: [2.104, 1.836, 0.0]
fit_rsq: 1.0
fit_intercept: 1.38222915025
except: True
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 1.31333333333
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [125828.0, 126364.0, 126902.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -9.52116305252
actual_array: [ 18.1 -119. ]
varray: [0.256, -2.383, 0.0]
fit_rsq: 1.0
fit_intercept: 8.79377478947
except: True
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -0.709
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [125828.0, 126364.0, 126902.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1293.43738095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -431.366047619
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 52 - Sample 52 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 52
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.645666666667
varray: [0.0, 0.0, 1.937]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [128089.0, 128690.0, 129291.0]
source_col: residuals
SAMP_residuals_STD | 1.11832747142
varray: [0.0, 0.0, 1.937]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [128089.0, 128690.0, 129291.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.393333333333
varray: [0.0, 0.0, -1.18]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [128089.0, 128690.0, 129291.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.681273317644
varray: [0.0, 0.0, -1.18]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [128089.0, 128690.0, 129291.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 431.449666667
varray: [0.0, 0.0, 1294.349]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [128089.0, 128690.0, 129291.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 747.292743575
varray: [0.0, 0.0, 1294.349]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [128089.0, 128690.0, 129291.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -1.45266666667
varray: [0.0, 0.0, -4.358]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [128089.0, 128690.0, 129291.0]
source_col: slope_shift
SAMP_slope_shift_STD | 2.51609247313
varray: [0.0, 0.0, -4.358]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [128089.0, 128690.0, 129291.0]
source_col: slope_shift
SAMP_d18O_STD | 1.89544093375
varray: [0.0, 0.0, 3.283]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [128089.0, 128690.0, 129291.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.709563480834
varray: [0.0, 0.0, 1.229]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [128089.0, 128690.0, 129291.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 26.3033333333
varray: [0.0, 0.0, 78.91]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [128089.0, 128690.0, 129291.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 45.5587097418
varray: [0.0, 0.0, 78.91]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [128089.0, 128690.0, 129291.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.645666666667
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.393333333333
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 431.449666667
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.45266666667
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.23571428571
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.23571428571
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -77.5917632088
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -84.4455714286
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [0.0, 0.0, 3.283]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [128089.0, 128690.0, 129291.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [0.0, 0.0, 1.229]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [128089.0, 128690.0, 129291.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.09433333333
varray: [0.0, 0.0, 3.283]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [128089.0, 128690.0, 129291.0]
source_col: d(18_16)Mean
dD_mean | 0.409666666667
varray: [0.0, 0.0, 1.229]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [128089.0, 128690.0, 129291.0]
source_col: d(D_H)Mean
d18O_stdev | 1.89544093375
varray: [0.0, 0.0, 3.283]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [128089.0, 128690.0, 129291.0]
source_col: d(18_16)Mean
dD_stdev | 0.709563480834
varray: [0.0, 0.0, 1.229]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [128089.0, 128690.0, 129291.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.09433333333
varray: [0.0, 0.0, 3.283]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [128089.0, 128690.0, 129291.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.409666666667
varray: [0.0, 0.0, 1.229]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [128089.0, 128690.0, 129291.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 1.89544093375
varray: [0.0, 0.0, 3.283]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [128089.0, 128690.0, 129291.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.709563480834
varray: [0.0, 0.0, 1.229]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [128089.0, 128690.0, 129291.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.311369857809
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 3.283]
fit_rsq: 1.0
fit_intercept: 1.38222915025
except: True
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 1.09433333333
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [128089.0, 128690.0, 129291.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -8.40033171949
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 1.229]
fit_rsq: 1.0
fit_intercept: 8.79377478947
except: True
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.409666666667
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [128089.0, 128690.0, 129291.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1294.34771429
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -863.543714286
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 53 - Sample 53 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 53
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.31633333333
varray: [2.197, 1.752, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [130415.0, 130952.0, 131488.0]
source_col: residuals
SAMP_residuals_STD | 1.1614888434
varray: [2.197, 1.752, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [130415.0, 130952.0, 131488.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.660333333333
varray: [-1.249, -0.732, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [130415.0, 130952.0, 131488.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.627576555755
varray: [-1.249, -0.732, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [130415.0, 130952.0, 131488.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 859.542666667
varray: [1304.709, 1273.919, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [130415.0, 130952.0, 131488.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 744.544963646
varray: [1304.709, 1273.919, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [130415.0, 130952.0, 131488.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -2.21833333333
varray: [-4.211, -2.444, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [130415.0, 130952.0, 131488.0]
source_col: slope_shift
SAMP_slope_shift_STD | 2.11455062208
varray: [-4.211, -2.444, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [130415.0, 130952.0, 131488.0]
source_col: slope_shift
SAMP_d18O_STD | 1.82199268202
varray: [3.322, 2.958, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [130415.0, 130952.0, 131488.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.29946500714
varray: [6.576, 3.764, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [130415.0, 130952.0, 131488.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 53.5473333333
varray: [90.126, 70.516, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [130415.0, 130952.0, 131488.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 47.3985833684
varray: [90.126, 70.516, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [130415.0, 130952.0, 131488.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.31633333333
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.660333333333
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 859.542666667
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.21833333333
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.56504761905
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.56504761905
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -54.3158868929
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -57.2015714286
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [3.322, 2.958, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [130415.0, 130952.0, 131488.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [6.576, 3.764, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [130415.0, 130952.0, 131488.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.09333333333
varray: [3.322, 2.958, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [130415.0, 130952.0, 131488.0]
source_col: d(18_16)Mean
dD_mean | 3.44666666667
varray: [6.576, 3.764, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [130415.0, 130952.0, 131488.0]
source_col: d(D_H)Mean
d18O_stdev | 1.82199268202
varray: [3.322, 2.958, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [130415.0, 130952.0, 131488.0]
source_col: d(18_16)Mean
dD_stdev | 3.29946500714
varray: [6.576, 3.764, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [130415.0, 130952.0, 131488.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.09333333333
varray: [3.322, 2.958, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [130415.0, 130952.0, 131488.0]
source_col: d(18_16)Mean
measured_dD_mean | 3.44666666667
varray: [6.576, 3.764, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [130415.0, 130952.0, 131488.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 1.82199268202
varray: [3.322, 2.958, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [130415.0, 130952.0, 131488.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.29946500714
varray: [6.576, 3.764, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [130415.0, 130952.0, 131488.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 0.769085187628
actual_array: [ 1.93 -15.84]
varray: [3.322, 2.958, 0.0]
fit_rsq: 1.0
fit_intercept: 1.38222915025
except: True
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 2.09333333333
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [130415.0, 130952.0, 131488.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -5.35745499863
actual_array: [ 18.1 -119. ]
varray: [6.576, 3.764, 0.0]
fit_rsq: 1.0
fit_intercept: 8.79377478947
except: True
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 3.44666666667
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [130415.0, 130952.0, 131488.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1293.67704762
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -435.450714286
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 54 - Sample 54 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 54
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.45966666667
varray: [2.428, 1.951, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [132676.0, 133213.0, 133749.0]
source_col: residuals
SAMP_residuals_STD | 1.28641063947
varray: [2.428, 1.951, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [132676.0, 133213.0, 133749.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.12966666667
varray: [-1.996, -1.393, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [132676.0, 133213.0, 133749.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 1.02372473514
varray: [-1.996, -1.393, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [132676.0, 133213.0, 133749.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 897.227333333
varray: [1366.384, 1325.298, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [132676.0, 133213.0, 133749.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 777.293175455
varray: [1366.384, 1325.298, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [132676.0, 133213.0, 133749.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -4.17033333333
varray: [-7.016, -5.495, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [132676.0, 133213.0, 133749.0]
source_col: slope_shift
SAMP_slope_shift_STD | 3.69081567317
varray: [-7.016, -5.495, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [132676.0, 133213.0, 133749.0]
source_col: slope_shift
SAMP_d18O_STD | 2.15028517489
varray: [3.907, 3.51, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [132676.0, 133213.0, 133749.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 4.27095227477
varray: [8.41, 5.5, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [132676.0, 133213.0, 133749.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 58.6673333333
varray: [97.032, 78.97, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [132676.0, 133213.0, 133749.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 51.603788246
varray: [97.032, 78.97, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [132676.0, 133213.0, 133749.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.45966666667
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.12966666667
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 897.227333333
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 4.17033333333
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.42171428571
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.42171428571
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -49.3414202846
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -52.0815714286
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [3.907, 3.51, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [132676.0, 133213.0, 133749.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [8.41, 5.5, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [132676.0, 133213.0, 133749.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.47233333333
varray: [3.907, 3.51, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [132676.0, 133213.0, 133749.0]
source_col: d(18_16)Mean
dD_mean | 4.63666666667
varray: [8.41, 5.5, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [132676.0, 133213.0, 133749.0]
source_col: d(D_H)Mean
d18O_stdev | 2.15028517489
varray: [3.907, 3.51, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [132676.0, 133213.0, 133749.0]
source_col: d(18_16)Mean
dD_stdev | 4.27095227477
varray: [8.41, 5.5, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [132676.0, 133213.0, 133749.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.47233333333
varray: [3.907, 3.51, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [132676.0, 133213.0, 133749.0]
source_col: d(18_16)Mean
measured_dD_mean | 4.63666666667
varray: [8.41, 5.5, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [132676.0, 133213.0, 133749.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 2.15028517489
varray: [3.907, 3.51, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [132676.0, 133213.0, 133749.0]
source_col: d(18_16)Mean
measured_dD_stdev | 4.27095227477
varray: [8.41, 5.5, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [132676.0, 133213.0, 133749.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.17898755221
actual_array: [ 1.93 -15.84]
varray: [3.907, 3.51, 0.0]
fit_rsq: 1.0
fit_intercept: 1.38222915025
except: True
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 2.47233333333
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [132676.0, 133213.0, 133749.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -4.16515229931
actual_array: [ 18.1 -119. ]
varray: [8.41, 5.5, 0.0]
fit_rsq: 1.0
fit_intercept: 8.79377478947
except: True
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 4.63666666667
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [132676.0, 133213.0, 133749.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1293.53371429
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -397.766047619
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 55 - Sample 55 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 55
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.733
varray: [0.0, 2.199, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [134937.0, 135538.0, 136074.0]
source_col: residuals
SAMP_residuals_STD | 1.26959324195
varray: [0.0, 2.199, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [134937.0, 135538.0, 136074.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0163333333333
varray: [0.0, 0.049, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [134937.0, 135538.0, 136074.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0282901631903
varray: [0.0, 0.049, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [134937.0, 135538.0, 136074.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 423.271
varray: [0.0, 1269.813, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [134937.0, 135538.0, 136074.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 733.12687737
varray: [0.0, 1269.813, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [134937.0, 135538.0, 136074.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.281
varray: [0.0, 0.843, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [134937.0, 135538.0, 136074.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.486706276927
varray: [0.0, 0.843, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [134937.0, 135538.0, 136074.0]
source_col: slope_shift
SAMP_d18O_STD | 1.07098474935
varray: [0.0, 1.855, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [134937.0, 135538.0, 136074.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 9.14118681208
varray: [0.0, 15.833, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [134937.0, 135538.0, 136074.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 29.3943333333
varray: [0.0, 88.183, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [134937.0, 135538.0, 136074.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 50.9124787879
varray: [0.0, 88.183, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [134937.0, 135538.0, 136074.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.733
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0163333333333
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 423.271
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.281
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.14838095238
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.14838095238
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -74.5608091358
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -81.3545714286
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [0.0, 1.855, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [134937.0, 135538.0, 136074.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [0.0, 15.833, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [134937.0, 135538.0, 136074.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.618333333333
varray: [0.0, 1.855, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [134937.0, 135538.0, 136074.0]
source_col: d(18_16)Mean
dD_mean | 5.27766666667
varray: [0.0, 15.833, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [134937.0, 135538.0, 136074.0]
source_col: d(D_H)Mean
d18O_stdev | 1.07098474935
varray: [0.0, 1.855, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [134937.0, 135538.0, 136074.0]
source_col: d(18_16)Mean
dD_stdev | 9.14118681208
varray: [0.0, 15.833, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [134937.0, 135538.0, 136074.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.618333333333
varray: [0.0, 1.855, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [134937.0, 135538.0, 136074.0]
source_col: d(18_16)Mean
measured_dD_mean | 5.27766666667
varray: [0.0, 15.833, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [134937.0, 135538.0, 136074.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 1.07098474935
varray: [0.0, 1.855, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [134937.0, 135538.0, 136074.0]
source_col: d(18_16)Mean
measured_dD_stdev | 9.14118681208
varray: [0.0, 15.833, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [134937.0, 135538.0, 136074.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.82618127085
actual_array: [ 1.93 -15.84]
varray: [0.0, 1.855, 0.0]
fit_rsq: 1.0
fit_intercept: 1.38222915025
except: True
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.618333333333
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [134937.0, 135538.0, 136074.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -3.52291193774
actual_array: [ 18.1 -119. ]
varray: [0.0, 15.833, 0.0]
fit_rsq: 1.0
fit_intercept: 8.79377478947
except: True
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 5.27766666667
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [134937.0, 135538.0, 136074.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1294.26038095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -871.722380952
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 56 - Sample 56 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 56
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 4.95466666667
varray: [5.18, 4.746, 4.938]
inst: mean_all
operation: mean_all(residuals)
tarray: [137263.0, 137799.0, 138336.0]
source_col: residuals
SAMP_residuals_STD | 0.21747950095
varray: [5.18, 4.746, 4.938]
inst: std_all
operation: std_all(residuals)
tarray: [137263.0, 137799.0, 138336.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.19433333333
varray: [-2.219, -2.18, -2.184]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [137263.0, 137799.0, 138336.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0214553800557
varray: [-2.219, -2.18, -2.184]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [137263.0, 137799.0, 138336.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1526.26333333
varray: [1541.346, 1512.745, 1524.699]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [137263.0, 137799.0, 138336.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 14.3645276405
varray: [1541.346, 1512.745, 1524.699]
inst: std_all
operation: std_all(baseline_shift)
tarray: [137263.0, 137799.0, 138336.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -14.277
varray: [-15.083, -13.144, -14.604]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [137263.0, 137799.0, 138336.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.01001336625
varray: [-15.083, -13.144, -14.604]
inst: std_all
operation: std_all(slope_shift)
tarray: [137263.0, 137799.0, 138336.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0650410127022
varray: [4.927, 4.822, 4.941]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [137263.0, 137799.0, 138336.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.725398511165
varray: [20.355, 18.915, 19.788]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [137263.0, 137799.0, 138336.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 184.458
varray: [191.81, 176.689, 184.875]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [137263.0, 137799.0, 138336.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 7.56911996206
varray: [191.81, 176.689, 184.875]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [137263.0, 137799.0, 138336.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 4.95466666667
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.19433333333
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1526.26333333
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 14.277
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 2.07328571429
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.07328571429
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 71.9545852683
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 73.7090952381
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [4.927, 4.822, 4.941]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [137263.0, 137799.0, 138336.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [20.355, 18.915, 19.788]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [137263.0, 137799.0, 138336.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 4.89666666667
varray: [4.927, 4.822, 4.941]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [137263.0, 137799.0, 138336.0]
source_col: d(18_16)Mean
dD_mean | 19.686
varray: [20.355, 18.915, 19.788]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [137263.0, 137799.0, 138336.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0650410127022
varray: [4.927, 4.822, 4.941]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [137263.0, 137799.0, 138336.0]
source_col: d(18_16)Mean
dD_stdev | 0.725398511165
varray: [20.355, 18.915, 19.788]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [137263.0, 137799.0, 138336.0]
source_col: d(D_H)Mean
measured_d18O_mean | 4.89666666667
varray: [4.927, 4.822, 4.941]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [137263.0, 137799.0, 138336.0]
source_col: d(18_16)Mean
measured_dD_mean | 19.686
varray: [20.355, 18.915, 19.788]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [137263.0, 137799.0, 138336.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0650410127022
varray: [4.927, 4.822, 4.941]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [137263.0, 137799.0, 138336.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.725398511165
varray: [20.355, 18.915, 19.788]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [137263.0, 137799.0, 138336.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 3.80099273922
actual_array: [ 1.93 -15.84]
varray: [4.927, 4.822, 4.941]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 4.89666666667
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [137263.0, 137799.0, 138336.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 10.9133021177
actual_array: [ 18.1 -119. ]
varray: [20.355, 18.915, 19.788]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 19.686
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [137263.0, 137799.0, 138336.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1290.03871429
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 231.269952381
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 57 - Sample 57 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 57
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.92366666667
varray: [1.307, 2.919, 1.545]
inst: mean_all
operation: mean_all(residuals)
tarray: [139459.0, 139995.0, 140532.0]
source_col: residuals
SAMP_residuals_STD | 0.870159372376
varray: [1.307, 2.919, 1.545]
inst: std_all
operation: std_all(residuals)
tarray: [139459.0, 139995.0, 140532.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.254
varray: [-1.198, -1.581, -0.983]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [139459.0, 139995.0, 140532.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.302907576663
varray: [-1.198, -1.581, -0.983]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [139459.0, 139995.0, 140532.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1289.84466667
varray: [1257.073, 1345.472, 1266.989]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [139459.0, 139995.0, 140532.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 48.4291433368
varray: [1257.073, 1345.472, 1266.989]
inst: std_all
operation: std_all(baseline_shift)
tarray: [139459.0, 139995.0, 140532.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -4.07666666667
varray: [-3.186, -6.057, -2.987]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [139459.0, 139995.0, 140532.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.71790288821
varray: [-3.186, -6.057, -2.987]
inst: std_all
operation: std_all(slope_shift)
tarray: [139459.0, 139995.0, 140532.0]
source_col: slope_shift
SAMP_d18O_STD | 0.495387054063
varray: [3.328, 4.203, 3.363]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [139459.0, 139995.0, 140532.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.70675388087
varray: [4.845, 11.478, 5.294]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [139459.0, 139995.0, 140532.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 77.5746666667
varray: [54.627, 114.485, 63.612]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [139459.0, 139995.0, 140532.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 32.2794375158
varray: [54.627, 114.485, 63.612]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [139459.0, 139995.0, 140532.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.92366666667
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.254
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1289.84466667
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 4.07666666667
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.957714285714
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.957714285714
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -33.2380307062
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -33.1742380952
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [3.328, 4.203, 3.363]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [139459.0, 139995.0, 140532.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [4.845, 11.478, 5.294]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [139459.0, 139995.0, 140532.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 3.63133333333
varray: [3.328, 4.203, 3.363]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [139459.0, 139995.0, 140532.0]
source_col: d(18_16)Mean
dD_mean | 7.20566666667
varray: [4.845, 11.478, 5.294]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [139459.0, 139995.0, 140532.0]
source_col: d(D_H)Mean
d18O_stdev | 0.495387054063
varray: [3.328, 4.203, 3.363]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [139459.0, 139995.0, 140532.0]
source_col: d(18_16)Mean
dD_stdev | 3.70675388087
varray: [4.845, 11.478, 5.294]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [139459.0, 139995.0, 140532.0]
source_col: d(D_H)Mean
measured_d18O_mean | 3.63133333333
varray: [3.328, 4.203, 3.363]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [139459.0, 139995.0, 140532.0]
source_col: d(18_16)Mean
measured_dD_mean | 7.20566666667
varray: [4.845, 11.478, 5.294]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [139459.0, 139995.0, 140532.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.495387054063
varray: [3.328, 4.203, 3.363]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [139459.0, 139995.0, 140532.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.70675388087
varray: [4.845, 11.478, 5.294]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [139459.0, 139995.0, 140532.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 2.43248845077
actual_array: [ 1.93 -15.84]
varray: [3.328, 4.203, 3.363]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 3.63133333333
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [139459.0, 139995.0, 140532.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -1.59118117784
actual_array: [ 18.1 -119. ]
varray: [4.845, 11.478, 5.294]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 7.20566666667
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [139459.0, 139995.0, 140532.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1293.06971429
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -5.14871428571
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 58 - Sample 58 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 58
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [141119.0, 141720.0, 142321.0, 142922.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [141119.0, 141720.0, 142321.0, 142922.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [141119.0, 141720.0, 142321.0, 142922.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [141119.0, 141720.0, 142321.0, 142922.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [141119.0, 141720.0, 142321.0, 142922.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [141119.0, 141720.0, 142321.0, 142922.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [141119.0, 141720.0, 142321.0, 142922.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [141119.0, 141720.0, 142321.0, 142922.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [141119.0, 141720.0, 142321.0, 142922.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [141119.0, 141720.0, 142321.0, 142922.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [141119.0, 141720.0, 142321.0, 142922.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [141119.0, 141720.0, 142321.0, 142922.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.88138095238
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.88138095238
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -110.748904762
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [141119.0, 141720.0, 142321.0, 142922.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [141119.0, 141720.0, 142321.0, 142922.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [141119.0, 141720.0, 142321.0, 142922.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [141119.0, 141720.0, 142321.0, 142922.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [141119.0, 141720.0, 142321.0, 142922.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [141119.0, 141720.0, 142321.0, 142922.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [141119.0, 141720.0, 142321.0, 142922.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [141119.0, 141720.0, 142321.0, 142922.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [141119.0, 141720.0, 142321.0, 142922.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [141119.0, 141720.0, 142321.0, 142922.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 1.38222915025
except: True
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [141119.0, 141720.0, 142321.0, 142922.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 8.79377478947
except: True
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [141119.0, 141720.0, 142321.0, 142922.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1294.99338095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1294.99338095
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 59 - Sample 59 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 59
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.49366666667
varray: [2.594, 0.0, 1.887]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [144111.0, 144648.0, 145249.0]
source_col: residuals
SAMP_residuals_STD | 1.34098558282
varray: [2.594, 0.0, 1.887]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [144111.0, 144648.0, 145249.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.467333333333
varray: [-0.609, 0.0, -0.793]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [144111.0, 144648.0, 145249.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.415047386853
varray: [-0.609, 0.0, -0.793]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [144111.0, 144648.0, 145249.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 859.036
varray: [1304.828, 0.0, 1272.28]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [144111.0, 144648.0, 145249.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 744.124976095
varray: [1304.828, 0.0, 1272.28]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [144111.0, 144648.0, 145249.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -1.223
varray: [-2.403, 0.0, -1.266]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [144111.0, 144648.0, 145249.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.20207695261
varray: [-2.403, 0.0, -1.266]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [144111.0, 144648.0, 145249.0]
source_col: slope_shift
SAMP_d18O_STD | 0.862746969472
varray: [1.636, 0.0, 1.293]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [144111.0, 144648.0, 145249.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 4.81887611378
varray: [-6.559, 0.0, -9.395]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [144111.0, 144648.0, 145249.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 60.4716666667
varray: [103.867, 0.0, 77.548]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [144111.0, 144648.0, 145249.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 53.9980489678
varray: [103.867, 0.0, 77.548]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [144111.0, 144648.0, 145249.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.49366666667
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.467333333333
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 859.036
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.223
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.38771428571
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.38771428571
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -48.161430531
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -50.2772380952
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [1.636, 0.0, 1.293]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [144111.0, 144648.0, 145249.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-6.559, 0.0, -9.395]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [144111.0, 144648.0, 145249.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.976333333333
varray: [1.636, 0.0, 1.293]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [144111.0, 144648.0, 145249.0]
source_col: d(18_16)Mean
dD_mean | -5.318
varray: [-6.559, 0.0, -9.395]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [144111.0, 144648.0, 145249.0]
source_col: d(D_H)Mean
d18O_stdev | 0.862746969472
varray: [1.636, 0.0, 1.293]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [144111.0, 144648.0, 145249.0]
source_col: d(18_16)Mean
dD_stdev | 4.81887611378
varray: [-6.559, 0.0, -9.395]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [144111.0, 144648.0, 145249.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.976333333333
varray: [1.636, 0.0, 1.293]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [144111.0, 144648.0, 145249.0]
source_col: d(18_16)Mean
measured_dD_mean | -5.318
varray: [-6.559, 0.0, -9.395]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [144111.0, 144648.0, 145249.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.862746969472
varray: [1.636, 0.0, 1.293]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [144111.0, 144648.0, 145249.0]
source_col: d(18_16)Mean
measured_dD_stdev | 4.81887611378
varray: [-6.559, 0.0, -9.395]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [144111.0, 144648.0, 145249.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.438991174487
actual_array: [ 1.93 -15.84]
varray: [1.636, 0.0, 1.293]
fit_rsq: 1.0
fit_intercept: 1.38222915025
except: True
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.976333333333
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [144111.0, 144648.0, 145249.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -14.1390816585
actual_array: [ 18.1 -119. ]
varray: [-6.559, 0.0, -9.395]
fit_rsq: 1.0
fit_intercept: 8.79377478947
except: True
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -5.318
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [144111.0, 144648.0, 145249.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1293.49971429
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -435.957380952
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 60 - Sample 60 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 60
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.69333333333
varray: [2.724, 2.356, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [146372.0, 146909.0, 147445.0]
source_col: residuals
SAMP_residuals_STD | 1.47796797439
varray: [2.724, 2.356, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [146372.0, 146909.0, 147445.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.558333333333
varray: [-0.911, -0.764, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [146372.0, 146909.0, 147445.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.489085200485
varray: [-0.911, -0.764, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [146372.0, 146909.0, 147445.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 870.438666667
varray: [1315.968, 1295.348, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [146372.0, 146909.0, 147445.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 753.892499247
varray: [1315.968, 1295.348, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [146372.0, 146909.0, 147445.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -1.97066666667
varray: [-3.362, -2.55, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [146372.0, 146909.0, 147445.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.75427515896
varray: [-3.362, -2.55, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [146372.0, 146909.0, 147445.0]
source_col: slope_shift
SAMP_d18O_STD | 1.1680295373
varray: [2.122, 1.907, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [146372.0, 146909.0, 147445.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.85106523952
varray: [-3.758, -5.593, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [146372.0, 146909.0, 147445.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 68.063
varray: [109.682, 94.507, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [146372.0, 146909.0, 147445.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 59.4306245382
varray: [109.682, 94.507, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [146372.0, 146909.0, 147445.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.69333333333
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.558333333333
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 870.438666667
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.97066666667
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.18804761905
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.18804761905
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -41.2318828604
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -42.6859047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [2.122, 1.907, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [146372.0, 146909.0, 147445.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-3.758, -5.593, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [146372.0, 146909.0, 147445.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.343
varray: [2.122, 1.907, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [146372.0, 146909.0, 147445.0]
source_col: d(18_16)Mean
dD_mean | -3.117
varray: [-3.758, -5.593, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [146372.0, 146909.0, 147445.0]
source_col: d(D_H)Mean
d18O_stdev | 1.1680295373
varray: [2.122, 1.907, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [146372.0, 146909.0, 147445.0]
source_col: d(18_16)Mean
dD_stdev | 2.85106523952
varray: [-3.758, -5.593, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [146372.0, 146909.0, 147445.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.343
varray: [2.122, 1.907, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [146372.0, 146909.0, 147445.0]
source_col: d(18_16)Mean
measured_dD_mean | -3.117
varray: [-3.758, -5.593, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [146372.0, 146909.0, 147445.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 1.1680295373
varray: [2.122, 1.907, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [146372.0, 146909.0, 147445.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.85106523952
varray: [-3.758, -5.593, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [146372.0, 146909.0, 147445.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.0424277610804
actual_array: [ 1.93 -15.84]
varray: [2.122, 1.907, 0.0]
fit_rsq: 1.0
fit_intercept: 1.38222915025
except: True
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 1.343
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [146372.0, 146909.0, 147445.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -11.9338226323
actual_array: [ 18.1 -119. ]
varray: [-3.758, -5.593, 0.0]
fit_rsq: 1.0
fit_intercept: 8.79377478947
except: True
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -3.117
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [146372.0, 146909.0, 147445.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1293.30004762
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -424.554714286
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 61 - Sample 61 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 61
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.708333333333
varray: [0.0, 0.0, 2.125]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [148635.0, 149235.0, 149836.0]
source_col: residuals
SAMP_residuals_STD | 1.22686932203
varray: [0.0, 0.0, 2.125]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [148635.0, 149235.0, 149836.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.165
varray: [0.0, 0.0, -0.495]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [148635.0, 149235.0, 149836.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.285788383249
varray: [0.0, 0.0, -0.495]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [148635.0, 149235.0, 149836.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 428.122
varray: [0.0, 0.0, 1284.366]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [148635.0, 149235.0, 149836.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 741.529055838
varray: [0.0, 0.0, 1284.366]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [148635.0, 149235.0, 149836.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -0.526333333333
varray: [0.0, 0.0, -1.579]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [148635.0, 149235.0, 149836.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.91163607505
varray: [0.0, 0.0, -1.579]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [148635.0, 149235.0, 149836.0]
source_col: slope_shift
SAMP_d18O_STD | 1.35792783313
varray: [0.0, 0.0, 2.352]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [148635.0, 149235.0, 149836.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.764411756407
varray: [0.0, 0.0, 1.324]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [148635.0, 149235.0, 149836.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 28.763
varray: [0.0, 0.0, 86.289]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [148635.0, 149235.0, 149836.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 49.8189773781
varray: [0.0, 0.0, 86.289]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [148635.0, 149235.0, 149836.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.708333333333
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.165
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 428.122
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.526333333333
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.17304761905
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.17304761905
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -75.4168801335
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -81.9859047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [0.0, 0.0, 2.352]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [148635.0, 149235.0, 149836.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [0.0, 0.0, 1.324]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [148635.0, 149235.0, 149836.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.784
varray: [0.0, 0.0, 2.352]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [148635.0, 149235.0, 149836.0]
source_col: d(18_16)Mean
dD_mean | 0.441333333333
varray: [0.0, 0.0, 1.324]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [148635.0, 149235.0, 149836.0]
source_col: d(D_H)Mean
d18O_stdev | 1.35792783313
varray: [0.0, 0.0, 2.352]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [148635.0, 149235.0, 149836.0]
source_col: d(18_16)Mean
dD_stdev | 0.764411756407
varray: [0.0, 0.0, 1.324]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [148635.0, 149235.0, 149836.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.784
varray: [0.0, 0.0, 2.352]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [148635.0, 149235.0, 149836.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.441333333333
varray: [0.0, 0.0, 1.324]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [148635.0, 149235.0, 149836.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 1.35792783313
varray: [0.0, 0.0, 2.352]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [148635.0, 149235.0, 149836.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.764411756407
varray: [0.0, 0.0, 1.324]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [148635.0, 149235.0, 149836.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.647006710429
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 2.352]
fit_rsq: 1.0
fit_intercept: 1.38222915025
except: True
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.784
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [148635.0, 149235.0, 149836.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -8.36860377651
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 1.324]
fit_rsq: 1.0
fit_intercept: 8.79377478947
except: True
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.441333333333
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [148635.0, 149235.0, 149836.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1294.28504762
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -866.871380952
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 62 - Sample 62 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 62
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.99266666667
varray: [2.55, 1.989, 1.439]
inst: mean_all
operation: mean_all(residuals)
tarray: [150961.0, 151497.0, 152034.0]
source_col: residuals
SAMP_residuals_STD | 0.555509075833
varray: [2.55, 1.989, 1.439]
inst: std_all
operation: std_all(residuals)
tarray: [150961.0, 151497.0, 152034.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.671
varray: [-0.939, -0.478, -0.596]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [150961.0, 151497.0, 152034.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.239476512418
varray: [-0.939, -0.478, -0.596]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [150961.0, 151497.0, 152034.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1284.34733333
varray: [1314.242, 1284.316, 1254.484]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [150961.0, 151497.0, 152034.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 29.8790123219
varray: [1314.242, 1284.316, 1254.484]
inst: std_all
operation: std_all(baseline_shift)
tarray: [150961.0, 151497.0, 152034.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -2.39666666667
varray: [-3.665, -1.88, -1.645]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [150961.0, 151497.0, 152034.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.10467566884
varray: [-3.665, -1.88, -1.645]
inst: std_all
operation: std_all(slope_shift)
tarray: [150961.0, 151497.0, 152034.0]
source_col: slope_shift
SAMP_d18O_STD | 0.302754576073
varray: [2.47, 2.205, 1.866]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [150961.0, 151497.0, 152034.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.59458558027
varray: [4.339, 1.781, -0.85]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [150961.0, 151497.0, 152034.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 81.4193333333
varray: [102.23, 81.422, 60.606]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [150961.0, 151497.0, 152034.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 20.8120001281
varray: [102.23, 81.422, 60.606]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [150961.0, 151497.0, 152034.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.99266666667
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.671
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1284.34733333
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.39666666667
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.888714285714
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.888714285714
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -30.843345618
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -29.3295714286
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [2.47, 2.205, 1.866]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [150961.0, 151497.0, 152034.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [4.339, 1.781, -0.85]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [150961.0, 151497.0, 152034.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.18033333333
varray: [2.47, 2.205, 1.866]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [150961.0, 151497.0, 152034.0]
source_col: d(18_16)Mean
dD_mean | 1.75666666667
varray: [4.339, 1.781, -0.85]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [150961.0, 151497.0, 152034.0]
source_col: d(D_H)Mean
d18O_stdev | 0.302754576073
varray: [2.47, 2.205, 1.866]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [150961.0, 151497.0, 152034.0]
source_col: d(18_16)Mean
dD_stdev | 2.59458558027
varray: [4.339, 1.781, -0.85]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [150961.0, 151497.0, 152034.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.18033333333
varray: [2.47, 2.205, 1.866]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [150961.0, 151497.0, 152034.0]
source_col: d(18_16)Mean
measured_dD_mean | 1.75666666667
varray: [4.339, 1.781, -0.85]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [150961.0, 151497.0, 152034.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.302754576073
varray: [2.47, 2.205, 1.866]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [150961.0, 151497.0, 152034.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.59458558027
varray: [4.339, 1.781, -0.85]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [150961.0, 151497.0, 152034.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 0.863178870264
actual_array: [ 1.93 -15.84]
varray: [2.47, 2.205, 1.866]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 2.18033333333
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [150961.0, 151497.0, 152034.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -7.05072521867
actual_array: [ 18.1 -119. ]
varray: [4.339, 1.781, -0.85]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 1.75666666667
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [150961.0, 151497.0, 152034.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1293.00071429
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.646047619
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 63 - Sample 63 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 63
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.045
varray: [2.561, 1.991, 1.583]
inst: mean_all
operation: mean_all(residuals)
tarray: [153159.0, 153695.0, 154231.0]
source_col: residuals
SAMP_residuals_STD | 0.491231106507
varray: [2.561, 1.991, 1.583]
inst: std_all
operation: std_all(residuals)
tarray: [153159.0, 153695.0, 154231.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.475666666667
varray: [-0.477, -0.529, -0.421]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [153159.0, 153695.0, 154231.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0540123442681
varray: [-0.477, -0.529, -0.421]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [153159.0, 153695.0, 154231.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1287.426
varray: [1315.157, 1285.086, 1262.035]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [153159.0, 153695.0, 154231.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 26.6381947774
varray: [1315.157, 1285.086, 1262.035]
inst: std_all
operation: std_all(baseline_shift)
tarray: [153159.0, 153695.0, 154231.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -1.42033333333
varray: [-2.006, -1.418, -0.837]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [153159.0, 153695.0, 154231.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.584503493004
varray: [-2.006, -1.418, -0.837]
inst: std_all
operation: std_all(slope_shift)
tarray: [153159.0, 153695.0, 154231.0]
source_col: slope_shift
SAMP_d18O_STD | 0.255370971987
varray: [1.6, 1.309, 1.091]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [153159.0, 153695.0, 154231.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.32408519637
varray: [2.792, 0.346, -1.854]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [153159.0, 153695.0, 154231.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 83.384
varray: [102.072, 82.226, 65.854]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [153159.0, 153695.0, 154231.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 18.1367473379
varray: [102.072, 82.226, 65.854]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [153159.0, 153695.0, 154231.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.045
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.475666666667
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1287.426
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.42033333333
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.836380952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.836380952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -29.0270868796
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -27.3649047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [1.6, 1.309, 1.091]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [153159.0, 153695.0, 154231.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [2.792, 0.346, -1.854]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [153159.0, 153695.0, 154231.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.33333333333
varray: [1.6, 1.309, 1.091]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [153159.0, 153695.0, 154231.0]
source_col: d(18_16)Mean
dD_mean | 0.428
varray: [2.792, 0.346, -1.854]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [153159.0, 153695.0, 154231.0]
source_col: d(D_H)Mean
d18O_stdev | 0.255370971987
varray: [1.6, 1.309, 1.091]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [153159.0, 153695.0, 154231.0]
source_col: d(18_16)Mean
dD_stdev | 2.32408519637
varray: [2.792, 0.346, -1.854]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [153159.0, 153695.0, 154231.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.33333333333
varray: [1.6, 1.309, 1.091]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [153159.0, 153695.0, 154231.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.428
varray: [2.792, 0.346, -1.854]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [153159.0, 153695.0, 154231.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.255370971987
varray: [1.6, 1.309, 1.091]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [153159.0, 153695.0, 154231.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.32408519637
varray: [2.792, 0.346, -1.854]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [153159.0, 153695.0, 154231.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.0528826147066
actual_array: [ 1.93 -15.84]
varray: [1.6, 1.309, 1.091]
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 1.33333333333
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [153159.0, 153695.0, 154231.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -8.3819629104
actual_array: [ 18.1 -119. ]
varray: [2.792, 0.346, -1.854]
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.428
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [153159.0, 153695.0, 154231.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1292.94838095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -7.56738095238
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 64 - Sample 64 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Sample Variables
Sample 64
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.6155
varray: [1.305, 0.0, 1.157, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [154818.0, 155355.0, 155956.0, 156493.0]
source_col: residuals
SAMP_residuals_STD | 0.713281851725
varray: [1.305, 0.0, 1.157, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [154818.0, 155355.0, 155956.0, 156493.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.034
varray: [0.096, 0.0, 0.04, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [154818.0, 155355.0, 155956.0, 156493.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0454312667664
varray: [0.096, 0.0, 0.04, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [154818.0, 155355.0, 155956.0, 156493.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 614.456
varray: [1231.873, 0.0, 1225.951, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [154818.0, 155355.0, 155956.0, 156493.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 709.516793033
varray: [1231.873, 0.0, 1225.951, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [154818.0, 155355.0, 155956.0, 156493.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.426
varray: [0.782, 0.0, 0.922, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [154818.0, 155355.0, 155956.0, 156493.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.495211739225
varray: [0.782, 0.0, 0.922, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [154818.0, 155355.0, 155956.0, 156493.0]
source_col: slope_shift
SAMP_d18O_STD | 4.23021894626
varray: [-6.829, 0.0, -7.765, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [154818.0, 155355.0, 155956.0, 156493.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 31.2601371878
varray: [-49.326, 0.0, -58.228, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [154818.0, 155355.0, 155956.0, 156493.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 25.7105
varray: [54.175, 0.0, 48.667, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [154818.0, 155355.0, 155956.0, 156493.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 29.7729646548
varray: [54.175, 0.0, 48.667, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [154818.0, 155355.0, 155956.0, 156493.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.6155
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.034
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 614.456
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.426
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.26588095238
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.26588095238
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -78.6387149019
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -85.0384047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.2636
varray: [-6.829, 0.0, -7.765, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [154818.0, 155355.0, 155956.0, 156493.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -109.9764
varray: [-49.326, 0.0, -58.228, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [154818.0, 155355.0, 155956.0, 156493.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.6485
varray: [-6.829, 0.0, -7.765, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [154818.0, 155355.0, 155956.0, 156493.0]
source_col: d(18_16)Mean
dD_mean | -26.8885
varray: [-49.326, 0.0, -58.228, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [154818.0, 155355.0, 155956.0, 156493.0]
source_col: d(D_H)Mean
d18O_stdev | 4.23021894626
varray: [-6.829, 0.0, -7.765, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [154818.0, 155355.0, 155956.0, 156493.0]
source_col: d(18_16)Mean
dD_stdev | 31.2601371878
varray: [-49.326, 0.0, -58.228, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [154818.0, 155355.0, 155956.0, 156493.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.6485
varray: [-6.829, 0.0, -7.765, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [154818.0, 155355.0, 155956.0, 156493.0]
source_col: d(18_16)Mean
measured_dD_mean | -26.8885
varray: [-49.326, 0.0, -58.228, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [154818.0, 155355.0, 155956.0, 156493.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 4.23021894626
varray: [-6.829, 0.0, -7.765, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [154818.0, 155355.0, 155956.0, 156493.0]
source_col: d(18_16)Mean
measured_dD_stdev | 31.2601371878
varray: [-49.326, 0.0, -58.228, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [154818.0, 155355.0, 155956.0, 156493.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.44091761023
actual_array: [ 1.93 -15.84]
varray: [-6.829, 0.0, -7.765, 0.0]
fit_rsq: 1.0
fit_intercept: 1.38222915025
except: True
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -3.6485
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: [154818.0, 155355.0, 155956.0, 156493.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -35.75132147
actual_array: [ 18.1 -119. ]
varray: [-49.326, 0.0, -58.228, 0.0]
fit_rsq: 1.0
fit_intercept: 8.79377478947
except: True
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -26.8885
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: [154818.0, 155355.0, 155956.0, 156493.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1294.37788095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -680.537380952
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 1 - Sample 1 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 1 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.107
operation: col(d(18_16)Mean)
inst: col
dD_actual | -54.516
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.09987529712
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -6.107
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -63.4322818275
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -54.516
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 1 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.837
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.248
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.889397002
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -6.837
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.1695683607
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -57.248
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 1 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.245
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.741
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.33066392746
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.245
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.6654573776
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -58.741
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 1 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.129
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.353
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.20520568395
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.129
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.2767065815
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -58.353
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 1 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.203
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.641
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.28523939102
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.203
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.5652638735
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -58.641
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 1 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.273
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.81
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.36094695176
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.273
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.7345908955
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -58.81
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 1 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.395
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.0
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.49289441476
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.395
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.9249585533
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -59.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 1 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.349
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.974
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.44314373199
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.349
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.8989082423
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -58.974
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 1 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.42
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.346
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.51993282931
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.42
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -68.2716280777
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -59.346
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 1 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.411
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.315
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.51019900008
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.411
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -68.2405680914
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -59.315
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 2 - Sample 2 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 2 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -12.113
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -92.388
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 3.16672727273
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 26.8588181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.01
operation: col(d(18_16)Mean)
inst: col
dD_actual | 28.804
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.84203030316
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 4.01
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.0489458256
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 28.804
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 2 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -12.113
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -92.388
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 3.16672727273
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 26.8588181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.535
operation: col(d(18_16)Mean)
inst: col
dD_actual | 36.832
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.40983700872
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 4.535
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 28.0924803383
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 36.832
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 2 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -12.113
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -92.388
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 3.16672727273
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 26.8588181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.691
operation: col(d(18_16)Mean)
inst: col
dD_actual | 38.754
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.57855671552
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 4.691
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 30.018199488
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 38.754
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 2 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -12.113
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -92.388
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 3.16672727273
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 26.8588181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.763
operation: col(d(18_16)Mean)
inst: col
dD_actual | 39.807
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.65642734942
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 4.763
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 31.0732370866
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 39.807
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 2 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -12.113
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -92.388
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 3.16672727273
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 26.8588181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.819
operation: col(d(18_16)Mean)
inst: col
dD_actual | 39.876
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.71699339802
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 4.819
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 31.1423706045
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 39.876
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 2 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -12.113
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -92.388
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 3.16672727273
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 26.8588181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.777
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.238
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.67156886157
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 4.777
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 31.5050710895
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 40.238
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 2 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -12.113
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -92.388
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 3.16672727273
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 26.8588181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.806
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.639
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.70293342245
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 4.806
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 31.9068470411
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 40.639
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 2 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -12.113
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -92.388
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 3.16672727273
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 26.8588181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.84
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.948
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.73970566624
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 4.84
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 32.2164449689
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 40.948
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 2 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -12.113
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -92.388
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 3.16672727273
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 26.8588181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.859
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.953
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.7602548613
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 4.859
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 32.2214546441
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 40.953
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 2 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -12.113
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -92.388
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 3.16672727273
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 26.8588181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.847
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.984
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.74727642231
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 4.847
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 32.2525146304
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 40.984
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 2 - 11
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -12.113
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -92.388
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 3.16672727273
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 26.8588181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.113
operation: col(d(18_16)Mean)
inst: col
dD_actual | -92.388
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.5955840087
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -12.113
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -101.377565717
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -92.388
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 3 - Sample 3 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 3 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.906
operation: col(d(18_16)Mean)
inst: col
dD_actual | -105.184
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.4532425183
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -12.906
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -114.198326507
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -105.184
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 3 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.147
operation: col(d(18_16)Mean)
inst: col
dD_actual | -107.874
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.7138928345
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -13.147
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -116.893531769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -107.874
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 3 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.225
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.468
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.7982526879
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -13.225
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -118.490616225
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -109.468
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 3 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.24
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.097
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.8144757367
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -13.24
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -119.120833366
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -110.097
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 3 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.302
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.614
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.8815310047
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -13.302
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -119.638833782
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -110.614
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 3 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.359
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.064
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9431785899
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -13.359
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.089704551
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -111.064
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 3 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.35
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.175
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9334447607
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -13.35
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.200919341
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -111.175
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 3 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.305
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.049
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.8847756145
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -13.305
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.074675525
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -111.049
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 3 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.351
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.485
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9345262973
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -13.351
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.511519204
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -111.485
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 3 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.451
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.754
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0426799555
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -13.451
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.78103973
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -111.754
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 4 - Sample 4 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 4 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.017
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.554
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1656101688
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -8.017
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.4935758564
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -66.554
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 4 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.77
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.362
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.89847063302
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.77
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.2934641627
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -62.362
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 4 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.704
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.345
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.82708921861
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.704
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.2744962255
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -61.345
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 4 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.772
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.072
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.90063370618
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.772
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.0009679592
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -61.072
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 4 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.641
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.439
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.75895241394
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.641
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.366743078
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -60.439
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 4 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.676
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.479
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.79680619431
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.676
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.4068204796
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -60.479
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 4 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.661
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.432
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.78058314558
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.661
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.3597295327
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -60.432
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 4 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.73
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.284
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.85520916974
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.73
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.2114431465
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -60.284
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 4 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.625
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.196
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.74164782863
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.625
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.1232728629
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -60.196
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 4 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.672
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.306
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.79248004798
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.672
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.2334857175
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -60.306
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 5 - Sample 5 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 5 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.791
operation: col(d(18_16)Mean)
inst: col
dD_actual | -82.99
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.83964631922
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -6.791
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -91.9613801976
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -82.99
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 5 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.874
operation: col(d(18_16)Mean)
inst: col
dD_actual | -85.569
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.92941385553
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -6.874
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -94.5453706695
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -85.569
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 5 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.827
operation: col(d(18_16)Mean)
inst: col
dD_actual | -86.014
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.87858163618
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -6.827
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -94.9912317629
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -86.014
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 5 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.816
operation: col(d(18_16)Mean)
inst: col
dD_actual | -86.273
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.86668473377
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -6.816
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -95.2507329386
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -86.273
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 6 - Sample 6 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 6 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.587
operation: col(d(18_16)Mean)
inst: col
dD_actual | -87.678
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.70054943851
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.587
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -96.6584516719
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -87.678
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 6 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.618
operation: col(d(18_16)Mean)
inst: col
dD_actual | -87.937
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.73407707255
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.618
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -96.9179528476
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -87.937
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 6 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.582
operation: col(d(18_16)Mean)
inst: col
dD_actual | -87.918
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.6951417556
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.582
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -96.8989160818
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -87.918
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 6 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.568
operation: col(d(18_16)Mean)
inst: col
dD_actual | -87.922
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.68000024345
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.568
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -96.902923822
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -87.922
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 7 - Sample 7 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 7 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.831
operation: col(d(18_16)Mean)
inst: col
dD_actual | -84.193
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.39368829039
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -0.831
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -93.1667080524
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -84.193
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 7 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.885
operation: col(d(18_16)Mean)
inst: col
dD_actual | -85.44
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.45209126582
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -0.885
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -94.4161210491
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -85.44
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 7 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.809
operation: col(d(18_16)Mean)
inst: col
dD_actual | -85.029
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.36989448559
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -0.809
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -94.0043257471
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -85.029
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 7 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.803
operation: col(d(18_16)Mean)
inst: col
dD_actual | -85.421
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.36340526609
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -0.803
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -94.3970842833
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -85.421
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 8 - Sample 8 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 8 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.058
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.454
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.43220226897
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.058
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.3585516063
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -48.454
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 8 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.34
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.661
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.12720895285
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.34
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.5562769526
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -43.661
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 8 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.344
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.446
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.12288280652
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.344
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.3389258773
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -42.446
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 8 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.413
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.307
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.04825678236
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.413
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.1996569065
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -42.307
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 9 - Sample 9 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 9 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.605
operation: col(d(18_16)Mean)
inst: col
dD_actual | -75.346
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.240934823409
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 1.605
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -84.3025887408
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -75.346
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 9 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.324
operation: col(d(18_16)Mean)
inst: col
dD_actual | -79.685
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.0629769561388
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 1.324
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -88.6499848856
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -79.685
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 9 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.195
operation: col(d(18_16)Mean)
inst: col
dD_actual | -80.885
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.202495175219
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 1.195
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -89.8523069353
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -80.885
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 9 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.149
operation: col(d(18_16)Mean)
inst: col
dD_actual | -81.261
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.252245857992
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 1.149
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.2290345109
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -81.261
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 10 - Sample 10 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 10 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.023
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.465
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.92748964221
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -5.023
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.4179489278
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -73.465
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 10 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.894
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.082
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.78797142313
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -4.894
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.0303377241
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -71.082
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 10 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.849
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.574
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.73930227694
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -4.849
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.521354723
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -70.574
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 10 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.884
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.58
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.77715605731
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -4.884
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.5273663333
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -70.58
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 11 - Sample 11 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 11 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.951
operation: col(d(18_16)Mean)
inst: col
dD_actual | -84.755
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0942287544
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.951
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -93.7297955457
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -84.755
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 11 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.137
operation: col(d(18_16)Mean)
inst: col
dD_actual | -86.056
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2953945586
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -8.137
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -95.0333130346
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -86.056
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 11 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.194
operation: col(d(18_16)Mean)
inst: col
dD_actual | -86.542
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3570421438
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -8.194
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -95.5202534648
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -86.542
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 11 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.151
operation: col(d(18_16)Mean)
inst: col
dD_actual | -86.406
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3105360708
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -8.151
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -95.3839902991
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -86.406
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 12 - Sample 12 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 12 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.016
operation: col(d(18_16)Mean)
inst: col
dD_actual | -90.222
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.51223597604
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -0.016
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -99.2073744173
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -90.222
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 12 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.13
operation: col(d(18_16)Mean)
inst: col
dD_actual | -92.735
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.63553114639
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -0.13
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -101.725237176
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -92.735
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 12 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.119
operation: col(d(18_16)Mean)
inst: col
dD_actual | -93.263
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.62363424399
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -0.119
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -102.254258878
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -93.263
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 12 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.002
operation: col(d(18_16)Mean)
inst: col
dD_actual | -92.877
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.4970944639
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -0.002
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -101.867511952
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -92.877
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 13 - Sample 13 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 13 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.754
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.852
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.88116604771
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.754
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.7921524988
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -66.852
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 13 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.667
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.719
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.78707236507
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.667
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.6511549725
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -62.719
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 13 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.63
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.621
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.74705551154
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.63
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.551030297
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -61.621
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 13 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.695
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.319
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.81735538937
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.695
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.2484459144
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -61.319
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 14 - Sample 14 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 14 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.643
operation: col(d(18_16)Mean)
inst: col
dD_actual | -91.026
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.7611154871
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.643
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -100.012930191
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -91.026
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 14 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.619
operation: col(d(18_16)Mean)
inst: col
dD_actual | -93.792
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.73515860914
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.619
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -102.784282515
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -93.792
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 14 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.651
operation: col(d(18_16)Mean)
inst: col
dD_actual | -94.497
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.76976777976
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.651
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -103.490646719
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -94.497
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 14 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.643
operation: col(d(18_16)Mean)
inst: col
dD_actual | -94.812
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.7611154871
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.643
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -103.806256258
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -94.812
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 15 - Sample 15 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 15 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.339
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.043
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.19785583865
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 4.339
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.989327216
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -70.043
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 15 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.344
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.171
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.20326352156
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 4.344
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.1156398599
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -69.171
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 15 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.448
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.672
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.31574332609
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 4.448
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.6176093156
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -69.672
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 15 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.423
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.411
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.28870491154
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 4.423
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.3561042698
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -69.411
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 16 - Sample 16 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 16 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.102
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.471
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.38461465937
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.102
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.4220254966
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -72.471
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 16 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.122
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.016
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.36298392773
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.122
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.9661450528
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -72.016
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 16 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.024
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.599
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.46897451276
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.024
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.5502731819
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -72.599
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 16 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.073
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.315
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.41597922024
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.073
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.2637885887
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -71.315
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 17 - Sample 17 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 17 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.845
operation: col(d(18_16)Mean)
inst: col
dD_actual | -82.982
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.81651271263
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -5.845
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -91.9533647172
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -82.982
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 17 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.865
operation: col(d(18_16)Mean)
inst: col
dD_actual | -82.614
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.83814344427
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -5.865
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -91.584652622
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -82.614
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 17 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.83
operation: col(d(18_16)Mean)
inst: col
dD_actual | -82.666
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.8002896639
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -5.83
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -91.6367532441
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -82.666
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 17 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.87
operation: col(d(18_16)Mean)
inst: col
dD_actual | -82.616
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.84355112718
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -5.87
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -91.5866564921
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -82.616
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 18 - Sample 18 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 18 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.535
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.918
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.39969979019
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -4.535
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.8544101286
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -64.918
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 18 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.437
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.862
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.29370920515
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -4.437
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.7944316834
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -62.862
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 18 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.37
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.185
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.22124625416
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -4.37
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.1161216603
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -62.185
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 18 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.347
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.894
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.19637091277
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -4.347
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.8245585633
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -61.894
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 19 - Sample 19 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 19 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.65
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.689
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.27946675106
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -1.65
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.6269020456
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -65.689
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 19 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.445
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.814
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.05775175175
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -1.445
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.7521439257
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -65.814
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 19 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.454
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.94
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.06748558099
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -1.454
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.878387741
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -65.94
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 19 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.386
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.684
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.99394109341
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -1.386
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.6218923703
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -65.684
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 20 - Sample 20 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 20 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.465
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.957
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.32399222945
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -4.465
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.9031608024
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -69.957
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 20 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.535
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.473
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.39969979019
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -4.535
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.4182242424
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -69.473
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 20 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.63
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.417
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.50244576548
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -4.63
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.3640509215
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -70.417
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 20 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.659
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.201
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.53381032636
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -4.659
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.1476329525
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -70.201
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 21 - Sample 21 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 21 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.0
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.544
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.57646797275
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -1.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.4545308843
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -51.544
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 21 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.902
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.713
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.47047738771
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -0.902
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.6199878234
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -49.713
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 21 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.791
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.627
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.35042682711
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -0.791
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.5318863684
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -48.627
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 21 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.805
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.561
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.36556833926
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -0.805
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.4657586557
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -48.561
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 22 - Sample 22 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 22 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.231
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.237
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.31552241531
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.231
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -68.1624171582
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -59.237
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 22 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.499
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.137
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.60537421929
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.499
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.0641586955
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -60.137
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 22 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.539
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.285
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.64863568257
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.539
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.2124450816
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -60.285
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 22 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.555
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.233
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.66594026789
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.555
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.1603444594
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -60.233
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 23 - Sample 23 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 23 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.197
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.284
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.27875017152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.197
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.2443388511
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -77.284
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 23 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.052
operation: col(d(18_16)Mean)
inst: col
dD_actual | -78.336
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.12192736713
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.052
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -87.2983745147
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -78.336
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 23 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.087
operation: col(d(18_16)Mean)
inst: col
dD_actual | -78.826
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.1597811475
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.087
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -87.789322685
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -78.826
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 23 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.062
operation: col(d(18_16)Mean)
inst: col
dD_actual | -78.845
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.13274273295
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.062
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -87.8083594508
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -78.845
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 24 - Sample 24 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 24 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.508
operation: col(d(18_16)Mean)
inst: col
dD_actual | -84.543
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.6966446306
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -8.508
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -93.5173853169
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -84.543
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 24 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.421
operation: col(d(18_16)Mean)
inst: col
dD_actual | -84.644
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.6025509479
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -8.421
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -93.6185807561
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -84.644
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 24 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.531
operation: col(d(18_16)Mean)
inst: col
dD_actual | -85.075
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.7215199719
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -8.531
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -94.050414759
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -85.075
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 24 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.559
operation: col(d(18_16)Mean)
inst: col
dD_actual | -85.022
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.7518029962
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -8.559
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -93.9973122018
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -85.022
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 25 - Sample 25 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 25 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.249
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.355
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.92730716369
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -2.249
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.2729053273
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -55.355
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 25 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.064
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.545
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.72722289602
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -2.064
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.4574678608
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -52.545
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 25 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.895
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.654
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.54444321366
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -1.895
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.5647437389
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -51.654
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 25 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.837
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.546
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.4817140919
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -1.837
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.4565347544
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -51.546
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 26 - Sample 26 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 26 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.395
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.595
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.32982125072
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -5.395
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.5443304003
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -71.595
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 26 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.575
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.689
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.52449783549
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -5.575
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.6423823771
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -73.689
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 26 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.542
operation: col(d(18_16)Mean)
inst: col
dD_actual | -74.04
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.48880712828
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -5.542
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.9940615767
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -74.04
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 26 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.601
operation: col(d(18_16)Mean)
inst: col
dD_actual | -74.139
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.55261778662
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -5.601
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.0932531458
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -74.139
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 27 - Sample 27 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 27 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.643
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.876
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.59804232307
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -5.643
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.8200690226
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -68.876
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 27 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.601
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.076
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.55261778662
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -5.601
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.0185209895
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -68.076
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 27 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.669
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.237
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.6261622742
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -5.669
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.1798325312
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -68.237
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 27 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.565
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.015
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.51368246967
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -5.565
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.957402952
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -68.015
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 28 - Sample 28 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 28 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.709
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.085
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.3432774094
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -1.709
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.9965777417
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -52.085
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 28 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.544
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.204
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.16482387337
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -1.544
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -59.1119379288
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -50.204
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 28 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.523
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.887
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.14211160515
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -1.523
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.7943245207
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -49.887
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 28 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.313
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.851
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.91498892292
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -1.313
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.7563198177
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -48.851
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 29 - Sample 29 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 29 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.7
operation: col(d(18_16)Mean)
inst: col
dD_actual | -46.99
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.33354358016
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -1.7
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.8917187056
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -46.99
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 29 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.411
operation: col(d(18_16)Mean)
inst: col
dD_actual | -45.341
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.02097950796
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -1.411
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.2395278222
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -45.341
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 29 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.628
operation: col(d(18_16)Mean)
inst: col
dD_actual | -45.977
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.25567294626
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -1.628
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.8767585086
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -45.977
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 29 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.676
operation: col(d(18_16)Mean)
inst: col
dD_actual | -46.385
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.3075867022
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -1.676
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.2855480055
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -46.385
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 30 - Sample 30 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 30 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.976
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.542
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.71358425882
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -2.976
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.4699423872
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -60.542
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 30 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.133
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.524
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.8833855022
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -3.133
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.4557776394
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -62.524
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 30 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.057
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.317
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.80118872196
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -3.057
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.2483770858
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -62.317
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 30 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.112
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.653
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.86067323398
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -3.112
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.5850272597
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -62.653
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 31 - Sample 31 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 31 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.339
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.525
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.43232836617
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.339
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.4529094915
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -60.525
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 31 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.539
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.537
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.64863568257
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.539
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.464932712
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -60.537
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 31 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.604
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.762
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.7189355604
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.604
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.6903680964
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -60.762
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 31 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.593
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.458
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.707038658
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.593
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.3857798438
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -60.458
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 32 - Sample 32 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 32 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.893
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.288
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.70535330453
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -3.893
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -68.2135158453
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -59.288
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 32 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.694
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.079
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.49012752471
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -3.694
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -68.0041114216
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -59.079
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 32 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.648
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.873
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.44037684194
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -3.648
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.7977128031
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -58.873
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 32 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.677
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.914
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.47174140282
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -3.677
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.8387921398
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -58.914
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 33 - Sample 33 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 33 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.865
operation: col(d(18_16)Mean)
inst: col
dD_actual | -99.378
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.3273629364
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -11.865
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -108.381091657
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -99.378
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 33 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.25
operation: col(d(18_16)Mean)
inst: col
dD_actual | -103.851
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.7437545205
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -12.25
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -112.862747097
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -103.851
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 33 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.307
operation: col(d(18_16)Mean)
inst: col
dD_actual | -104.771
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.8054021056
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -12.307
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -113.784527335
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -104.771
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 33 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.318
operation: col(d(18_16)Mean)
inst: col
dD_actual | -104.988
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.817299008
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -12.318
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -114.001947239
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -104.988
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 34 - Sample 34 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 34 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.217
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.551
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.26023795243
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.217
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.4692845954
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -55.551
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 34 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.873
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.327
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.550749954629
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.873
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -59.2351759389
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -50.327
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 34 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.973
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.624
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.442596296427
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.973
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.5308156048
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -49.624
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 34 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.072
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.749
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.335524174808
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 1.072
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.6541224435
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -48.749
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 35 - Sample 35 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 35 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.117
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.937
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.70300775285
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -1.117
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.8444212727
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -49.937
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 35 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.198
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.531
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.79061221599
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -1.198
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.4376356459
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -49.531
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 35 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.257
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.982
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.85442287433
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -1.257
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.8895083496
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -49.982
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 35 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.254
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.604
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.85117826458
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -1.254
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.5107769039
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -49.604
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 36 - Sample 36 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 36 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.54
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.09
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.24203430906
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -2.54
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.0035224584
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -53.09
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 36 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.51
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.959
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.2095882116
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -2.51
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.872268968
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -52.959
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 36 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.531
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.269
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.23230047982
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -2.531
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.1828688308
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -53.269
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 36 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.583
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.508
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.28854038209
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -2.583
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.4223313057
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -53.508
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 37 - Sample 37 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 37 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.924
operation: col(d(18_16)Mean)
inst: col
dD_actual | -75.113
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.82041752059
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -4.924
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -84.0691378761
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -75.113
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 37 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.092
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.434
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.00211566637
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -5.092
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.3946291073
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -77.434
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 37 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.056
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.498
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.96318034942
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -5.056
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.45875295
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -77.498
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 37 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.07
operation: col(d(18_16)Mean)
inst: col
dD_actual | -78.085
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.97832186157
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -5.07
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -87.0468888193
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -78.085
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 38 - Sample 38 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 38 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.533
operation: col(d(18_16)Mean)
inst: col
dD_actual | -114.157
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.2129025372
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -14.533
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -123.188689634
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -114.157
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 38 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.933
operation: col(d(18_16)Mean)
inst: col
dD_actual | -117.82
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.64551717
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -14.933
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -126.858777691
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -117.82
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 38 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -15.019
operation: col(d(18_16)Mean)
inst: col
dD_actual | -118.827
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.7385293161
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -15.019
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -127.867726278
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -118.827
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 38 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -15.147
operation: col(d(18_16)Mean)
inst: col
dD_actual | -119.149
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.8769659986
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -15.147
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -128.190349361
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -119.149
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 39 - Sample 39 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 39 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.197
operation: col(d(18_16)Mean)
inst: col
dD_actual | -78.995
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.11567700749
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -5.197
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -87.958649707
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -78.995
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 39 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.674
operation: col(d(18_16)Mean)
inst: col
dD_actual | -74.782
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.55003337509
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -4.674
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.7374973774
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -74.782
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 39 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.61
operation: col(d(18_16)Mean)
inst: col
dD_actual | -74.184
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.48081503384
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -4.61
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.1383402226
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -74.184
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 39 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.516
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.578
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.37915059513
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -4.516
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.5311675875
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -73.578
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 40 - Sample 40 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 40 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.527
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.874
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.63565724359
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.527
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.8064549039
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -62.874
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 40 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.597
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.285
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.71136480433
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.597
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.214380123
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -61.285
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 40 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.641
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.009
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.75895241394
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.641
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.9378460516
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -61.009
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 40 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.619
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.934
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.73515860914
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.619
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.8627009235
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -60.934
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 41 - Sample 41 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 41 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.543
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.795
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.32681550083
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -3.543
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.7195618698
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -58.795
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 41 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.889
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.487
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.61949057619
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -2.889
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -65.4070957942
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -56.487
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 41 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.997
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.922
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.73629652704
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -2.997
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -65.8429375372
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -56.922
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 41 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.071
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.549
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.81633023411
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -3.071
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.4711508082
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -57.549
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 42 - Sample 42 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 42 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.064
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.207
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.645686314
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -1.064
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.1168787754
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -51.207
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 42 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.019
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.628
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.59701716781
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -1.019
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -59.5367583864
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -50.628
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 42 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.033
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.643
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.61215867996
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -1.033
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -59.551787412
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -50.643
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 42 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.035
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.541
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.61432175312
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -1.035
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -59.4495900378
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -50.541
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 43 - Sample 43 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 43 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -10.578
operation: col(d(18_16)Mean)
inst: col
dD_actual | -94.53
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.9354253553
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -10.578
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -103.523710576
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -94.53
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 43 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.034
operation: col(d(18_16)Mean)
inst: col
dD_actual | -98.613
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -13.4286060367
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -11.034
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -107.61461135
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -98.613
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 43 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.166
operation: col(d(18_16)Mean)
inst: col
dD_actual | -100.015
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -13.5713688656
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -11.166
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -109.019324278
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -100.015
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 43 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.22
operation: col(d(18_16)Mean)
inst: col
dD_actual | -100.453
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -13.629771841
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -11.22
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -109.458171826
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -100.453
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 44 - Sample 44 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 44 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.302
operation: col(d(18_16)Mean)
inst: col
dD_actual | -102.242
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -13.7184578407
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -11.302
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -111.250633615
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -102.242
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 44 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.191
operation: col(d(18_16)Mean)
inst: col
dD_actual | -101.691
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -13.5984072801
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -11.191
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -110.698567408
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -101.691
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 44 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 44 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 45 - Sample 45 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 45 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 45 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 45 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 45 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 46 - Sample 46 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 46 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.541
operation: col(d(18_16)Mean)
inst: col
dD_actual | -78.666
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.56926217372
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -6.541
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -87.6290130784
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -78.666
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 46 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 46 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 47 - Sample 47 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 47 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.313
operation: col(d(18_16)Mean)
inst: col
dD_actual | 6.697
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.0881993055
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 3.313
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -2.10083213559
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 6.697
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 47 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.393
operation: col(d(18_16)Mean)
inst: col
dD_actual | 7.723
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.17472223206
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 3.393
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -1.07284678307
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 7.723
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 47 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 48 - Sample 48 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 48 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.091
operation: col(d(18_16)Mean)
inst: col
dD_actual | 24.543
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.84809818429
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 3.091
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 15.779700614
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 24.543
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 48 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.228
operation: col(d(18_16)Mean)
inst: col
dD_actual | 17.608
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.914732114006
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 2.228
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 8.83128110161
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 17.608
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 48 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.748
operation: col(d(18_16)Mean)
inst: col
dD_actual | 21.605
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.47713113666
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 2.748
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 12.8360154623
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 21.605
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 49 - Sample 49 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 49 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 49 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 49 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 49 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 50 - Sample 50 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 50 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.552
operation: col(d(18_16)Mean)
inst: col
dD_actual | -4.699
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.183613384562
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 1.552
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -13.5188838679
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -4.699
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 50 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 50 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.58
operation: col(d(18_16)Mean)
inst: col
dD_actual | -4.464
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.213896408858
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 1.58
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -13.2834291331
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -4.464
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 51 - Sample 51 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 51 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.104
operation: col(d(18_16)Mean)
inst: col
dD_actual | 0.256
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.780621577836
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 2.104
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.55429573753
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.256
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 51 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.836
operation: col(d(18_16)Mean)
inst: col
dD_actual | -2.383
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.490769773855
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 1.836
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -11.1984023119
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -2.383
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 51 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 52 - Sample 52 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 52 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 52 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 52 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.283
operation: col(d(18_16)Mean)
inst: col
dD_actual | 1.229
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.05575320804
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 3.283
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -7.5794129422
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 1.229
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 53 - Sample 53 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 53 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.322
operation: col(d(18_16)Mean)
inst: col
dD_actual | 6.576
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.09793313474
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 3.322
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -2.2220662756
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 6.576
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 53 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.958
operation: col(d(18_16)Mean)
inst: col
dD_actual | 3.764
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.70425381888
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 2.958
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -5.03950761214
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 3.764
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 53 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 54 - Sample 54 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 54 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.907
operation: col(d(18_16)Mean)
inst: col
dD_actual | 8.41
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.73063203522
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 3.907
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.384517409593
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 8.41
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 54 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.51
operation: col(d(18_16)Mean)
inst: col
dD_actual | 5.5
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.30126201215
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 3.51
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -3.3001483802
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 5.5
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 54 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 55 - Sample 55 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 55 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 55 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.855
operation: col(d(18_16)Mean)
inst: col
dD_actual | 15.833
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.511318968913
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 1.855
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 7.05284640305
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 15.833
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 55 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 56 - Sample 56 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 56 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.927
operation: col(d(18_16)Mean)
inst: col
dD_actual | 20.355
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.83379934888
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 4.927
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 11.5835966605
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 20.355
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 56 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.822
operation: col(d(18_16)Mean)
inst: col
dD_actual | 18.915
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.72023800776
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 4.822
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 10.1408102008
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 18.915
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 56 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.941
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.788
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.84894086102
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 4.941
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 11.015499492
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 19.788
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 57 - Sample 57 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 57 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.328
operation: col(d(18_16)Mean)
inst: col
dD_actual | 4.845
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.10442235423
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 3.328
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -3.95641583234
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 4.845
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 57 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.203
operation: col(d(18_16)Mean)
inst: col
dD_actual | 11.478
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.05076686349
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 4.203
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 2.68941929756
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 11.478
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 57 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.363
operation: col(d(18_16)Mean)
inst: col
dD_actual | 5.294
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.1422761346
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 3.363
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -3.50654699873
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 5.294
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 58 - Sample 58 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 58 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 58 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 58 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 58 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 59 - Sample 59 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 59 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.636
operation: col(d(18_16)Mean)
inst: col
dD_actual | -6.559
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.274462457451
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 1.636
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -15.382483045
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -6.559
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 59 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 59 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.293
operation: col(d(18_16)Mean)
inst: col
dD_actual | -9.395
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.0965045901814
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 1.293
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.2239708225
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -9.395
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 60 - Sample 60 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 60 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.122
operation: col(d(18_16)Mean)
inst: col
dD_actual | -3.758
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.800089236312
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 2.122
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -12.5760629939
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -3.758
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 60 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.907
operation: col(d(18_16)Mean)
inst: col
dD_actual | -5.593
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.567558871178
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 1.907
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -14.4146137949
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -5.593
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 60 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 61 - Sample 61 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 61 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 61 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 61 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.352
operation: col(d(18_16)Mean)
inst: col
dD_actual | 1.324
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.04884265018
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 2.352
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -7.48422911327
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 1.324
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 62 - Sample 62 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 62 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.47
operation: col(d(18_16)Mean)
inst: col
dD_actual | 4.339
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.17646396686
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 2.47
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -4.46339496331
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 4.339
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 62 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.205
operation: col(d(18_16)Mean)
inst: col
dD_actual | 1.781
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.88985677262
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 2.205
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -7.02634479933
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 1.781
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 62 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.866
operation: col(d(18_16)Mean)
inst: col
dD_actual | -0.85
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.523215871316
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 1.866
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -9.66243589336
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -0.85
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 63 - Sample 63 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 63 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.6
operation: col(d(18_16)Mean)
inst: col
dD_actual | 2.792
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.235527140499
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 1.6
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -6.01338847243
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 2.792
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 63 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.309
operation: col(d(18_16)Mean)
inst: col
dD_actual | 0.346
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.079200004869
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 1.309
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.4641215838
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.346
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 63 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.091
operation: col(d(18_16)Mean)
inst: col
dD_actual | -1.854
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.314974979749
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 1.091
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -10.668378675
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -1.854
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean

## 64 - Sample 64 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 12:55:49
Detail Variables
Detail 64 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.829
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.326
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.88074470934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -6.829
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.2322389624
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -49.326
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 64 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 64 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.765
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.228
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.89306295011
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: -7.765
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.1514647013
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: -58.228
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean
Detail 64 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 3.16672727273
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 26.8588181818
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.451
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.754
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.2636
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -109.9764
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.49493139073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.38222915025
measure_array: [ 3.16672727 -13.2636 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.924610426152
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.81079110814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 8.79377478947
measure_array: [ 26.85881818 -109.9764 ]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.99806869571
tarray: []
source_col: d(D_H)Mean