## Summary
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
221201_22-284_HIDS2046
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -10.4842 | -76.3602 |  | 74.9475 |  | 0.25553 | 2.451 |  | -0.23815
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.21736 | 3.4506
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.32352 | 5.0035
4 |  | EV | -10.3299 | -74.9985 |  |  |  | 0.093657 | 1.5429
5 |  | UNDE6-D-2 | -6.5764 | -59.4353 |  |  |  | 0.11115 | 0.93877
6 | * | UNDE6-E-1 | -0.96238 | -0.75035 |  |  |  | 0 | 0
7 |  | UNDE6-C-3 | -5.2071 | -61.2984 |  |  |  | 0.084727 | 0.27979
8 |  | UNDE6-010-2 | -7.3558 | -57.8958 |  |  |  | 0.060139 | 0.34887
9 |  | UNDE6-E-2 | -4.2397 | -51.2129 |  |  |  | 0.046313 | 0.22232
10 |  | UNDE6-2030-3 | -7.736 | -61.9207 |  |  |  | 0.090521 | 0.50291
11 |  | UNDE6-E-3 | -3.6742 | -57.7821 |  |  |  | 0.10608 | 0.21449
12 |  | UNDE6-2030-2 | -6.8258 | -56.7489 |  |  |  | 0.071444 | 0.11371
13 |  | EV | -10.1691 | -72.6004 |  |  |  | 0.047181 | 0.6463
14 |  | UNDE6-2030-1 | -6.6128 | -58.7496 |  |  |  | 0.090524 | 0.68683
15 |  | UNDE6-C-1 | 0.46722 | -58.26 | -4.4145 |  | 36.1617 | 0.031309 | 0.43683 |  | 1.5997
16 |  | UNDE6-D-3 | -4.2302 | -56.0088 |  |  | 7.5604 | 0.048724 | 0.71272 | 32.8788 | 3.8249
17 |  | UNDE6-D-5 | -4.8676 | -64.1931 | -1.6402 |  |  | 0.050296 | 0.53094
18 |  | UNDE6-B-2 | -3.4009 | -59.1151 | -1.4447 |  |  | 0.068956 | 0.43649
19 |  | UNDE6-C-2 | -6.5321 | -70.7229 |  |  |  | 0.075716 | 0.6975
20 |  | UNDE6-A-1 | -4.5996 | -59.8898 | -1.4512 | 113.4795 |  | 0.023274 | 0.54899 |  | -6.0718
21 |  | UNDE6-B-4 | -4.8191 | -61.2505 |  |  |  | 0.02181 | 0.20358
22 |  | EV | -10.107 | -72.763 |  |  |  | 0.1224 | 0.64297
23 |  | UNDE6-A-5 | -4.3649 | -60.9761 | -1.4322 |  |  | 0.10787 | 0.59568
24 |  | UNDE6-D-4 | -2.4894 | -67.6903 |  |  |  | 0.072334 | 0.2834
25 |  | UNDE6-C-4 | -6.0728 | -67.3742 | -1.367 | 293.8063 |  | 0.16039 | 0.23523 |  | -9.5784
26 |  | UNDE6-3040-1 | -6.3384 | -55.2677 |  |  |  | 0.021156 | 0.7052
27 |  | UNDE6-1020-1 | -4.1539 | -51.8706 |  |  |  | 0.042669 | 0.20917
28 |  | UNDE6-B-3 | -3.8736 | -61.3005 | -2.26 | 130.1583 |  | 0.060056 | 0.65826 |  | -7.5606
29 |  | UNDE6-3040-2 | -6.9881 | -61.0737 |  |  |  | 0.047479 | 0.27466
30 |  | UNDE6-C-5 | -2.9704 | -62.2808 |  |  |  | 0.11166 | 0.27282
31 |  | EV | -10.0619 | -72.8207 |  |  |  | 0.14976 | 0.51989
32 |  | UNDE6-B-1 | -4.7758 | -61.5489 |  |  |  | 0.13937 | 0.66944
33 |  | UNDE6-E-4 | -1.6446 | -49.3441 |  |  |  | 0.094067 | 0.74614
34 |  | UNDE6-A-2 | -5.7121 | -65.4494 |  |  |  | 0.13416 | 0.99823
35 |  | UNDE6-3040-3 | -5.8298 | -53.5169 |  |  |  | 0.032816 | 0.63182
36 |  | UNDE6-B-5 | -3.4609 | -49.2187 |  |  |  | 0.035351 | 0.24795
37 |  | UNDE6-A-4 | -3.9089 | -68.485 | -2.1522 |  |  | 0.066365 | 1.1034
38 |  | UNDE6-D-1 | -6.4687 | -64.8469 |  |  |  | 0.04435 | 0.24586
39 |  | UNDE6-E-5 | -5.1471 | -63.567 |  |  |  | 0.042177 | 0.10238
40 |  | EV | -10.114 | -72.7165 |  |  |  | 0.14733 | 0.50335
41 |  | UNDE6-1020-2 | -6.857 | -56.1342 |  |  |  | 0.099321 | 0.9816
42 |  | UNDE6-A-3 | -4.5574 | -60.4577 |  |  |  | 0.11122 | 0.43894
43 |  | UNDE6-1020-3 | -5.2416 | -54.1006 |  |  |  | 0.0093986 | 0.3424
44 |  | UNDE6-010-1 | -3.9308 | -49.7679 |  |  |  | 0.034191 | 0.20599
45 |  | UNDE6-010-3 | -3.3506 | -49.5851 |  |  |  | 0.052265 | 0.15805
46 |  | UNDE6-deep-1 | -1.3429 | -47.2379 |  |  |  | 0.08543 | 0.29736
47 |  | UNDE6-deep-2 | -4.6707 | -54.9598 |  |  |  | 0.098882 | 0.48036
48 |  | UNDE6-deep-3 | -3.3365 | -46.1641 |  |  |  | 0.11884 | 0.67022
49 |  | EV | -10.04 | -71.6852 |  |  |  | 0.16262 | 1.261

## Detail
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
1 |  | EV | -9.875 | -75.572 | -11.1388 | -82.9679 | -1.159 | 322.041 | 2022/12/01 11:08:44
1 |  | EV | -9.474 | -71.771 | -10.7256 | -78.8326 | -1.196 | 334.503 | 2022/12/01 11:17:40
1 |  | EV | -9.258 | -69.755 | -10.503 | -76.6393 | -1.072 | 157.32 | 2022/12/01 11:26:37
1 |  | EV | -9.183 | -68.948 | -10.4257 | -75.7614 | -0.996 | 104.251 | 2022/12/01 11:35:32
1 |  | EV | -9.176 | -68.891 | -10.4185 | -75.6994 | -0.944 | 130.925 | 2022/12/01 11:44:28
1 |  | EV | -9.107 | -68.321 | -10.3474 | -75.0792 | -0.988 | 73.958 | 2022/12/01 11:53:25
1 |  | EV | -9.126 | -68.064 | -10.367 | -74.7996 | -0.966 | 68.608 | 2022/12/01 12:02:21
1 |  | EV | -9.074 | -67.978 | -10.3134 | -74.7061 | -0.974 | 65.311 | 2022/12/01 12:11:17
1 |  | EV | -9.043 | -67.716 | -10.2814 | -74.421 | -0.793 | 63.311 | 2022/12/01 12:20:13
1 |  | EV | -9.082 | -67.968 | -10.3216 | -74.6952 | -0.996 | 61.662 | 2022/12/01 12:29:09
Sample 1 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -10.4842 | -76.3602 |  | 74.9475 |  | 0.25553 | 2.451 |  | -0.23815
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
2 |  | PZ | 2.239 | 8.309 | 1.345 | 8.2893 | -1.028 | 56.725 | 2022/12/01 12:38:57
2 |  | PZ | 2.687 | 15.198 | 1.8066 | 15.7841 | -1.026 | 61.151 | 2022/12/01 12:47:53
2 |  | PZ | 2.769 | 16.918 | 1.8911 | 17.6554 | -1.118 | 62.387 | 2022/12/01 12:56:49
2 |  | PZ | 2.883 | 18.143 | 2.0086 | 18.9881 | -1.221 | 61.23 | 2022/12/01 13:05:45
2 |  | PZ | 2.846 | 18.353 | 1.9705 | 19.2165 | -0.906 | 60.994 | 2022/12/01 13:14:41
2 |  | PZ | 2.873 | 18.732 | 1.9983 | 19.6289 | -1.009 | 61.493 | 2022/12/01 13:23:38
2 |  | PZ | 2.926 | 19.289 | 2.0529 | 20.2349 | -1.069 | 60.963 | 2022/12/01 13:32:34
2 |  | PZ | 2.923 | 19.124 | 2.0499 | 20.0553 | -1.063 | 61.254 | 2022/12/01 13:41:31
2 |  | PZ | 2.947 | 19.488 | 2.0746 | 20.4514 | -1.118 | 61.744 | 2022/12/01 13:50:28
2 |  | PZ | 2.974 | 19.713 | 2.1024 | 20.6961 | -1.014 | 61.811 | 2022/12/01 13:59:23
Sample 2 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.21736 | 3.4506
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
3 |  | UT2 | -13.561 | -95.448 | -14.9374 | -104.5917 | -0.887 | 57.918 | 2022/12/01 14:09:12
3 |  | UT2 | -14.282 | -105.739 | -15.6804 | -115.7877 | -0.998 | 64.733 | 2022/12/01 14:18:08
3 |  | UT2 | -14.497 | -108.726 | -15.9019 | -119.0373 | -1.112 | 66.678 | 2022/12/01 14:27:05
3 |  | UT2 | -14.482 | -109.713 | -15.8865 | -120.1111 | -0.792 | 66.492 | 2022/12/01 14:36:01
3 |  | UT2 | -14.55 | -110.198 | -15.9566 | -120.6388 | -0.88 | 66.397 | 2022/12/01 14:44:57
3 |  | UT2 | -14.588 | -110.9 | -15.9957 | -121.4025 | -0.905 | 65.23 | 2022/12/01 14:53:53
3 |  | UT2 | -14.593 | -111.092 | -16.0009 | -121.6114 | -0.997 | 64.931 | 2022/12/01 15:02:50
3 |  | UT2 | -14.611 | -111.537 | -16.0194 | -122.0955 | -0.815 | 65.122 | 2022/12/01 15:11:47
3 |  | UT2 | -14.604 | -111.624 | -16.0122 | -122.1902 | -0.895 | 65.529 | 2022/12/01 15:20:43
3 |  | UT2 | -14.601 | -111.94 | -16.0091 | -122.5339 | -1.094 | 72.047 | 2022/12/01 15:29:40
Sample 3 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.32352 | 5.0035
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
4 |  | EV | -9.327 | -72.374 | -10.5741 | -79.4887 | -1.07 | 56.074 | 2022/12/01 15:39:27
4 |  | EV | -9.156 | -69.081 | -10.3979 | -75.9061 | -0.924 | 61.077 | 2022/12/01 15:48:23
4 |  | EV | -9.052 | -68.062 | -10.2907 | -74.7975 | -1.178 | 62.713 | 2022/12/01 15:57:19
4 |  | EV | -9.074 | -67.733 | -10.3134 | -74.4395 | -1.018 | 64.375 | 2022/12/01 16:06:16
4 |  | EV | -9.101 | -68.034 | -10.3412 | -74.767 | -0.849 | 65.207 | 2022/12/01 16:15:13
4 |  | EV | -9.048 | -67.499 | -10.2866 | -74.185 | -0.98 | 64.828 | 2022/12/01 16:24:09
4 |  | EV | -9.08 | -67.604 | -10.3196 | -74.2992 | -0.89 | 66.553 | 2022/12/01 16:33:05
4 |  | EV | -9.035 | -67.439 | -10.2732 | -74.1197 | -1.117 | 65.658 | 2022/12/01 16:42:02
4 |  | EV | -9.012 | -67.402 | -10.2495 | -74.0794 | -0.926 | 64.68 | 2022/12/01 16:50:58
4 |  | EV | -9.015 | -67.24 | -10.2526 | -73.9032 | -1.143 | 63.787 | 2022/12/01 16:59:54
Sample 4 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
4 |  | EV | -10.3299 | -74.9985 |  |  |  | 0.093657 | 1.5429
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
5 |  | UNDE6-D-2 | -5.591 | -55.307 | -6.7241 | -60.9208 | -0.88 | 54.057 | 2022/12/01 17:09:44
5 |  | UNDE6-D-2 | -5.456 | -53.775 | -6.5849 | -59.2541 | -0.893 | 55.258 | 2022/12/01 17:18:41
5 |  | UNDE6-D-2 | -5.322 | -53.214 | -6.4468 | -58.6438 | -1.103 | 54.797 | 2022/12/01 17:27:38
5 |  | UNDE6-D-2 | -5.422 | -53.47 | -6.5499 | -58.9223 | -1.038 | 54.37 | 2022/12/01 17:36:34
Sample 5 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
5 |  | UNDE6-D-2 | -6.5764 | -59.4353 |  |  |  | 0.11115 | 0.93877
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
6 |  | UNDE6-E-1 |  |  | -0.96238 | -0.75035 |  |  | 2022/12/01 17:46:24
6 |  | UNDE6-E-1 |  |  | -0.96238 | -0.75035 |  |  | 2022/12/01 17:56:23
6 |  | UNDE6-E-1 |  |  | -0.96238 | -0.75035 |  |  | 2022/12/01 18:06:21
6 |  | UNDE6-E-1 |  |  | -0.96238 | -0.75035 |  |  | 2022/12/01 18:16:20
Sample 6 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
6 | * | UNDE6-E-1 | -0.96238 | -0.75035 |  |  |  | 0 | 0
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
7 |  | UNDE6-C-3 | -4.239 | -56.017 | -5.3308 | -61.6933 | -0.963 | 53.68 | 2022/12/01 18:27:12
7 |  | UNDE6-C-3 | -4.083 | -55.354 | -5.17 | -60.972 | -1.016 | 54.553 | 2022/12/01 18:36:09
7 |  | UNDE6-C-3 | -4.043 | -55.548 | -5.1288 | -61.183 | -1.099 | 54.58 | 2022/12/01 18:45:05
7 |  | UNDE6-C-3 | -4.111 | -55.697 | -5.1989 | -61.3451 | -1.086 | 54.182 | 2022/12/01 18:54:01
Sample 7 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
7 |  | UNDE6-C-3 | -5.2071 | -61.2984 |  |  |  | 0.084727 | 0.27979
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
8 |  | UNDE6-010-2 | -6.127 | -53.044 | -7.2764 | -58.4588 | -0.897 | 53.616 | 2022/12/01 19:03:48
8 |  | UNDE6-010-2 | -6.186 | -52.282 | -7.3372 | -57.6298 | -0.762 | 53.404 | 2022/12/01 19:12:44
8 |  | UNDE6-010-2 | -6.245 | -52.402 | -7.398 | -57.7604 | -1.258 | 54.342 | 2022/12/01 19:21:40
8 |  | UNDE6-010-2 | -6.258 | -52.378 | -7.4114 | -57.7343 | -0.829 | 53.527 | 2022/12/01 19:30:36
Sample 8 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
8 |  | UNDE6-010-2 | -7.3558 | -57.8958 |  |  |  | 0.060139 | 0.34887
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
9 |  | UNDE6-E-2 | -3.245 | -46.704 | -4.3064 | -51.5613 | -1.15 | 54.324 | 2022/12/01 19:40:23
9 |  | UNDE6-E-2 | -3.169 | -46.343 | -4.2281 | -51.1686 | -1.115 | 53.329 | 2022/12/01 19:49:20
9 |  | UNDE6-E-2 | -3.172 | -46.194 | -4.2312 | -51.0065 | -1.035 | 53.282 | 2022/12/01 19:58:17
9 |  | UNDE6-E-2 | -3.135 | -46.294 | -4.1931 | -51.1153 | -1.114 | 53.357 | 2022/12/01 20:07:13
Sample 9 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
9 |  | UNDE6-E-2 | -4.2397 | -51.2129 |  |  |  | 0.046313 | 0.22232
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
10 |  | UNDE6-2030-3 | -6.441 | -55.51 | -7.6 | -61.1417 | -1.079 | 53.055 | 2022/12/01 20:17:02
10 |  | UNDE6-2030-3 | -6.587 | -56.316 | -7.7505 | -62.0186 | -0.857 | 53.155 | 2022/12/01 20:25:59
10 |  | UNDE6-2030-3 | -6.632 | -56.394 | -7.7968 | -62.1034 | -1.019 | 53.388 | 2022/12/01 20:34:55
10 |  | UNDE6-2030-3 | -6.632 | -56.684 | -7.7968 | -62.4189 | -1.035 | 53.244 | 2022/12/01 20:43:52
Sample 10 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
10 |  | UNDE6-2030-3 | -7.736 | -61.9207 |  |  |  | 0.090521 | 0.50291
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
11 |  | UNDE6-E-3 | -2.788 | -52.644 | -3.8355 | -58.0237 | -1.054 | 51.967 | 2022/12/01 20:53:40
11 |  | UNDE6-E-3 | -2.559 | -52.147 | -3.5995 | -57.483 | -1.1 | 53.143 | 2022/12/01 21:02:37
11 |  | UNDE6-E-3 | -2.605 | -52.525 | -3.6469 | -57.8942 | -0.91 | 52.916 | 2022/12/01 21:11:33
11 |  | UNDE6-E-3 | -2.574 | -52.372 | -3.615 | -57.7278 | -0.978 | 52.883 | 2022/12/01 21:20:30
Sample 11 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
11 |  | UNDE6-E-3 | -3.6742 | -57.7821 |  |  |  | 0.10608 | 0.21449
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
12 |  | UNDE6-2030-2 | -5.588 | -51.622 | -6.721 | -56.9118 | -1.016 | 52.182 | 2022/12/01 21:30:19
12 |  | UNDE6-2030-2 | -5.692 | -51.495 | -6.8281 | -56.7736 | -0.954 | 52.439 | 2022/12/01 21:39:15
12 |  | UNDE6-2030-2 | -5.741 | -51.364 | -6.8786 | -56.6311 | -1.179 | 52.462 | 2022/12/01 21:48:11
12 |  | UNDE6-2030-2 | -5.738 | -51.408 | -6.8755 | -56.679 | -1.115 | 52.435 | 2022/12/01 21:57:07
Sample 12 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
12 |  | UNDE6-2030-2 | -6.8258 | -56.7489 |  |  |  | 0.071444 | 0.11371
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
13 |  | EV | -8.869 | -65.125 | -10.1021 | -71.6022 | -1.002 | 53.661 | 2022/12/01 22:06:57
13 |  | EV | -8.941 | -66.111 | -10.1763 | -72.6749 | -0.969 | 52.867 | 2022/12/01 22:15:53
13 |  | EV | -8.982 | -66.619 | -10.2186 | -73.2276 | -0.949 | 53.019 | 2022/12/01 22:24:49
13 |  | EV | -8.944 | -66.315 | -10.1794 | -72.8968 | -0.948 | 52.812 | 2022/12/01 22:33:44
Sample 13 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
13 |  | EV | -10.1691 | -72.6004 |  |  |  | 0.047181 | 0.6463
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
14 |  | UNDE6-2030-1 | -5.597 | -54.193 | -6.7302 | -59.7089 | -1.008 | 52.107 | 2022/12/01 22:43:34
14 |  | UNDE6-2030-1 | -5.497 | -53.342 | -6.6272 | -58.783 | -0.943 | 52.833 | 2022/12/01 22:52:31
14 |  | UNDE6-2030-1 | -5.459 | -53.187 | -6.588 | -58.6144 | -0.94 | 52.951 | 2022/12/01 23:01:27
14 |  | UNDE6-2030-1 | -5.379 | -52.523 | -6.5056 | -57.892 | -0.991 | 52.81 | 2022/12/01 23:10:24
Sample 14 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
14 |  | UNDE6-2030-1 | -6.6128 | -58.7496 |  |  |  | 0.090524 | 0.68683
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
15 |  | UNDE6-C-1 | 1.341 | -52.228 | 0.41956 | -57.5711 | -4.481 | 64.989 | 2022/12/01 23:20:10
15 |  | UNDE6-C-1 | 1.406 | -52.915 | 0.48654 | -58.3185 | -4.383 | 65.559 | 2022/12/01 23:29:07
15 |  | UNDE6-C-1 | 1.407 | -53.169 | 0.48757 | -58.5948 | -4.316 | 65.082 | 2022/12/01 23:38:03
15 |  | UNDE6-C-1 | 1.395 | -53.133 | 0.4752 | -58.5557 | -4.478 | 64.728 | 2022/12/01 23:46:59
Sample 15 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
15 |  | UNDE6-C-1 | 0.46722 | -58.26 | -4.4145 |  | 36.1617 | 0.031309 | 0.43683 |  | 1.5997
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
16 |  | UNDE6-D-3 | -3.241 | -51.852 | -4.3023 | -57.162 | -1.473 | 65.541 | 2022/12/01 23:56:45
16 |  | UNDE6-D-3 | -3.166 | -50.439 | -4.225 | -55.6248 | -1.303 | 53.744 | 2022/12/02 00:05:41
16 |  | UNDE6-D-3 | -3.145 | -50.552 | -4.2034 | -55.7477 | -1.172 | 53.721 | 2022/12/02 00:14:38
16 |  | UNDE6-D-3 | -3.132 | -50.325 | -4.19 | -55.5007 | -1.224 | 53.639 | 2022/12/02 00:23:34
Sample 16 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
16 |  | UNDE6-D-3 | -4.2302 | -56.0088 |  |  | 7.5604 | 0.048724 | 0.71272 | 32.8788 | 3.8249
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
17 |  | UNDE6-D-5 | -3.725 | -57.535 | -4.8011 | -63.3448 | -1.71 | 55.354 | 2022/12/02 00:33:21
17 |  | UNDE6-D-5 | -3.8 | -58.425 | -4.8784 | -64.313 | -1.709 | 55.402 | 2022/12/02 00:42:17
17 |  | UNDE6-D-5 | -3.847 | -58.676 | -4.9268 | -64.5861 | -1.545 | 55.452 | 2022/12/02 00:51:13
17 |  | UNDE6-D-5 | -3.786 | -58.623 | -4.864 | -64.5284 | -1.597 | 55.971 | 2022/12/02 01:00:09
Sample 17 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
17 |  | UNDE6-D-5 | -4.8676 | -64.1931 | -1.6402 |  |  | 0.050296 | 0.53094
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
18 |  | UNDE6-B-2 | -2.46 | -54.253 | -3.4975 | -59.7742 | -1.318 | 54.678 | 2022/12/02 01:09:57
18 |  | UNDE6-B-2 | -2.343 | -53.593 | -3.3769 | -59.0561 | -1.388 | 54.863 | 2022/12/02 01:18:54
18 |  | UNDE6-B-2 | -2.296 | -53.214 | -3.3285 | -58.6438 | -1.415 | 54.765 | 2022/12/02 01:27:50
18 |  | UNDE6-B-2 | -2.366 | -53.529 | -3.4006 | -58.9865 | -1.658 | 55.383 | 2022/12/02 01:36:46
Sample 18 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
18 |  | UNDE6-B-2 | -3.4009 | -59.1151 | -1.4447 |  |  | 0.068956 | 0.43649
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
19 |  | UNDE6-C-2 | -5.292 | -63.28 | -6.4159 | -69.595 | -0.967 | 51.104 | 2022/12/02 01:46:34
19 |  | UNDE6-C-2 | -5.438 | -64.535 | -6.5664 | -70.9603 | -0.879 | 50.963 | 2022/12/02 01:55:31
19 |  | UNDE6-C-2 | -5.434 | -64.694 | -6.5623 | -71.1333 | -1.011 | 51.182 | 2022/12/02 02:04:28
19 |  | UNDE6-C-2 | -5.455 | -64.758 | -6.5839 | -71.2029 | -0.831 | 51.066 | 2022/12/02 02:13:24
Sample 19 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
19 |  | UNDE6-C-2 | -6.5321 | -70.7229 |  |  |  | 0.075716 | 0.6975
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
20 |  | UNDE6-A-1 | -3.562 | -55.128 | -4.6331 | -60.7261 | -1.231 | 169.825 | 2022/12/02 02:23:12
20 |  | UNDE6-A-1 | -3.522 | -54.363 | -4.5919 | -59.8938 | -1.612 | 177.781 | 2022/12/02 02:32:09
20 |  | UNDE6-A-1 | -3.527 | -54.059 | -4.5971 | -59.5631 | -1.394 | 178.95 | 2022/12/02 02:41:05
20 |  | UNDE6-A-1 | -3.507 | -53.887 | -4.5764 | -59.376 | -1.568 | 180.328 | 2022/12/02 02:50:02
Sample 20 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
20 |  | UNDE6-A-1 | -4.5996 | -59.8898 | -1.4512 | 113.4795 |  | 0.023274 | 0.54899 |  | -6.0718
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
21 |  | UNDE6-B-4 | -3.75 | -55.353 | -4.8269 | -60.9709 | -1.118 | 53.735 | 2022/12/02 02:59:50
21 |  | UNDE6-B-4 | -3.726 | -55.565 | -4.8021 | -61.2015 | -0.96 | 52.167 | 2022/12/02 03:08:46
21 |  | UNDE6-B-4 | -3.77 | -55.688 | -4.8475 | -61.3353 | -1.01 | 52.293 | 2022/12/02 03:17:42
21 |  | UNDE6-B-4 | -3.724 | -55.834 | -4.8001 | -61.4942 | -1.195 | 52.416 | 2022/12/02 03:26:38
Sample 21 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
21 |  | UNDE6-B-4 | -4.8191 | -61.2505 |  |  |  | 0.02181 | 0.20358
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
22 |  | EV | -8.7 | -65.266 | -9.928 | -71.7556 | -0.999 | 53.693 | 2022/12/02 03:36:23
22 |  | EV | -8.893 | -66.285 | -10.1269 | -72.8642 | -0.886 | 52.802 | 2022/12/02 03:45:19
22 |  | EV | -8.916 | -66.492 | -10.1506 | -73.0894 | -0.959 | 52.921 | 2022/12/02 03:54:15
22 |  | EV | -8.986 | -66.725 | -10.2227 | -73.3429 | -0.968 | 52.875 | 2022/12/02 04:03:11
Sample 22 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
22 |  | EV | -10.107 | -72.763 |  |  |  | 0.1224 | 0.64297
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
23 |  | UNDE6-A-5 | -3.455 | -56.143 | -4.5229 | -61.8304 | -1.475 | 54.072 | 2022/12/02 04:12:57
23 |  | UNDE6-A-5 | -3.296 | -55.427 | -4.359 | -61.0514 | -1.376 | 53.908 | 2022/12/02 04:21:54
23 |  | UNDE6-A-5 | -3.212 | -54.733 | -4.2724 | -60.2964 | -1.454 | 54.311 | 2022/12/02 04:30:49
23 |  | UNDE6-A-5 | -3.244 | -55.128 | -4.3054 | -60.7261 | -1.424 | 54.203 | 2022/12/02 04:39:45
Sample 23 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
23 |  | UNDE6-A-5 | -4.3649 | -60.9761 | -1.4322 |  |  | 0.10787 | 0.59568
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
24 |  | UNDE6-D-4 | -1.585 | -61.114 | -2.5958 | -67.2385 | -1.502 | 54.075 | 2022/12/02 04:49:32
24 |  | UNDE6-D-4 | -1.478 | -61.639 | -2.4855 | -67.8097 | -1.27 | 53.626 | 2022/12/02 04:58:28
24 |  | UNDE6-D-4 | -1.438 | -61.612 | -2.4443 | -67.7803 | -1.135 | 54.004 | 2022/12/02 05:07:24
24 |  | UNDE6-D-4 | -1.426 | -61.752 | -2.4319 | -67.9326 | -1.164 | 53.505 | 2022/12/02 05:16:21
Sample 24 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
24 |  | UNDE6-D-4 | -2.4894 | -67.6903 |  |  |  | 0.072334 | 0.2834
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
25 |  | UNDE6-C-4 | -4.721 | -60.91 | -5.8275 | -67.0166 | -1.488 | 353.262 | 2022/12/02 05:26:08
25 |  | UNDE6-C-4 | -5.031 | -61.468 | -6.147 | -67.6236 | -1.363 | 359.865 | 2022/12/02 05:35:03
25 |  | UNDE6-C-4 | -5.014 | -61.274 | -6.1294 | -67.4126 | -1.315 | 357.967 | 2022/12/02 05:44:00
25 |  | UNDE6-C-4 | -5.07 | -61.303 | -6.1872 | -67.4441 | -1.302 | 357.097 | 2022/12/02 05:52:56
Sample 25 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
25 |  | UNDE6-C-4 | -6.0728 | -67.3742 | -1.367 | 293.8063 |  | 0.16039 | 0.23523 |  | -9.5784
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
26 |  | UNDE6-3040-1 | -5.244 | -51.089 | -6.3665 | -56.3319 | -1.114 | 57.025 | 2022/12/02 06:02:43
26 |  | UNDE6-3040-1 | -5.223 | -50.093 | -6.3448 | -55.2483 | -1.231 | 54.164 | 2022/12/02 06:11:39
26 |  | UNDE6-3040-1 | -5.2 | -49.822 | -6.3211 | -54.9535 | -0.82 | 53.376 | 2022/12/02 06:20:34
26 |  | UNDE6-3040-1 | -5.2 | -49.439 | -6.3211 | -54.5368 | -1.001 | 53.454 | 2022/12/02 06:29:31
Sample 26 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
26 |  | UNDE6-3040-1 | -6.3384 | -55.2677 |  |  |  | 0.021156 | 0.7052
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
27 |  | UNDE6-1020-1 | -3.142 | -47.208 | -4.2003 | -52.1096 | -0.976 | 52.849 | 2022/12/02 06:39:17
27 |  | UNDE6-1020-1 | -3.121 | -47.125 | -4.1787 | -52.0193 | -0.757 | 52.818 | 2022/12/02 06:48:13
27 |  | UNDE6-1020-1 | -3.047 | -46.791 | -4.1024 | -51.656 | -0.912 | 52.826 | 2022/12/02 06:57:09
27 |  | UNDE6-1020-1 | -3.078 | -46.829 | -4.1343 | -51.6973 | -1.001 | 52.467 | 2022/12/02 07:06:05
Sample 27 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
27 |  | UNDE6-1020-1 | -4.1539 | -51.8706 |  |  |  | 0.042669 | 0.20917
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
28 |  | UNDE6-B-3 | -2.742 | -54.752 | -3.7881 | -60.317 | -2.346 | 192.856 | 2022/12/02 07:15:52
28 |  | UNDE6-B-3 | -2.828 | -55.596 | -3.8767 | -61.2353 | -2.141 | 191.34 | 2022/12/02 07:24:49
28 |  | UNDE6-B-3 | -2.846 | -56.053 | -3.8953 | -61.7324 | -2.306 | 193.739 | 2022/12/02 07:33:45
28 |  | UNDE6-B-3 | -2.884 | -56.223 | -3.9344 | -61.9174 | -2.247 | 195.664 | 2022/12/02 07:42:41
Sample 28 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
28 |  | UNDE6-B-3 | -3.8736 | -61.3005 | -2.26 | 130.1583 |  | 0.060056 | 0.65826 |  | -7.5606
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
29 |  | UNDE6-3040-2 | -5.827 | -55.832 | -6.9673 | -61.492 | -1.05 | 54.868 | 2022/12/02 07:52:25
29 |  | UNDE6-3040-2 | -5.828 | -55.204 | -6.9683 | -60.8088 | -1.091 | 53.01 | 2022/12/02 08:01:21
29 |  | UNDE6-3040-2 | -5.816 | -55.309 | -6.9559 | -60.923 | -0.973 | 53.703 | 2022/12/02 08:10:16
29 |  | UNDE6-3040-2 | -5.918 | -55.445 | -7.061 | -61.071 | -0.769 | 53.497 | 2022/12/02 08:19:11
Sample 29 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
29 |  | UNDE6-3040-2 | -6.9881 | -61.0737 |  |  |  | 0.047479 | 0.27466
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
30 |  | UNDE6-C-5 | -2.102 | -56.55 | -3.1285 | -62.2731 | -1.112 | 53.836 | 2022/12/02 08:28:55
30 |  | UNDE6-C-5 | -1.949 | -56.768 | -2.9709 | -62.5103 | -1.299 | 54.642 | 2022/12/02 08:37:51
30 |  | UNDE6-C-5 | -1.84 | -56.174 | -2.8586 | -61.8641 | -1.142 | 54.117 | 2022/12/02 08:46:46
30 |  | UNDE6-C-5 | -1.903 | -56.736 | -2.9235 | -62.4755 | -1.139 | 54.967 | 2022/12/02 08:55:41
Sample 30 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
30 |  | UNDE6-C-5 | -2.9704 | -62.2808 |  |  |  | 0.11166 | 0.27282
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
31 |  | EV | -8.618 | -65.528 | -9.8435 | -72.0406 | -1.067 | 53.419 | 2022/12/02 09:05:27
31 |  | EV | -8.832 | -66.347 | -10.064 | -72.9317 | -0.861 | 52.891 | 2022/12/02 09:14:23
31 |  | EV | -8.947 | -66.773 | -10.1825 | -73.3951 | -0.994 | 52.736 | 2022/12/02 09:23:19
31 |  | EV | -8.923 | -66.332 | -10.1578 | -72.9153 | -0.954 | 52.638 | 2022/12/02 09:32:15
Sample 31 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
31 |  | EV | -10.0619 | -72.8207 |  |  |  | 0.14976 | 0.51989
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
32 |  | UNDE6-B-1 | -3.907 | -56.875 | -4.9887 | -62.6267 | -0.946 | 51.534 | 2022/12/02 09:42:01
32 |  | UNDE6-B-1 | -3.661 | -55.672 | -4.7351 | -61.3179 | -1.213 | 52.883 | 2022/12/02 09:50:56
32 |  | UNDE6-B-1 | -3.625 | -55.581 | -4.698 | -61.2189 | -1.143 | 52.627 | 2022/12/02 09:59:52
32 |  | UNDE6-B-1 | -3.609 | -55.409 | -4.6816 | -61.0318 | -1.142 | 52.922 | 2022/12/02 10:08:47
Sample 32 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
32 |  | UNDE6-B-1 | -4.7758 | -61.5489 |  |  |  | 0.13937 | 0.66944
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
33 |  | UNDE6-E-4 | -0.803 | -45.759 | -1.7899 | -50.5332 | -1.048 | 49.006 | 2022/12/02 10:18:33
33 |  | UNDE6-E-4 | -0.618 | -44.528 | -1.5992 | -49.194 | -0.823 | 51.283 | 2022/12/02 10:27:28
33 |  | UNDE6-E-4 | -0.617 | -44.204 | -1.5982 | -48.8415 | -0.957 | 51.785 | 2022/12/02 10:36:23
33 |  | UNDE6-E-4 | -0.61 | -44.173 | -1.591 | -48.8078 | -1.363 | 52.246 | 2022/12/02 10:45:19
Sample 33 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
33 |  | UNDE6-E-4 | -1.6446 | -49.3441 |  |  |  | 0.094067 | 0.74614
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
34 |  | UNDE6-A-2 | -4.408 | -57.993 | -5.5049 | -63.843 | -1.225 | 52.666 | 2022/12/02 10:55:06
34 |  | UNDE6-A-2 | -4.667 | -59.737 | -5.7719 | -65.7404 | -0.888 | 52.447 | 2022/12/02 11:04:02
34 |  | UNDE6-A-2 | -4.682 | -60.015 | -5.7873 | -66.0428 | -0.935 | 52.217 | 2022/12/02 11:12:59
34 |  | UNDE6-A-2 | -4.679 | -60.133 | -5.7842 | -66.1712 | -1.286 | 52.359 | 2022/12/02 11:21:55
Sample 34 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
34 |  | UNDE6-A-2 | -5.7121 | -65.4494 |  |  |  | 0.13416 | 0.99823
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
35 |  | UNDE6-3040-3 | -4.742 | -49.42 | -5.8491 | -54.5162 | -1.281 | 52.379 | 2022/12/02 11:31:40
35 |  | UNDE6-3040-3 | -4.678 | -48.205 | -5.7832 | -53.1943 | -0.936 | 52.773 | 2022/12/02 11:40:35
35 |  | UNDE6-3040-3 | -4.721 | -48 | -5.8275 | -52.9713 | -1.05 | 52.649 | 2022/12/02 11:49:31
35 |  | UNDE6-3040-3 | -4.752 | -48.381 | -5.8594 | -53.3858 | -0.99 | 52.261 | 2022/12/02 11:58:27
Sample 35 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
35 |  | UNDE6-3040-3 | -5.8298 | -53.5169 |  |  |  | 0.032816 | 0.63182
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
36 |  | UNDE6-B-5 | -2.466 | -44.616 | -3.5037 | -49.2897 | -1.217 | 53.005 | 2022/12/02 12:08:10
36 |  | UNDE6-B-5 | -2.44 | -44.852 | -3.4769 | -49.5465 | -1.014 | 52.476 | 2022/12/02 12:17:06
36 |  | UNDE6-B-5 | -2.387 | -44.473 | -3.4222 | -49.1341 | -1.222 | 52.489 | 2022/12/02 12:26:01
36 |  | UNDE6-B-5 | -2.405 | -44.262 | -3.4408 | -48.9046 | -1.172 | 51.792 | 2022/12/02 12:34:57
Sample 36 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
36 |  | UNDE6-B-5 | -3.4609 | -49.2187 |  |  |  | 0.035351 | 0.24795
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
37 |  | UNDE6-A-4 | -2.761 | -60.688 | -3.8077 | -66.775 | -2.194 | 62.273 | 2022/12/02 12:44:41
37 |  | UNDE6-A-4 | -2.877 | -62.313 | -3.9272 | -68.5429 | -2.09 | 63.466 | 2022/12/02 12:53:36
37 |  | UNDE6-A-4 | -2.901 | -62.904 | -3.9519 | -69.1859 | -2.191 | 63.753 | 2022/12/02 13:02:32
37 |  | UNDE6-A-4 | -2.898 | -63.134 | -3.9488 | -69.4361 | -2.134 | 63.58 | 2022/12/02 13:11:28
Sample 37 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
37 |  | UNDE6-A-4 | -3.9089 | -68.485 | -2.1522 |  |  | 0.066365 | 1.1034
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
38 |  | UNDE6-D-1 | -5.278 | -59.153 | -6.4015 | -65.105 | -0.87 | 52.056 | 2022/12/02 13:21:11
38 |  | UNDE6-D-1 | -5.353 | -59.096 | -6.4788 | -65.043 | -1.066 | 52.969 | 2022/12/02 13:30:07
38 |  | UNDE6-D-1 | -5.369 | -58.655 | -6.4953 | -64.5633 | -1.212 | 52.415 | 2022/12/02 13:39:03
38 |  | UNDE6-D-1 | -5.373 | -58.759 | -6.4994 | -64.6764 | -1.001 | 52.56 | 2022/12/02 13:47:59
Sample 38 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
38 |  | UNDE6-D-1 | -6.4687 | -64.8469 |  |  |  | 0.04435 | 0.24586
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
39 |  | UNDE6-E-5 | -4.107 | -57.613 | -5.1948 | -63.4296 | -1.043 | 51.273 | 2022/12/02 13:57:43
39 |  | UNDE6-E-5 | -4.084 | -57.826 | -5.1711 | -63.6614 | -1.028 | 51.409 | 2022/12/02 14:06:39
39 |  | UNDE6-E-5 | -4.037 | -57.819 | -5.1226 | -63.6537 | -1.061 | 51.353 | 2022/12/02 14:15:34
39 |  | UNDE6-E-5 | -4.015 | -57.699 | -5.0999 | -63.5232 | -1.103 | 51.591 | 2022/12/02 14:24:30
Sample 39 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
39 |  | UNDE6-E-5 | -5.1471 | -63.567 |  |  |  | 0.042177 | 0.10238
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
40 |  | EV | -8.67 | -65.503 | -9.8971 | -72.0134 | -1.123 | 53.379 | 2022/12/02 14:34:14
40 |  | EV | -8.916 | -66.067 | -10.1506 | -72.627 | -0.941 | 52.343 | 2022/12/02 14:43:09
40 |  | EV | -8.922 | -66.326 | -10.1567 | -72.9088 | -1.054 | 52.619 | 2022/12/02 14:52:04
40 |  | EV | -9.014 | -66.701 | -10.2516 | -73.3168 | -1.111 | 52.461 | 2022/12/02 15:01:00
Sample 40 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
40 |  | EV | -10.114 | -72.7165 |  |  |  | 0.14733 | 0.50335
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
41 |  | UNDE6-1020-2 | -5.863 | -52.308 | -7.0044 | -57.6581 | -0.782 | 50.663 | 2022/12/02 15:10:45
41 |  | UNDE6-1020-2 | -5.69 | -50.715 | -6.8261 | -55.925 | -0.888 | 50.943 | 2022/12/02 15:19:41
41 |  | UNDE6-1020-2 | -5.694 | -50.587 | -6.8302 | -55.7858 | -1.092 | 50.999 | 2022/12/02 15:28:38
41 |  | UNDE6-1020-2 | -5.633 | -50.019 | -6.7673 | -55.1678 | -0.893 | 51.199 | 2022/12/02 15:37:33
Sample 41 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
41 |  | UNDE6-1020-2 | -6.857 | -56.1342 |  |  |  | 0.099321 | 0.9816
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
42 |  | UNDE6-A-3 | -3.6 | -54.468 | -4.6723 | -60.0081 | -1.159 | 52.181 | 2022/12/02 15:47:17
42 |  | UNDE6-A-3 | -3.337 | -54.536 | -4.4013 | -60.082 | -0.853 | 53.213 | 2022/12/02 15:56:13
42 |  | UNDE6-A-3 | -3.487 | -55.247 | -4.5558 | -60.8556 | -0.983 | 51.902 | 2022/12/02 16:05:09
42 |  | UNDE6-A-3 | -3.53 | -55.274 | -4.6001 | -60.8849 | -1.11 | 52.239 | 2022/12/02 16:14:05
Sample 42 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
42 |  | UNDE6-A-3 | -4.5574 | -60.4577 |  |  |  | 0.11122 | 0.43894
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
43 |  | UNDE6-1020-3 | -4.149 | -49.546 | -5.238 | -54.6532 | -0.906 | 51.923 | 2022/12/02 16:23:49
43 |  | UNDE6-1020-3 | -4.162 | -48.912 | -5.2514 | -53.9635 | -1.119 | 52.8 | 2022/12/02 16:32:45
43 |  | UNDE6-1020-3 | -4.141 | -48.896 | -5.2298 | -53.9461 | -1.107 | 52.198 | 2022/12/02 16:41:41
43 |  | UNDE6-1020-3 | -4.158 | -48.798 | -5.2473 | -53.8395 | -1.009 | 52.406 | 2022/12/02 16:50:37
Sample 43 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
43 |  | UNDE6-1020-3 | -5.2416 | -54.1006 |  |  |  | 0.0093986 | 0.3424
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
44 |  | UNDE6-010-1 | -2.907 | -45.2 | -3.9581 | -49.9251 | -0.831 | 51.185 | 2022/12/02 17:00:21
44 |  | UNDE6-010-1 | -2.913 | -45.263 | -3.9643 | -49.9936 | -1.142 | 51.654 | 2022/12/02 17:09:17
44 |  | UNDE6-010-1 | -2.849 | -44.853 | -3.8984 | -49.5476 | -1.047 | 51.251 | 2022/12/02 17:18:12
44 |  | UNDE6-010-1 | -2.853 | -44.906 | -3.9025 | -49.6052 | -1.108 | 51.358 | 2022/12/02 17:27:07
Sample 44 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
44 |  | UNDE6-010-1 | -3.9308 | -49.7679 |  |  |  | 0.034191 | 0.20599
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
45 |  | UNDE6-010-3 | -2.362 | -45.058 | -3.3965 | -49.7706 | -1.123 | 51.182 | 2022/12/02 17:36:51
45 |  | UNDE6-010-3 | -2.318 | -44.958 | -3.3511 | -49.6618 | -1.085 | 51.863 | 2022/12/02 17:45:46
45 |  | UNDE6-010-3 | -2.346 | -44.844 | -3.38 | -49.5378 | -1.118 | 51.878 | 2022/12/02 17:54:42
45 |  | UNDE6-010-3 | -2.244 | -44.69 | -3.2749 | -49.3702 | -0.961 | 51.771 | 2022/12/02 18:03:37
Sample 45 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
45 |  | UNDE6-010-3 | -3.3506 | -49.5851 |  |  |  | 0.052265 | 0.15805
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
46 |  | UNDE6-deep-1 | -0.49 | -43.088 | -1.4673 | -47.6273 | -1.108 | 52.829 | 2022/12/02 18:13:23
46 |  | UNDE6-deep-1 | -0.336 | -42.66 | -1.3086 | -47.1617 | -1.135 | 52.311 | 2022/12/02 18:22:19
46 |  | UNDE6-deep-1 | -0.36 | -42.799 | -1.3334 | -47.3129 | -1.141 | 52.176 | 2022/12/02 18:31:13
46 |  | UNDE6-deep-1 | -0.291 | -42.373 | -1.2623 | -46.8495 | -1.059 | 52.179 | 2022/12/02 18:40:09
Sample 46 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
46 |  | UNDE6-deep-1 | -1.3429 | -47.2379 |  |  |  | 0.08543 | 0.29736
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
47 |  | UNDE6-deep-2 | -3.453 | -49.154 | -4.5208 | -54.2268 | -0.874 | 52.127 | 2022/12/02 18:49:53
47 |  | UNDE6-deep-2 | -3.636 | -49.829 | -4.7094 | -54.9611 | -1.089 | 52.341 | 2022/12/02 18:58:49
47 |  | UNDE6-deep-2 | -3.631 | -50.242 | -4.7042 | -55.4104 | -1.053 | 51.662 | 2022/12/02 19:07:45
47 |  | UNDE6-deep-2 | -3.674 | -50.086 | -4.7485 | -55.2407 | -0.941 | 51.857 | 2022/12/02 19:16:40
Sample 47 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
47 |  | UNDE6-deep-2 | -4.6707 | -54.9598 |  |  |  | 0.098882 | 0.48036
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
48 |  | UNDE6-deep-3 | -2.397 | -42.692 | -3.4326 | -47.1965 | -1.017 | 52.117 | 2022/12/02 19:26:25
48 |  | UNDE6-deep-3 | -2.133 | -41.206 | -3.1605 | -45.5799 | -0.983 | 53.678 | 2022/12/02 19:35:21
48 |  | UNDE6-deep-3 | -2.315 | -41.345 | -3.3481 | -45.7311 | -1.158 | 52.073 | 2022/12/02 19:44:15
48 |  | UNDE6-deep-3 | -2.37 | -41.729 | -3.4047 | -46.1488 | -1.229 | 53.334 | 2022/12/02 19:53:10
Sample 48 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
48 |  | UNDE6-deep-3 | -3.3365 | -46.1641 |  |  |  | 0.11884 | 0.67022
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
49 |  | EV | -8.57 | -63.388 | -9.794 | -69.7125 | -0.903 | 53.059 | 2022/12/02 20:02:56
49 |  | EV | -8.842 | -65.386 | -10.0743 | -71.8862 | -1.018 | 52.413 | 2022/12/02 20:11:51
49 |  | EV | -8.919 | -65.768 | -10.1537 | -72.3017 | -1.029 | 52.529 | 2022/12/02 20:20:47
49 |  | EV | -8.904 | -66.263 | -10.1382 | -72.8403 | -0.996 | 52.316 | 2022/12/02 20:29:42
Sample 49 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
49 |  | EV | -10.04 | -71.6852 |  |  |  | 0.16262 | 1.261

## Source
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Line | Analysis | Time Code | Port | Inj Nr | d(18_16)Mean | d(D_H)Mean | H2O_Mean | Ignore | Good | Identifier 1 | Identifier 2 | Gas Configuration | Timestamp Mean | d(18_16)_SD | d(D_H)_SD | H2O_SD | d(18_16)_Sl | d(D_H)_Sl | H2O_Sl | baseline_shift | slope_shift | residuals | baseline_curvature | interval | ch4_ppm | h16od_adjust | h16od_shift | n2_flag | DAS Temp | Tray | Sample | Job | Method | Error Code
1 | P-22639 | 2022/12/01 11:08:44 | 1-01 | 1 | -9.875 | -75.572 | 14843 | -1 | 1 | EV | standard | H2O | 1669918429.62 | 0.202 | 0.819 | 65.922 | 0 | -0.001 | 1.608 | 322.041 | 26.997 | 2.523 | 67.677 | 0.857 | -1.159 | 0 | 0 | 1 | 43.438 | 1 | 1 | 1 | GJB_Rinse | 0
2 | P-22639 | 2022/12/01 11:17:40 | 1-01 | 2 | -9.474 | -71.771 | 14046 | -1 | 1 | EV | standard | H2O | 1669918965.83 | 0.197 | 0.774 | 64.994 | 0 | -0.004 | 1.583 | 334.503 | 24.129 | 2.402 | 64.709 | 0.857 | -1.196 | 0 | 0 | 1 | 42.562 | 1 | 1 | 1 | GJB_Rinse | 0
3 | P-22639 | 2022/12/01 11:26:37 | 1-01 | 3 | -9.258 | -69.755 | 13235 | -1 | 1 | EV | standard | H2O | 1669919501.9 | 0.22 | 0.73 | 64.901 | 0.001 | 0 | 1.594 | 157.32 | 27.73 | 2.253 | 63.666 | 0.856 | -1.072 | 0 | 0 | 1 | 42.125 | 1 | 1 | 1 | GJB_Rinse | 0
4 | P-22639 | 2022/12/01 11:35:32 | 1-01 | 4 | -9.183 | -68.948 | 13272 | 0 | 1 | EV | standard | H2O | 1669920037.89 | 0.199 | 0.837 | 60.703 | 0.001 | -0.001 | 1.49 | 104.251 | 29.732 | 2.253 | 64.403 | 0.856 | -0.996 | 0 | 0 | 1 | 42.125 | 1 | 1 | 1 | GJB_Rinse | 0
5 | P-22639 | 2022/12/01 11:44:28 | 1-01 | 5 | -9.176 | -68.891 | 13252 | 0 | 1 | EV | standard | H2O | 1669920574.06 | 0.209 | 0.789 | 61.657 | 0.001 | -0.002 | 1.509 | 130.925 | 28.367 | 2.238 | 64.702 | 0.856 | -0.944 | 0 | 0 | 1 | 42.125 | 1 | 1 | 1 | GJB_Rinse | 0
6 | P-22639 | 2022/12/01 11:53:25 | 1-01 | 6 | -9.107 | -68.321 | 13046 | 0 | 1 | EV | standard | H2O | 1669921110.37 | 0.215 | 0.841 | 66.401 | 0 | 0.002 | 1.633 | 73.958 | 29.175 | 2.215 | 65.36 | 0.855 | -0.988 | 0 | 0 | 1 | 42.312 | 1 | 1 | 1 | GJB_Rinse | 0
7 | P-22639 | 2022/12/01 12:02:21 | 1-01 | 7 | -9.126 | -68.064 | 13088 | 0 | 1 | EV | standard | H2O | 1669921646.16 | 0.189 | 0.79 | 65.804 | 0.001 | -0.001 | 1.615 | 68.608 | 29.965 | 2.22 | 65.276 | 0.855 | -0.966 | 0 | 0 | 1 | 42.375 | 1 | 1 | 1 | GJB_Rinse | 0
8 | P-22639 | 2022/12/01 12:11:17 | 1-01 | 8 | -9.074 | -67.978 | 13038 | 0 | 1 | EV | standard | H2O | 1669922182.36 | 0.224 | 0.837 | 61.922 | 0.002 | -0.001 | 1.516 | 65.311 | 29.905 | 2.213 | 65.028 | 0.856 | -0.974 | 0 | 0 | 1 | 42.312 | 1 | 1 | 1 | GJB_Rinse | 0
9 | P-22639 | 2022/12/01 12:20:13 | 1-01 | 9 | -9.043 | -67.716 | 13014 | 0 | 1 | EV | standard | H2O | 1669922718.32 | 0.197 | 0.675 | 62.906 | 0.001 | -0.002 | 1.544 | 63.311 | 30.081 | 2.208 | 65.258 | 0.856 | -0.793 | 0 | 0 | 1 | 42.25 | 1 | 1 | 1 | GJB_Rinse | 0
10 | P-22639 | 2022/12/01 12:29:09 | 1-01 | 10 | -9.082 | -67.968 | 12662 | 0 | 1 | EV | standard | H2O | 1669923254.41 | 0.189 | 0.742 | 63.3 | 0.001 | -0.002 | 1.558 | 61.662 | 28.873 | 2.146 | 63.183 | 0.856 | -0.996 | 0 | 0 | 1 | 42.312 | 1 | 1 | 1 | GJB_Rinse | 0
11 | P-22640 | 2022/12/01 12:38:57 | 1-02 | 1 | 2.239 | 8.309 | 13087 | -1 | 1 | PZ | standard | H2O | 1669923842.37 | 0.206 | 0.777 | 63.597 | 0 | -0.002 | 1.562 | 56.725 | 30.507 | 2.346 | 68.285 | 0.856 | -1.028 | 0 | 0 | 1 | 42.25 | 1 | 2 | 1 | GJB_Rinse | 0
12 | P-22640 | 2022/12/01 12:47:53 | 1-02 | 2 | 2.687 | 15.198 | 12879 | -1 | 1 | PZ | standard | H2O | 1669924377.86 | 0.204 | 0.834 | 61.803 | 0 | -0.007 | 1.525 | 61.151 | 29.724 | 2.316 | 67.105 | 0.856 | -1.026 | 0 | 0 | 1 | 42.312 | 1 | 2 | 1 | GJB_Rinse | 0
13 | P-22640 | 2022/12/01 12:56:49 | 1-02 | 3 | 2.769 | 16.918 | 12736 | -1 | 1 | PZ | standard | H2O | 1669924914.39 | 0.191 | 0.854 | 62.685 | 0.001 | -0.005 | 1.54 | 62.387 | 29.303 | 2.288 | 66.257 | 0.856 | -1.118 | 0 | 0 | 1 | 42.5 | 1 | 2 | 1 | GJB_Rinse | 0
14 | P-22640 | 2022/12/01 13:05:45 | 1-02 | 4 | 2.883 | 18.143 | 12741 | 0 | 1 | PZ | standard | H2O | 1669925450.61 | 0.231 | 0.908 | 62.403 | 0.001 | -0.007 | 1.528 | 61.23 | 29.227 | 2.303 | 65.436 | 0.856 | -1.221 | 0 | 0 | 1 | 42.562 | 1 | 2 | 1 | GJB_Rinse | 0
15 | P-22640 | 2022/12/01 13:14:41 | 1-02 | 5 | 2.846 | 18.353 | 12776 | 0 | 1 | PZ | standard | H2O | 1669925986.08 | 0.215 | 0.907 | 60.823 | 0.001 | -0.006 | 1.504 | 60.994 | 29.218 | 2.306 | 66.605 | 0.856 | -0.906 | 0 | 0 | 1 | 42.625 | 1 | 2 | 1 | GJB_Rinse | 0
16 | P-22640 | 2022/12/01 13:23:38 | 1-02 | 6 | 2.873 | 18.732 | 13002 | 0 | 1 | PZ | standard | H2O | 1669926523.79 | 0.223 | 0.923 | 59.273 | 0 | -0.004 | 1.449 | 61.493 | 30.082 | 2.347 | 67.42 | 0.857 | -1.009 | 0 | 0 | 1 | 42.625 | 1 | 2 | 1 | GJB_Rinse | 0
17 | P-22640 | 2022/12/01 13:32:34 | 1-02 | 7 | 2.926 | 19.289 | 12778 | 0 | 1 | PZ | standard | H2O | 1669927060.12 | 0.189 | 0.769 | 62.923 | 0 | -0.005 | 1.555 | 60.963 | 29.477 | 2.325 | 66.577 | 0.856 | -1.069 | 0 | 0 | 1 | 42.625 | 1 | 2 | 1 | GJB_Rinse | 0
18 | P-22640 | 2022/12/01 13:41:31 | 1-02 | 8 | 2.923 | 19.124 | 12753 | 0 | 1 | PZ | standard | H2O | 1669927596.17 | 0.207 | 0.792 | 57.641 | 0 | -0.003 | 1.425 | 61.254 | 29.334 | 2.303 | 66.221 | 0.856 | -1.063 | 0 | 0 | 1 | 42.5 | 1 | 2 | 1 | GJB_Rinse | 0
19 | P-22640 | 2022/12/01 13:50:28 | 1-02 | 9 | 2.947 | 19.488 | 12808 | 0 | 1 | PZ | standard | H2O | 1669928133.1 | 0.196 | 0.805 | 63.748 | 0 | -0.002 | 1.576 | 61.744 | 29.503 | 2.306 | 66.162 | 0.856 | -1.118 | 0 | 0 | 1 | 42.5 | 1 | 2 | 1 | GJB_Rinse | 0
20 | P-22640 | 2022/12/01 13:59:23 | 1-02 | 10 | 2.974 | 19.713 | 12818 | 0 | 1 | PZ | standard | H2O | 1669928668.81 | 0.216 | 0.821 | 61.368 | 0.002 | -0.003 | 1.51 | 61.811 | 29.644 | 2.318 | 66.331 | 0.856 | -1.014 | 0 | 0 | 1 | 42.562 | 1 | 2 | 1 | GJB_Rinse | 0
21 | P-22641 | 2022/12/01 14:09:12 | 1-03 | 1 | -13.561 | -95.448 | 13000 | -1 | 1 | UT2 | standard | H2O | 1669929257.55 | 0.198 | 0.669 | 62.805 | 0.002 | -0.002 | 1.541 | 57.918 | 29.915 | 2.16 | 64.692 | 0.855 | -0.887 | 0 | 0 | 1 | 42.688 | 1 | 3 | 1 | GJB_Rinse | 0
22 | P-22641 | 2022/12/01 14:18:08 | 1-03 | 2 | -14.282 | -105.739 | 12735 | -1 | 1 | UT2 | standard | H2O | 1669929793.79 | 0.206 | 0.736 | 58.65 | 0.002 | 0.004 | 1.437 | 64.733 | 29.066 | 2.099 | 62.328 | 0.856 | -0.998 | 0 | 0 | 1 | 42.562 | 1 | 3 | 1 | GJB_Rinse | 0
23 | P-22641 | 2022/12/01 14:27:05 | 1-03 | 3 | -14.497 | -108.726 | 13049 | -1 | 1 | UT2 | standard | H2O | 1669930330.2 | 0.207 | 0.708 | 61.025 | 0.001 | 0.007 | 1.493 | 66.678 | 29.778 | 2.143 | 63.115 | 0.856 | -1.112 | 0 | 0 | 1 | 42.562 | 1 | 3 | 1 | GJB_Rinse | 0
24 | P-22641 | 2022/12/01 14:36:01 | 1-03 | 4 | -14.482 | -109.713 | 13068 | 0 | 1 | UT2 | standard | H2O | 1669930866.22 | 0.206 | 0.764 | 61.977 | 0.001 | 0.005 | 1.519 | 66.492 | 29.882 | 2.157 | 64.694 | 0.856 | -0.792 | 0 | 0 | 1 | 42.625 | 1 | 3 | 1 | GJB_Rinse | 0
25 | P-22641 | 2022/12/01 14:44:57 | 1-03 | 5 | -14.55 | -110.198 | 13018 | 0 | 1 | UT2 | standard | H2O | 1669931402.75 | 0.172 | 0.777 | 62.077 | 0.001 | 0.003 | 1.511 | 66.397 | 29.478 | 2.153 | 64.466 | 0.856 | -0.88 | 0 | 0 | 1 | 42.625 | 1 | 3 | 1 | GJB_Rinse | 0
26 | P-22641 | 2022/12/01 14:53:53 | 1-03 | 6 | -14.588 | -110.9 | 12926 | 0 | 1 | UT2 | standard | H2O | 1669931938.82 | 0.214 | 0.748 | 61.491 | 0.002 | 0.004 | 1.502 | 65.23 | 29.348 | 2.133 | 63.507 | 0.856 | -0.905 | 0 | 0 | 1 | 42.625 | 1 | 3 | 1 | GJB_Rinse | 0
27 | P-22641 | 2022/12/01 15:02:50 | 1-03 | 7 | -14.593 | -111.092 | 13023 | 0 | 1 | UT2 | standard | H2O | 1669932475.25 | 0.205 | 0.774 | 63.046 | 0 | 0.005 | 1.551 | 64.931 | 29.739 | 2.138 | 63.664 | 0.856 | -0.997 | 0 | 0 | 1 | 42.688 | 1 | 3 | 1 | GJB_Rinse | 0
28 | P-22641 | 2022/12/01 15:11:47 | 1-03 | 8 | -14.611 | -111.537 | 13071 | 0 | 1 | UT2 | standard | H2O | 1669933012.28 | 0.19 | 0.727 | 61.973 | 0.001 | 0 | 1.521 | 65.122 | 29.936 | 2.155 | 64.359 | 0.856 | -0.815 | 0 | 0 | 1 | 42.562 | 1 | 3 | 1 | GJB_Rinse | 0
29 | P-22641 | 2022/12/01 15:20:43 | 1-03 | 9 | -14.604 | -111.624 | 12930 | 0 | 1 | UT2 | standard | H2O | 1669933548.63 | 0.19 | 0.705 | 62.514 | 0.001 | 0.003 | 1.541 | 65.529 | 29.514 | 2.127 | 62.912 | 0.856 | -0.895 | 0 | 0 | 1 | 42.5 | 1 | 3 | 1 | GJB_Rinse | 0
30 | P-22641 | 2022/12/01 15:29:40 | 1-03 | 10 | -14.601 | -111.94 | 13036 | 0 | 1 | UT2 | standard | H2O | 1669934084.94 | 0.186 | 0.805 | 63.643 | 0 | 0.004 | 1.563 | 72.047 | 29.72 | 2.138 | 63.151 | 0.856 | -1.094 | 0 | 0 | 1 | 42.562 | 1 | 3 | 1 | GJB_Rinse | 0
31 | P-22642 | 2022/12/01 15:39:27 | 1-04 | 1 | -9.327 | -72.374 | 13120 | -1 | 1 | EV | standard | H2O | 1669934672.79 | 0.198 | 0.724 | 62.528 | 0.001 | 0.003 | 1.545 | 56.074 | 30.534 | 2.21 | 65.645 | 0.856 | -1.07 | 0 | 0 | 1 | 42.625 | 1 | 4 | 1 | GJB_Rinse | 0
32 | P-22642 | 2022/12/01 15:48:23 | 1-04 | 2 | -9.156 | -69.081 | 12884 | -1 | 1 | EV | standard | H2O | 1669935209.03 | 0.204 | 0.887 | 61.938 | 0 | 0.001 | 1.526 | 61.077 | 29.425 | 2.179 | 64.757 | 0.856 | -0.924 | 0 | 0 | 1 | 42.5 | 1 | 4 | 1 | GJB_Rinse | 0
33 | P-22642 | 2022/12/01 15:57:19 | 1-04 | 3 | -9.052 | -68.062 | 12962 | -1 | 1 | EV | standard | H2O | 1669935744.53 | 0.21 | 0.834 | 62.82 | 0.001 | -0.001 | 1.539 | 62.713 | 29.997 | 2.179 | 63.263 | 0.856 | -1.178 | 0 | 0 | 1 | 42.438 | 1 | 4 | 1 | GJB_Rinse | 0
34 | P-22642 | 2022/12/01 16:06:16 | 1-04 | 4 | -9.074 | -67.733 | 12924 | 0 | 1 | EV | standard | H2O | 1669936280.97 | 0.219 | 0.662 | 61.487 | 0.001 | 0.002 | 1.518 | 64.375 | 29.529 | 2.199 | 65.139 | 0.856 | -1.018 | 0 | 0 | 1 | 42.5 | 1 | 4 | 1 | GJB_Rinse | 0
35 | P-22642 | 2022/12/01 16:15:13 | 1-04 | 5 | -9.101 | -68.034 | 12707 | 0 | 1 | EV | standard | H2O | 1669936818.14 | 0.197 | 0.811 | 63.111 | 0.001 | 0.001 | 1.556 | 65.207 | 28.985 | 2.154 | 63.903 | 0.856 | -0.849 | 0 | 0 | 1 | 42.5 | 1 | 4 | 1 | GJB_Rinse | 0
36 | P-22642 | 2022/12/01 16:24:09 | 1-04 | 6 | -9.048 | -67.499 | 12698 | 0 | 1 | EV | standard | H2O | 1669937354.56 | 0.198 | 0.77 | 65.634 | 0.001 | 0.003 | 1.612 | 64.828 | 28.812 | 2.155 | 63.751 | 0.856 | -0.98 | 0 | 0 | 1 | 42.625 | 1 | 4 | 1 | GJB_Rinse | 0
37 | P-22642 | 2022/12/01 16:33:05 | 1-04 | 7 | -9.08 | -67.604 | 12755 | 0 | 1 | EV | standard | H2O | 1669937890.68 | 0.203 | 0.779 | 63.629 | 0.001 | -0.001 | 1.56 | 66.553 | 28.933 | 2.159 | 63.862 | 0.856 | -0.89 | 0 | 0 | 1 | 42.75 | 1 | 4 | 1 | GJB_Rinse | 0
38 | P-22642 | 2022/12/01 16:42:02 | 1-04 | 8 | -9.035 | -67.439 | 12757 | 0 | 1 | EV | standard | H2O | 1669938427.58 | 0.207 | 0.743 | 65.398 | 0.001 | 0.001 | 1.616 | 65.658 | 28.979 | 2.161 | 63.72 | 0.856 | -1.117 | 0 | 0 | 1 | 42.562 | 1 | 4 | 1 | GJB_Rinse | 0
39 | P-22642 | 2022/12/01 16:50:58 | 1-04 | 9 | -9.012 | -67.402 | 12754 | 0 | 1 | EV | standard | H2O | 1669938963.89 | 0.179 | 0.756 | 64.295 | 0.001 | -0.002 | 1.576 | 64.68 | 29.017 | 2.168 | 63.824 | 0.856 | -0.926 | 0 | 0 | 1 | 42.188 | 1 | 4 | 1 | GJB_Rinse | 0
40 | P-22642 | 2022/12/01 16:59:54 | 1-04 | 10 | -9.015 | -67.24 | 12749 | 0 | 1 | EV | standard | H2O | 1669939499.67 | 0.198 | 0.74 | 62.847 | 0.001 | 0.002 | 1.544 | 63.787 | 29.178 | 2.148 | 63.592 | 0.856 | -1.143 | 0 | 0 | 1 | 42.188 | 1 | 4 | 1 | GJB_Rinse | 0
41 | P-22643 | 2022/12/01 17:09:44 | 1-05 | 1 | -5.591 | -55.307 | 12810 | -1 | 1 | UNDE6-D-2 | 22-284 | H2O | 1669940089.73 | 0.192 | 0.742 | 60.352 | 0.001 | 0 | 1.487 | 54.057 | 29.224 | 2.201 | 65.197 | 0.856 | -0.88 | 0 | 0 | 1 | 42.25 | 1 | 5 | 1 | GJB_Rinse | 0
42 | P-22643 | 2022/12/01 17:18:41 | 1-05 | 2 | -5.456 | -53.775 | 12681 | -1 | 1 | UNDE6-D-2 | 22-284 | H2O | 1669940626.93 | 0.212 | 0.808 | 63.952 | 0 | 0 | 1.574 | 55.258 | 28.971 | 2.176 | 65.141 | 0.856 | -0.893 | 0 | 0 | 1 | 42.25 | 1 | 5 | 1 | GJB_Rinse | 0
43 | P-22643 | 2022/12/01 17:27:38 | 1-05 | 3 | -5.322 | -53.214 | 12649 | -1 | 1 | UNDE6-D-2 | 22-284 | H2O | 1669941163.73 | 0.183 | 0.828 | 62.078 | 0.001 | -0.003 | 1.522 | 54.797 | 29.306 | 2.154 | 63.143 | 0.856 | -1.103 | 0 | 0 | 1 | 42.312 | 1 | 5 | 1 | GJB_Rinse | 0
44 | P-22643 | 2022/12/01 17:36:34 | 1-05 | 4 | -5.422 | -53.47 | 12622 | 0 | 1 | UNDE6-D-2 | 22-284 | H2O | 1669941699.72 | 0.213 | 0.757 | 61.992 | 0.001 | 0.002 | 1.53 | 54.37 | 29.128 | 2.166 | 63.758 | 0.856 | -1.038 | 0 | 0 | 1 | 42.375 | 1 | 5 | 1 | GJB_Rinse | 0
45 | P-22644 | 2022/12/01 17:46:24 | 1-06 | 1 | [5 empty] | UNDE6-E-1 | 22-284 | H2O | [17 empty] | 1 | 6 | 1 | GJB_Rinse | 0
46 | P-22644 | 2022/12/01 17:56:23 | 1-06 | 2 | [5 empty] | UNDE6-E-1 | 22-284 | H2O | [17 empty] | 1 | 6 | 1 | GJB_Rinse | 0
47 | P-22644 | 2022/12/01 18:06:21 | 1-06 | 3 | [5 empty] | UNDE6-E-1 | 22-284 | H2O | [17 empty] | 1 | 6 | 1 | GJB_Rinse | 0
48 | P-22644 | 2022/12/01 18:16:20 | 1-06 | 4 | [5 empty] | UNDE6-E-1 | 22-284 | H2O | [17 empty] | 1 | 6 | 1 | GJB_Rinse | 0
49 | P-22645 | 2022/12/01 18:27:12 | 1-07 | 1 | -4.239 | -56.017 | 12354 | -1 | 1 | UNDE6-C-3 | 22-284 | H2O | 1669944738.28 | 0.216 | 0.819 | 60.877 | 0.001 | 0 | 1.5 | 53.68 | 28.111 | 2.13 | 63.152 | 0.855 | -0.963 | 0 | 0 | 1 | 42.938 | 1 | 7 | 1 | GJB_Rinse | 0
50 | P-22645 | 2022/12/01 18:36:09 | 1-07 | 2 | -4.083 | -55.354 | 12563 | -1 | 1 | UNDE6-C-3 | 22-284 | H2O | 1669945274.18 | 0.218 | 0.847 | 54.069 | 0.002 | 0.002 | 1.337 | 54.553 | 28.304 | 2.175 | 64.534 | 0.856 | -1.016 | 0 | 0 | 1 | 42.688 | 1 | 7 | 1 | GJB_Rinse | 0
51 | P-22645 | 2022/12/01 18:45:05 | 1-07 | 3 | -4.043 | -55.548 | 12604 | -1 | 1 | UNDE6-C-3 | 22-284 | H2O | 1669945810.64 | 0.205 | 0.748 | 57.3 | 0.001 | -0.003 | 1.407 | 54.58 | 28.323 | 2.172 | 64.887 | 0.856 | -1.099 | 0 | 0 | 1 | 42.375 | 1 | 7 | 1 | GJB_Rinse | 0
52 | P-22645 | 2022/12/01 18:54:01 | 1-07 | 4 | -4.111 | -55.697 | 12507 | 0 | 1 | UNDE6-C-3 | 22-284 | H2O | 1669946346.96 | 0.208 | 0.722 | 59.283 | 0.001 | 0.001 | 1.465 | 54.182 | 28.605 | 2.158 | 63.353 | 0.856 | -1.086 | 0 | 0 | 1 | 42.25 | 1 | 7 | 1 | GJB_Rinse | 0
53 | P-22646 | 2022/12/01 19:03:48 | 1-08 | 1 | -6.127 | -53.044 | 12803 | -1 | 1 | UNDE6-010-2 | 22-284 | H2O | 1669946934.06 | 0.21 | 0.858 | 56.905 | 0.001 | -0.005 | 1.405 | 53.616 | 29.552 | 2.198 | 64.616 | 0.856 | -0.897 | 0 | 0 | 1 | 42.438 | 1 | 8 | 1 | GJB_Rinse | 0
54 | P-22646 | 2022/12/01 19:12:44 | 1-08 | 2 | -6.186 | -52.282 | 12625 | -1 | 1 | UNDE6-010-2 | 22-284 | H2O | 1669947470.1 | 0.213 | 0.822 | 63.017 | 0.001 | 0.001 | 1.555 | 53.404 | 28.918 | 2.175 | 64.197 | 0.856 | -0.762 | 0 | 0 | 1 | 42.438 | 1 | 8 | 1 | GJB_Rinse | 0
55 | P-22646 | 2022/12/01 19:21:40 | 1-08 | 3 | -6.245 | -52.402 | 13005 | -1 | 1 | UNDE6-010-2 | 22-284 | H2O | 1669948006.26 | 0.205 | 0.727 | 58.547 | 0 | 0 | 1.437 | 54.342 | 30.043 | 2.217 | 65.03 | 0.856 | -1.258 | 0 | 0 | 1 | 42.438 | 1 | 8 | 1 | GJB_Rinse | 0
56 | P-22646 | 2022/12/01 19:30:36 | 1-08 | 4 | -6.258 | -52.378 | 12722 | 0 | 1 | UNDE6-010-2 | 22-284 | H2O | 1669948541.98 | 0.19 | 0.793 | 65.098 | 0.001 | 0.001 | 1.602 | 53.527 | 29.019 | 2.18 | 65.796 | 0.856 | -0.829 | 0 | 0 | 1 | 42.562 | 1 | 8 | 1 | GJB_Rinse | 0
57 | P-22647 | 2022/12/01 19:40:23 | 1-09 | 1 | -3.245 | -46.704 | 13077 | -1 | 1 | UNDE6-E-2 | 22-284 | H2O | 1669949128.71 | 0.187 | 0.773 | 61.204 | 0 | 0 | 1.513 | 54.324 | 30.319 | 2.255 | 65.634 | 0.856 | -1.15 | 0 | 0 | 1 | 42.562 | 1 | 9 | 1 | GJB_Rinse | 0
58 | P-22647 | 2022/12/01 19:49:20 | 1-09 | 2 | -3.169 | -46.343 | 12543 | -1 | 1 | UNDE6-E-2 | 22-284 | H2O | 1669949666.17 | 0.194 | 0.852 | 61.033 | 0.001 | 0.001 | 1.5 | 53.329 | 28.714 | 2.162 | 63.591 | 0.856 | -1.115 | 0 | 0 | 1 | 42.562 | 1 | 9 | 1 | GJB_Rinse | 0
59 | P-22647 | 2022/12/01 19:58:17 | 1-09 | 3 | -3.172 | -46.194 | 12539 | -1 | 1 | UNDE6-E-2 | 22-284 | H2O | 1669950202.72 | 0.194 | 0.81 | 61.792 | 0 | 0 | 1.527 | 53.282 | 28.638 | 2.172 | 64.018 | 0.856 | -1.035 | 0 | 0 | 1 | 42.562 | 1 | 9 | 1 | GJB_Rinse | 0
60 | P-22647 | 2022/12/01 20:07:13 | 1-09 | 4 | -3.135 | -46.294 | 12565 | 0 | 1 | UNDE6-E-2 | 22-284 | H2O | 1669950738.68 | 0.2 | 0.793 | 66.986 | 0 | 0.003 | 1.659 | 53.357 | 28.936 | 2.176 | 63.694 | 0.856 | -1.114 | 0 | 0 | 1 | 42.5 | 1 | 9 | 1 | GJB_Rinse | 0
61 | P-22648 | 2022/12/01 20:17:02 | 1-10 | 1 | -6.441 | -55.51 | 12518 | -1 | 1 | UNDE6-2030-3 | 22-284 | H2O | 1669951327.59 | 0.212 | 0.673 | 61.584 | 0.002 | 0.002 | 1.52 | 53.055 | 28.719 | 2.155 | 63.647 | 0.856 | -1.079 | 0 | 0 | 1 | 42.312 | 1 | 10 | 1 | GJB_Rinse | 0
62 | P-22648 | 2022/12/01 20:25:59 | 1-10 | 2 | -6.587 | -56.316 | 12565 | -1 | 1 | UNDE6-2030-3 | 22-284 | H2O | 1669951864.22 | 0.201 | 0.871 | 64.183 | 0.001 | -0.001 | 1.588 | 53.155 | 28.781 | 2.15 | 63.728 | 0.856 | -0.857 | 0 | 0 | 1 | 42.438 | 1 | 10 | 1 | GJB_Rinse | 0
63 | P-22648 | 2022/12/01 20:34:55 | 1-10 | 3 | -6.632 | -56.394 | 12623 | -1 | 1 | UNDE6-2030-3 | 22-284 | H2O | 1669952401.08 | 0.2 | 0.823 | 60.488 | 0.001 | -0.002 | 1.483 | 53.388 | 28.969 | 2.153 | 63.079 | 0.856 | -1.019 | 0 | 0 | 1 | 42.438 | 1 | 10 | 1 | GJB_Rinse | 0
64 | P-22648 | 2022/12/01 20:43:52 | 1-10 | 4 | -6.632 | -56.684 | 12557 | 0 | 1 | UNDE6-2030-3 | 22-284 | H2O | 1669952936.91 | 0.215 | 0.748 | 61.784 | 0.002 | -0.001 | 1.528 | 53.244 | 28.809 | 2.134 | 63.009 | 0.855 | -1.035 | 0 | 0 | 1 | 42.625 | 1 | 10 | 1 | GJB_Rinse | 0
65 | P-22649 | 2022/12/01 20:53:40 | 1-11 | 1 | -2.788 | -52.644 | 11880 | -1 | 1 | UNDE6-E-3 | 22-284 | H2O | 1669953526.32 | 0.201 | 0.829 | 60.39 | 0.001 | 0.001 | 1.501 | 51.967 | 27.083 | 2.059 | 60.048 | 0.856 | -1.054 | 0 | 0 | 1 | 42.625 | 1 | 11 | 1 | GJB_Rinse | 0
66 | P-22649 | 2022/12/01 21:02:37 | 1-11 | 2 | -2.559 | -52.147 | 12378 | -1 | 1 | UNDE6-E-3 | 22-284 | H2O | 1669954062.41 | 0.201 | 0.842 | 58.803 | 0.001 | 0.002 | 1.46 | 53.143 | 28.177 | 2.132 | 62.688 | 0.856 | -1.1 | 0 | 0 | 1 | 42.688 | 1 | 11 | 1 | GJB_Rinse | 0
67 | P-22649 | 2022/12/01 21:11:33 | 1-11 | 3 | -2.605 | -52.525 | 12384 | -1 | 1 | UNDE6-E-3 | 22-284 | H2O | 1669954599.18 | 0.194 | 0.851 | 60.124 | 0 | 0.003 | 1.497 | 52.916 | 28.354 | 2.148 | 62.633 | 0.856 | -0.91 | 0 | 0 | 1 | 42.812 | 1 | 11 | 1 | GJB_Rinse | 0
68 | P-22649 | 2022/12/01 21:20:30 | 1-11 | 4 | -2.574 | -52.372 | 12344 | 0 | 1 | UNDE6-E-3 | 22-284 | H2O | 1669955135.73 | 0.177 | 0.862 | 63.376 | 0.001 | 0.002 | 1.568 | 52.883 | 28.025 | 2.145 | 62.943 | 0.855 | -0.978 | 0 | 0 | 1 | 42.688 | 1 | 11 | 1 | GJB_Rinse | 0
69 | P-22650 | 2022/12/01 21:30:19 | 1-12 | 1 | -5.588 | -51.622 | 12027 | -1 | 1 | UNDE6-2030-2 | 22-284 | H2O | 1669955724.42 | 0.209 | 0.761 | 58.942 | 0.001 | 0.003 | 1.46 | 52.182 | 27.225 | 2.075 | 62.273 | 0.856 | -1.016 | 0 | 0 | 1 | 42.562 | 1 | 12 | 1 | GJB_Rinse | 0
70 | P-22650 | 2022/12/01 21:39:15 | 1-12 | 2 | -5.692 | -51.495 | 12109 | -1 | 1 | UNDE6-2030-2 | 22-284 | H2O | 1669956260.12 | 0.186 | 0.723 | 60.34 | 0.001 | 0.001 | 1.492 | 52.439 | 27.551 | 2.089 | 61.755 | 0.855 | -0.954 | 0 | 0 | 1 | 42.625 | 1 | 12 | 1 | GJB_Rinse | 0
71 | P-22650 | 2022/12/01 21:48:11 | 1-12 | 3 | -5.741 | -51.364 | 12063 | -1 | 1 | UNDE6-2030-2 | 22-284 | H2O | 1669956796.63 | 0.206 | 0.732 | 61.55 | 0.001 | -0.001 | 1.531 | 52.462 | 27.12 | 2.07 | 61.47 | 0.856 | -1.179 | 0 | 0 | 1 | 42.5 | 1 | 12 | 1 | GJB_Rinse | 0
72 | P-22650 | 2022/12/01 21:57:07 | 1-12 | 4 | -5.738 | -51.408 | 12088 | 0 | 1 | UNDE6-2030-2 | 22-284 | H2O | 1669957333.73 | 0.216 | 0.707 | 61.739 | 0 | 0.001 | 1.532 | 52.435 | 27.415 | 2.082 | 60.922 | 0.856 | -1.115 | 0 | 0 | 1 | 42.5 | 1 | 12 | 1 | GJB_Rinse | 0
73 | P-22651 | 2022/12/01 22:06:57 | 1-13 | 1 | -8.869 | -65.125 | 12907 | -1 | 1 | EV | standard | H2O | 1669957922.38 | 0.207 | 0.747 | 61.218 | 0.001 | -0.003 | 1.503 | 53.661 | 29.978 | 2.193 | 64.449 | 0.856 | -1.002 | 0 | 0 | 1 | 42.625 | 1 | 13 | 1 | GJB_Rinse | 0
74 | P-22651 | 2022/12/01 22:15:53 | 1-13 | 2 | -8.941 | -66.111 | 12479 | -1 | 1 | EV | standard | H2O | 1669958458.93 | 0.222 | 0.695 | 60.464 | 0.001 | 0.001 | 1.491 | 52.867 | 28.509 | 2.125 | 63.486 | 0.856 | -0.969 | 0 | 0 | 1 | 42.75 | 1 | 13 | 1 | GJB_Rinse | 0
75 | P-22651 | 2022/12/01 22:24:49 | 1-13 | 3 | -8.982 | -66.619 | 12519 | -1 | 1 | EV | standard | H2O | 1669958994.2 | 0.187 | 0.748 | 60.775 | 0 | 0 | 1.499 | 53.019 | 28.482 | 2.119 | 63.73 | 0.856 | -0.949 | 0 | 0 | 1 | 42.75 | 1 | 13 | 1 | GJB_Rinse | 0
76 | P-22651 | 2022/12/01 22:33:44 | 1-13 | 4 | -8.944 | -66.315 | 12454 | 0 | 1 | EV | standard | H2O | 1669959530.15 | 0.192 | 0.648 | 58.629 | 0 | 0 | 1.448 | 52.812 | 28.491 | 2.124 | 62.827 | 0.856 | -0.948 | 0 | 0 | 1 | 42.625 | 1 | 13 | 1 | GJB_Rinse | 0
77 | P-22652 | 2022/12/01 22:43:34 | 1-14 | 1 | -5.597 | -54.193 | 11848 | -1 | 1 | UNDE6-2030-1 | 22-284 | H2O | 1669960120.19 | 0.202 | 0.792 | 57.105 | 0.001 | -0.001 | 1.421 | 52.107 | 26.426 | 2.043 | 61.411 | 0.855 | -1.008 | 0 | 0 | 1 | 42.5 | 1 | 14 | 1 | GJB_Rinse | 0
78 | P-22652 | 2022/12/01 22:52:31 | 1-14 | 2 | -5.497 | -53.342 | 12287 | -1 | 1 | UNDE6-2030-1 | 22-284 | H2O | 1669960656.56 | 0.185 | 0.906 | 62.229 | 0.001 | 0 | 1.537 | 52.833 | 27.953 | 2.119 | 62.981 | 0.856 | -0.943 | 0 | 0 | 1 | 42.688 | 1 | 14 | 1 | GJB_Rinse | 0
79 | P-22652 | 2022/12/01 23:01:27 | 1-14 | 3 | -5.459 | -53.187 | 12298 | -1 | 1 | UNDE6-2030-1 | 22-284 | H2O | 1669961193.01 | 0.203 | 0.748 | 63.035 | 0.001 | -0.001 | 1.561 | 52.951 | 27.735 | 2.116 | 63.403 | 0.855 | -0.94 | 0 | 0 | 1 | 42.75 | 1 | 14 | 1 | GJB_Rinse | 0
80 | P-22652 | 2022/12/01 23:10:24 | 1-14 | 4 | -5.379 | -52.523 | 12264 | 0 | 1 | UNDE6-2030-1 | 22-284 | H2O | 1669961729.36 | 0.184 | 0.811 | 57.188 | -0.001 | -0.001 | 1.415 | 52.81 | 27.848 | 2.117 | 62.746 | 0.856 | -0.991 | 0 | 0 | 1 | 42.812 | 1 | 14 | 1 | GJB_Rinse | 0
81 | P-22653 | 2022/12/01 23:20:10 | 1-15 | 1 | 1.341 | -52.228 | 12342 | -1 | 1 | UNDE6-C-1 | 22-284 | H2O | 1669962316.49 | 0.192 | 0.845 | 61.654 | 0.001 | 0 | 1.523 | 64.989 | 20.914 | 3.015 | 66.126 | 0.856 | -4.481 | 0 | 0 | 1 | 42.812 | 1 | 15 | 1 | GJB_Rinse | 0
82 | P-22653 | 2022/12/01 23:29:07 | 1-15 | 2 | 1.406 | -52.915 | 12543 | -1 | 1 | UNDE6-C-1 | 22-284 | H2O | 1669962852.17 | 0.214 | 0.921 | 63.266 | -0.001 | 0 | 1.565 | 65.559 | 21.296 | 3.086 | 68.229 | 0.856 | -4.383 | 0 | 0 | 1 | 42.625 | 1 | 15 | 1 | GJB_Rinse | 0
83 | P-22653 | 2022/12/01 23:38:03 | 1-15 | 3 | 1.407 | -53.169 | 12341 | -1 | 1 | UNDE6-C-1 | 22-284 | H2O | 1669963388.78 | 0.203 | 0.927 | 61.434 | 0 | 0.002 | 1.514 | 65.082 | 20.695 | 3.041 | 67.291 | 0.855 | -4.316 | 0 | 0 | 1 | 42.5 | 1 | 15 | 1 | GJB_Rinse | 0
84 | P-22653 | 2022/12/01 23:46:59 | 1-15 | 4 | 1.395 | -53.133 | 12180 | 0 | 1 | UNDE6-C-1 | 22-284 | H2O | 1669963924.56 | 0.203 | 0.931 | 59.011 | 0.001 | 0 | 1.456 | 64.728 | 20.296 | 2.993 | 65.41 | 0.856 | -4.478 | 0 | 0 | 1 | 42.625 | 1 | 15 | 1 | GJB_Rinse | 0
85 | P-22654 | 2022/12/01 23:56:45 | 1-16 | 1 | -3.241 | -51.852 | 18075 | -1 | 1 | UNDE6-D-3 | 22-284 | H2O | 1669964510.75 | 0.209 | 0.802 | 71.316 | 0.001 | 0.002 | 1.683 | 65.541 | 44.152 | 3.037 | 85.79 | 0.86 | -1.473 | 0 | 0 | 1 | 42.688 | 1 | 16 | 1 | GJB_Rinse | 0
86 | P-22654 | 2022/12/02 00:05:41 | 1-16 | 2 | -3.166 | -50.439 | 12683 | -1 | 1 | UNDE6-D-3 | 22-284 | H2O | 1669965047.17 | 0.21 | 0.692 | 56.051 | 0 | 0.001 | 1.372 | 53.744 | 29.283 | 2.184 | 62.859 | 0.856 | -1.303 | 0 | 0 | 1 | 42.688 | 1 | 16 | 1 | GJB_Rinse | 0
87 | P-22654 | 2022/12/02 00:14:38 | 1-16 | 3 | -3.145 | -50.552 | 12715 | -1 | 1 | UNDE6-D-3 | 22-284 | H2O | 1669965583.39 | 0.203 | 0.847 | 62.43 | 0.001 | 0 | 1.543 | 53.721 | 29.304 | 2.193 | 63.928 | 0.856 | -1.172 | 0 | 0 | 1 | 42.812 | 1 | 16 | 1 | GJB_Rinse | 0
88 | P-22654 | 2022/12/02 00:23:34 | 1-16 | 4 | -3.132 | -50.325 | 12602 | 0 | 1 | UNDE6-D-3 | 22-284 | H2O | 1669966119.95 | 0.188 | 0.742 | 61.392 | 0 | 0.002 | 1.519 | 53.639 | 28.776 | 2.172 | 63.38 | 0.855 | -1.224 | 0 | 0 | 1 | 42.562 | 1 | 16 | 1 | GJB_Rinse | 0
89 | P-22655 | 2022/12/02 00:33:21 | 1-17 | 1 | -3.725 | -57.535 | 12279 | -1 | 1 | UNDE6-D-5 | 22-284 | H2O | 1669966706.73 | 0.207 | 0.858 | 60.031 | 0 | 0.001 | 1.481 | 55.354 | 26.35 | 2.259 | 63.476 | 0.855 | -1.71 | 0 | 0 | 1 | 42.562 | 1 | 17 | 1 | GJB_Rinse | 0
90 | P-22655 | 2022/12/02 00:42:17 | 1-17 | 2 | -3.8 | -58.425 | 12299 | -1 | 1 | UNDE6-D-5 | 22-284 | H2O | 1669967242.85 | 0.187 | 0.797 | 65.78 | 0.001 | 0.005 | 1.628 | 55.402 | 26.472 | 2.261 | 63.174 | 0.856 | -1.709 | 0 | 0 | 1 | 42.5 | 1 | 17 | 1 | GJB_Rinse | 0
91 | P-22655 | 2022/12/02 00:51:13 | 1-17 | 3 | -3.847 | -58.676 | 12352 | -1 | 1 | UNDE6-D-5 | 22-284 | H2O | 1669967778.69 | 0.209 | 0.859 | 62.278 | 0 | -0.001 | 1.539 | 55.452 | 26.613 | 2.279 | 64.203 | 0.855 | -1.545 | 0 | 0 | 1 | 42.438 | 1 | 17 | 1 | GJB_Rinse | 0
92 | P-22655 | 2022/12/02 01:00:09 | 1-17 | 4 | -3.786 | -58.623 | 12591 | 0 | 1 | UNDE6-D-5 | 22-284 | H2O | 1669968315.17 | 0.21 | 0.808 | 59.74 | 0.001 | 0.002 | 1.477 | 55.971 | 26.983 | 2.322 | 65.8 | 0.856 | -1.597 | 0 | 0 | 1 | 42.375 | 1 | 17 | 1 | GJB_Rinse | 0
93 | P-22656 | 2022/12/02 01:09:57 | 1-18 | 1 | -2.46 | -54.253 | 12342 | -1 | 1 | UNDE6-B-2 | 22-284 | H2O | 1669968902.78 | 0.207 | 0.754 | 62.557 | 0.001 | 0.002 | 1.54 | 54.678 | 26.944 | 2.225 | 64.128 | 0.856 | -1.318 | 0 | 0 | 1 | 42.25 | 1 | 18 | 1 | GJB_Rinse | 0
94 | P-22656 | 2022/12/02 01:18:54 | 1-18 | 2 | -2.343 | -53.593 | 12429 | -1 | 1 | UNDE6-B-2 | 22-284 | H2O | 1669969439.51 | 0.193 | 0.703 | 64.311 | 0.001 | 0.003 | 1.593 | 54.863 | 27.446 | 2.233 | 64.504 | 0.855 | -1.388 | 0 | 0 | 1 | 42.312 | 1 | 18 | 1 | GJB_Rinse | 0
95 | P-22656 | 2022/12/02 01:27:50 | 1-18 | 3 | -2.296 | -53.214 | 12351 | -1 | 1 | UNDE6-B-2 | 22-284 | H2O | 1669969975.64 | 0.182 | 0.743 | 60.807 | 0.001 | 0 | 1.514 | 54.765 | 27.213 | 2.215 | 63.496 | 0.855 | -1.415 | 0 | 0 | 1 | 42.438 | 1 | 18 | 1 | GJB_Rinse | 0
96 | P-22656 | 2022/12/02 01:36:46 | 1-18 | 4 | -2.366 | -53.529 | 12590 | 0 | 1 | UNDE6-B-2 | 22-284 | H2O | 1669970511.4 | 0.201 | 0.817 | 65.288 | 0 | 0.002 | 1.618 | 55.383 | 27.912 | 2.259 | 63.536 | 0.856 | -1.658 | 0 | 0 | 1 | 42.438 | 1 | 18 | 1 | GJB_Rinse | 0
97 | P-22657 | 2022/12/02 01:46:34 | 1-19 | 1 | -5.292 | -63.28 | 11456 | -1 | 1 | UNDE6-C-2 | 22-284 | H2O | 1669971099.97 | 0.193 | 0.874 | 59.76 | 0.001 | -0.003 | 1.497 | 51.104 | 25.486 | 1.977 | 59.504 | 0.855 | -0.967 | 0 | 0 | 1 | 42.375 | 1 | 19 | 1 | GJB_Rinse | 0
98 | P-22657 | 2022/12/02 01:55:31 | 1-19 | 2 | -5.438 | -64.535 | 11408 | -1 | 1 | UNDE6-C-2 | 22-284 | H2O | 1669971637.4 | 0.2 | 0.806 | 57.003 | 0.001 | 0.002 | 1.425 | 50.963 | 25.277 | 1.965 | 59.514 | 0.855 | -0.879 | 0 | 0 | 1 | 42.312 | 1 | 19 | 1 | GJB_Rinse | 0
99 | P-22657 | 2022/12/02 02:04:28 | 1-19 | 3 | -5.434 | -64.694 | 11510 | -1 | 1 | UNDE6-C-2 | 22-284 | H2O | 1669972174.15 | 0.201 | 0.871 | 56.255 | 0 | 0.004 | 1.406 | 51.182 | 25.898 | 1.964 | 58.936 | 0.855 | -1.011 | 0 | 0 | 1 | 42.375 | 1 | 19 | 1 | GJB_Rinse | 0
100 | P-22657 | 2022/12/02 02:13:24 | 1-19 | 4 | -5.455 | -64.758 | 11467 | 0 | 1 | UNDE6-C-2 | 22-284 | H2O | 1669972710.45 | 0.208 | 0.769 | 58.189 | 0.001 | 0 | 1.459 | 51.066 | 25.42 | 1.973 | 60.263 | 0.855 | -0.831 | 0 | 0 | 1 | 42.188 | 1 | 19 | 1 | GJB_Rinse | 0
101 | P-22658 | 2022/12/02 02:23:12 | 1-20 | 1 | -3.562 | -55.128 | 11427 | -1 | 1 | UNDE6-A-1 | 22-284 | H2O | 1669973298.35 | 0.195 | 0.861 | 56.965 | 0.001 | 0 | 1.424 | 169.825 | 21.153 | 2.058 | 57.557 | 0.856 | -1.231 | 0 | 0 | 1 | 42.312 | 1 | 20 | 1 | GJB_Rinse | 0
102 | P-22658 | 2022/12/02 02:32:09 | 1-20 | 2 | -3.522 | -54.363 | 11980 | -1 | 1 | UNDE6-A-1 | 22-284 | H2O | 1669973834.73 | 0.211 | 0.872 | 59.252 | 0 | 0 | 1.472 | 177.781 | 22.571 | 2.136 | 58.975 | 0.856 | -1.612 | 0 | 0 | 1 | 42.438 | 1 | 20 | 1 | GJB_Rinse | 0
103 | P-22658 | 2022/12/02 02:41:05 | 1-20 | 3 | -3.527 | -54.059 | 12038 | -1 | 1 | UNDE6-A-1 | 22-284 | H2O | 1669974370.61 | 0.218 | 0.819 | 54.225 | 0.001 | -0.002 | 1.348 | 178.95 | 22.32 | 2.154 | 60.152 | 0.855 | -1.394 | 0 | 0 | 1 | 42.188 | 1 | 20 | 1 | GJB_Rinse | 0
104 | P-22658 | 2022/12/02 02:50:02 | 1-20 | 4 | -3.507 | -53.887 | 12137 | 0 | 1 | UNDE6-A-1 | 22-284 | H2O | 1669974907.7 | 0.201 | 0.856 | 58.503 | 0.001 | -0.001 | 1.453 | 180.328 | 22.883 | 2.164 | 59.686 | 0.856 | -1.568 | 0 | 0 | 1 | 42.188 | 1 | 20 | 1 | GJB_Rinse | 0
105 | P-22659 | 2022/12/02 02:59:50 | 1-21 | 1 | -3.75 | -55.353 | 11816 | -1 | 1 | UNDE6-B-4 | 22-284 | H2O | 1669975496.52 | 0.198 | 0.743 | 61.587 | 0.001 | 0.001 | 1.532 | 53.735 | 26.221 | 2.042 | 60.554 | 0.855 | -1.118 | 0 | 0 | 1 | 42.375 | 1 | 21 | 1 | GJB_Rinse | 0
106 | P-22659 | 2022/12/02 03:08:46 | 1-21 | 2 | -3.726 | -55.565 | 11599 | -1 | 1 | UNDE6-B-4 | 22-284 | H2O | 1669976032.38 | 0.209 | 0.865 | 57.049 | 0 | 0 | 1.438 | 52.167 | 25.672 | 2.007 | 60.666 | 0.856 | -0.96 | 0 | 0 | 1 | 42.438 | 1 | 21 | 1 | GJB_Rinse | 0
107 | P-22659 | 2022/12/02 03:17:42 | 1-21 | 3 | -3.77 | -55.688 | 11724 | -1 | 1 | UNDE6-B-4 | 22-284 | H2O | 1669976568.45 | 0.194 | 0.845 | 56.918 | 0.001 | -0.002 | 1.417 | 52.293 | 26.318 | 2.014 | 59.984 | 0.855 | -1.01 | 0 | 0 | 1 | 42.375 | 1 | 21 | 1 | GJB_Rinse | 0
108 | P-22659 | 2022/12/02 03:26:38 | 1-21 | 4 | -3.724 | -55.834 | 11834 | 0 | 1 | UNDE6-B-4 | 22-284 | H2O | 1669977104.21 | 0.173 | 0.81 | 58.766 | 0.001 | 0.002 | 1.463 | 52.416 | 26.623 | 2.054 | 60.806 | 0.856 | -1.195 | 0 | 0 | 1 | 42.25 | 1 | 21 | 1 | GJB_Rinse | 0
109 | P-22660 | 2022/12/02 03:36:23 | 1-22 | 1 | -8.7 | -65.266 | 12789 | -1 | 1 | EV | standard | H2O | 1669977688.96 | 0.203 | 0.783 | 60.988 | 0.001 | 0.004 | 1.507 | 53.693 | 29.268 | 2.169 | 64.548 | 0.856 | -0.999 | 0 | 0 | 1 | 42.375 | 1 | 22 | 1 | GJB_Rinse | 0
110 | P-22660 | 2022/12/02 03:45:19 | 1-22 | 2 | -8.893 | -66.285 | 12449 | -1 | 1 | EV | standard | H2O | 1669978224.47 | 0.227 | 0.779 | 59.877 | 0.002 | 0.003 | 1.492 | 52.802 | 28.743 | 2.121 | 62.85 | 0.855 | -0.886 | 0 | 0 | 1 | 42.438 | 1 | 22 | 1 | GJB_Rinse | 0
111 | P-22660 | 2022/12/02 03:54:15 | 1-22 | 3 | -8.916 | -66.492 | 12472 | -1 | 1 | EV | standard | H2O | 1669978760.71 | 0.207 | 0.713 | 57.508 | 0.001 | 0 | 1.412 | 52.921 | 28.702 | 2.109 | 62.692 | 0.856 | -0.959 | 0 | 0 | 1 | 42.312 | 1 | 22 | 1 | GJB_Rinse | 0
112 | P-22660 | 2022/12/02 04:03:11 | 1-22 | 4 | -8.986 | -66.725 | 12454 | 0 | 1 | EV | standard | H2O | 1669979296.89 | 0.219 | 0.791 | 60.243 | 0.001 | -0.001 | 1.488 | 52.875 | 28.584 | 2.11 | 63.003 | 0.855 | -0.968 | 0 | 0 | 1 | 42.438 | 1 | 22 | 1 | GJB_Rinse | 0
113 | P-22661 | 2022/12/02 04:12:57 | 1-23 | 1 | -3.455 | -56.143 | 12082 | -1 | 1 | UNDE6-A-5 | 22-284 | H2O | 1669979883.3 | 0.201 | 0.826 | 60.945 | 0.001 | 0.001 | 1.507 | 54.072 | 26.606 | 2.159 | 61.785 | 0.856 | -1.475 | 0 | 0 | 1 | 42.375 | 1 | 23 | 1 | GJB_Rinse | 0
114 | P-22661 | 2022/12/02 04:21:54 | 1-23 | 2 | -3.296 | -55.427 | 12023 | -1 | 1 | UNDE6-A-5 | 22-284 | H2O | 1669980419.4 | 0.199 | 0.728 | 60.041 | 0.001 | -0.002 | 1.494 | 53.908 | 26.52 | 2.141 | 61.803 | 0.855 | -1.376 | 0 | 0 | 1 | 42.312 | 1 | 23 | 1 | GJB_Rinse | 0
115 | P-22661 | 2022/12/02 04:30:49 | 1-23 | 3 | -3.212 | -54.733 | 12130 | -1 | 1 | UNDE6-A-5 | 22-284 | H2O | 1669980954.58 | 0.195 | 0.742 | 58.556 | 0 | 0 | 1.457 | 54.311 | 26.435 | 2.164 | 62.499 | 0.856 | -1.454 | 0 | 0 | 1 | 42.375 | 1 | 23 | 1 | GJB_Rinse | 0
116 | P-22661 | 2022/12/02 04:39:45 | 1-23 | 4 | -3.244 | -55.128 | 12046 | 0 | 1 | UNDE6-A-5 | 22-284 | H2O | 1669981491.3 | 0.179 | 0.895 | 64.628 | 0.001 | 0 | 1.607 | 54.203 | 26.024 | 2.145 | 62.306 | 0.856 | -1.424 | 0 | 0 | 1 | 42.438 | 1 | 23 | 1 | GJB_Rinse | 0
117 | P-22662 | 2022/12/02 04:49:32 | 1-24 | 1 | -1.585 | -61.114 | 12526 | -1 | 1 | UNDE6-D-4 | 22-284 | H2O | 1669982077.33 | 0.197 | 0.769 | 58.03 | 0.001 | 0.002 | 1.434 | 54.075 | 28.504 | 2.161 | 61.731 | 0.856 | -1.502 | 0 | 0 | 1 | 42.312 | 1 | 24 | 1 | GJB_Rinse | 0
118 | P-22662 | 2022/12/02 04:58:28 | 1-24 | 2 | -1.478 | -61.639 | 12389 | -1 | 1 | UNDE6-D-4 | 22-284 | H2O | 1669982613.89 | 0.219 | 0.978 | 60.138 | 0.001 | 0 | 1.478 | 53.626 | 27.898 | 2.161 | 62.866 | 0.856 | -1.27 | 0 | 0 | 1 | 42.312 | 1 | 24 | 1 | GJB_Rinse | 0
119 | P-22662 | 2022/12/02 05:07:24 | 1-24 | 3 | -1.438 | -61.612 | 12532 | -1 | 1 | UNDE6-D-4 | 22-284 | H2O | 1669983150.09 | 0.208 | 0.81 | 62.706 | 0 | 0.002 | 1.554 | 54.004 | 27.925 | 2.183 | 64.34 | 0.855 | -1.135 | 0 | 0 | 1 | 42.438 | 1 | 24 | 1 | GJB_Rinse | 0
120 | P-22662 | 2022/12/02 05:16:21 | 1-24 | 4 | -1.426 | -61.752 | 12375 | 0 | 1 | UNDE6-D-4 | 22-284 | H2O | 1669983686.39 | 0.189 | 0.813 | 59.073 | 0.001 | -0.002 | 1.462 | 53.505 | 28.04 | 2.15 | 63.121 | 0.856 | -1.164 | 0 | 0 | 1 | 42.5 | 1 | 24 | 1 | GJB_Rinse | 0
121 | P-22663 | 2022/12/02 05:26:08 | 1-25 | 1 | -4.721 | -60.91 | 11949 | -1 | 1 | UNDE6-C-4 | 22-284 | H2O | 1669984273.49 | 0.216 | 0.952 | 55.584 | 0 | 0.001 | 1.381 | 353.262 | 17.145 | 2.092 | 55.625 | 0.856 | -1.488 | 0 | 0 | 1 | 42.438 | 1 | 25 | 1 | GJB_Rinse | 0
122 | P-22663 | 2022/12/02 05:35:03 | 1-25 | 2 | -5.031 | -61.468 | 12052 | -1 | 1 | UNDE6-C-4 | 22-284 | H2O | 1669984809.09 | 0.239 | 0.899 | 60.958 | 0.001 | 0.004 | 1.504 | 359.865 | 17.423 | 2.117 | 55.542 | 0.857 | -1.363 | 0 | 0 | 1 | 42.375 | 1 | 25 | 1 | GJB_Rinse | 0
123 | P-22663 | 2022/12/02 05:44:00 | 1-25 | 3 | -5.014 | -61.274 | 11963 | -1 | 1 | UNDE6-C-4 | 22-284 | H2O | 1669985346.45 | 0.235 | 0.986 | 58.93 | 0.001 | 0.003 | 1.465 | 357.967 | 17.423 | 2.128 | 55.195 | 0.856 | -1.315 | 0 | 0 | 1 | 42.375 | 1 | 25 | 1 | GJB_Rinse | 0
124 | P-22663 | 2022/12/02 05:52:56 | 1-25 | 4 | -5.07 | -61.303 | 11911 | 0 | 1 | UNDE6-C-4 | 22-284 | H2O | 1669985882.45 | 0.229 | 0.887 | 61.159 | 0.001 | 0.005 | 1.519 | 357.097 | 16.755 | 2.112 | 55.982 | 0.856 | -1.302 | 0 | 0 | 1 | 42.375 | 1 | 25 | 1 | GJB_Rinse | 0
125 | P-22664 | 2022/12/02 06:02:43 | 1-26 | 1 | -5.244 | -51.089 | 12464 | -1 | 1 | UNDE6-3040-1 | 22-284 | H2O | 1669986469.11 | 0.197 | 0.828 | 62.092 | 0 | 0.002 | 1.537 | 57.025 | 28.113 | 2.153 | 63.559 | 0.856 | -1.114 | 0 | 0 | 1 | 42.562 | 1 | 26 | 1 | GJB_Rinse | 0
126 | P-22664 | 2022/12/02 06:11:39 | 1-26 | 2 | -5.223 | -50.093 | 12063 | -1 | 1 | UNDE6-3040-1 | 22-284 | H2O | 1669987004.59 | 0.214 | 0.845 | 61.563 | 0.001 | -0.003 | 1.527 | 54.164 | 27.535 | 2.079 | 60.413 | 0.856 | -1.231 | 0 | 0 | 1 | 42.438 | 1 | 26 | 1 | GJB_Rinse | 0
127 | P-22664 | 2022/12/02 06:20:34 | 1-26 | 3 | -5.2 | -49.822 | 12046 | -1 | 1 | UNDE6-3040-1 | 22-284 | H2O | 1669987539.46 | 0.208 | 0.798 | 59.405 | 0.001 | -0.001 | 1.48 | 53.376 | 27.097 | 2.08 | 62.734 | 0.855 | -0.82 | 0 | 0 | 1 | 42.438 | 1 | 26 | 1 | GJB_Rinse | 0
128 | P-22664 | 2022/12/02 06:29:31 | 1-26 | 4 | -5.2 | -49.439 | 12150 | 0 | 1 | UNDE6-3040-1 | 22-284 | H2O | 1669988076.26 | 0.193 | 0.829 | 62.016 | 0.001 | 0.004 | 1.545 | 53.454 | 27.347 | 2.091 | 61.755 | 0.856 | -1.001 | 0 | 0 | 1 | 42.5 | 1 | 26 | 1 | GJB_Rinse | 0
129 | P-22665 | 2022/12/02 06:39:17 | 1-27 | 1 | -3.142 | -47.208 | 12229 | -1 | 1 | UNDE6-1020-1 | 22-284 | H2O | 1669988662.97 | 0.205 | 0.834 | 65.318 | 0.002 | 0.002 | 1.614 | 52.849 | 27.702 | 2.119 | 62.414 | 0.856 | -0.976 | 0 | 0 | 1 | 42.5 | 1 | 27 | 1 | GJB_Rinse | 0
130 | P-22665 | 2022/12/02 06:48:13 | 1-27 | 2 | -3.121 | -47.125 | 12185 | -1 | 1 | UNDE6-1020-1 | 22-284 | H2O | 1669989199.31 | 0.199 | 0.799 | 64.736 | 0.001 | 0.002 | 1.601 | 52.818 | 27.327 | 2.126 | 64.063 | 0.856 | -0.757 | 0 | 0 | 1 | 42.5 | 1 | 27 | 1 | GJB_Rinse | 0
131 | P-22665 | 2022/12/02 06:57:09 | 1-27 | 3 | -3.047 | -46.791 | 12205 | -1 | 1 | UNDE6-1020-1 | 22-284 | H2O | 1669989735.21 | 0.19 | 0.818 | 59.917 | 0 | 0.001 | 1.474 | 52.826 | 27.614 | 2.109 | 63.196 | 0.855 | -0.912 | 0 | 0 | 1 | 42.562 | 1 | 27 | 1 | GJB_Rinse | 0
132 | P-22665 | 2022/12/02 07:06:05 | 1-27 | 4 | -3.078 | -46.829 | 12027 | 0 | 1 | UNDE6-1020-1 | 22-284 | H2O | 1669990271.04 | 0.212 | 0.831 | 60.874 | 0.001 | 0.002 | 1.505 | 52.467 | 26.932 | 2.085 | 62.277 | 0.856 | -1.001 | 0 | 0 | 1 | 42.562 | 1 | 27 | 1 | GJB_Rinse | 0
133 | P-22666 | 2022/12/02 07:15:52 | 1-28 | 1 | -2.742 | -54.752 | 11552 | -1 | 1 | UNDE6-B-3 | 22-284 | H2O | 1669990858.54 | 0.203 | 0.829 | 54.904 | 0 | -0.003 | 1.366 | 192.856 | 19.352 | 2.217 | 57.361 | 0.856 | -2.346 | 0 | 0 | 1 | 42.438 | 1 | 28 | 1 | GJB_Rinse | 0
134 | P-22666 | 2022/12/02 07:24:49 | 1-28 | 2 | -2.828 | -55.596 | 11339 | -1 | 1 | UNDE6-B-3 | 22-284 | H2O | 1669991395.43 | 0.208 | 0.923 | 56.74 | 0 | 0.001 | 1.42 | 191.34 | 18.354 | 2.194 | 57.84 | 0.855 | -2.141 | 0 | 0 | 1 | 42.438 | 1 | 28 | 1 | GJB_Rinse | 0
135 | P-22666 | 2022/12/02 07:33:45 | 1-28 | 3 | -2.846 | -56.053 | 11470 | -1 | 1 | UNDE6-B-3 | 22-284 | H2O | 1669991931.17 | 0.212 | 0.804 | 57.067 | 0.001 | -0.003 | 1.428 | 193.739 | 18.974 | 2.208 | 57.403 | 0.855 | -2.306 | 0 | 0 | 1 | 42.562 | 1 | 28 | 1 | GJB_Rinse | 0
136 | P-22666 | 2022/12/02 07:42:41 | 1-28 | 4 | -2.884 | -56.223 | 11574 | 0 | 1 | UNDE6-B-3 | 22-284 | H2O | 1669992467.25 | 0.207 | 0.85 | 54.457 | 0.001 | -0.003 | 1.359 | 195.664 | 18.906 | 2.232 | 57.811 | 0.857 | -2.247 | 0 | 0 | 1 | 42.562 | 1 | 28 | 1 | GJB_Rinse | 0
137 | P-22667 | 2022/12/02 07:52:25 | 1-29 | 1 | -5.827 | -55.832 | 12260 | -1 | 1 | UNDE6-3040-2 | 22-284 | H2O | 1669993051.24 | 0.186 | 0.883 | 60.858 | 0.001 | 0.001 | 1.5 | 54.868 | 27.829 | 2.105 | 62.143 | 0.855 | -1.05 | 0 | 0 | 1 | 42.5 | 1 | 29 | 1 | GJB_Rinse | 0
138 | P-22667 | 2022/12/02 08:01:21 | 1-29 | 2 | -5.828 | -55.204 | 11840 | -1 | 1 | UNDE6-3040-2 | 22-284 | H2O | 1669993586.6 | 0.21 | 0.783 | 57.47 | 0.001 | -0.002 | 1.427 | 53.01 | 26.646 | 2.026 | 60.269 | 0.855 | -1.091 | 0 | 0 | 1 | 42.5 | 1 | 29 | 1 | GJB_Rinse | 0
139 | P-22667 | 2022/12/02 08:10:16 | 1-29 | 3 | -5.816 | -55.309 | 12342 | -1 | 1 | UNDE6-3040-2 | 22-284 | H2O | 1669994121.65 | 0.225 | 0.858 | 57.666 | 0.001 | 0.002 | 1.427 | 53.703 | 27.909 | 2.12 | 62.773 | 0.856 | -0.973 | 0 | 0 | 1 | 42.438 | 1 | 29 | 1 | GJB_Rinse | 0
140 | P-22667 | 2022/12/02 08:19:11 | 1-29 | 4 | -5.918 | -55.445 | 12416 | 0 | 1 | UNDE6-3040-2 | 22-284 | H2O | 1669994656.63 | 0.195 | 0.756 | 61.777 | 0.001 | 0 | 1.519 | 53.497 | 28.116 | 2.154 | 64.461 | 0.856 | -0.769 | 0 | 0 | 1 | 42.438 | 1 | 29 | 1 | GJB_Rinse | 0
141 | P-22668 | 2022/12/02 08:28:55 | 1-30 | 1 | -2.102 | -56.55 | 11832 | -1 | 1 | UNDE6-C-5 | 22-284 | H2O | 1669995241.14 | 0.196 | 0.803 | 58.983 | 0.001 | 0 | 1.469 | 53.836 | 26.319 | 2.055 | 61.127 | 0.855 | -1.112 | 0 | 0 | 1 | 42.438 | 1 | 30 | 1 | GJB_Rinse | 0
142 | P-22668 | 2022/12/02 08:37:51 | 1-30 | 2 | -1.949 | -56.768 | 12135 | -1 | 1 | UNDE6-C-5 | 22-284 | H2O | 1669995777.05 | 0.21 | 0.842 | 60.647 | 0.001 | -0.001 | 1.491 | 54.642 | 27.215 | 2.108 | 61.648 | 0.855 | -1.299 | 0 | 0 | 1 | 42.438 | 1 | 30 | 1 | GJB_Rinse | 0
143 | P-22668 | 2022/12/02 08:46:46 | 1-30 | 3 | -1.84 | -56.174 | 11904 | -1 | 1 | UNDE6-C-5 | 22-284 | H2O | 1669996312.13 | 0.202 | 0.688 | 54.453 | 0 | -0.003 | 1.35 | 54.117 | 26.476 | 2.073 | 61.225 | 0.855 | -1.142 | 0 | 0 | 1 | 42.5 | 1 | 30 | 1 | GJB_Rinse | 0
144 | P-22668 | 2022/12/02 08:55:41 | 1-30 | 4 | -1.903 | -56.736 | 12372 | 0 | 1 | UNDE6-C-5 | 22-284 | H2O | 1669996847.23 | 0.196 | 0.836 | 62.426 | 0.001 | 0.002 | 1.544 | 54.967 | 27.745 | 2.149 | 63.516 | 0.856 | -1.139 | 0 | 0 | 1 | 42.375 | 1 | 30 | 1 | GJB_Rinse | 0
145 | P-22669 | 2022/12/02 09:05:27 | 1-31 | 1 | -8.618 | -65.528 | 12659 | -1 | 1 | EV | standard | H2O | 1669997432.6 | 0.198 | 0.766 | 63.811 | 0.002 | -0.001 | 1.568 | 53.419 | 29.131 | 2.146 | 63.59 | 0.856 | -1.067 | 0 | 0 | 1 | 42.375 | 1 | 31 | 1 | GJB_Rinse | 0
146 | P-22669 | 2022/12/02 09:14:23 | 1-31 | 2 | -8.832 | -66.347 | 12438 | -1 | 1 | EV | standard | H2O | 1669997968.02 | 0.206 | 0.821 | 62.306 | 0.001 | 0.002 | 1.541 | 52.891 | 28.173 | 2.119 | 64.038 | 0.856 | -0.861 | 0 | 0 | 1 | 42.375 | 1 | 31 | 1 | GJB_Rinse | 0
147 | P-22669 | 2022/12/02 09:23:19 | 1-31 | 3 | -8.947 | -66.773 | 12352 | -1 | 1 | EV | standard | H2O | 1669998504.01 | 0.203 | 0.779 | 63.283 | 0.002 | 0.001 | 1.555 | 52.736 | 28.515 | 2.09 | 61.557 | 0.856 | -0.994 | 0 | 0 | 1 | 42.375 | 1 | 31 | 1 | GJB_Rinse | 0
148 | P-22669 | 2022/12/02 09:32:15 | 1-31 | 4 | -8.923 | -66.332 | 12335 | 0 | 1 | EV | standard | H2O | 1669999040.58 | 0.207 | 0.694 | 63.26 | 0.001 | 0.001 | 1.56 | 52.638 | 28.224 | 2.092 | 62.766 | 0.855 | -0.954 | 0 | 0 | 1 | 42.312 | 1 | 31 | 1 | GJB_Rinse | 0
149 | P-22670 | 2022/12/02 09:42:01 | 1-32 | 1 | -3.907 | -56.875 | 11470 | -1 | 1 | UNDE6-B-1 | 22-284 | H2O | 1669999626.69 | 0.188 | 0.764 | 59.793 | 0.001 | 0 | 1.499 | 51.534 | 25.658 | 1.991 | 59.221 | 0.855 | -0.946 | 0 | 0 | 1 | 42.062 | 1 | 32 | 1 | GJB_Rinse | 0
150 | P-22670 | 2022/12/02 09:50:56 | 1-32 | 2 | -3.661 | -55.672 | 12026 | -1 | 1 | UNDE6-B-1 | 22-284 | H2O | 1670000161.61 | 0.203 | 0.766 | 59.41 | 0 | 0.001 | 1.471 | 52.883 | 26.913 | 2.068 | 61.171 | 0.855 | -1.213 | 0 | 0 | 1 | 42 | 1 | 32 | 1 | GJB_Rinse | 0
151 | P-22670 | 2022/12/02 09:59:52 | 1-32 | 3 | -3.625 | -55.581 | 11946 | -1 | 1 | UNDE6-B-1 | 22-284 | H2O | 1670000697.45 | 0.204 | 0.771 | 58.897 | 0 | -0.001 | 1.466 | 52.627 | 26.908 | 2.07 | 61.208 | 0.855 | -1.143 | 0 | 0 | 1 | 42.062 | 1 | 32 | 1 | GJB_Rinse | 0
152 | P-22670 | 2022/12/02 10:08:47 | 1-32 | 4 | -3.609 | -55.409 | 12086 | 0 | 1 | UNDE6-B-1 | 22-284 | H2O | 1670001232.63 | 0.209 | 0.838 | 62.447 | 0.001 | 0.001 | 1.546 | 52.922 | 27.039 | 2.094 | 62.141 | 0.855 | -1.142 | 0 | 0 | 1 | 42.25 | 1 | 32 | 1 | GJB_Rinse | 0
153 | P-22671 | 2022/12/02 10:18:33 | 1-33 | 1 | -0.803 | -45.759 | 10307 | -1 | 1 | UNDE6-E-4 | 22-284 | H2O | 1670001819.91 | 0.205 | 0.796 | 52.104 | 0.001 | 0 | 1.328 | 49.006 | 22.267 | 1.791 | 54.365 | 0.855 | -1.048 | 0 | 0 | 1 | 42.25 | 1 | 33 | 1 | GJB_Rinse | 0
154 | P-22671 | 2022/12/02 10:27:28 | 1-33 | 2 | -0.618 | -44.528 | 11519 | -1 | 1 | UNDE6-E-4 | 22-284 | H2O | 1670002354.22 | 0.211 | 0.776 | 55.522 | 0.001 | 0 | 1.377 | 51.283 | 25.191 | 2.012 | 61.243 | 0.855 | -0.823 | 0 | 0 | 1 | 42 | 1 | 33 | 1 | GJB_Rinse | 0
155 | P-22671 | 2022/12/02 10:36:23 | 1-33 | 3 | -0.617 | -44.204 | 11782 | -1 | 1 | UNDE6-E-4 | 22-284 | H2O | 1670002889.28 | 0.204 | 0.828 | 61.52 | 0 | -0.001 | 1.531 | 51.785 | 26.067 | 2.054 | 62.327 | 0.855 | -0.957 | 0 | 0 | 1 | 41.938 | 1 | 33 | 1 | GJB_Rinse | 0
156 | P-22671 | 2022/12/02 10:45:19 | 1-33 | 4 | -0.61 | -44.173 | 11928 | 0 | 1 | UNDE6-E-4 | 22-284 | H2O | 1670003424.96 | 0.215 | 0.842 | 57.77 | 0.001 | -0.001 | 1.436 | 52.246 | 27.206 | 2.056 | 60.036 | 0.856 | -1.363 | 0 | 0 | 1 | 42 | 1 | 33 | 1 | GJB_Rinse | 0
157 | P-22672 | 2022/12/02 10:55:06 | 1-34 | 1 | -4.408 | -57.993 | 12160 | -1 | 1 | UNDE6-A-2 | 22-284 | H2O | 1670004011.97 | 0.2 | 0.787 | 54.787 | 0 | 0 | 1.362 | 52.666 | 27.668 | 2.082 | 61.381 | 0.855 | -1.225 | 0 | 0 | 1 | 42.125 | 1 | 34 | 1 | GJB_Rinse | 0
158 | P-22672 | 2022/12/02 11:04:02 | 1-34 | 2 | -4.667 | -59.737 | 12047 | -1 | 1 | UNDE6-A-2 | 22-284 | H2O | 1670004547.88 | 0.213 | 0.843 | 59.061 | 0.001 | 0 | 1.472 | 52.447 | 26.627 | 2.086 | 62.921 | 0.856 | -0.888 | 0 | 0 | 1 | 42.125 | 1 | 34 | 1 | GJB_Rinse | 0
159 | P-22672 | 2022/12/02 11:12:59 | 1-34 | 3 | -4.682 | -60.015 | 11933 | -1 | 1 | UNDE6-A-2 | 22-284 | H2O | 1670005084.04 | 0.21 | 0.875 | 60.717 | 0.001 | 0.003 | 1.51 | 52.217 | 26.355 | 2.046 | 62.161 | 0.856 | -0.935 | 0 | 0 | 1 | 41.875 | 1 | 34 | 1 | GJB_Rinse | 0
160 | P-22672 | 2022/12/02 11:21:55 | 1-34 | 4 | -4.679 | -60.133 | 12024 | 0 | 1 | UNDE6-A-2 | 22-284 | H2O | 1670005620.65 | 0.218 | 0.752 | 59.612 | 0.001 | -0.002 | 1.472 | 52.359 | 27.508 | 2.06 | 60.234 | 0.856 | -1.286 | 0 | 0 | 1 | 41.938 | 1 | 34 | 1 | GJB_Rinse | 0
161 | P-22673 | 2022/12/02 11:31:40 | 1-35 | 1 | -4.742 | -49.42 | 12009 | -1 | 1 | UNDE6-3040-3 | 22-284 | H2O | 1670006206.32 | 0.203 | 0.732 | 60.985 | 0.001 | 0.004 | 1.518 | 52.379 | 27.539 | 2.062 | 59.66 | 0.855 | -1.281 | 0 | 0 | 1 | 42.125 | 1 | 35 | 1 | GJB_Rinse | 0
162 | P-22673 | 2022/12/02 11:40:35 | 1-35 | 2 | -4.678 | -48.205 | 12301 | -1 | 1 | UNDE6-3040-3 | 22-284 | H2O | 1670006740.93 | 0.198 | 0.807 | 61.51 | 0.001 | 0 | 1.52 | 52.773 | 28.119 | 2.122 | 62.844 | 0.856 | -0.936 | 0 | 0 | 1 | 42.188 | 1 | 35 | 1 | GJB_Rinse | 0
163 | P-22673 | 2022/12/02 11:49:31 | 1-35 | 3 | -4.721 | -48 | 12179 | -1 | 1 | UNDE6-3040-3 | 22-284 | H2O | 1670007276.87 | 0.189 | 0.807 | 60.335 | 0.001 | 0.004 | 1.494 | 52.649 | 27.585 | 2.111 | 62.188 | 0.856 | -1.05 | 0 | 0 | 1 | 42.062 | 1 | 35 | 1 | GJB_Rinse | 0
164 | P-22673 | 2022/12/02 11:58:27 | 1-35 | 4 | -4.752 | -48.381 | 12017 | 0 | 1 | UNDE6-3040-3 | 22-284 | H2O | 1670007812.88 | 0.193 | 0.798 | 58.235 | 0.001 | 0 | 1.434 | 52.261 | 26.9 | 2.069 | 62.357 | 0.856 | -0.99 | 0 | 0 | 1 | 42.062 | 1 | 35 | 1 | GJB_Rinse | 0
165 | P-22674 | 2022/12/02 12:08:10 | 1-36 | 1 | -2.466 | -44.616 | 12092 | -1 | 1 | UNDE6-B-5 | 22-284 | H2O | 1670008396.14 | 0.201 | 0.811 | 59.052 | 0.001 | 0.002 | 1.47 | 53.005 | 27.111 | 2.124 | 62.501 | 0.855 | -1.217 | 0 | 0 | 1 | 42.188 | 1 | 36 | 1 | GJB_Rinse | 0
166 | P-22674 | 2022/12/02 12:17:06 | 1-36 | 2 | -2.44 | -44.852 | 11814 | -1 | 1 | UNDE6-B-5 | 22-284 | H2O | 1670008932.01 | 0.201 | 0.824 | 56.511 | 0.001 | -0.001 | 1.404 | 52.476 | 26.171 | 2.069 | 61.541 | 0.856 | -1.014 | 0 | 0 | 1 | 42.312 | 1 | 36 | 1 | GJB_Rinse | 0
167 | P-22674 | 2022/12/02 12:26:01 | 1-36 | 3 | -2.387 | -44.473 | 11791 | -1 | 1 | UNDE6-B-5 | 22-284 | H2O | 1670009467.64 | 0.198 | 0.796 | 62.131 | 0.001 | -0.001 | 1.546 | 52.489 | 26.366 | 2.05 | 60.386 | 0.855 | -1.222 | 0 | 0 | 1 | 42.312 | 1 | 36 | 1 | GJB_Rinse | 0
168 | P-22674 | 2022/12/02 12:34:57 | 1-36 | 4 | -2.405 | -44.262 | 11447 | 0 | 1 | UNDE6-B-5 | 22-284 | H2O | 1670010003.27 | 0.202 | 0.848 | 58.527 | 0.001 | -0.002 | 1.457 | 51.792 | 25.309 | 1.999 | 59.586 | 0.854 | -1.172 | 0 | 0 | 1 | 42.062 | 1 | 36 | 1 | GJB_Rinse | 0
169 | P-22675 | 2022/12/02 12:44:41 | 1-37 | 1 | -2.761 | -60.688 | 11865 | -1 | 1 | UNDE6-A-4 | 22-284 | H2O | 1670010586.29 | 0.202 | 0.749 | 54.922 | 0.001 | -0.001 | 1.368 | 62.273 | 24.063 | 2.267 | 62.098 | 0.856 | -2.194 | 0 | 0 | 1 | 42.125 | 1 | 37 | 1 | GJB_Rinse | 0
170 | P-22675 | 2022/12/02 12:53:36 | 1-37 | 2 | -2.877 | -62.313 | 12284 | -1 | 1 | UNDE6-A-4 | 22-284 | H2O | 1670011122.02 | 0.208 | 0.889 | 58.354 | 0.001 | 0.003 | 1.452 | 63.466 | 24.886 | 2.362 | 64.899 | 0.856 | -2.09 | 0 | 0 | 1 | 42.25 | 1 | 37 | 1 | GJB_Rinse | 0
171 | P-22675 | 2022/12/02 13:02:32 | 1-37 | 3 | -2.901 | -62.904 | 12338 | -1 | 1 | UNDE6-A-4 | 22-284 | H2O | 1670011658.03 | 0.181 | 0.782 | 61.361 | 0.001 | 0.003 | 1.517 | 63.753 | 25.074 | 2.355 | 64.121 | 0.856 | -2.191 | 0 | 0 | 1 | 42.312 | 1 | 37 | 1 | GJB_Rinse | 0
172 | P-22675 | 2022/12/02 13:11:28 | 1-37 | 4 | -2.898 | -63.134 | 12303 | 0 | 1 | UNDE6-A-4 | 22-284 | H2O | 1670012193.42 | 0.207 | 0.843 | 60.006 | 0.001 | 0 | 1.474 | 63.58 | 25.346 | 2.349 | 63.54 | 0.856 | -2.134 | 0 | 0 | 1 | 42.312 | 1 | 37 | 1 | GJB_Rinse | 0
173 | P-22676 | 2022/12/02 13:21:11 | 1-38 | 1 | -5.278 | -59.153 | 11861 | -1 | 1 | UNDE6-D-1 | 22-284 | H2O | 1670012776.74 | 0.192 | 0.794 | 60.167 | 0.001 | 0.005 | 1.497 | 52.056 | 26.205 | 2.052 | 62.152 | 0.855 | -0.87 | 0 | 0 | 1 | 42.125 | 1 | 38 | 1 | GJB_Rinse | 0
174 | P-22676 | 2022/12/02 13:30:07 | 1-38 | 2 | -5.353 | -59.096 | 12371 | -1 | 1 | UNDE6-D-1 | 22-284 | H2O | 1670013313.01 | 0.192 | 0.753 | 58.324 | 0.001 | 0 | 1.436 | 52.969 | 28.223 | 2.112 | 62.569 | 0.855 | -1.066 | 0 | 0 | 1 | 42.188 | 1 | 38 | 1 | GJB_Rinse | 0
175 | P-22676 | 2022/12/02 13:39:03 | 1-38 | 3 | -5.369 | -58.655 | 12059 | -1 | 1 | UNDE6-D-1 | 22-284 | H2O | 1670013848.77 | 0.204 | 0.756 | 62.174 | 0.001 | 0.002 | 1.546 | 52.415 | 27.488 | 2.073 | 60.423 | 0.856 | -1.212 | 0 | 0 | 1 | 42.25 | 1 | 38 | 1 | GJB_Rinse | 0
176 | P-22676 | 2022/12/02 13:47:59 | 1-38 | 4 | -5.373 | -58.759 | 12208 | 0 | 1 | UNDE6-D-1 | 22-284 | H2O | 1670014384.92 | 0.184 | 0.856 | 60.721 | 0 | 0.002 | 1.503 | 52.56 | 27.569 | 2.104 | 62.648 | 0.856 | -1.001 | 0 | 0 | 1 | 42.312 | 1 | 38 | 1 | GJB_Rinse | 0
177 | P-22677 | 2022/12/02 13:57:43 | 1-39 | 1 | -4.107 | -57.613 | 11536 | -1 | 1 | UNDE6-E-5 | 22-284 | H2O | 1670014968.95 | 0.203 | 0.863 | 58.364 | 0.001 | 0.001 | 1.449 | 51.273 | 25.899 | 1.972 | 58.753 | 0.856 | -1.043 | 0 | 0 | 1 | 42.375 | 1 | 39 | 1 | GJB_Rinse | 0
178 | P-22677 | 2022/12/02 14:06:39 | 1-39 | 2 | -4.084 | -57.826 | 11575 | -1 | 1 | UNDE6-E-5 | 22-284 | H2O | 1670015504.7 | 0.191 | 0.769 | 59.279 | 0.001 | 0.002 | 1.476 | 51.409 | 25.896 | 1.988 | 58.558 | 0.855 | -1.028 | 0 | 0 | 1 | 42.438 | 1 | 39 | 1 | GJB_Rinse | 0
179 | P-22677 | 2022/12/02 14:15:34 | 1-39 | 3 | -4.037 | -57.819 | 11619 | -1 | 1 | UNDE6-E-5 | 22-284 | H2O | 1670016040.24 | 0.208 | 0.736 | 57.107 | 0.001 | 0.002 | 1.42 | 51.353 | 26.232 | 1.992 | 59.378 | 0.856 | -1.061 | 0 | 0 | 1 | 42.312 | 1 | 39 | 1 | GJB_Rinse | 0
180 | P-22677 | 2022/12/02 14:24:30 | 1-39 | 4 | -4.015 | -57.699 | 11690 | 0 | 1 | UNDE6-E-5 | 22-284 | H2O | 1670016575.67 | 0.203 | 0.837 | 59.437 | 0.001 | 0.002 | 1.477 | 51.591 | 26.185 | 2.001 | 59.979 | 0.856 | -1.103 | 0 | 0 | 1 | 42.25 | 1 | 39 | 1 | GJB_Rinse | 0
181 | P-22678 | 2022/12/02 14:34:14 | 1-40 | 1 | -8.67 | -65.503 | 12674 | -1 | 1 | EV | standard | H2O | 1670017159.24 | 0.192 | 0.787 | 61.027 | 0.001 | 0.001 | 1.498 | 53.379 | 29.04 | 2.146 | 63.173 | 0.855 | -1.123 | 0 | 0 | 1 | 42.375 | 1 | 40 | 1 | GJB_Rinse | 0
182 | P-22678 | 2022/12/02 14:43:09 | 1-40 | 2 | -8.916 | -66.067 | 12254 | -1 | 1 | EV | standard | H2O | 1670017694.7 | 0.198 | 0.782 | 62.549 | 0.002 | -0.001 | 1.546 | 52.343 | 27.877 | 2.084 | 62.799 | 0.856 | -0.941 | 0 | 0 | 1 | 42.375 | 1 | 40 | 1 | GJB_Rinse | 0
183 | P-22678 | 2022/12/02 14:52:04 | 1-40 | 3 | -8.922 | -66.326 | 12301 | -1 | 1 | EV | standard | H2O | 1670018230.34 | 0.209 | 0.74 | 62.686 | 0.001 | 0 | 1.549 | 52.619 | 27.715 | 2.081 | 62.718 | 0.856 | -1.054 | 0 | 0 | 1 | 42.375 | 1 | 40 | 1 | GJB_Rinse | 0
184 | P-22678 | 2022/12/02 15:01:00 | 1-40 | 4 | -9.014 | -66.701 | 12240 | 0 | 1 | EV | standard | H2O | 1670018766.26 | 0.2 | 0.798 | 60.006 | 0.002 | 0.001 | 1.485 | 52.461 | 27.835 | 2.078 | 61.595 | 0.855 | -1.111 | 0 | 0 | 1 | 42.438 | 1 | 40 | 1 | GJB_Rinse | 0
185 | P-22679 | 2022/12/02 15:10:45 | 1-41 | 1 | -5.863 | -52.308 | 11253 | -1 | 1 | UNDE6-1020-2 | 22-284 | H2O | 1670019351.56 | 0.22 | 0.849 | 56.855 | 0 | 0.001 | 1.423 | 50.663 | 24.807 | 1.946 | 58.811 | 0.856 | -0.782 | 0 | 0 | 1 | 42.375 | 1 | 41 | 1 | GJB_Rinse | 0
186 | P-22679 | 2022/12/02 15:19:41 | 1-41 | 2 | -5.69 | -50.715 | 11352 | -1 | 1 | UNDE6-1020-2 | 22-284 | H2O | 1670019887.28 | 0.215 | 0.738 | 59.119 | 0.001 | -0.001 | 1.479 | 50.943 | 24.889 | 1.954 | 60.124 | 0.855 | -0.888 | 0 | 0 | 1 | 42.375 | 1 | 41 | 1 | GJB_Rinse | 0
187 | P-22679 | 2022/12/02 15:28:38 | 1-41 | 3 | -5.694 | -50.587 | 11398 | -1 | 1 | UNDE6-1020-2 | 22-284 | H2O | 1670020423.73 | 0.189 | 0.732 | 59.65 | 0.001 | 0 | 1.486 | 50.999 | 25.599 | 1.96 | 58.546 | 0.855 | -1.092 | 0 | 0 | 1 | 42.5 | 1 | 41 | 1 | GJB_Rinse | 0
188 | P-22679 | 2022/12/02 15:37:33 | 1-41 | 4 | -5.633 | -50.019 | 11557 | 0 | 1 | UNDE6-1020-2 | 22-284 | H2O | 1670020958.85 | 0.216 | 0.827 | 55.704 | 0.001 | -0.005 | 1.393 | 51.199 | 25.922 | 1.997 | 60.079 | 0.855 | -0.893 | 0 | 0 | 1 | 42.562 | 1 | 41 | 1 | GJB_Rinse | 0
189 | P-22680 | 2022/12/02 15:47:17 | 1-42 | 1 | -3.6 | -54.468 | 11775 | -1 | 1 | UNDE6-A-3 | 22-284 | H2O | 1670021542.97 | 0.203 | 0.758 | 55.014 | 0.001 | 0 | 1.366 | 52.181 | 26.248 | 2.032 | 60.392 | 0.856 | -1.159 | 0 | 0 | 1 | 42.438 | 1 | 42 | 1 | GJB_Rinse | 0
190 | P-22680 | 2022/12/02 15:56:13 | 1-42 | 2 | -3.337 | -54.536 | 12380 | -1 | 1 | UNDE6-A-3 | 22-284 | H2O | 1670022079.38 | 0.188 | 0.797 | 59.673 | 0.001 | -0.001 | 1.474 | 53.213 | 28.144 | 2.143 | 63.665 | 0.856 | -0.853 | 0 | 0 | 1 | 42.375 | 1 | 42 | 1 | GJB_Rinse | 0
191 | P-22680 | 2022/12/02 16:05:09 | 1-42 | 3 | -3.487 | -55.247 | 11638 | -1 | 1 | UNDE6-A-3 | 22-284 | H2O | 1670022615.28 | 0.2 | 0.738 | 54.481 | 0 | -0.001 | 1.359 | 51.902 | 25.716 | 2.02 | 60.595 | 0.856 | -0.983 | 0 | 0 | 1 | 42.438 | 1 | 42 | 1 | GJB_Rinse | 0
192 | P-22680 | 2022/12/02 16:14:05 | 1-42 | 4 | -3.53 | -55.274 | 11832 | 0 | 1 | UNDE6-A-3 | 22-284 | H2O | 1670023151.04 | 0.213 | 0.78 | 57.65 | 0.001 | -0.001 | 1.432 | 52.239 | 26.692 | 2.047 | 60.341 | 0.856 | -1.11 | 0 | 0 | 1 | 42.5 | 1 | 42 | 1 | GJB_Rinse | 0
193 | P-22681 | 2022/12/02 16:23:49 | 1-43 | 1 | -4.149 | -49.546 | 11898 | -1 | 1 | UNDE6-1020-3 | 22-284 | H2O | 1670023734.72 | 0.187 | 0.89 | 59.643 | 0 | -0.003 | 1.484 | 51.923 | 26.593 | 2.054 | 61.794 | 0.855 | -0.906 | 0 | 0 | 1 | 42.625 | 1 | 43 | 1 | GJB_Rinse | 0
194 | P-22681 | 2022/12/02 16:32:45 | 1-43 | 2 | -4.162 | -48.912 | 12310 | -1 | 1 | UNDE6-1020-3 | 22-284 | H2O | 1670024271.24 | 0.206 | 0.736 | 61.559 | 0.001 | 0.002 | 1.522 | 52.8 | 28.201 | 2.119 | 62.111 | 0.856 | -1.119 | 0 | 0 | 1 | 42.562 | 1 | 43 | 1 | GJB_Rinse | 0
195 | P-22681 | 2022/12/02 16:41:41 | 1-43 | 3 | -4.141 | -48.896 | 11996 | -1 | 1 | UNDE6-1020-3 | 22-284 | H2O | 1670024806.58 | 0.202 | 0.796 | 62.172 | 0.002 | -0.002 | 1.547 | 52.198 | 27.282 | 2.07 | 60.894 | 0.856 | -1.107 | 0 | 0 | 1 | 42.25 | 1 | 43 | 1 | GJB_Rinse | 0
196 | P-22681 | 2022/12/02 16:50:37 | 1-43 | 4 | -4.158 | -48.798 | 12126 | 0 | 1 | UNDE6-1020-3 | 22-284 | H2O | 1670025342.56 | 0.208 | 0.66 | 56.704 | 0.001 | -0.001 | 1.409 | 52.406 | 27.428 | 2.11 | 62.198 | 0.856 | -1.009 | 0 | 0 | 1 | 42.25 | 1 | 43 | 1 | GJB_Rinse | 0
197 | P-22682 | 2022/12/02 17:00:21 | 1-44 | 1 | -2.907 | -45.2 | 11630 | -1 | 1 | UNDE6-010-1 | 22-284 | H2O | 1670025927.31 | 0.2 | 0.831 | 55.389 | 0.001 | -0.004 | 1.363 | 51.185 | 26.027 | 2.018 | 61.044 | 0.855 | -0.831 | 0 | 0 | 1 | 42.5 | 1 | 44 | 1 | GJB_Rinse | 0
198 | P-22682 | 2022/12/02 17:09:17 | 1-44 | 2 | -2.913 | -45.263 | 11784 | -1 | 1 | UNDE6-010-1 | 22-284 | H2O | 1670026462.69 | 0.201 | 0.828 | 59.628 | 0.002 | -0.001 | 1.484 | 51.654 | 26.416 | 2.024 | 60.436 | 0.856 | -1.142 | 0 | 0 | 1 | 42.562 | 1 | 44 | 1 | GJB_Rinse | 0
199 | P-22682 | 2022/12/02 17:18:12 | 1-44 | 3 | -2.849 | -44.853 | 11618 | -1 | 1 | UNDE6-010-1 | 22-284 | H2O | 1670026997.6 | 0.178 | 0.708 | 58.338 | 0.001 | 0.001 | 1.449 | 51.251 | 26.054 | 2.005 | 59.789 | 0.855 | -1.047 | 0 | 0 | 1 | 42.438 | 1 | 44 | 1 | GJB_Rinse | 0
200 | P-22682 | 2022/12/02 17:27:07 | 1-44 | 4 | -2.853 | -44.906 | 11650 | 0 | 1 | UNDE6-010-1 | 22-284 | H2O | 1670027532.82 | 0.208 | 0.797 | 60.397 | 0.001 | 0 | 1.516 | 51.358 | 26.268 | 2.019 | 59.528 | 0.855 | -1.108 | 0 | 0 | 1 | 42.312 | 1 | 44 | 1 | GJB_Rinse | 0
201 | P-22683 | 2022/12/02 17:36:51 | 1-45 | 1 | -2.362 | -45.058 | 11497 | -1 | 1 | UNDE6-010-3 | 22-284 | H2O | 1670028116.82 | 0.198 | 0.694 | 55.832 | 0.001 | -0.004 | 1.398 | 51.182 | 25.47 | 1.989 | 59.638 | 0.856 | -1.123 | 0 | 0 | 1 | 42.375 | 1 | 45 | 1 | GJB_Rinse | 0
202 | P-22683 | 2022/12/02 17:45:46 | 1-45 | 2 | -2.318 | -44.958 | 11890 | -1 | 1 | UNDE6-010-3 | 22-284 | H2O | 1670028652.46 | 0.196 | 0.807 | 58.125 | 0 | 0.001 | 1.444 | 51.863 | 26.874 | 2.054 | 60.805 | 0.855 | -1.085 | 0 | 0 | 1 | 42.5 | 1 | 45 | 1 | GJB_Rinse | 0
203 | P-22683 | 2022/12/02 17:54:42 | 1-45 | 3 | -2.346 | -44.844 | 11848 | -1 | 1 | UNDE6-010-3 | 22-284 | H2O | 1670029187.37 | 0.186 | 0.926 | 61.264 | 0 | 0 | 1.524 | 51.878 | 26.704 | 2.043 | 59.996 | 0.856 | -1.118 | 0 | 0 | 1 | 42.562 | 1 | 45 | 1 | GJB_Rinse | 0
204 | P-22683 | 2022/12/02 18:03:37 | 1-45 | 4 | -2.244 | -44.69 | 11917 | 0 | 1 | UNDE6-010-3 | 22-284 | H2O | 1670029722.86 | 0.21 | 0.765 | 58.257 | 0.001 | 0 | 1.446 | 51.771 | 26.831 | 2.082 | 62.3 | 0.856 | -0.961 | 0 | 0 | 1 | 42.438 | 1 | 45 | 1 | GJB_Rinse | 0
205 | P-22684 | 2022/12/02 18:13:23 | 1-46 | 1 | -0.49 | -43.088 | 11934 | -1 | 1 | UNDE6-deep-1 | 22-284 | H2O | 1670030308.33 | 0.201 | 0.854 | 60.058 | 0.001 | 0.001 | 1.492 | 52.829 | 26.586 | 2.084 | 61.966 | 0.856 | -1.108 | 0 | 0 | 1 | 42.25 | 1 | 46 | 1 | GJB_Rinse | 0
206 | P-22684 | 2022/12/02 18:22:19 | 1-46 | 2 | -0.336 | -42.66 | 11692 | -1 | 1 | UNDE6-deep-1 | 22-284 | H2O | 1670030844.14 | 0.184 | 0.72 | 57.735 | 0 | 0.001 | 1.438 | 52.311 | 26.105 | 2.047 | 61.288 | 0.855 | -1.135 | 0 | 0 | 1 | 42.375 | 1 | 46 | 1 | GJB_Rinse | 0
207 | P-22684 | 2022/12/02 18:31:13 | 1-46 | 3 | -0.36 | -42.799 | 11612 | -1 | 1 | UNDE6-deep-1 | 22-284 | H2O | 1670031378.8 | 0.19 | 0.863 | 54.63 | 0.001 | -0.001 | 1.365 | 52.176 | 26.088 | 2.036 | 59.152 | 0.856 | -1.141 | 0 | 0 | 1 | 42 | 1 | 46 | 1 | GJB_Rinse | 0
208 | P-22684 | 2022/12/02 18:40:09 | 1-46 | 4 | -0.291 | -42.373 | 11690 | 0 | 1 | UNDE6-deep-1 | 22-284 | H2O | 1670031914.37 | 0.199 | 0.883 | 61.272 | 0.001 | 0.003 | 1.525 | 52.179 | 26.542 | 2.042 | 60.178 | 0.855 | -1.059 | 0 | 0 | 1 | 42 | 1 | 46 | 1 | GJB_Rinse | 0
209 | P-22685 | 2022/12/02 18:49:53 | 1-47 | 1 | -3.453 | -49.154 | 11824 | -1 | 1 | UNDE6-deep-2 | 22-284 | H2O | 1670032499.63 | 0.204 | 0.832 | 61.165 | 0.001 | 0.001 | 1.52 | 52.127 | 26.49 | 2.055 | 62.037 | 0.855 | -0.874 | 0 | 0 | 1 | 42.062 | 1 | 47 | 1 | GJB_Rinse | 0
210 | P-22685 | 2022/12/02 18:58:49 | 1-47 | 2 | -3.636 | -49.829 | 11845 | -1 | 1 | UNDE6-deep-2 | 22-284 | H2O | 1670033035.47 | 0.203 | 0.833 | 57.046 | 0.001 | 0.001 | 1.418 | 52.341 | 26.387 | 2.051 | 60.938 | 0.856 | -1.089 | 0 | 0 | 1 | 42 | 1 | 47 | 1 | GJB_Rinse | 0
211 | P-22685 | 2022/12/02 19:07:45 | 1-47 | 3 | -3.631 | -50.242 | 11557 | -1 | 1 | UNDE6-deep-2 | 22-284 | H2O | 1670033570.65 | 0.209 | 0.814 | 56.286 | 0.001 | 0.003 | 1.402 | 51.662 | 25.826 | 2.01 | 59.564 | 0.856 | -1.053 | 0 | 0 | 1 | 42.062 | 1 | 47 | 1 | GJB_Rinse | 0
212 | P-22685 | 2022/12/02 19:16:40 | 1-47 | 4 | -3.674 | -50.086 | 11620 | 0 | 1 | UNDE6-deep-2 | 22-284 | H2O | 1670034105.96 | 0.19 | 0.769 | 55.709 | 0.001 | 0.002 | 1.391 | 51.857 | 25.618 | 2.02 | 60.794 | 0.856 | -0.941 | 0 | 0 | 1 | 42.25 | 1 | 47 | 1 | GJB_Rinse | 0
213 | P-22686 | 2022/12/02 19:26:25 | 1-48 | 1 | -2.397 | -42.692 | 11725 | -1 | 1 | UNDE6-deep-3 | 22-284 | H2O | 1670034691.5 | 0.216 | 0.772 | 57.34 | 0.001 | 0.004 | 1.427 | 52.117 | 26.167 | 2.038 | 61.229 | 0.856 | -1.017 | 0 | 0 | 1 | 42.438 | 1 | 48 | 1 | GJB_Rinse | 0
214 | P-22686 | 2022/12/02 19:35:21 | 1-48 | 2 | -2.133 | -41.206 | 12555 | -1 | 1 | UNDE6-deep-3 | 22-284 | H2O | 1670035226.4 | 0.224 | 0.799 | 59.133 | 0.001 | 0 | 1.454 | 53.678 | 28.734 | 2.194 | 64.858 | 0.856 | -0.983 | 0 | 0 | 1 | 42.25 | 1 | 48 | 1 | GJB_Rinse | 0
215 | P-22686 | 2022/12/02 19:44:15 | 1-48 | 3 | -2.315 | -41.345 | 11689 | -1 | 1 | UNDE6-deep-3 | 22-284 | H2O | 1670035761.39 | 0.225 | 0.747 | 56.68 | 0.002 | 0.004 | 1.41 | 52.073 | 26.326 | 2.04 | 60.285 | 0.855 | -1.158 | 0 | 0 | 1 | 42.062 | 1 | 48 | 1 | GJB_Rinse | 0
216 | P-22686 | 2022/12/02 19:53:10 | 1-48 | 4 | -2.37 | -41.729 | 12309 | 0 | 1 | UNDE6-deep-3 | 22-284 | H2O | 1670036295.74 | 0.214 | 0.799 | 58.966 | 0.002 | 0.002 | 1.458 | 53.334 | 27.939 | 2.132 | 63.228 | 0.856 | -1.229 | 0 | 0 | 1 | 42.125 | 1 | 48 | 1 | GJB_Rinse | 0
217 | P-22687 | 2022/12/02 20:02:56 | 1-49 | 1 | -8.57 | -63.388 | 12615 | -1 | 1 | EV | standard | H2O | 1670036881.11 | 0.187 | 0.747 | 59.673 | 0.001 | -0.002 | 1.471 | 53.059 | 28.958 | 2.149 | 64.019 | 0.856 | -0.903 | 0 | 0 | 1 | 42.062 | 1 | 49 | 1 | GJB_Rinse | 0
218 | P-22687 | 2022/12/02 20:11:51 | 1-49 | 2 | -8.842 | -65.386 | 12218 | -1 | 1 | EV | standard | H2O | 1670037416.58 | 0.213 | 0.793 | 57.43 | 0.001 | 0.001 | 1.427 | 52.413 | 27.509 | 2.076 | 62.191 | 0.856 | -1.018 | 0 | 0 | 1 | 42.062 | 1 | 49 | 1 | GJB_Rinse | 0
219 | P-22687 | 2022/12/02 20:20:47 | 1-49 | 3 | -8.919 | -65.768 | 12265 | -1 | 1 | EV | standard | H2O | 1670037952.47 | 0.196 | 0.81 | 60.982 | 0.001 | -0.001 | 1.517 | 52.529 | 27.65 | 2.078 | 62.038 | 0.855 | -1.029 | 0 | 0 | 1 | 42.25 | 1 | 49 | 1 | GJB_Rinse | 0
220 | P-22687 | 2022/12/02 20:29:42 | 1-49 | 4 | -8.904 | -66.263 | 12241 | 0 | 1 | EV | standard | H2O | 1670038488.13 | 0.19 | 0.783 | 64.439 | 0.001 | 0.004 | 1.594 | 52.316 | 28.016 | 2.071 | 61.735 | 0.857 | -0.996 | 0 | 0 | 1 | 42.312 | 1 | 49 | 1 | GJB_Rinse | 0

## Instructions
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Variable | Instruction | Comment
StandardsFile | standards file.csv
OutputBase | ChemCorrect
CorrectionType | M
timecode | Time Code
identifier | Identifier 1
sample | Sample
injection | Inj Nr | Inj Nr
line | Line
job | Job
Variable | Instruction | Comment
baseline_shift_MEAN | mean_standards(baseline_shift) | The mean of the organic baseline of the standards.
baseline_shift_STD | std_standards(baseline_shift) | The standard deviation of the organic baseline of the standards.
ch4_ppm_MEAN | mean_standards(ch4_ppm) | The mean methane concentration of the standards.
ch4_ppm_STD | std_standards(ch4_ppm) | The standard deviation of the methane concentrations of the standards.
residuals_MEAN | mean_standards(residuals) | The mean difference between the measured and modeled spectra of the standards.
residuals_STD | std_standards(residuals) | The standard deviation of the difference between measured and modeled spectra of the standards.
slope_shift_MEAN | mean_standards(slope_shift) | The mean organic slope of the standards.
slope_shift_STD | std_standards(slope_shift) | The standard deviation of the organic slope of the standards.
baseline_curvature_MEAN | mean_standards(baseline_curvature) | The mean of the baseline curvature of the standards.
baseline_curvature_STD | std_standards(baseline_curvature) | The standard deviation of the baseline curvature of the standards.
residuals_MEAN_ABS | abs(residuals_MEAN) | The absolute value of the mean difference between the measured and modeled spectra of the standards.
baseline_shift_MEAN_ABS | abs(baseline_shift_MEAN) | The absolute value of the mean of the organic baseline of the standards.
ch4_ppm_MEAN_ABS | abs(ch4_ppm_MEAN) | The absolute value of the mean methane concentration of the standards.
slope_shift_MEAN_ABS | abs(slope_shift_MEAN) | The absolute value of the mean organic slope of the standards.
residuals_15SIGMA | evaluation(residuals_STD*1.5) | One and half times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_15SIGMA | evaluation(residuals_MEAN+residuals_15SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_2SIGMA | evaluation(residuals_STD*2) | Twice the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_2SIGMA | evaluation(residuals_MEAN+residuals_2SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_3SIGMA | evaluation(residuals_STD*3) | Three times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_3SIGMA | evaluation(residuals_MEAN+residuals_3SIGMA) | The mean difference between measured and modeled spectra of the standards plus three standard deviation.
residuals_10PCT | evaluation(residuals_MEAN*0.1) | 30% of the difference between measured and modeled spectra of the standards.
ch4_ppm_3SIGMA | evaluation(ch4_ppm_STD*3) | Three times the standard deviation of the methanol amplitudes of the standards.
ch4_ppm_PLUS_3SIGMA | evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA) | The mean methanol amplitude of the standards plus three standard deviation.
HALF | float(0.5) | A fixed value of 0.5 or half.
baseline_shift_HALFSIGMA | evaluation(baseline_shift_STD*HALF) | Half the standard deviation of the organic baseline of the standards.
baseline_shift_PLUS_HALFSIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA) | The absolute value of the mean organic baseline of the standards plus one-half standard deviation.
baseline_shift_PLUS_1SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD) | The absolute value of the mean organic baseline of the standards plus one standard deviation.
baseline_shift_3SIGMA | evaluation(baseline_shift_STD*3) | Three times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_3SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA) | The absolute value of the mean organic baseline of the standards plus three standard deviation.
baseline_shift_18SIGMA | evaluation(baseline_shift_STD*18) | Eighteen times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_18SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA) | The absolute value of the mean organic baseline of the standards plus eighteen standard deviation.
slope_shift_5SIGMA | evaluation(slope_shift_STD*5) | Five times the standard deviation of the organic slope of the standards.
slope_shift_PLUS_5SIGMA | evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA) | The absolute value of the mean organic slope of the standards plus five standard deviation.
baseline_curvature_3SIGMA | evaluation(baseline_curvature_STD*3) | Three times the standard deviation of the baseline curvature of the standards.
baseline_curvature_PLUS_3SIGMA | evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
TRUE | evaluation(1>0)
FALSE | evaluation(1<0)
PROC_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the uncalibrated d18O value in order to calibrate it.
PROC_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the uncalibrated dD value in order to calibrate it.
Variable | Instruction | Comment
IS_STANDARD | is_standard_sample() | Standard recognition.
IS_NOT_STANDARD | bool_xor(IS_STANDARD,TRUE)
SAMP_residuals_MEAN | mean_all(residuals) | The mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_residuals_STD | std_all(residuals) | One standard deviation of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN | mean_all(ch4_ppm) | The mean methane concentration of a sample.
SAMP_ch4_ppm_STD | std_all(ch4_ppm) | The standard deviation of the methane concentrations of a sample.
SAMP_baseline_shift_MEAN | mean_all(baseline_shift) | The mean of the organic baseline of a sample.
SAMP_baseline_shift_STD | std_all(baseline_shift) | The standard deviation of the organic baseline of a sample.
SAMP_slope_shift_MEAN | mean_all(slope_shift) | The mean organic slope of a sample.
SAMP_slope_shift_STD | std_all(slope_shift) | The standard deviation of the organic slope of a sample.
SAMP_d18O_STD | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
SAMP_dD_STD | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_baseline_curvature_MEAN | mean_all(baseline_curvature) | The mean of the baseline curvature of the ssamples
SAMP_baseline_curvature_STD | std_all(baseline_curvature) | The standard deviation of the baseline curvature of the samples
SAMP_residuals_MEAN_ABS | abs(SAMP_residuals_MEAN) | Absolute value of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN_ABS | abs(SAMP_ch4_ppm_MEAN) | Absolute value of the mean methane concentration of a sample.
SAMP_baseline_shift_MEAN_ABS | abs(SAMP_baseline_shift_MEAN) | Absolute value of the mean organic baseline of a sample.
SAMP_slope_shift_MEAN_ABS | abs(SAMP_slope_shift_MEAN) | Absolute value of the mean organic slope of a sample.
SAMP_residuals_DIFFERENTIAL | evaluation(SAMP_residuals_MEAN-residuals_MEAN)
SAMP_residuals_DIFFERENTIAL_ABS | abs(SAMP_residuals_DIFFERENTIAL) | Absolute value of the difference between the mean of the sample residual and the mean of all standards residuals
SAMP_residuals_DIFFERENTIAL_PCT | evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
SAMP_baseline_curvature_DIFFERENTIAL | evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
SAMP_STD_DELTA_18_16O_CURR_MEAN | current_standards_mean(d(18_16)Mean) | The d18O mean of the last run standard.
SAMP_STD_DELTA_DH_CURR_MEAN | current_standards_mean(d(D_H)Mean) | The dD mean of the last run standard.
known_dD | known_standard(d(D_H)Mean) | The dD value of the known Standard.
known_d18O | known_standard(d(18_16)Mean) | The d18O value of the known Standard.
d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
measured_d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
measured_dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
measured_d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
measured_dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the mean uncalibrated d18O value in order to calibrate it.
SAMP_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the mean uncalibrated dD value in order to calibrate it.
SAMP_TEST_VALUE | evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
residuals_GT_3SIGMA | evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA) | Threshold of red trigger. If the mean difference between the measured spectra of the sample and the modeled spectra of the standard is greater than that of the mean difference of the standard plus three standard deviations, this will trigger.
residuals_GE_2SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GE_15SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GT_10PCT | evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
ch4_ppm_GE_3SIGMA | evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA) | True if the mean methane concentration of the sample is greater than or equal to that of the standard plus three standard deviations, otherwise false.
baseline_shift_GT_1SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus one standard deviation, otherwise false.
baseline_shift_GT_3SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus three standard deviation, otherwise false. Sets yellow trigger for "other", trigger threshold is approximately sample to sample dD precision specification
baseline_shift_GT_18SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus eighteen standard deviation, otherwise false. Sets red trigger for "other", trigger threshold is approximately 24 hour dD drift specification
slope_shift_GT_5SIGMA | evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA) | True if the mean organic slope of the sample is greater than that of the standard plus five standard deviations, otherwise false.
baseline_curvature_GT_3SIGMA | evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
RED | truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA)) | Red trigger. If the sample is not a standard and the threshold is triggered (measured-modeled > standard+3stdev's), then RED.
NOT_RED | bool_xor(RED,TRUE) | If the sample is determined to be "red" than it can not be "not red."
YELLOW | truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA))) | Truth evaluation used to determine whether or not a given sample is yellow.
NOT_YELLOW | bool_xor(YELLOW,TRUE) | If sample is determined to be "yellow" than it can not be "not yellow."
GREEN | truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW)) | This is the pre-requisite for a sample to be colored green. It must not be a standard, must not be yellow, and must not be red.
NOT_GREEN | bool_xor(GREEN,TRUE) | Pre-requisite for being "not-green."
CYAN | truth_evaluation((IS_STANDARD)) | This is the pre-requisite for a sample to be a standard and therefore cyan colored.
NOT_CYAN | bool_xor(CYAN,TRUE) | Pre-requisite for being not a standard.
METHANE | truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA)) | This is the threshold for amounts of methane being included in the main ChemCorrect spreadsheet. Methane will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured methane concentration must be greater than or equal to that of the standard plus three standard deviations.
OTHER1 | truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA)) | This is the threshold for amounts of "OTHER1" being included in the main ChemCorrect spreadsheet. This column (OTHER1) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic baseline must be greater than that of the standard plus one standard deviation.
OTHER2 | truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA)) | This is the threshold for amounts of "OTHER2" being included in the main ChemCorrect spreadsheet. This column (OTHER2) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic slope must be greater than that of the standard plus three standard deviations.
OTHER3 | truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | truth_evaluation(RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN)) | Organic baseline differential showing the difference between the measured baseline and the standard baseline.
Variable | Instruction | Comment
DTL_TEST | math_evaluation((1 + 2)) | Test evaluation of 1+2 being equal to 3.
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample | FALSE | FALSE | FALSE | FALSE | Testing Sample
B | identifier | TRUE | Name | RED | YELLOW | GREEN | CYAN | Testing Name
C | SAMP_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d<sup>18</sup>O Mean | RED | YELLOW | GREEN | CYAN | Testing18 16 Mean
D | SAMP_DELTA_DH_CALIBRATE | TRUE | Calibrated d<sup>2</sup>H Mean | RED | YELLOW | GREEN | CYAN | This is a test to see that I can use tool tips
E | SAMP_ch4_ppm_MEAN | METHANE | CH<sub>4</sub> | RED | YELLOW | FALSE | FALSE
F | baseline_shift_DIFF | OTHER1 | C<sub>2+</sub> alcohols | RED | YELLOW | FALSE | FALSE
G | SAMP_residuals_DIFFERENTIAL_PCT | OTHER3 | relative %deviation | RED | YELLOW | FALSE | FALSE
H | SAMP_d18O_STD | TRUE | Uncalibrated d<sup>18</sup>O precision | RED | YELLOW | GREEN | CYAN
I | SAMP_dD_STD | TRUE | Uncalibrated d<sup>2</sup>H precision | RED | YELLOW | GREEN | CYAN
J | SAMP_slope_shift_MEAN | OTHER2 | slope | RED | YELLOW | FALSE | FALSE
K | SAMP_baseline_curvature_DIFFERENTIAL | OTHER4 | curvature | RED | YELLOW | FALSE | FALSE
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample
B | identifier | TRUE | Name
C | col(d(18_16)Mean) | TRUE | Measured d(18_16) Mean
D | col(d(D_H)Mean) | TRUE | Measured d(D_H) Mean
E | DTL_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d(18_16) Mean
F | DTL_DELTA_DH_CALIBRATE | TRUE | Calibrated d(D_H) Mean
G | col(ch4_ppm) | METHANE | CH<sub>4</sub>
H | col(baseline_shift) | OTHER1 | Other
I | col(Time Code) | TRUE | Time Code
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | d18O_mean | TRUE | d(18_16) Mean | RED | YELLOW | GREEN | CYAN
B | dD_mean | TRUE | d(D_H) Mean | RED | YELLOW | GREEN | CYAN
C | XY_X(d18O_mean) XY_Y(dD_mean) | TRUE | d18O x dD | gmwl | best_fit

## Process Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Process Variables
Variable | Value | Operations
baseline_shift_MEAN | 63.24145
varray: [56.725, 61.151, 62.387, 61.23, 60.994, 61.493, 60.963, 61.254, 61.744, 61.811, 57.918, 64.733, 66.678, 66.492, 66.397, 65.23, 64.931, 65.122, 65.529, 72.047]
inst: mean_standards
operation: mean_standards(baseline_shift)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0, 10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: baseline_shift
baseline_shift_STD | 3.47291266314
varray: [56.725, 61.151, 62.387, 61.23, 60.994, 61.493, 60.963, 61.254, 61.744, 61.811, 57.918, 64.733, 66.678, 66.492, 66.397, 65.23, 64.931, 65.122, 65.529, 72.047]
inst: std_standards
operation: std_standards(baseline_shift)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0, 10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: baseline_shift
ch4_ppm_MEAN | -0.99735
varray: [-1.028, -1.026, -1.118, -1.221, -0.906, -1.009, -1.069, -1.063, -1.118, -1.014, -0.887, -0.998, -1.112, -0.792, -0.88, -0.905, -0.997, -0.815, -0.895, -1.094]
inst: mean_standards
operation: mean_standards(ch4_ppm)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0, 10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: ch4_ppm
ch4_ppm_STD | 0.112846553943
varray: [-1.028, -1.026, -1.118, -1.221, -0.906, -1.009, -1.069, -1.063, -1.118, -1.014, -0.887, -0.998, -1.112, -0.792, -0.88, -0.905, -0.997, -0.815, -0.895, -1.094]
inst: std_standards
operation: std_standards(ch4_ppm)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0, 10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: ch4_ppm
residuals_MEAN | 2.22805
varray: [2.346, 2.316, 2.288, 2.303, 2.306, 2.347, 2.325, 2.303, 2.306, 2.318, 2.16, 2.099, 2.143, 2.157, 2.153, 2.133, 2.138, 2.155, 2.127, 2.138]
inst: mean_standards
operation: mean_standards(residuals)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0, 10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: residuals
residuals_STD | 0.0918428148065
varray: [2.346, 2.316, 2.288, 2.303, 2.306, 2.347, 2.325, 2.303, 2.306, 2.318, 2.16, 2.099, 2.143, 2.157, 2.153, 2.133, 2.138, 2.155, 2.127, 2.138]
inst: std_standards
operation: std_standards(residuals)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0, 10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: residuals
slope_shift_MEAN | 29.61975
varray: [30.507, 29.724, 29.303, 29.227, 29.218, 30.082, 29.477, 29.334, 29.503, 29.644, 29.915, 29.066, 29.778, 29.882, 29.478, 29.348, 29.739, 29.936, 29.514, 29.72]
inst: mean_standards
operation: mean_standards(slope_shift)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0, 10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: slope_shift
slope_shift_STD | 0.344966035704
varray: [30.507, 29.724, 29.303, 29.227, 29.218, 30.082, 29.477, 29.334, 29.503, 29.644, 29.915, 29.066, 29.778, 29.882, 29.478, 29.348, 29.739, 29.936, 29.514, 29.72]
inst: std_standards
operation: std_standards(slope_shift)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0, 10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: slope_shift
baseline_curvature_MEAN | 65.16435
varray: [68.285, 67.105, 66.257, 65.436, 66.605, 67.42, 66.577, 66.221, 66.162, 66.331, 64.692, 62.328, 63.115, 64.694, 64.466, 63.507, 63.664, 64.359, 62.912, 63.151]
inst: mean_standards
operation: mean_standards(baseline_curvature)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0, 10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: baseline_curvature
baseline_curvature_STD | 1.70685942057
varray: [68.285, 67.105, 66.257, 65.436, 66.605, 67.42, 66.577, 66.221, 66.162, 66.331, 64.692, 62.328, 63.115, 64.694, 64.466, 63.507, 63.664, 64.359, 62.912, 63.151]
inst: std_standards
operation: std_standards(baseline_curvature)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0, 10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: baseline_curvature
residuals_MEAN_ABS | 2.22805
operation: abs(residuals_MEAN)
inst: abs
source_col: residuals_MEAN
baseline_shift_MEAN_ABS | 63.24145
operation: abs(baseline_shift_MEAN)
inst: abs
source_col: baseline_shift_MEAN
ch4_ppm_MEAN_ABS | 0.99735
operation: abs(ch4_ppm_MEAN)
inst: abs
source_col: ch4_ppm_MEAN
slope_shift_MEAN_ABS | 29.61975
operation: abs(slope_shift_MEAN)
inst: abs
source_col: slope_shift_MEAN
residuals_15SIGMA | 0.13776422221
operation: evaluation(residuals_STD*1.5)
inst: evaluation
factors: residuals_STD*1.5
residuals_PLUS_15SIGMA | 2.36581422221
operation: evaluation(residuals_MEAN+residuals_15SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_15SIGMA
residuals_2SIGMA | 0.183685629613
operation: evaluation(residuals_STD*2)
inst: evaluation
factors: residuals_STD*2
residuals_PLUS_2SIGMA | 2.41173562961
operation: evaluation(residuals_MEAN+residuals_2SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_2SIGMA
residuals_3SIGMA | 0.275528444419
operation: evaluation(residuals_STD*3)
inst: evaluation
factors: residuals_STD*3
residuals_PLUS_3SIGMA | 2.50357844442
operation: evaluation(residuals_MEAN+residuals_3SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_3SIGMA
residuals_10PCT | 0.222805
operation: evaluation(residuals_MEAN*0.1)
inst: evaluation
factors: residuals_MEAN*0.1
ch4_ppm_3SIGMA | 0.338539661829
operation: evaluation(ch4_ppm_STD*3)
inst: evaluation
factors: ch4_ppm_STD*3
ch4_ppm_PLUS_3SIGMA | 1.33588966183
operation: evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA)
inst: evaluation
factors: ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA
HALF | 0.5
operation: float(0.5)
inst: float
baseline_shift_HALFSIGMA | 1.73645633157
operation: evaluation(baseline_shift_STD*HALF)
inst: evaluation
factors: baseline_shift_STD*HALF
baseline_shift_PLUS_HALFSIGMA | 64.9779063316
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA
baseline_shift_PLUS_1SIGMA | 66.7143626631
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_STD
baseline_shift_3SIGMA | 10.4187379894
operation: evaluation(baseline_shift_STD*3)
inst: evaluation
factors: baseline_shift_STD*3
baseline_shift_PLUS_3SIGMA | 73.6601879894
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_3SIGMA
baseline_shift_18SIGMA | 62.5124279365
operation: evaluation(baseline_shift_STD*18)
inst: evaluation
factors: baseline_shift_STD*18
baseline_shift_PLUS_18SIGMA | 125.753877936
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_18SIGMA
slope_shift_5SIGMA | 1.72483017852
operation: evaluation(slope_shift_STD*5)
inst: evaluation
factors: slope_shift_STD*5
slope_shift_PLUS_5SIGMA | 31.3445801785
operation: evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA)
inst: evaluation
factors: slope_shift_MEAN_ABS+slope_shift_5SIGMA
baseline_curvature_3SIGMA | 5.1205782617
operation: evaluation(baseline_curvature_STD*3)
inst: evaluation
factors: baseline_curvature_STD*3
baseline_curvature_PLUS_3SIGMA | 70.2849282617
operation: evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
inst: evaluation
factors: baseline_curvature_3SIGMA+baseline_curvature_MEAN
TRUE | True
operation: evaluation(1>0)
inst: evaluation
factors: 1>0
FALSE | False
operation: evaluation(1<0)
inst: evaluation
factors: 1<0
PROC_DELTA_18_16O_CALIBRATE | -5.85385350876
actual_array: [ 1.93 -15.84]
varray: [-9.875, -9.474, -9.258, -9.183, -9.176, -9.107, -9.126, -9.074, -9.043, -9.082, 2.239, 2.687, 2.769, 2.883, 2.846, 2.873, 2.926, 2.923, 2.947, 2.974, -13.561, -14.282, -14.497, -14.482, -14.55, -14.588, -14.593, -14.611, -14.604, -14.601, -9.327, -9.156, -9.052, -9.074, -9.101, -9.048, -9.08, -9.035, -9.012, -9.015, -5.591, -5.456, -5.322, -5.422, 0.0, 0.0, 0.0, 0.0, -4.239, -4.083, -4.043, -4.111, -6.127, -6.186, -6.245, -6.258, -3.245, -3.169, -3.172, -3.135, -6.441, -6.587, -6.632, -6.632, -2.788, -2.559, -2.605, -2.574, -5.588, -5.692, -5.741, -5.738, -8.869, -8.941, -8.982, -8.944, -5.597, -5.497, -5.459, -5.379, 1.341, 1.406, 1.407, 1.395, -3.241, -3.166, -3.145, -3.132, -3.725, -3.8, -3.847, -3.786, -2.46, -2.343, -2.296, -2.366, -5.292, -5.438, -5.434, -5.455, -3.562, -3.522, -3.527, -3.507, -3.75, -3.726, -3.77, -3.724, -8.7, -8.893, -8.916, -8.986, -3.455, -3.296, -3.212, -3.244, -1.585, -1.478, -1.438, -1.426, -4.721, -5.031, -5.014, -5.07, -5.244, -5.223, -5.2, -5.2, -3.142, -3.121, -3.047, -3.078, -2.742, -2.828, -2.846, -2.884, -5.827, -5.828, -5.816, -5.918, -2.102, -1.949, -1.84, -1.903, -8.618, -8.832, -8.947, -8.923, -3.907, -3.661, -3.625, -3.609, -0.803, -0.618, -0.617, -0.61, -4.408, -4.667, -4.682, -4.679, -4.742, -4.678, -4.721, -4.752, -2.466, -2.44, -2.387, -2.405, -2.761, -2.877, -2.901, -2.898, -5.278, -5.353, -5.369, -5.373, -4.107, -4.084, -4.037, -4.015, -8.67, -8.916, -8.922, -9.014, -5.863, -5.69, -5.694, -5.633, -3.6, -3.337, -3.487, -3.53, -4.149, -4.162, -4.141, -4.158, -2.907, -2.913, -2.849, -2.853, -2.362, -2.318, -2.346, -2.244, -0.49, -0.336, -0.36, -0.291, -3.453, -3.636, -3.631, -3.674, -2.397, -2.133, -2.315, -2.37, -8.57, -8.842, -8.919, -8.904]
fit_rsq: 1.0
fit_intercept: 0.933872312887
except: True
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -4.74657272727
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [1.0, 537.0, 1074.0, 1609.0, 2145.0, 2682.0, 3218.0, 3754.0, 4290.0, 4826.0, 5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0, 10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0, 16244.0, 16780.0, 17316.0, 17853.0, 18390.0, 18926.0, 19462.0, 19999.0, 20535.0, 21071.0, 21661.0, 22198.0, 22735.0, 23271.0, 23861.0, 24460.0, 25058.0, 25657.0, 26309.0, 26846.0, 27382.0, 27918.0, 28505.0, 29041.0, 29577.0, 30113.0, 30700.0, 31237.0, 31774.0, 32310.0, 32899.0, 33436.0, 33972.0, 34509.0, 35097.0, 35634.0, 36170.0, 36707.0, 37296.0, 37832.0, 38368.0, 38904.0, 39494.0, 40030.0, 40566.0, 41101.0, 41691.0, 42228.0, 42764.0, 43301.0, 43887.0, 44424.0, 44960.0, 45496.0, 46082.0, 46618.0, 47155.0, 47691.0, 48278.0, 48814.0, 49350.0, 49886.0, 50474.0, 51011.0, 51547.0, 52083.0, 52671.0, 53208.0, 53745.0, 54281.0, 54869.0, 55406.0, 55942.0, 56479.0, 57067.0, 57603.0, 58139.0, 58675.0, 59260.0, 59796.0, 60332.0, 60868.0, 61454.0, 61991.0, 62526.0, 63062.0, 63649.0, 64185.0, 64721.0, 65258.0, 65845.0, 66380.0, 66917.0, 67453.0, 68040.0, 68576.0, 69111.0, 69648.0, 70234.0, 70770.0, 71306.0, 71842.0, 72429.0, 72966.0, 73502.0, 74038.0, 74622.0, 75158.0, 75693.0, 76228.0, 76812.0, 77348.0, 77883.0, 78418.0, 79004.0, 79540.0, 80076.0, 80612.0, 81198.0, 81733.0, 82269.0, 82804.0, 83390.0, 83925.0, 84460.0, 84996.0, 85583.0, 86119.0, 86656.0, 87192.0, 87777.0, 88312.0, 88848.0, 89384.0, 89967.0, 90503.0, 91038.0, 91574.0, 92158.0, 92693.0, 93229.0, 93765.0, 94348.0, 94884.0, 95420.0, 95956.0, 96540.0, 97076.0, 97611.0, 98147.0, 98731.0, 99266.0, 99801.0, 100337.0, 100922.0, 101458.0, 101995.0, 102530.0, 103114.0, 103650.0, 104186.0, 104722.0, 105306.0, 105842.0, 106378.0, 106914.0, 107498.0, 108034.0, 108569.0, 109104.0, 109688.0, 110223.0, 110759.0, 111294.0, 111880.0, 112416.0, 112950.0, 113486.0, 114070.0, 114606.0, 115142.0, 115677.0, 116262.0, 116798.0, 117332.0, 117867.0, 118453.0, 118988.0, 119524.0, 120059.0]
source_col: d(18_16)Mean
PROC_DELTA_DH_CALIBRATE | -59.4669896726
actual_array: [ 18.1 -119. ]
varray: [-75.572, -71.771, -69.755, -68.948, -68.891, -68.321, -68.064, -67.978, -67.716, -67.968, 8.309, 15.198, 16.918, 18.143, 18.353, 18.732, 19.289, 19.124, 19.488, 19.713, -95.448, -105.739, -108.726, -109.713, -110.198, -110.9, -111.092, -111.537, -111.624, -111.94, -72.374, -69.081, -68.062, -67.733, -68.034, -67.499, -67.604, -67.439, -67.402, -67.24, -55.307, -53.775, -53.214, -53.47, 0.0, 0.0, 0.0, 0.0, -56.017, -55.354, -55.548, -55.697, -53.044, -52.282, -52.402, -52.378, -46.704, -46.343, -46.194, -46.294, -55.51, -56.316, -56.394, -56.684, -52.644, -52.147, -52.525, -52.372, -51.622, -51.495, -51.364, -51.408, -65.125, -66.111, -66.619, -66.315, -54.193, -53.342, -53.187, -52.523, -52.228, -52.915, -53.169, -53.133, -51.852, -50.439, -50.552, -50.325, -57.535, -58.425, -58.676, -58.623, -54.253, -53.593, -53.214, -53.529, -63.28, -64.535, -64.694, -64.758, -55.128, -54.363, -54.059, -53.887, -55.353, -55.565, -55.688, -55.834, -65.266, -66.285, -66.492, -66.725, -56.143, -55.427, -54.733, -55.128, -61.114, -61.639, -61.612, -61.752, -60.91, -61.468, -61.274, -61.303, -51.089, -50.093, -49.822, -49.439, -47.208, -47.125, -46.791, -46.829, -54.752, -55.596, -56.053, -56.223, -55.832, -55.204, -55.309, -55.445, -56.55, -56.768, -56.174, -56.736, -65.528, -66.347, -66.773, -66.332, -56.875, -55.672, -55.581, -55.409, -45.759, -44.528, -44.204, -44.173, -57.993, -59.737, -60.015, -60.133, -49.42, -48.205, -48.0, -48.381, -44.616, -44.852, -44.473, -44.262, -60.688, -62.313, -62.904, -63.134, -59.153, -59.096, -58.655, -58.759, -57.613, -57.826, -57.819, -57.699, -65.503, -66.067, -66.326, -66.701, -52.308, -50.715, -50.587, -50.019, -54.468, -54.536, -55.247, -55.274, -49.546, -48.912, -48.896, -48.798, -45.2, -45.263, -44.853, -44.906, -45.058, -44.958, -44.844, -44.69, -43.088, -42.66, -42.799, -42.373, -49.154, -49.829, -50.242, -50.086, -42.692, -41.206, -41.345, -41.729, -63.388, -65.386, -65.768, -66.263]
fit_rsq: 1.0
fit_intercept: 0.689697520058
except: True
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -53.9706590909
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [1.0, 537.0, 1074.0, 1609.0, 2145.0, 2682.0, 3218.0, 3754.0, 4290.0, 4826.0, 5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0, 10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0, 16244.0, 16780.0, 17316.0, 17853.0, 18390.0, 18926.0, 19462.0, 19999.0, 20535.0, 21071.0, 21661.0, 22198.0, 22735.0, 23271.0, 23861.0, 24460.0, 25058.0, 25657.0, 26309.0, 26846.0, 27382.0, 27918.0, 28505.0, 29041.0, 29577.0, 30113.0, 30700.0, 31237.0, 31774.0, 32310.0, 32899.0, 33436.0, 33972.0, 34509.0, 35097.0, 35634.0, 36170.0, 36707.0, 37296.0, 37832.0, 38368.0, 38904.0, 39494.0, 40030.0, 40566.0, 41101.0, 41691.0, 42228.0, 42764.0, 43301.0, 43887.0, 44424.0, 44960.0, 45496.0, 46082.0, 46618.0, 47155.0, 47691.0, 48278.0, 48814.0, 49350.0, 49886.0, 50474.0, 51011.0, 51547.0, 52083.0, 52671.0, 53208.0, 53745.0, 54281.0, 54869.0, 55406.0, 55942.0, 56479.0, 57067.0, 57603.0, 58139.0, 58675.0, 59260.0, 59796.0, 60332.0, 60868.0, 61454.0, 61991.0, 62526.0, 63062.0, 63649.0, 64185.0, 64721.0, 65258.0, 65845.0, 66380.0, 66917.0, 67453.0, 68040.0, 68576.0, 69111.0, 69648.0, 70234.0, 70770.0, 71306.0, 71842.0, 72429.0, 72966.0, 73502.0, 74038.0, 74622.0, 75158.0, 75693.0, 76228.0, 76812.0, 77348.0, 77883.0, 78418.0, 79004.0, 79540.0, 80076.0, 80612.0, 81198.0, 81733.0, 82269.0, 82804.0, 83390.0, 83925.0, 84460.0, 84996.0, 85583.0, 86119.0, 86656.0, 87192.0, 87777.0, 88312.0, 88848.0, 89384.0, 89967.0, 90503.0, 91038.0, 91574.0, 92158.0, 92693.0, 93229.0, 93765.0, 94348.0, 94884.0, 95420.0, 95956.0, 96540.0, 97076.0, 97611.0, 98147.0, 98731.0, 99266.0, 99801.0, 100337.0, 100922.0, 101458.0, 101995.0, 102530.0, 103114.0, 103650.0, 104186.0, 104722.0, 105306.0, 105842.0, 106378.0, 106914.0, 107498.0, 108034.0, 108569.0, 109104.0, 109688.0, 110223.0, 110759.0, 111294.0, 111880.0, 112416.0, 112950.0, 113486.0, 114070.0, 114606.0, 115142.0, 115677.0, 116262.0, 116798.0, 117332.0, 117867.0, 118453.0, 118988.0, 119524.0, 120059.0]
source_col: d(D_H)Mean

## 1 - Sample 1 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 1
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.2671
varray: [2.523, 2.402, 2.253, 2.253, 2.238, 2.215, 2.22, 2.213, 2.208, 2.146]
inst: mean_all
operation: mean_all(residuals)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2145.0, 2682.0, 3218.0, 3754.0, 4290.0, 4826.0]
source_col: residuals
SAMP_residuals_STD | 0.111027974453
varray: [2.523, 2.402, 2.253, 2.253, 2.238, 2.215, 2.22, 2.213, 2.208, 2.146]
inst: std_all
operation: std_all(residuals)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2145.0, 2682.0, 3218.0, 3754.0, 4290.0, 4826.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0084
varray: [-1.159, -1.196, -1.072, -0.996, -0.944, -0.988, -0.966, -0.974, -0.793, -0.996]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2145.0, 2682.0, 3218.0, 3754.0, 4290.0, 4826.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.113611619124
varray: [-1.159, -1.196, -1.072, -0.996, -0.944, -0.988, -0.966, -0.974, -0.793, -0.996]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2145.0, 2682.0, 3218.0, 3754.0, 4290.0, 4826.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 138.189
varray: [322.041, 334.503, 157.32, 104.251, 130.925, 73.958, 68.608, 65.311, 63.311, 61.662]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2145.0, 2682.0, 3218.0, 3754.0, 4290.0, 4826.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 105.226068602
varray: [322.041, 334.503, 157.32, 104.251, 130.925, 73.958, 68.608, 65.311, 63.311, 61.662]
inst: std_all
operation: std_all(baseline_shift)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2145.0, 2682.0, 3218.0, 3754.0, 4290.0, 4826.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.4954
varray: [26.997, 24.129, 27.73, 29.732, 28.367, 29.175, 29.965, 29.905, 30.081, 28.873]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2145.0, 2682.0, 3218.0, 3754.0, 4290.0, 4826.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.84714242969
varray: [26.997, 24.129, 27.73, 29.732, 28.367, 29.175, 29.965, 29.905, 30.081, 28.873]
inst: std_all
operation: std_all(slope_shift)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2145.0, 2682.0, 3218.0, 3754.0, 4290.0, 4826.0]
source_col: slope_shift
SAMP_d18O_STD | 0.255530733094
varray: [-9.875, -9.474, -9.258, -9.183, -9.176, -9.107, -9.126, -9.074, -9.043, -9.082]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2145.0, 2682.0, 3218.0, 3754.0, 4290.0, 4826.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.45097509847
varray: [-75.572, -71.771, -69.755, -68.948, -68.891, -68.321, -68.064, -67.978, -67.716, -67.968]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2145.0, 2682.0, 3218.0, 3754.0, 4290.0, 4826.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.9262
varray: [67.677, 64.709, 63.666, 64.403, 64.702, 65.36, 65.276, 65.028, 65.258, 63.183]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2145.0, 2682.0, 3218.0, 3754.0, 4290.0, 4826.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.20166216365
varray: [67.677, 64.709, 63.666, 64.403, 64.702, 65.36, 65.276, 65.028, 65.258, 63.183]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2145.0, 2682.0, 3218.0, 3754.0, 4290.0, 4826.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.2671
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0084
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 138.189
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.4954
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.03905
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.03905
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.75265366576
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.23815
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 0
varray: [-9.875, -9.474, -9.258, -9.183, -9.176, -9.107, -9.126, -9.074, -9.043, -9.082]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2145.0, 2682.0, 3218.0, 3754.0, 4290.0, 4826.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 0
varray: [-75.572, -71.771, -69.755, -68.948, -68.891, -68.321, -68.064, -67.978, -67.716, -67.968]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2145.0, 2682.0, 3218.0, 3754.0, 4290.0, 4826.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.2398
varray: [-9.875, -9.474, -9.258, -9.183, -9.176, -9.107, -9.126, -9.074, -9.043, -9.082]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2145.0, 2682.0, 3218.0, 3754.0, 4290.0, 4826.0]
source_col: d(18_16)Mean
dD_mean | -69.4984
varray: [-75.572, -71.771, -69.755, -68.948, -68.891, -68.321, -68.064, -67.978, -67.716, -67.968]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2145.0, 2682.0, 3218.0, 3754.0, 4290.0, 4826.0]
source_col: d(D_H)Mean
d18O_stdev | 0.255530733094
varray: [-9.875, -9.474, -9.258, -9.183, -9.176, -9.107, -9.126, -9.074, -9.043, -9.082]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2145.0, 2682.0, 3218.0, 3754.0, 4290.0, 4826.0]
source_col: d(18_16)Mean
dD_stdev | 2.45097509847
varray: [-75.572, -71.771, -69.755, -68.948, -68.891, -68.321, -68.064, -67.978, -67.716, -67.968]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2145.0, 2682.0, 3218.0, 3754.0, 4290.0, 4826.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.2398
varray: [-9.875, -9.474, -9.258, -9.183, -9.176, -9.107, -9.126, -9.074, -9.043, -9.082]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2145.0, 2682.0, 3218.0, 3754.0, 4290.0, 4826.0]
source_col: d(18_16)Mean
measured_dD_mean | -69.4984
varray: [-75.572, -71.771, -69.755, -68.948, -68.891, -68.321, -68.064, -67.978, -67.716, -67.968]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2145.0, 2682.0, 3218.0, 3754.0, 4290.0, 4826.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.255530733094
varray: [-9.875, -9.474, -9.258, -9.183, -9.176, -9.107, -9.126, -9.074, -9.043, -9.082]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2145.0, 2682.0, 3218.0, 3754.0, 4290.0, 4826.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.45097509847
varray: [-75.572, -71.771, -69.755, -68.948, -68.891, -68.321, -68.064, -67.978, -67.716, -67.968]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2145.0, 2682.0, 3218.0, 3754.0, 4290.0, 4826.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.484246735
actual_array: [ 1.93 -15.84]
varray: [-9.875, -9.474, -9.258, -9.183, -9.176, -9.107, -9.126, -9.074, -9.043, -9.082]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -9.2398
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [1.0, 537.0, 1074.0, 1609.0, 2145.0, 2682.0, 3218.0, 3754.0, 4290.0, 4826.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -76.3601836716
actual_array: [ 18.1 -119. ]
varray: [-75.572, -71.771, -69.755, -68.948, -68.891, -68.321, -68.064, -67.978, -67.716, -67.968]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -69.4984
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [1.0, 537.0, 1074.0, 1609.0, 2145.0, 2682.0, 3218.0, 3754.0, 4290.0, 4826.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 60.97435
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 74.94755
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 2 - Sample 2 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 2
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.3158
varray: [2.346, 2.316, 2.288, 2.303, 2.306, 2.347, 2.325, 2.303, 2.306, 2.318]
inst: mean_all
operation: mean_all(residuals)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0]
source_col: residuals
SAMP_residuals_STD | 0.0190543083492
varray: [2.346, 2.316, 2.288, 2.303, 2.306, 2.347, 2.325, 2.303, 2.306, 2.318]
inst: std_all
operation: std_all(residuals)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0572
varray: [-1.028, -1.026, -1.118, -1.221, -0.906, -1.009, -1.069, -1.063, -1.118, -1.014]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0838474805823
varray: [-1.028, -1.026, -1.118, -1.221, -0.906, -1.009, -1.069, -1.063, -1.118, -1.014]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 60.9752
varray: [56.725, 61.151, 62.387, 61.23, 60.994, 61.493, 60.963, 61.254, 61.744, 61.811]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.55596242457
varray: [56.725, 61.151, 62.387, 61.23, 60.994, 61.493, 60.963, 61.254, 61.744, 61.811]
inst: std_all
operation: std_all(baseline_shift)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.6019
varray: [30.507, 29.724, 29.303, 29.227, 29.218, 30.082, 29.477, 29.334, 29.503, 29.644]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.413763875229
varray: [30.507, 29.724, 29.303, 29.227, 29.218, 30.082, 29.477, 29.334, 29.503, 29.644]
inst: std_all
operation: std_all(slope_shift)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0]
source_col: slope_shift
SAMP_d18O_STD | 0.217364542749
varray: [2.239, 2.687, 2.769, 2.883, 2.846, 2.873, 2.926, 2.923, 2.947, 2.974]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.45057760221
varray: [8.309, 15.198, 16.918, 18.143, 18.353, 18.732, 19.289, 19.124, 19.488, 19.713]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 66.6399
varray: [68.285, 67.105, 66.257, 65.436, 66.605, 67.42, 66.577, 66.221, 66.162, 66.331]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.79103834013
varray: [68.285, 67.105, 66.257, 65.436, 66.605, 67.42, 66.577, 66.221, 66.162, 66.331]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.3158
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0572
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 60.9752
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.6019
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.08775
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.08775
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 3.93842148964
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.47555
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 2.8067
varray: [2.239, 2.687, 2.769, 2.883, 2.846, 2.873, 2.926, 2.923, 2.947, 2.974]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 17.3267
varray: [8.309, 15.198, 16.918, 18.143, 18.353, 18.732, 19.289, 19.124, 19.488, 19.713]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0]
source_col: d(D_H)Mean
known_dD | 18.1
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 1.93
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.8067
varray: [2.239, 2.687, 2.769, 2.883, 2.846, 2.873, 2.926, 2.923, 2.947, 2.974]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0]
source_col: d(18_16)Mean
dD_mean | 17.3267
varray: [8.309, 15.198, 16.918, 18.143, 18.353, 18.732, 19.289, 19.124, 19.488, 19.713]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0]
source_col: d(D_H)Mean
d18O_stdev | 0.217364542749
varray: [2.239, 2.687, 2.769, 2.883, 2.846, 2.873, 2.926, 2.923, 2.947, 2.974]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0]
source_col: d(18_16)Mean
dD_stdev | 3.45057760221
varray: [8.309, 15.198, 16.918, 18.143, 18.353, 18.732, 19.289, 19.124, 19.488, 19.713]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.8067
varray: [2.239, 2.687, 2.769, 2.883, 2.846, 2.873, 2.926, 2.923, 2.947, 2.974]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0]
source_col: d(18_16)Mean
measured_dD_mean | 17.3267
varray: [8.309, 15.198, 16.918, 18.143, 18.353, 18.732, 19.289, 19.124, 19.488, 19.713]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.217364542749
varray: [2.239, 2.687, 2.769, 2.883, 2.846, 2.873, 2.926, 2.923, 2.947, 2.974]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.45057760221
varray: [8.309, 15.198, 16.918, 18.143, 18.353, 18.732, 19.289, 19.124, 19.488, 19.713]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.93
actual_array: [ 1.93 -15.84]
varray: [2.239, 2.687, 2.769, 2.883, 2.846, 2.873, 2.926, 2.923, 2.947, 2.974]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: 2.8067
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 18.1
actual_array: [ 18.1 -119. ]
varray: [8.309, 15.198, 16.918, 18.143, 18.353, 18.732, 19.289, 19.124, 19.488, 19.713]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: 17.3267
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [5414.0, 5950.0, 6486.0, 7022.0, 7558.0, 8095.0, 8631.0, 9168.0, 9705.0, 10240.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 60.92565
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -2.26625
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 3 - Sample 3 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 3
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1403
varray: [2.16, 2.099, 2.143, 2.157, 2.153, 2.133, 2.138, 2.155, 2.127, 2.138]
inst: mean_all
operation: mean_all(residuals)
tarray: [10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: residuals
SAMP_residuals_STD | 0.0182516361519
varray: [2.16, 2.099, 2.143, 2.157, 2.153, 2.133, 2.138, 2.155, 2.127, 2.138]
inst: std_all
operation: std_all(residuals)
tarray: [10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9375
varray: [-0.887, -0.998, -1.112, -0.792, -0.88, -0.905, -0.997, -0.815, -0.895, -1.094]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.109055796117
varray: [-0.887, -0.998, -1.112, -0.792, -0.88, -0.905, -0.997, -0.815, -0.895, -1.094]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 65.5077
varray: [57.918, 64.733, 66.678, 66.492, 66.397, 65.23, 64.931, 65.122, 65.529, 72.047]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.4100239833
varray: [57.918, 64.733, 66.678, 66.492, 66.397, 65.23, 64.931, 65.122, 65.529, 72.047]
inst: std_all
operation: std_all(baseline_shift)
tarray: [10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.6376
varray: [29.915, 29.066, 29.778, 29.882, 29.478, 29.348, 29.739, 29.936, 29.514, 29.72]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.281632621217
varray: [29.915, 29.066, 29.778, 29.882, 29.478, 29.348, 29.739, 29.936, 29.514, 29.72]
inst: std_all
operation: std_all(slope_shift)
tarray: [10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: slope_shift
SAMP_d18O_STD | 0.32352106234
varray: [-13.561, -14.282, -14.497, -14.482, -14.55, -14.588, -14.593, -14.611, -14.604, -14.601]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 5.00345207609
varray: [-95.448, -105.739, -108.726, -109.713, -110.198, -110.9, -111.092, -111.537, -111.624, -111.94]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.6888
varray: [64.692, 62.328, 63.115, 64.694, 64.466, 63.507, 63.664, 64.359, 62.912, 63.151]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.828477961888
varray: [64.692, 62.328, 63.115, 64.694, 64.466, 63.507, 63.664, 64.359, 62.912, 63.151]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1403
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9375
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 65.5077
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.6376
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.08775
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.08775
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.93842148964
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.47555
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-13.561, -14.282, -14.497, -14.482, -14.55, -14.588, -14.593, -14.611, -14.604, -14.601]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-95.448, -105.739, -108.726, -109.713, -110.198, -110.9, -111.092, -111.537, -111.624, -111.94]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: d(D_H)Mean
known_dD | -119.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | -15.84
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -14.4369
varray: [-13.561, -14.282, -14.497, -14.482, -14.55, -14.588, -14.593, -14.611, -14.604, -14.601]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: d(18_16)Mean
dD_mean | -108.6917
varray: [-95.448, -105.739, -108.726, -109.713, -110.198, -110.9, -111.092, -111.537, -111.624, -111.94]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: d(D_H)Mean
d18O_stdev | 0.32352106234
varray: [-13.561, -14.282, -14.497, -14.482, -14.55, -14.588, -14.593, -14.611, -14.604, -14.601]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: d(18_16)Mean
dD_stdev | 5.00345207609
varray: [-95.448, -105.739, -108.726, -109.713, -110.198, -110.9, -111.092, -111.537, -111.624, -111.94]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: d(D_H)Mean
measured_d18O_mean | -14.4369
varray: [-13.561, -14.282, -14.497, -14.482, -14.55, -14.588, -14.593, -14.611, -14.604, -14.601]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: d(18_16)Mean
measured_dD_mean | -108.6917
varray: [-95.448, -105.739, -108.726, -109.713, -110.198, -110.9, -111.092, -111.537, -111.624, -111.94]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.32352106234
varray: [-13.561, -14.282, -14.497, -14.482, -14.55, -14.588, -14.593, -14.611, -14.604, -14.601]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: d(18_16)Mean
measured_dD_stdev | 5.00345207609
varray: [-95.448, -105.739, -108.726, -109.713, -110.198, -110.9, -111.092, -111.537, -111.624, -111.94]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -15.84
actual_array: [ 1.93 -15.84]
varray: [-13.561, -14.282, -14.497, -14.482, -14.55, -14.588, -14.593, -14.611, -14.604, -14.601]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -14.4369
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -119.0
actual_array: [ 18.1 -119. ]
varray: [-95.448, -105.739, -108.726, -109.713, -110.198, -110.9, -111.092, -111.537, -111.624, -111.94]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -108.6917
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [10829.0, 11365.0, 11902.0, 12438.0, 12974.0, 13510.0, 14047.0, 14584.0, 15120.0, 15657.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.10115
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 2.26625
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 4 - Sample 4 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 4
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1712
varray: [2.21, 2.179, 2.179, 2.199, 2.154, 2.155, 2.159, 2.161, 2.168, 2.148]
inst: mean_all
operation: mean_all(residuals)
tarray: [16244.0, 16780.0, 17316.0, 17853.0, 18390.0, 18926.0, 19462.0, 19999.0, 20535.0, 21071.0]
source_col: residuals
SAMP_residuals_STD | 0.0204385257133
varray: [2.21, 2.179, 2.179, 2.199, 2.154, 2.155, 2.159, 2.161, 2.168, 2.148]
inst: std_all
operation: std_all(residuals)
tarray: [16244.0, 16780.0, 17316.0, 17853.0, 18390.0, 18926.0, 19462.0, 19999.0, 20535.0, 21071.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0095
varray: [-1.07, -0.924, -1.178, -1.018, -0.849, -0.98, -0.89, -1.117, -0.926, -1.143]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [16244.0, 16780.0, 17316.0, 17853.0, 18390.0, 18926.0, 19462.0, 19999.0, 20535.0, 21071.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.113869564756
varray: [-1.07, -0.924, -1.178, -1.018, -0.849, -0.98, -0.89, -1.117, -0.926, -1.143]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [16244.0, 16780.0, 17316.0, 17853.0, 18390.0, 18926.0, 19462.0, 19999.0, 20535.0, 21071.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 63.4952
varray: [56.074, 61.077, 62.713, 64.375, 65.207, 64.828, 66.553, 65.658, 64.68, 63.787]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [16244.0, 16780.0, 17316.0, 17853.0, 18390.0, 18926.0, 19462.0, 19999.0, 20535.0, 21071.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.02822176643
varray: [56.074, 61.077, 62.713, 64.375, 65.207, 64.828, 66.553, 65.658, 64.68, 63.787]
inst: std_all
operation: std_all(baseline_shift)
tarray: [16244.0, 16780.0, 17316.0, 17853.0, 18390.0, 18926.0, 19462.0, 19999.0, 20535.0, 21071.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.3389
varray: [30.534, 29.425, 29.997, 29.529, 28.985, 28.812, 28.933, 28.979, 29.017, 29.178]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [16244.0, 16780.0, 17316.0, 17853.0, 18390.0, 18926.0, 19462.0, 19999.0, 20535.0, 21071.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.550915490595
varray: [30.534, 29.425, 29.997, 29.529, 28.985, 28.812, 28.933, 28.979, 29.017, 29.178]
inst: std_all
operation: std_all(slope_shift)
tarray: [16244.0, 16780.0, 17316.0, 17853.0, 18390.0, 18926.0, 19462.0, 19999.0, 20535.0, 21071.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0936565830871
varray: [-9.327, -9.156, -9.052, -9.074, -9.101, -9.048, -9.08, -9.035, -9.012, -9.015]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16244.0, 16780.0, 17316.0, 17853.0, 18390.0, 18926.0, 19462.0, 19999.0, 20535.0, 21071.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.54291807228
varray: [-72.374, -69.081, -68.062, -67.733, -68.034, -67.499, -67.604, -67.439, -67.402, -67.24]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16244.0, 16780.0, 17316.0, 17853.0, 18390.0, 18926.0, 19462.0, 19999.0, 20535.0, 21071.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.1456
varray: [65.645, 64.757, 63.263, 65.139, 63.903, 63.751, 63.862, 63.72, 63.824, 63.592]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [16244.0, 16780.0, 17316.0, 17853.0, 18390.0, 18926.0, 19462.0, 19999.0, 20535.0, 21071.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.765496527316
varray: [65.645, 64.757, 63.263, 65.139, 63.903, 63.751, 63.862, 63.72, 63.824, 63.592]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [16244.0, 16780.0, 17316.0, 17853.0, 18390.0, 18926.0, 19462.0, 19999.0, 20535.0, 21071.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1712
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0095
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 63.4952
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.3389
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.05685
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.05685
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.55155853773
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.01875
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-9.327, -9.156, -9.052, -9.074, -9.101, -9.048, -9.08, -9.035, -9.012, -9.015]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [16244.0, 16780.0, 17316.0, 17853.0, 18390.0, 18926.0, 19462.0, 19999.0, 20535.0, 21071.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-72.374, -69.081, -68.062, -67.733, -68.034, -67.499, -67.604, -67.439, -67.402, -67.24]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [16244.0, 16780.0, 17316.0, 17853.0, 18390.0, 18926.0, 19462.0, 19999.0, 20535.0, 21071.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.09
varray: [-9.327, -9.156, -9.052, -9.074, -9.101, -9.048, -9.08, -9.035, -9.012, -9.015]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16244.0, 16780.0, 17316.0, 17853.0, 18390.0, 18926.0, 19462.0, 19999.0, 20535.0, 21071.0]
source_col: d(18_16)Mean
dD_mean | -68.2468
varray: [-72.374, -69.081, -68.062, -67.733, -68.034, -67.499, -67.604, -67.439, -67.402, -67.24]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16244.0, 16780.0, 17316.0, 17853.0, 18390.0, 18926.0, 19462.0, 19999.0, 20535.0, 21071.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0936565830871
varray: [-9.327, -9.156, -9.052, -9.074, -9.101, -9.048, -9.08, -9.035, -9.012, -9.015]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16244.0, 16780.0, 17316.0, 17853.0, 18390.0, 18926.0, 19462.0, 19999.0, 20535.0, 21071.0]
source_col: d(18_16)Mean
dD_stdev | 1.54291807228
varray: [-72.374, -69.081, -68.062, -67.733, -68.034, -67.499, -67.604, -67.439, -67.402, -67.24]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16244.0, 16780.0, 17316.0, 17853.0, 18390.0, 18926.0, 19462.0, 19999.0, 20535.0, 21071.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.09
varray: [-9.327, -9.156, -9.052, -9.074, -9.101, -9.048, -9.08, -9.035, -9.012, -9.015]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16244.0, 16780.0, 17316.0, 17853.0, 18390.0, 18926.0, 19462.0, 19999.0, 20535.0, 21071.0]
source_col: d(18_16)Mean
measured_dD_mean | -68.2468
varray: [-72.374, -69.081, -68.062, -67.733, -68.034, -67.499, -67.604, -67.439, -67.402, -67.24]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16244.0, 16780.0, 17316.0, 17853.0, 18390.0, 18926.0, 19462.0, 19999.0, 20535.0, 21071.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0936565830871
varray: [-9.327, -9.156, -9.052, -9.074, -9.101, -9.048, -9.08, -9.035, -9.012, -9.015]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16244.0, 16780.0, 17316.0, 17853.0, 18390.0, 18926.0, 19462.0, 19999.0, 20535.0, 21071.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.54291807228
varray: [-72.374, -69.081, -68.062, -67.733, -68.034, -67.499, -67.604, -67.439, -67.402, -67.24]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16244.0, 16780.0, 17316.0, 17853.0, 18390.0, 18926.0, 19462.0, 19999.0, 20535.0, 21071.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.3298737503
actual_array: [ 1.93 -15.84]
varray: [-9.327, -9.156, -9.052, -9.074, -9.101, -9.048, -9.08, -9.035, -9.012, -9.015]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -9.09
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [16244.0, 16780.0, 17316.0, 17853.0, 18390.0, 18926.0, 19462.0, 19999.0, 20535.0, 21071.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -74.9985225173
actual_array: [ 18.1 -119. ]
varray: [-72.374, -69.081, -68.062, -67.733, -68.034, -67.499, -67.604, -67.439, -67.402, -67.24]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -68.2468
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [16244.0, 16780.0, 17316.0, 17853.0, 18390.0, 18926.0, 19462.0, 19999.0, 20535.0, 21071.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.07025
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.25375
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 5 - Sample 5 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 5
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.17425
varray: [2.201, 2.176, 2.154, 2.166]
inst: mean_all
operation: mean_all(residuals)
tarray: [21661.0, 22198.0, 22735.0, 23271.0]
source_col: residuals
SAMP_residuals_STD | 0.0199728983041
varray: [2.201, 2.176, 2.154, 2.166]
inst: std_all
operation: std_all(residuals)
tarray: [21661.0, 22198.0, 22735.0, 23271.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9785
varray: [-0.88, -0.893, -1.103, -1.038]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [21661.0, 22198.0, 22735.0, 23271.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.109625118776
varray: [-0.88, -0.893, -1.103, -1.038]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [21661.0, 22198.0, 22735.0, 23271.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 54.6205
varray: [54.057, 55.258, 54.797, 54.37]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [21661.0, 22198.0, 22735.0, 23271.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.522124187016
varray: [54.057, 55.258, 54.797, 54.37]
inst: std_all
operation: std_all(baseline_shift)
tarray: [21661.0, 22198.0, 22735.0, 23271.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.15725
varray: [29.224, 28.971, 29.306, 29.128]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [21661.0, 22198.0, 22735.0, 23271.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.143905929922
varray: [29.224, 28.971, 29.306, 29.128]
inst: std_all
operation: std_all(slope_shift)
tarray: [21661.0, 22198.0, 22735.0, 23271.0]
source_col: slope_shift
SAMP_d18O_STD | 0.111152672782
varray: [-5.591, -5.456, -5.322, -5.422]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21661.0, 22198.0, 22735.0, 23271.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.938772425387
varray: [-55.307, -53.775, -53.214, -53.47]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21661.0, 22198.0, 22735.0, 23271.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.30975
varray: [65.197, 65.141, 63.143, 63.758]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [21661.0, 22198.0, 22735.0, 23271.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.02370613459
varray: [65.197, 65.141, 63.143, 63.758]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [21661.0, 22198.0, 22735.0, 23271.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.17425
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9785
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 54.6205
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.15725
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0538
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0538
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.41466753439
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.8546
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-5.591, -5.456, -5.322, -5.422]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [21661.0, 22198.0, 22735.0, 23271.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-55.307, -53.775, -53.214, -53.47]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [21661.0, 22198.0, 22735.0, 23271.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.44775
varray: [-5.591, -5.456, -5.322, -5.422]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21661.0, 22198.0, 22735.0, 23271.0]
source_col: d(18_16)Mean
dD_mean | -53.9415
varray: [-55.307, -53.775, -53.214, -53.47]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21661.0, 22198.0, 22735.0, 23271.0]
source_col: d(D_H)Mean
d18O_stdev | 0.111152672782
varray: [-5.591, -5.456, -5.322, -5.422]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21661.0, 22198.0, 22735.0, 23271.0]
source_col: d(18_16)Mean
dD_stdev | 0.938772425387
varray: [-55.307, -53.775, -53.214, -53.47]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21661.0, 22198.0, 22735.0, 23271.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.44775
varray: [-5.591, -5.456, -5.322, -5.422]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21661.0, 22198.0, 22735.0, 23271.0]
source_col: d(18_16)Mean
measured_dD_mean | -53.9415
varray: [-55.307, -53.775, -53.214, -53.47]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21661.0, 22198.0, 22735.0, 23271.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.111152672782
varray: [-5.591, -5.456, -5.322, -5.422]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21661.0, 22198.0, 22735.0, 23271.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.938772425387
varray: [-55.307, -53.775, -53.214, -53.47]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21661.0, 22198.0, 22735.0, 23271.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.57643580807
actual_array: [ 1.93 -15.84]
varray: [-5.591, -5.456, -5.322, -5.422]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.44775
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [21661.0, 22198.0, 22735.0, 23271.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -59.4352664373
actual_array: [ 18.1 -119. ]
varray: [-55.307, -53.775, -53.214, -53.47]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -53.9415
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [21661.0, 22198.0, 22735.0, 23271.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.0672
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -8.62095
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 6 - Sample 6 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 6
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [23861.0, 24460.0, 25058.0, 25657.0]
source_col: residuals
SAMP_residuals_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [23861.0, 24460.0, 25058.0, 25657.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [23861.0, 24460.0, 25058.0, 25657.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [23861.0, 24460.0, 25058.0, 25657.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [23861.0, 24460.0, 25058.0, 25657.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [23861.0, 24460.0, 25058.0, 25657.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [23861.0, 24460.0, 25058.0, 25657.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [23861.0, 24460.0, 25058.0, 25657.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23861.0, 24460.0, 25058.0, 25657.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23861.0, 24460.0, 25058.0, 25657.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [23861.0, 24460.0, 25058.0, 25657.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [23861.0, 24460.0, 25058.0, 25657.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 0.0
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.0
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 0.0
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.0
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -2.22805
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.22805
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -100.0
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -65.16435
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [23861.0, 24460.0, 25058.0, 25657.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [23861.0, 24460.0, 25058.0, 25657.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23861.0, 24460.0, 25058.0, 25657.0]
source_col: d(18_16)Mean
dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23861.0, 24460.0, 25058.0, 25657.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23861.0, 24460.0, 25058.0, 25657.0]
source_col: d(18_16)Mean
dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23861.0, 24460.0, 25058.0, 25657.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23861.0, 24460.0, 25058.0, 25657.0]
source_col: d(18_16)Mean
measured_dD_mean | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23861.0, 24460.0, 25058.0, 25657.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23861.0, 24460.0, 25058.0, 25657.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0
varray: [0.0, 0.0, 0.0, 0.0]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23861.0, 24460.0, 25058.0, 25657.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.962380883342
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.933872312887
except: True
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [23861.0, 24460.0, 25058.0, 25657.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -0.750347012817
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 0.0, 0.0]
fit_rsq: 1.0
fit_intercept: 0.689697520058
except: True
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [23861.0, 24460.0, 25058.0, 25657.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 63.24145
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -63.24145
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 7 - Sample 7 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 7
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.15875
varray: [2.13, 2.175, 2.172, 2.158]
inst: mean_all
operation: mean_all(residuals)
tarray: [26309.0, 26846.0, 27382.0, 27918.0]
source_col: residuals
SAMP_residuals_STD | 0.020548722588
varray: [2.13, 2.175, 2.172, 2.158]
inst: std_all
operation: std_all(residuals)
tarray: [26309.0, 26846.0, 27382.0, 27918.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.041
varray: [-0.963, -1.016, -1.099, -1.086]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [26309.0, 26846.0, 27382.0, 27918.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0635032807551
varray: [-0.963, -1.016, -1.099, -1.086]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [26309.0, 26846.0, 27382.0, 27918.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 54.24875
varray: [53.68, 54.553, 54.58, 54.182]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [26309.0, 26846.0, 27382.0, 27918.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.420407243039
varray: [53.68, 54.553, 54.58, 54.182]
inst: std_all
operation: std_all(baseline_shift)
tarray: [26309.0, 26846.0, 27382.0, 27918.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.33575
varray: [28.111, 28.304, 28.323, 28.605]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [26309.0, 26846.0, 27382.0, 27918.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.203452492407
varray: [28.111, 28.304, 28.323, 28.605]
inst: std_all
operation: std_all(slope_shift)
tarray: [26309.0, 26846.0, 27382.0, 27918.0]
source_col: slope_shift
SAMP_d18O_STD | 0.084727012615
varray: [-4.239, -4.083, -4.043, -4.111]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26309.0, 26846.0, 27382.0, 27918.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.27979397182
varray: [-56.017, -55.354, -55.548, -55.697]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26309.0, 26846.0, 27382.0, 27918.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.9815
varray: [63.152, 64.534, 64.887, 63.353]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [26309.0, 26846.0, 27382.0, 27918.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.857956681113
varray: [63.152, 64.534, 64.887, 63.353]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [26309.0, 26846.0, 27382.0, 27918.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.15875
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.041
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 54.24875
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.33575
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0693
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0693
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.11034312515
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.18285
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-4.239, -4.083, -4.043, -4.111]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [26309.0, 26846.0, 27382.0, 27918.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-56.017, -55.354, -55.548, -55.697]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [26309.0, 26846.0, 27382.0, 27918.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.119
varray: [-4.239, -4.083, -4.043, -4.111]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26309.0, 26846.0, 27382.0, 27918.0]
source_col: d(18_16)Mean
dD_mean | -55.654
varray: [-56.017, -55.354, -55.548, -55.697]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26309.0, 26846.0, 27382.0, 27918.0]
source_col: d(D_H)Mean
d18O_stdev | 0.084727012615
varray: [-4.239, -4.083, -4.043, -4.111]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26309.0, 26846.0, 27382.0, 27918.0]
source_col: d(18_16)Mean
dD_stdev | 0.27979397182
varray: [-56.017, -55.354, -55.548, -55.697]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26309.0, 26846.0, 27382.0, 27918.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.119
varray: [-4.239, -4.083, -4.043, -4.111]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26309.0, 26846.0, 27382.0, 27918.0]
source_col: d(18_16)Mean
measured_dD_mean | -55.654
varray: [-56.017, -55.354, -55.548, -55.697]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26309.0, 26846.0, 27382.0, 27918.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.084727012615
varray: [-4.239, -4.083, -4.043, -4.111]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26309.0, 26846.0, 27382.0, 27918.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.27979397182
varray: [-56.017, -55.354, -55.548, -55.697]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26309.0, 26846.0, 27382.0, 27918.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.2071227006
actual_array: [ 1.93 -15.84]
varray: [-4.239, -4.083, -4.043, -4.111]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -4.119
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [26309.0, 26846.0, 27382.0, 27918.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -61.2983574621
actual_array: [ 18.1 -119. ]
varray: [-56.017, -55.354, -55.548, -55.697]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -55.654
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [26309.0, 26846.0, 27382.0, 27918.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.0827
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -8.9927
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 8 - Sample 8 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 8
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1925
varray: [2.198, 2.175, 2.217, 2.18]
inst: mean_all
operation: mean_all(residuals)
tarray: [28505.0, 29041.0, 29577.0, 30113.0]
source_col: residuals
SAMP_residuals_STD | 0.0190875177363
varray: [2.198, 2.175, 2.217, 2.18]
inst: std_all
operation: std_all(residuals)
tarray: [28505.0, 29041.0, 29577.0, 30113.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9365
varray: [-0.897, -0.762, -1.258, -0.829]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [28505.0, 29041.0, 29577.0, 30113.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.221305972204
varray: [-0.897, -0.762, -1.258, -0.829]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [28505.0, 29041.0, 29577.0, 30113.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 53.72225
varray: [53.616, 53.404, 54.342, 53.527]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [28505.0, 29041.0, 29577.0, 30113.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.422210354365
varray: [53.616, 53.404, 54.342, 53.527]
inst: std_all
operation: std_all(baseline_shift)
tarray: [28505.0, 29041.0, 29577.0, 30113.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.383
varray: [29.552, 28.918, 30.043, 29.019]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [28505.0, 29041.0, 29577.0, 30113.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.5205388234
varray: [29.552, 28.918, 30.043, 29.019]
inst: std_all
operation: std_all(slope_shift)
tarray: [28505.0, 29041.0, 29577.0, 30113.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0601387285089
varray: [-6.127, -6.186, -6.245, -6.258]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28505.0, 29041.0, 29577.0, 30113.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.348873902721
varray: [-53.044, -52.282, -52.402, -52.378]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28505.0, 29041.0, 29577.0, 30113.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.90975
varray: [64.616, 64.197, 65.03, 65.796]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [28505.0, 29041.0, 29577.0, 30113.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.681713710683
varray: [64.616, 64.197, 65.03, 65.796]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [28505.0, 29041.0, 29577.0, 30113.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1925
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9365
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 53.72225
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.383
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.03555
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.03555
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.59556562914
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.2546
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-6.127, -6.186, -6.245, -6.258]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [28505.0, 29041.0, 29577.0, 30113.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-53.044, -52.282, -52.402, -52.378]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [28505.0, 29041.0, 29577.0, 30113.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.204
varray: [-6.127, -6.186, -6.245, -6.258]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28505.0, 29041.0, 29577.0, 30113.0]
source_col: d(18_16)Mean
dD_mean | -52.5265
varray: [-53.044, -52.282, -52.402, -52.378]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28505.0, 29041.0, 29577.0, 30113.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0601387285089
varray: [-6.127, -6.186, -6.245, -6.258]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28505.0, 29041.0, 29577.0, 30113.0]
source_col: d(18_16)Mean
dD_stdev | 0.348873902721
varray: [-53.044, -52.282, -52.402, -52.378]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28505.0, 29041.0, 29577.0, 30113.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.204
varray: [-6.127, -6.186, -6.245, -6.258]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28505.0, 29041.0, 29577.0, 30113.0]
source_col: d(18_16)Mean
measured_dD_mean | -52.5265
varray: [-53.044, -52.282, -52.402, -52.378]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28505.0, 29041.0, 29577.0, 30113.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0601387285089
varray: [-6.127, -6.186, -6.245, -6.258]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28505.0, 29041.0, 29577.0, 30113.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.348873902721
varray: [-53.044, -52.282, -52.402, -52.378]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28505.0, 29041.0, 29577.0, 30113.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.35577205456
actual_array: [ 1.93 -15.84]
varray: [-6.127, -6.186, -6.245, -6.258]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -6.204
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [28505.0, 29041.0, 29577.0, 30113.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -57.895836481
actual_array: [ 18.1 -119. ]
varray: [-53.044, -52.282, -52.402, -52.378]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -52.5265
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [28505.0, 29041.0, 29577.0, 30113.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.04895
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -9.5192
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 9 - Sample 9 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 9
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.19125
varray: [2.255, 2.162, 2.172, 2.176]
inst: mean_all
operation: mean_all(residuals)
tarray: [30700.0, 31237.0, 31774.0, 32310.0]
source_col: residuals
SAMP_residuals_STD | 0.0429059047995
varray: [2.255, 2.162, 2.172, 2.176]
inst: std_all
operation: std_all(residuals)
tarray: [30700.0, 31237.0, 31774.0, 32310.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.1035
varray: [-1.15, -1.115, -1.035, -1.114]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [30700.0, 31237.0, 31774.0, 32310.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0486381194812
varray: [-1.15, -1.115, -1.035, -1.114]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [30700.0, 31237.0, 31774.0, 32310.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 53.573
varray: [54.324, 53.329, 53.282, 53.357]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [30700.0, 31237.0, 31774.0, 32310.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.501622035667
varray: [54.324, 53.329, 53.282, 53.357]
inst: std_all
operation: std_all(baseline_shift)
tarray: [30700.0, 31237.0, 31774.0, 32310.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.15175
varray: [30.319, 28.714, 28.638, 28.936]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [30700.0, 31237.0, 31774.0, 32310.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.788370629844
varray: [30.319, 28.714, 28.638, 28.936]
inst: std_all
operation: std_all(slope_shift)
tarray: [30700.0, 31237.0, 31774.0, 32310.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0463132450457
varray: [-3.245, -3.169, -3.172, -3.135]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30700.0, 31237.0, 31774.0, 32310.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.222321651367
varray: [-46.704, -46.343, -46.194, -46.294]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30700.0, 31237.0, 31774.0, 32310.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.23425
varray: [65.634, 63.591, 64.018, 63.694]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [30700.0, 31237.0, 31774.0, 32310.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.950737389258
varray: [65.634, 63.591, 64.018, 63.694]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [30700.0, 31237.0, 31774.0, 32310.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.19125
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.1035
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 53.573
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.15175
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0368
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0368
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.65166849936
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.9301
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-3.245, -3.169, -3.172, -3.135]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [30700.0, 31237.0, 31774.0, 32310.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-46.704, -46.343, -46.194, -46.294]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [30700.0, 31237.0, 31774.0, 32310.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.18025
varray: [-3.245, -3.169, -3.172, -3.135]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30700.0, 31237.0, 31774.0, 32310.0]
source_col: d(18_16)Mean
dD_mean | -46.38375
varray: [-46.704, -46.343, -46.194, -46.294]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30700.0, 31237.0, 31774.0, 32310.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0463132450457
varray: [-3.245, -3.169, -3.172, -3.135]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30700.0, 31237.0, 31774.0, 32310.0]
source_col: d(18_16)Mean
dD_stdev | 0.222321651367
varray: [-46.704, -46.343, -46.194, -46.294]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30700.0, 31237.0, 31774.0, 32310.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.18025
varray: [-3.245, -3.169, -3.172, -3.135]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30700.0, 31237.0, 31774.0, 32310.0]
source_col: d(18_16)Mean
measured_dD_mean | -46.38375
varray: [-46.704, -46.343, -46.194, -46.294]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30700.0, 31237.0, 31774.0, 32310.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0463132450457
varray: [-3.245, -3.169, -3.172, -3.135]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30700.0, 31237.0, 31774.0, 32310.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.222321651367
varray: [-46.704, -46.343, -46.194, -46.294]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30700.0, 31237.0, 31774.0, 32310.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.23971522768
actual_array: [ 1.93 -15.84]
varray: [-3.245, -3.169, -3.172, -3.135]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.18025
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [30700.0, 31237.0, 31774.0, 32310.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -51.2129153758
actual_array: [ 18.1 -119. ]
varray: [-46.704, -46.343, -46.194, -46.294]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -46.38375
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [30700.0, 31237.0, 31774.0, 32310.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.0502
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -9.66845
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 10 - Sample 10 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 10
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.148
varray: [2.155, 2.15, 2.153, 2.134]
inst: mean_all
operation: mean_all(residuals)
tarray: [32899.0, 33436.0, 33972.0, 34509.0]
source_col: residuals
SAMP_residuals_STD | 0.00955684745789
varray: [2.155, 2.15, 2.153, 2.134]
inst: std_all
operation: std_all(residuals)
tarray: [32899.0, 33436.0, 33972.0, 34509.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9975
varray: [-1.079, -0.857, -1.019, -1.035]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [32899.0, 33436.0, 33972.0, 34509.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0970412283517
varray: [-1.079, -0.857, -1.019, -1.035]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [32899.0, 33436.0, 33972.0, 34509.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 53.2105
varray: [53.055, 53.155, 53.388, 53.244]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [32899.0, 33436.0, 33972.0, 34509.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.141290480925
varray: [53.055, 53.155, 53.388, 53.244]
inst: std_all
operation: std_all(baseline_shift)
tarray: [32899.0, 33436.0, 33972.0, 34509.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.8195
varray: [28.719, 28.781, 28.969, 28.809]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [32899.0, 33436.0, 33972.0, 34509.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.106525427324
varray: [28.719, 28.781, 28.969, 28.809]
inst: std_all
operation: std_all(slope_shift)
tarray: [32899.0, 33436.0, 33972.0, 34509.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0905207158611
varray: [-6.441, -6.587, -6.632, -6.632]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32899.0, 33436.0, 33972.0, 34509.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.502906220549
varray: [-55.51, -56.316, -56.394, -56.684]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32899.0, 33436.0, 33972.0, 34509.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.36575
varray: [63.647, 63.728, 63.079, 63.009]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [32899.0, 33436.0, 33972.0, 34509.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.37408677692
varray: [63.647, 63.728, 63.079, 63.009]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [32899.0, 33436.0, 33972.0, 34509.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.148
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9975
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 53.2105
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.8195
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.08005
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.08005
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.59282780907
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.7986
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-6.441, -6.587, -6.632, -6.632]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [32899.0, 33436.0, 33972.0, 34509.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-55.51, -56.316, -56.394, -56.684]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [32899.0, 33436.0, 33972.0, 34509.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.573
varray: [-6.441, -6.587, -6.632, -6.632]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32899.0, 33436.0, 33972.0, 34509.0]
source_col: d(18_16)Mean
dD_mean | -56.226
varray: [-55.51, -56.316, -56.394, -56.684]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32899.0, 33436.0, 33972.0, 34509.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0905207158611
varray: [-6.441, -6.587, -6.632, -6.632]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32899.0, 33436.0, 33972.0, 34509.0]
source_col: d(18_16)Mean
dD_stdev | 0.502906220549
varray: [-55.51, -56.316, -56.394, -56.684]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32899.0, 33436.0, 33972.0, 34509.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.573
varray: [-6.441, -6.587, -6.632, -6.632]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32899.0, 33436.0, 33972.0, 34509.0]
source_col: d(18_16)Mean
measured_dD_mean | -56.226
varray: [-55.51, -56.316, -56.394, -56.684]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32899.0, 33436.0, 33972.0, 34509.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0905207158611
varray: [-6.441, -6.587, -6.632, -6.632]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32899.0, 33436.0, 33972.0, 34509.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.502906220549
varray: [-55.51, -56.316, -56.394, -56.684]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32899.0, 33436.0, 33972.0, 34509.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.73603661648
actual_array: [ 1.93 -15.84]
varray: [-6.441, -6.587, -6.632, -6.632]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -6.573
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [32899.0, 33436.0, 33972.0, 34509.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -61.9206570628
actual_array: [ 18.1 -119. ]
varray: [-55.51, -56.316, -56.394, -56.684]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -56.226
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [32899.0, 33436.0, 33972.0, 34509.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.09345
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.03095
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 11 - Sample 11 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 11
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.121
varray: [2.059, 2.132, 2.148, 2.145]
inst: mean_all
operation: mean_all(residuals)
tarray: [35097.0, 35634.0, 36170.0, 36707.0]
source_col: residuals
SAMP_residuals_STD | 0.0419126074907
varray: [2.059, 2.132, 2.148, 2.145]
inst: std_all
operation: std_all(residuals)
tarray: [35097.0, 35634.0, 36170.0, 36707.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0105
varray: [-1.054, -1.1, -0.91, -0.978]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [35097.0, 35634.0, 36170.0, 36707.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.083783451031
varray: [-1.054, -1.1, -0.91, -0.978]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [35097.0, 35634.0, 36170.0, 36707.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 52.72725
varray: [51.967, 53.143, 52.916, 52.883]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [35097.0, 35634.0, 36170.0, 36707.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.51984380667
varray: [51.967, 53.143, 52.916, 52.883]
inst: std_all
operation: std_all(baseline_shift)
tarray: [35097.0, 35634.0, 36170.0, 36707.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.90975
varray: [27.083, 28.177, 28.354, 28.025]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [35097.0, 35634.0, 36170.0, 36707.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.567326698943
varray: [27.083, 28.177, 28.354, 28.025]
inst: std_all
operation: std_all(slope_shift)
tarray: [35097.0, 35634.0, 36170.0, 36707.0]
source_col: slope_shift
SAMP_d18O_STD | 0.106077016047
varray: [-2.788, -2.559, -2.605, -2.574]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [35097.0, 35634.0, 36170.0, 36707.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.214490093011
varray: [-52.644, -52.147, -52.525, -52.372]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [35097.0, 35634.0, 36170.0, 36707.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.078
varray: [60.048, 62.688, 62.633, 62.943]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [35097.0, 35634.0, 36170.0, 36707.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.36005514594
varray: [60.048, 62.688, 62.633, 62.943]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [35097.0, 35634.0, 36170.0, 36707.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.121
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0105
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 52.72725
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.90975
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.10705
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.10705
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.80464980588
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.08635
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-2.788, -2.559, -2.605, -2.574]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [35097.0, 35634.0, 36170.0, 36707.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-52.644, -52.147, -52.525, -52.372]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [35097.0, 35634.0, 36170.0, 36707.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.6315
varray: [-2.788, -2.559, -2.605, -2.574]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [35097.0, 35634.0, 36170.0, 36707.0]
source_col: d(18_16)Mean
dD_mean | -52.422
varray: [-52.644, -52.147, -52.525, -52.372]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [35097.0, 35634.0, 36170.0, 36707.0]
source_col: d(D_H)Mean
d18O_stdev | 0.106077016047
varray: [-2.788, -2.559, -2.605, -2.574]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [35097.0, 35634.0, 36170.0, 36707.0]
source_col: d(18_16)Mean
dD_stdev | 0.214490093011
varray: [-52.644, -52.147, -52.525, -52.372]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [35097.0, 35634.0, 36170.0, 36707.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.6315
varray: [-2.788, -2.559, -2.605, -2.574]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [35097.0, 35634.0, 36170.0, 36707.0]
source_col: d(18_16)Mean
measured_dD_mean | -52.422
varray: [-52.644, -52.147, -52.525, -52.372]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [35097.0, 35634.0, 36170.0, 36707.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.106077016047
varray: [-2.788, -2.559, -2.605, -2.574]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [35097.0, 35634.0, 36170.0, 36707.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.214490093011
varray: [-52.644, -52.147, -52.525, -52.372]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [35097.0, 35634.0, 36170.0, 36707.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.67421338932
actual_array: [ 1.93 -15.84]
varray: [-2.788, -2.559, -2.605, -2.574]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.6315
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [35097.0, 35634.0, 36170.0, 36707.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -57.7821471309
actual_array: [ 18.1 -119. ]
varray: [-52.644, -52.147, -52.525, -52.372]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -52.422
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [35097.0, 35634.0, 36170.0, 36707.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.12045
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.5142
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 12 - Sample 12 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 12
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.079
varray: [2.075, 2.089, 2.07, 2.082]
inst: mean_all
operation: mean_all(residuals)
tarray: [37296.0, 37832.0, 38368.0, 38904.0]
source_col: residuals
SAMP_residuals_STD | 0.0082865352631
varray: [2.075, 2.089, 2.07, 2.082]
inst: std_all
operation: std_all(residuals)
tarray: [37296.0, 37832.0, 38368.0, 38904.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.066
varray: [-1.016, -0.954, -1.179, -1.115]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [37296.0, 37832.0, 38368.0, 38904.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.10035603287
varray: [-1.016, -0.954, -1.179, -1.115]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [37296.0, 37832.0, 38368.0, 38904.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 52.3795
varray: [52.182, 52.439, 52.462, 52.435]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [37296.0, 37832.0, 38368.0, 38904.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.132203126539
varray: [52.182, 52.439, 52.462, 52.435]
inst: std_all
operation: std_all(baseline_shift)
tarray: [37296.0, 37832.0, 38368.0, 38904.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.32775
varray: [27.225, 27.551, 27.12, 27.415]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [37296.0, 37832.0, 38368.0, 38904.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.192501731594
varray: [27.225, 27.551, 27.12, 27.415]
inst: std_all
operation: std_all(slope_shift)
tarray: [37296.0, 37832.0, 38368.0, 38904.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0714440340406
varray: [-5.588, -5.692, -5.741, -5.738]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37296.0, 37832.0, 38368.0, 38904.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.113708325699
varray: [-51.622, -51.495, -51.364, -51.408]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37296.0, 37832.0, 38368.0, 38904.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.605
varray: [62.273, 61.755, 61.47, 60.922]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [37296.0, 37832.0, 38368.0, 38904.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.563748762009
varray: [62.273, 61.755, 61.47, 60.922]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [37296.0, 37832.0, 38368.0, 38904.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.079
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.066
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 52.3795
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.32775
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.14905
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.14905
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.68970624537
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.55935
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-5.588, -5.692, -5.741, -5.738]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [37296.0, 37832.0, 38368.0, 38904.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-51.622, -51.495, -51.364, -51.408]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [37296.0, 37832.0, 38368.0, 38904.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.68975
varray: [-5.588, -5.692, -5.741, -5.738]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37296.0, 37832.0, 38368.0, 38904.0]
source_col: d(18_16)Mean
dD_mean | -51.47225
varray: [-51.622, -51.495, -51.364, -51.408]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37296.0, 37832.0, 38368.0, 38904.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0714440340406
varray: [-5.588, -5.692, -5.741, -5.738]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37296.0, 37832.0, 38368.0, 38904.0]
source_col: d(18_16)Mean
dD_stdev | 0.113708325699
varray: [-51.622, -51.495, -51.364, -51.408]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37296.0, 37832.0, 38368.0, 38904.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.68975
varray: [-5.588, -5.692, -5.741, -5.738]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37296.0, 37832.0, 38368.0, 38904.0]
source_col: d(18_16)Mean
measured_dD_mean | -51.47225
varray: [-51.622, -51.495, -51.364, -51.408]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37296.0, 37832.0, 38368.0, 38904.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0714440340406
varray: [-5.588, -5.692, -5.741, -5.738]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37296.0, 37832.0, 38368.0, 38904.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.113708325699
varray: [-51.622, -51.495, -51.364, -51.408]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37296.0, 37832.0, 38368.0, 38904.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.82582340695
actual_array: [ 1.93 -15.84]
varray: [-5.588, -5.692, -5.741, -5.738]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.68975
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [37296.0, 37832.0, 38368.0, 38904.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -56.7488795684
actual_array: [ 18.1 -119. ]
varray: [-51.622, -51.495, -51.364, -51.408]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -51.47225
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [37296.0, 37832.0, 38368.0, 38904.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.16245
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.86195
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 13 - Sample 13 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 13
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.14025
varray: [2.193, 2.125, 2.119, 2.124]
inst: mean_all
operation: mean_all(residuals)
tarray: [39494.0, 40030.0, 40566.0, 41101.0]
source_col: residuals
SAMP_residuals_STD | 0.0352644769326
varray: [2.193, 2.125, 2.119, 2.124]
inst: std_all
operation: std_all(residuals)
tarray: [39494.0, 40030.0, 40566.0, 41101.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.967
varray: [-1.002, -0.969, -0.949, -0.948]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [39494.0, 40030.0, 40566.0, 41101.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0252586618806
varray: [-1.002, -0.969, -0.949, -0.948]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [39494.0, 40030.0, 40566.0, 41101.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 53.08975
varray: [53.661, 52.867, 53.019, 52.812]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [39494.0, 40030.0, 40566.0, 41101.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.390766234468
varray: [53.661, 52.867, 53.019, 52.812]
inst: std_all
operation: std_all(baseline_shift)
tarray: [39494.0, 40030.0, 40566.0, 41101.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.865
varray: [29.978, 28.509, 28.482, 28.491]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [39494.0, 40030.0, 40566.0, 41101.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.742084900803
varray: [29.978, 28.509, 28.482, 28.491]
inst: std_all
operation: std_all(slope_shift)
tarray: [39494.0, 40030.0, 40566.0, 41101.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0471805044483
varray: [-8.869, -8.941, -8.982, -8.944]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39494.0, 40030.0, 40566.0, 41101.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.646298950435
varray: [-65.125, -66.111, -66.619, -66.315]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39494.0, 40030.0, 40566.0, 41101.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.623
varray: [64.449, 63.486, 63.73, 62.827]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [39494.0, 40030.0, 40566.0, 41101.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.669853217752
varray: [64.449, 63.486, 63.73, 62.827]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [39494.0, 40030.0, 40566.0, 41101.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.14025
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.967
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 53.08975
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.865
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0878
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0878
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.94066560445
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.54135
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-8.869, -8.941, -8.982, -8.944]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [39494.0, 40030.0, 40566.0, 41101.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-65.125, -66.111, -66.619, -66.315]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [39494.0, 40030.0, 40566.0, 41101.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.934
varray: [-8.869, -8.941, -8.982, -8.944]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39494.0, 40030.0, 40566.0, 41101.0]
source_col: d(18_16)Mean
dD_mean | -66.0425
varray: [-65.125, -66.111, -66.619, -66.315]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39494.0, 40030.0, 40566.0, 41101.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0471805044483
varray: [-8.869, -8.941, -8.982, -8.944]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39494.0, 40030.0, 40566.0, 41101.0]
source_col: d(18_16)Mean
dD_stdev | 0.646298950435
varray: [-65.125, -66.111, -66.619, -66.315]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39494.0, 40030.0, 40566.0, 41101.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.934
varray: [-8.869, -8.941, -8.982, -8.944]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39494.0, 40030.0, 40566.0, 41101.0]
source_col: d(18_16)Mean
measured_dD_mean | -66.0425
varray: [-65.125, -66.111, -66.619, -66.315]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39494.0, 40030.0, 40566.0, 41101.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0471805044483
varray: [-8.869, -8.941, -8.982, -8.944]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39494.0, 40030.0, 40566.0, 41101.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.646298950435
varray: [-65.125, -66.111, -66.619, -66.315]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39494.0, 40030.0, 40566.0, 41101.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.1691114964
actual_array: [ 1.93 -15.84]
varray: [-8.869, -8.941, -8.982, -8.944]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -8.934
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [39494.0, 40030.0, 40566.0, 41101.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -72.6003843883
actual_array: [ 18.1 -119. ]
varray: [-65.125, -66.111, -66.619, -66.315]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -66.0425
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [39494.0, 40030.0, 40566.0, 41101.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.1012
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.1517
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 14 - Sample 14 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 14
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.09875
varray: [2.043, 2.119, 2.116, 2.117]
inst: mean_all
operation: mean_all(residuals)
tarray: [41691.0, 42228.0, 42764.0, 43301.0]
source_col: residuals
SAMP_residuals_STD | 0.0371875875349
varray: [2.043, 2.119, 2.116, 2.117]
inst: std_all
operation: std_all(residuals)
tarray: [41691.0, 42228.0, 42764.0, 43301.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9705
varray: [-1.008, -0.943, -0.94, -0.991]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [41691.0, 42228.0, 42764.0, 43301.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0342198772645
varray: [-1.008, -0.943, -0.94, -0.991]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [41691.0, 42228.0, 42764.0, 43301.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 52.67525
varray: [52.107, 52.833, 52.951, 52.81]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [41691.0, 42228.0, 42764.0, 43301.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.38383535967
varray: [52.107, 52.833, 52.951, 52.81]
inst: std_all
operation: std_all(baseline_shift)
tarray: [41691.0, 42228.0, 42764.0, 43301.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.4905
varray: [26.426, 27.953, 27.735, 27.848]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [41691.0, 42228.0, 42764.0, 43301.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.715227935696
varray: [26.426, 27.953, 27.735, 27.848]
inst: std_all
operation: std_all(slope_shift)
tarray: [41691.0, 42228.0, 42764.0, 43301.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0905243981845
varray: [-5.597, -5.497, -5.459, -5.379]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41691.0, 42228.0, 42764.0, 43301.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.686829612544
varray: [-54.193, -53.342, -53.187, -52.523]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41691.0, 42228.0, 42764.0, 43301.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.63525
varray: [61.411, 62.981, 63.403, 62.746]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [41691.0, 42228.0, 42764.0, 43301.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.860239646843
varray: [61.411, 62.981, 63.403, 62.746]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [41691.0, 42228.0, 42764.0, 43301.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.09875
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9705
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 52.67525
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.4905
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1293
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1293
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.80328089585
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.5291
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-5.597, -5.497, -5.459, -5.379]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [41691.0, 42228.0, 42764.0, 43301.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-54.193, -53.342, -53.187, -52.523]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [41691.0, 42228.0, 42764.0, 43301.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.483
varray: [-5.597, -5.497, -5.459, -5.379]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [41691.0, 42228.0, 42764.0, 43301.0]
source_col: d(18_16)Mean
dD_mean | -53.31125
varray: [-54.193, -53.342, -53.187, -52.523]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [41691.0, 42228.0, 42764.0, 43301.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0905243981845
varray: [-5.597, -5.497, -5.459, -5.379]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41691.0, 42228.0, 42764.0, 43301.0]
source_col: d(18_16)Mean
dD_stdev | 0.686829612544
varray: [-54.193, -53.342, -53.187, -52.523]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41691.0, 42228.0, 42764.0, 43301.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.483
varray: [-5.597, -5.497, -5.459, -5.379]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [41691.0, 42228.0, 42764.0, 43301.0]
source_col: d(18_16)Mean
measured_dD_mean | -53.31125
varray: [-54.193, -53.342, -53.187, -52.523]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [41691.0, 42228.0, 42764.0, 43301.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0905243981845
varray: [-5.597, -5.497, -5.459, -5.379]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41691.0, 42228.0, 42764.0, 43301.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.686829612544
varray: [-54.193, -53.342, -53.187, -52.523]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41691.0, 42228.0, 42764.0, 43301.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.61276189427
actual_array: [ 1.93 -15.84]
varray: [-5.597, -5.497, -5.459, -5.379]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.483
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [41691.0, 42228.0, 42764.0, 43301.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -58.7495945433
actual_array: [ 18.1 -119. ]
varray: [-54.193, -53.342, -53.187, -52.523]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -53.31125
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [41691.0, 42228.0, 42764.0, 43301.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.1427
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.5662
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 15 - Sample 15 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 15
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.03375
varray: [3.015, 3.086, 3.041, 2.993]
inst: mean_all
operation: mean_all(residuals)
tarray: [43887.0, 44424.0, 44960.0, 45496.0]
source_col: residuals
SAMP_residuals_STD | 0.0399781190153
varray: [3.015, 3.086, 3.041, 2.993]
inst: std_all
operation: std_all(residuals)
tarray: [43887.0, 44424.0, 44960.0, 45496.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -4.4145
varray: [-4.481, -4.383, -4.316, -4.478]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [43887.0, 44424.0, 44960.0, 45496.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0798936793495
varray: [-4.481, -4.383, -4.316, -4.478]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [43887.0, 44424.0, 44960.0, 45496.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 65.0895
varray: [64.989, 65.559, 65.082, 64.728]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [43887.0, 44424.0, 44960.0, 45496.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.347020172324
varray: [64.989, 65.559, 65.082, 64.728]
inst: std_all
operation: std_all(baseline_shift)
tarray: [43887.0, 44424.0, 44960.0, 45496.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 20.80025
varray: [20.914, 21.296, 20.695, 20.296]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [43887.0, 44424.0, 44960.0, 45496.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.417952449448
varray: [20.914, 21.296, 20.695, 20.296]
inst: std_all
operation: std_all(slope_shift)
tarray: [43887.0, 44424.0, 44960.0, 45496.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0313089444089
varray: [1.341, 1.406, 1.407, 1.395]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43887.0, 44424.0, 44960.0, 45496.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.436826720031
varray: [-52.228, -52.915, -53.169, -53.133]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43887.0, 44424.0, 44960.0, 45496.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 66.764
varray: [66.126, 68.229, 67.291, 65.41]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [43887.0, 44424.0, 44960.0, 45496.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.24690470098
varray: [66.126, 68.229, 67.291, 65.41]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [43887.0, 44424.0, 44960.0, 45496.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.03375
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 4.4145
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 65.0895
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 20.80025
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.8057
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.8057
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 36.1616660308
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.59965
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [1.341, 1.406, 1.407, 1.395]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [43887.0, 44424.0, 44960.0, 45496.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-52.228, -52.915, -53.169, -53.133]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [43887.0, 44424.0, 44960.0, 45496.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.38725
varray: [1.341, 1.406, 1.407, 1.395]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [43887.0, 44424.0, 44960.0, 45496.0]
source_col: d(18_16)Mean
dD_mean | -52.86125
varray: [-52.228, -52.915, -53.169, -53.133]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [43887.0, 44424.0, 44960.0, 45496.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0313089444089
varray: [1.341, 1.406, 1.407, 1.395]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43887.0, 44424.0, 44960.0, 45496.0]
source_col: d(18_16)Mean
dD_stdev | 0.436826720031
varray: [-52.228, -52.915, -53.169, -53.133]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43887.0, 44424.0, 44960.0, 45496.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.38725
varray: [1.341, 1.406, 1.407, 1.395]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [43887.0, 44424.0, 44960.0, 45496.0]
source_col: d(18_16)Mean
measured_dD_mean | -52.86125
varray: [-52.228, -52.915, -53.169, -53.133]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [43887.0, 44424.0, 44960.0, 45496.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0313089444089
varray: [1.341, 1.406, 1.407, 1.395]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43887.0, 44424.0, 44960.0, 45496.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.436826720031
varray: [-52.228, -52.915, -53.169, -53.133]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43887.0, 44424.0, 44960.0, 45496.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 0.467218069313
actual_array: [ 1.93 -15.84]
varray: [1.341, 1.406, 1.407, 1.395]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: 1.38725
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [43887.0, 44424.0, 44960.0, 45496.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -58.2600231792
actual_array: [ 18.1 -119. ]
varray: [-52.228, -52.915, -53.169, -53.133]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -52.86125
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [43887.0, 44424.0, 44960.0, 45496.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 60.2077
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1.84805
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 16 - Sample 16 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 16
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.3965
varray: [3.037, 2.184, 2.193, 2.172]
inst: mean_all
operation: mean_all(residuals)
tarray: [46082.0, 46618.0, 47155.0, 47691.0]
source_col: residuals
SAMP_residuals_STD | 0.427086642264
varray: [3.037, 2.184, 2.193, 2.172]
inst: std_all
operation: std_all(residuals)
tarray: [46082.0, 46618.0, 47155.0, 47691.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.293
varray: [-1.473, -1.303, -1.172, -1.224]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [46082.0, 46618.0, 47155.0, 47691.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.131531998642
varray: [-1.473, -1.303, -1.172, -1.224]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [46082.0, 46618.0, 47155.0, 47691.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 56.66125
varray: [65.541, 53.744, 53.721, 53.639]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [46082.0, 46618.0, 47155.0, 47691.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 5.92000486346
varray: [65.541, 53.744, 53.721, 53.639]
inst: std_all
operation: std_all(baseline_shift)
tarray: [46082.0, 46618.0, 47155.0, 47691.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 32.87875
varray: [44.152, 29.283, 29.304, 28.776]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [46082.0, 46618.0, 47155.0, 47691.0]
source_col: slope_shift
SAMP_slope_shift_STD | 7.51946316235
varray: [44.152, 29.283, 29.304, 28.776]
inst: std_all
operation: std_all(slope_shift)
tarray: [46082.0, 46618.0, 47155.0, 47691.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0487237108603
varray: [-3.241, -3.166, -3.145, -3.132]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [46082.0, 46618.0, 47155.0, 47691.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.712717335274
varray: [-51.852, -50.439, -50.552, -50.325]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [46082.0, 46618.0, 47155.0, 47691.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 68.98925
varray: [85.79, 62.859, 63.928, 63.38]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [46082.0, 46618.0, 47155.0, 47691.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 11.2090008884
varray: [85.79, 62.859, 63.928, 63.38]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [46082.0, 46618.0, 47155.0, 47691.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.3965
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.293
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 56.66125
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 32.87875
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.16845
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.16845
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 7.56042279123
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 3.8249
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-3.241, -3.166, -3.145, -3.132]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [46082.0, 46618.0, 47155.0, 47691.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-51.852, -50.439, -50.552, -50.325]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [46082.0, 46618.0, 47155.0, 47691.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.171
varray: [-3.241, -3.166, -3.145, -3.132]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [46082.0, 46618.0, 47155.0, 47691.0]
source_col: d(18_16)Mean
dD_mean | -50.792
varray: [-51.852, -50.439, -50.552, -50.325]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [46082.0, 46618.0, 47155.0, 47691.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0487237108603
varray: [-3.241, -3.166, -3.145, -3.132]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [46082.0, 46618.0, 47155.0, 47691.0]
source_col: d(18_16)Mean
dD_stdev | 0.712717335274
varray: [-51.852, -50.439, -50.552, -50.325]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [46082.0, 46618.0, 47155.0, 47691.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.171
varray: [-3.241, -3.166, -3.145, -3.132]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [46082.0, 46618.0, 47155.0, 47691.0]
source_col: d(18_16)Mean
measured_dD_mean | -50.792
varray: [-51.852, -50.439, -50.552, -50.325]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [46082.0, 46618.0, 47155.0, 47691.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0487237108603
varray: [-3.241, -3.166, -3.145, -3.132]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [46082.0, 46618.0, 47155.0, 47691.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.712717335274
varray: [-51.852, -50.439, -50.552, -50.325]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [46082.0, 46618.0, 47155.0, 47691.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.23018285045
actual_array: [ 1.93 -15.84]
varray: [-3.241, -3.166, -3.145, -3.132]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.171
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [46082.0, 46618.0, 47155.0, 47691.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -56.0088108562
actual_array: [ 18.1 -119. ]
varray: [-51.852, -50.439, -50.552, -50.325]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -50.792
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [46082.0, 46618.0, 47155.0, 47691.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 60.84495
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -6.5802
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 17 - Sample 17 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 17
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.28025
varray: [2.259, 2.261, 2.279, 2.322]
inst: mean_all
operation: mean_all(residuals)
tarray: [48278.0, 48814.0, 49350.0, 49886.0]
source_col: residuals
SAMP_residuals_STD | 0.0292503561232
varray: [2.259, 2.261, 2.279, 2.322]
inst: std_all
operation: std_all(residuals)
tarray: [48278.0, 48814.0, 49350.0, 49886.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.64025
varray: [-1.71, -1.709, -1.545, -1.597]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [48278.0, 48814.0, 49350.0, 49886.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.082734011547
varray: [-1.71, -1.709, -1.545, -1.597]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [48278.0, 48814.0, 49350.0, 49886.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 55.54475
varray: [55.354, 55.402, 55.452, 55.971]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [48278.0, 48814.0, 49350.0, 49886.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.286969655771
varray: [55.354, 55.402, 55.452, 55.971]
inst: std_all
operation: std_all(baseline_shift)
tarray: [48278.0, 48814.0, 49350.0, 49886.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.6045
varray: [26.35, 26.472, 26.613, 26.983]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [48278.0, 48814.0, 49350.0, 49886.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.274263255529
varray: [26.35, 26.472, 26.613, 26.983]
inst: std_all
operation: std_all(slope_shift)
tarray: [48278.0, 48814.0, 49350.0, 49886.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0502957917391
varray: [-3.725, -3.8, -3.847, -3.786]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48278.0, 48814.0, 49350.0, 49886.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.530937770993
varray: [-57.535, -58.425, -58.676, -58.623]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48278.0, 48814.0, 49350.0, 49886.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.16325
varray: [63.476, 63.174, 64.203, 65.8]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [48278.0, 48814.0, 49350.0, 49886.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.17352158764
varray: [63.476, 63.174, 64.203, 65.8]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [48278.0, 48814.0, 49350.0, 49886.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.28025
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.64025
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 55.54475
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.6045
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0522
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0522
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.34285586051
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.0011
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-3.725, -3.8, -3.847, -3.786]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [48278.0, 48814.0, 49350.0, 49886.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-57.535, -58.425, -58.676, -58.623]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [48278.0, 48814.0, 49350.0, 49886.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.7895
varray: [-3.725, -3.8, -3.847, -3.786]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48278.0, 48814.0, 49350.0, 49886.0]
source_col: d(18_16)Mean
dD_mean | -58.31475
varray: [-57.535, -58.425, -58.676, -58.623]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48278.0, 48814.0, 49350.0, 49886.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0502957917391
varray: [-3.725, -3.8, -3.847, -3.786]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48278.0, 48814.0, 49350.0, 49886.0]
source_col: d(18_16)Mean
dD_stdev | 0.530937770993
varray: [-57.535, -58.425, -58.676, -58.623]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48278.0, 48814.0, 49350.0, 49886.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.7895
varray: [-3.725, -3.8, -3.847, -3.786]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48278.0, 48814.0, 49350.0, 49886.0]
source_col: d(18_16)Mean
measured_dD_mean | -58.31475
varray: [-57.535, -58.425, -58.676, -58.623]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48278.0, 48814.0, 49350.0, 49886.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0502957917391
varray: [-3.725, -3.8, -3.847, -3.786]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48278.0, 48814.0, 49350.0, 49886.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.530937770993
varray: [-57.535, -58.425, -58.676, -58.623]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48278.0, 48814.0, 49350.0, 49886.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.86756396576
actual_array: [ 1.93 -15.84]
varray: [-3.725, -3.8, -3.847, -3.786]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.7895
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [48278.0, 48814.0, 49350.0, 49886.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -64.1930841449
actual_array: [ 18.1 -119. ]
varray: [-57.535, -58.425, -58.676, -58.623]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -58.31475
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [48278.0, 48814.0, 49350.0, 49886.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 60.9612
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -7.6967
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 18 - Sample 18 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 18
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.233
varray: [2.225, 2.233, 2.215, 2.259]
inst: mean_all
operation: mean_all(residuals)
tarray: [50474.0, 51011.0, 51547.0, 52083.0]
source_col: residuals
SAMP_residuals_STD | 0.0188325958558
varray: [2.225, 2.233, 2.215, 2.259]
inst: std_all
operation: std_all(residuals)
tarray: [50474.0, 51011.0, 51547.0, 52083.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.44475
varray: [-1.318, -1.388, -1.415, -1.658]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [50474.0, 51011.0, 51547.0, 52083.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.14792650202
varray: [-1.318, -1.388, -1.415, -1.658]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [50474.0, 51011.0, 51547.0, 52083.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 54.92225
varray: [54.678, 54.863, 54.765, 55.383]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [50474.0, 51011.0, 51547.0, 52083.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.316326176596
varray: [54.678, 54.863, 54.765, 55.383]
inst: std_all
operation: std_all(baseline_shift)
tarray: [50474.0, 51011.0, 51547.0, 52083.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.37875
varray: [26.944, 27.446, 27.213, 27.912]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [50474.0, 51011.0, 51547.0, 52083.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.410430160523
varray: [26.944, 27.446, 27.213, 27.912]
inst: std_all
operation: std_all(slope_shift)
tarray: [50474.0, 51011.0, 51547.0, 52083.0]
source_col: slope_shift
SAMP_d18O_STD | 0.068955903784
varray: [-2.46, -2.343, -2.296, -2.366]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50474.0, 51011.0, 51547.0, 52083.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.436487781425
varray: [-54.253, -53.593, -53.214, -53.529]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50474.0, 51011.0, 51547.0, 52083.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.916
varray: [64.128, 64.504, 63.496, 63.536]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [50474.0, 51011.0, 51547.0, 52083.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.486993497561
varray: [64.128, 64.504, 63.496, 63.536]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [50474.0, 51011.0, 51547.0, 52083.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.233
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.44475
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 54.92225
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.37875
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.00495
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.00495
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.222167366082
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.24835
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-2.46, -2.343, -2.296, -2.366]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [50474.0, 51011.0, 51547.0, 52083.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-54.253, -53.593, -53.214, -53.529]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [50474.0, 51011.0, 51547.0, 52083.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.36625
varray: [-2.46, -2.343, -2.296, -2.366]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50474.0, 51011.0, 51547.0, 52083.0]
source_col: d(18_16)Mean
dD_mean | -53.64725
varray: [-54.253, -53.593, -53.214, -53.529]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50474.0, 51011.0, 51547.0, 52083.0]
source_col: d(D_H)Mean
d18O_stdev | 0.068955903784
varray: [-2.46, -2.343, -2.296, -2.366]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50474.0, 51011.0, 51547.0, 52083.0]
source_col: d(18_16)Mean
dD_stdev | 0.436487781425
varray: [-54.253, -53.593, -53.214, -53.529]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50474.0, 51011.0, 51547.0, 52083.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.36625
varray: [-2.46, -2.343, -2.296, -2.366]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50474.0, 51011.0, 51547.0, 52083.0]
source_col: d(18_16)Mean
measured_dD_mean | -53.64725
varray: [-54.253, -53.593, -53.214, -53.529]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50474.0, 51011.0, 51547.0, 52083.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.068955903784
varray: [-2.46, -2.343, -2.296, -2.366]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50474.0, 51011.0, 51547.0, 52083.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.436487781425
varray: [-54.253, -53.593, -53.214, -53.529]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50474.0, 51011.0, 51547.0, 52083.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.40086603146
actual_array: [ 1.93 -15.84]
varray: [-2.46, -2.343, -2.296, -2.366]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.36625
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [50474.0, 51011.0, 51547.0, 52083.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -59.1151411619
actual_array: [ 18.1 -119. ]
varray: [-54.253, -53.593, -53.214, -53.529]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -53.64725
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [50474.0, 51011.0, 51547.0, 52083.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.00845
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -8.3192
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 19 - Sample 19 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 19
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.96975
varray: [1.977, 1.965, 1.964, 1.973]
inst: mean_all
operation: mean_all(residuals)
tarray: [52671.0, 53208.0, 53745.0, 54281.0]
source_col: residuals
SAMP_residuals_STD | 0.00629152869606
varray: [1.977, 1.965, 1.964, 1.973]
inst: std_all
operation: std_all(residuals)
tarray: [52671.0, 53208.0, 53745.0, 54281.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.922
varray: [-0.967, -0.879, -1.011, -0.831]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [52671.0, 53208.0, 53745.0, 54281.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0818046453449
varray: [-0.967, -0.879, -1.011, -0.831]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [52671.0, 53208.0, 53745.0, 54281.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 51.07875
varray: [51.104, 50.963, 51.182, 51.066]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [52671.0, 53208.0, 53745.0, 54281.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0910288415833
varray: [51.104, 50.963, 51.182, 51.066]
inst: std_all
operation: std_all(baseline_shift)
tarray: [52671.0, 53208.0, 53745.0, 54281.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 25.52025
varray: [25.486, 25.277, 25.898, 25.42]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [52671.0, 53208.0, 53745.0, 54281.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.266513758244
varray: [25.486, 25.277, 25.898, 25.42]
inst: std_all
operation: std_all(slope_shift)
tarray: [52671.0, 53208.0, 53745.0, 54281.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0757160264849
varray: [-5.292, -5.438, -5.434, -5.455]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52671.0, 53208.0, 53745.0, 54281.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.697496176812
varray: [-63.28, -64.535, -64.694, -64.758]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52671.0, 53208.0, 53745.0, 54281.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 59.55425
varray: [59.504, 59.514, 58.936, 60.263]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [52671.0, 53208.0, 53745.0, 54281.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.544274670242
varray: [59.504, 59.514, 58.936, 60.263]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [52671.0, 53208.0, 53745.0, 54281.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.96975
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.922
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 51.07875
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 25.52025
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.2583
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.2583
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -11.5930971028
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.6101
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-5.292, -5.438, -5.434, -5.455]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [52671.0, 53208.0, 53745.0, 54281.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-63.28, -64.535, -64.694, -64.758]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [52671.0, 53208.0, 53745.0, 54281.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.40475
varray: [-5.292, -5.438, -5.434, -5.455]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [52671.0, 53208.0, 53745.0, 54281.0]
source_col: d(18_16)Mean
dD_mean | -64.31675
varray: [-63.28, -64.535, -64.694, -64.758]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [52671.0, 53208.0, 53745.0, 54281.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0757160264849
varray: [-5.292, -5.438, -5.434, -5.455]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52671.0, 53208.0, 53745.0, 54281.0]
source_col: d(18_16)Mean
dD_stdev | 0.697496176812
varray: [-63.28, -64.535, -64.694, -64.758]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52671.0, 53208.0, 53745.0, 54281.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.40475
varray: [-5.292, -5.438, -5.434, -5.455]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [52671.0, 53208.0, 53745.0, 54281.0]
source_col: d(18_16)Mean
measured_dD_mean | -64.31675
varray: [-63.28, -64.535, -64.694, -64.758]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [52671.0, 53208.0, 53745.0, 54281.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0757160264849
varray: [-5.292, -5.438, -5.434, -5.455]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52671.0, 53208.0, 53745.0, 54281.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.697496176812
varray: [-63.28, -64.535, -64.694, -64.758]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52671.0, 53208.0, 53745.0, 54281.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.53212313554
actual_array: [ 1.93 -15.84]
varray: [-5.292, -5.438, -5.434, -5.455]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.40475
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [52671.0, 53208.0, 53745.0, 54281.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -70.7228782067
actual_array: [ 18.1 -119. ]
varray: [-63.28, -64.535, -64.694, -64.758]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -64.31675
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [52671.0, 53208.0, 53745.0, 54281.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.2717
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -12.1627
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 20 - Sample 20 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 20
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.128
varray: [2.058, 2.136, 2.154, 2.164]
inst: mean_all
operation: mean_all(residuals)
tarray: [54869.0, 55406.0, 55942.0, 56479.0]
source_col: residuals
SAMP_residuals_STD | 0.0480832611207
varray: [2.058, 2.136, 2.154, 2.164]
inst: std_all
operation: std_all(residuals)
tarray: [54869.0, 55406.0, 55942.0, 56479.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.45125
varray: [-1.231, -1.612, -1.394, -1.568]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [54869.0, 55406.0, 55942.0, 56479.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.17441210776
varray: [-1.231, -1.612, -1.394, -1.568]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [54869.0, 55406.0, 55942.0, 56479.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 176.721
varray: [169.825, 177.781, 178.95, 180.328]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [54869.0, 55406.0, 55942.0, 56479.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 4.71371424675
varray: [169.825, 177.781, 178.95, 180.328]
inst: std_all
operation: std_all(baseline_shift)
tarray: [54869.0, 55406.0, 55942.0, 56479.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 22.23175
varray: [21.153, 22.571, 22.32, 22.883]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [54869.0, 55406.0, 55942.0, 56479.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.755139446283
varray: [21.153, 22.571, 22.32, 22.883]
inst: std_all
operation: std_all(slope_shift)
tarray: [54869.0, 55406.0, 55942.0, 56479.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0232737334063
varray: [-3.562, -3.522, -3.527, -3.507]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54869.0, 55406.0, 55942.0, 56479.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.548987173499
varray: [-55.128, -54.363, -54.059, -53.887]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54869.0, 55406.0, 55942.0, 56479.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 59.0925
varray: [57.557, 58.975, 60.152, 59.686]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [54869.0, 55406.0, 55942.0, 56479.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.13230575965
varray: [57.557, 58.975, 60.152, 59.686]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [54869.0, 55406.0, 55942.0, 56479.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.128
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.45125
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 176.721
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 22.23175
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.10005
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.10005
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.49047373264
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -6.07185
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-3.562, -3.522, -3.527, -3.507]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [54869.0, 55406.0, 55942.0, 56479.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-55.128, -54.363, -54.059, -53.887]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [54869.0, 55406.0, 55942.0, 56479.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.5295
varray: [-3.562, -3.522, -3.527, -3.507]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [54869.0, 55406.0, 55942.0, 56479.0]
source_col: d(18_16)Mean
dD_mean | -54.35925
varray: [-55.128, -54.363, -54.059, -53.887]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [54869.0, 55406.0, 55942.0, 56479.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0232737334063
varray: [-3.562, -3.522, -3.527, -3.507]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54869.0, 55406.0, 55942.0, 56479.0]
source_col: d(18_16)Mean
dD_stdev | 0.548987173499
varray: [-55.128, -54.363, -54.059, -53.887]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54869.0, 55406.0, 55942.0, 56479.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.5295
varray: [-3.562, -3.522, -3.527, -3.507]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [54869.0, 55406.0, 55942.0, 56479.0]
source_col: d(18_16)Mean
measured_dD_mean | -54.35925
varray: [-55.128, -54.363, -54.059, -53.887]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [54869.0, 55406.0, 55942.0, 56479.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0232737334063
varray: [-3.562, -3.522, -3.527, -3.507]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54869.0, 55406.0, 55942.0, 56479.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.548987173499
varray: [-55.128, -54.363, -54.059, -53.887]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54869.0, 55406.0, 55942.0, 56479.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.59962687606
actual_array: [ 1.93 -15.84]
varray: [-3.562, -3.522, -3.527, -3.507]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.5295
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [54869.0, 55406.0, 55942.0, 56479.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -59.8897518537
actual_array: [ 18.1 -119. ]
varray: [-55.128, -54.363, -54.059, -53.887]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -54.35925
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [54869.0, 55406.0, 55942.0, 56479.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.11345
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 113.47955
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 21 - Sample 21 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 21
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.02925
varray: [2.042, 2.007, 2.014, 2.054]
inst: mean_all
operation: mean_all(residuals)
tarray: [57067.0, 57603.0, 58139.0, 58675.0]
source_col: residuals
SAMP_residuals_STD | 0.0223811676788
varray: [2.042, 2.007, 2.014, 2.054]
inst: std_all
operation: std_all(residuals)
tarray: [57067.0, 57603.0, 58139.0, 58675.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.07075
varray: [-1.118, -0.96, -1.01, -1.195]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [57067.0, 57603.0, 58139.0, 58675.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.105872171351
varray: [-1.118, -0.96, -1.01, -1.195]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [57067.0, 57603.0, 58139.0, 58675.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 52.65275
varray: [53.735, 52.167, 52.293, 52.416]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [57067.0, 57603.0, 58139.0, 58675.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.728626275947
varray: [53.735, 52.167, 52.293, 52.416]
inst: std_all
operation: std_all(baseline_shift)
tarray: [57067.0, 57603.0, 58139.0, 58675.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.2085
varray: [26.221, 25.672, 26.318, 26.623]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [57067.0, 57603.0, 58139.0, 58675.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.396563992314
varray: [26.221, 25.672, 26.318, 26.623]
inst: std_all
operation: std_all(slope_shift)
tarray: [57067.0, 57603.0, 58139.0, 58675.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0218097837373
varray: [-3.75, -3.726, -3.77, -3.724]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [57067.0, 57603.0, 58139.0, 58675.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.203579632249
varray: [-55.353, -55.565, -55.688, -55.834]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [57067.0, 57603.0, 58139.0, 58675.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 60.5025
varray: [60.554, 60.666, 59.984, 60.806]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [57067.0, 57603.0, 58139.0, 58675.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.360711796314
varray: [60.554, 60.666, 59.984, 60.806]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [57067.0, 57603.0, 58139.0, 58675.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.02925
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.07075
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 52.65275
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.2085
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1988
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1988
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -8.92260048024
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.66185
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-3.75, -3.726, -3.77, -3.724]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [57067.0, 57603.0, 58139.0, 58675.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-55.353, -55.565, -55.688, -55.834]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [57067.0, 57603.0, 58139.0, 58675.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.7425
varray: [-3.75, -3.726, -3.77, -3.724]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [57067.0, 57603.0, 58139.0, 58675.0]
source_col: d(18_16)Mean
dD_mean | -55.61
varray: [-55.353, -55.565, -55.688, -55.834]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [57067.0, 57603.0, 58139.0, 58675.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0218097837373
varray: [-3.75, -3.726, -3.77, -3.724]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [57067.0, 57603.0, 58139.0, 58675.0]
source_col: d(18_16)Mean
dD_stdev | 0.203579632249
varray: [-55.353, -55.565, -55.688, -55.834]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [57067.0, 57603.0, 58139.0, 58675.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.7425
varray: [-3.75, -3.726, -3.77, -3.724]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [57067.0, 57603.0, 58139.0, 58675.0]
source_col: d(18_16)Mean
measured_dD_mean | -55.61
varray: [-55.353, -55.565, -55.688, -55.834]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [57067.0, 57603.0, 58139.0, 58675.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0218097837373
varray: [-3.75, -3.726, -3.77, -3.724]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [57067.0, 57603.0, 58139.0, 58675.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.203579632249
varray: [-55.353, -55.565, -55.688, -55.834]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [57067.0, 57603.0, 58139.0, 58675.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.81912918416
actual_array: [ 1.93 -15.84]
varray: [-3.75, -3.726, -3.77, -3.724]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.7425
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [57067.0, 57603.0, 58139.0, 58675.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -61.250488262
actual_array: [ 18.1 -119. ]
varray: [-55.353, -55.565, -55.688, -55.834]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -55.61
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [57067.0, 57603.0, 58139.0, 58675.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.2122
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.5887
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 22 - Sample 22 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 22
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.12725
varray: [2.169, 2.121, 2.109, 2.11]
inst: mean_all
operation: mean_all(residuals)
tarray: [59260.0, 59796.0, 60332.0, 60868.0]
source_col: residuals
SAMP_residuals_STD | 0.028359301825
varray: [2.169, 2.121, 2.109, 2.11]
inst: std_all
operation: std_all(residuals)
tarray: [59260.0, 59796.0, 60332.0, 60868.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.953
varray: [-0.999, -0.886, -0.959, -0.968]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [59260.0, 59796.0, 60332.0, 60868.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0478400111483
varray: [-0.999, -0.886, -0.959, -0.968]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [59260.0, 59796.0, 60332.0, 60868.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 53.07275
varray: [53.693, 52.802, 52.921, 52.875]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [59260.0, 59796.0, 60332.0, 60868.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.416392743293
varray: [53.693, 52.802, 52.921, 52.875]
inst: std_all
operation: std_all(baseline_shift)
tarray: [59260.0, 59796.0, 60332.0, 60868.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.82425
varray: [29.268, 28.743, 28.702, 28.584]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [59260.0, 59796.0, 60332.0, 60868.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.30341432069
varray: [29.268, 28.743, 28.702, 28.584]
inst: std_all
operation: std_all(slope_shift)
tarray: [59260.0, 59796.0, 60332.0, 60868.0]
source_col: slope_shift
SAMP_d18O_STD | 0.12239927832
varray: [-8.7, -8.893, -8.916, -8.986]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59260.0, 59796.0, 60332.0, 60868.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.642965525255
varray: [-65.266, -66.285, -66.492, -66.725]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59260.0, 59796.0, 60332.0, 60868.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.27325
varray: [64.548, 62.85, 62.692, 63.003]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [59260.0, 59796.0, 60332.0, 60868.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.859266111283
varray: [64.548, 62.85, 62.692, 63.003]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [59260.0, 59796.0, 60332.0, 60868.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.12725
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.953
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 53.07275
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.82425
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1008
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1008
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.52413545477
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.8911
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-8.7, -8.893, -8.916, -8.986]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [59260.0, 59796.0, 60332.0, 60868.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-65.266, -66.285, -66.492, -66.725]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [59260.0, 59796.0, 60332.0, 60868.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.87375
varray: [-8.7, -8.893, -8.916, -8.986]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [59260.0, 59796.0, 60332.0, 60868.0]
source_col: d(18_16)Mean
dD_mean | -66.192
varray: [-65.266, -66.285, -66.492, -66.725]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [59260.0, 59796.0, 60332.0, 60868.0]
source_col: d(D_H)Mean
d18O_stdev | 0.12239927832
varray: [-8.7, -8.893, -8.916, -8.986]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59260.0, 59796.0, 60332.0, 60868.0]
source_col: d(18_16)Mean
dD_stdev | 0.642965525255
varray: [-65.266, -66.285, -66.492, -66.725]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59260.0, 59796.0, 60332.0, 60868.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.87375
varray: [-8.7, -8.893, -8.916, -8.986]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [59260.0, 59796.0, 60332.0, 60868.0]
source_col: d(18_16)Mean
measured_dD_mean | -66.192
varray: [-65.266, -66.285, -66.492, -66.725]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [59260.0, 59796.0, 60332.0, 60868.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.12239927832
varray: [-8.7, -8.893, -8.916, -8.986]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59260.0, 59796.0, 60332.0, 60868.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.642965525255
varray: [-65.266, -66.285, -66.492, -66.725]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59260.0, 59796.0, 60332.0, 60868.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.1070222285
actual_array: [ 1.93 -15.84]
varray: [-8.7, -8.893, -8.916, -8.986]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -8.87375
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [59260.0, 59796.0, 60332.0, 60868.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -72.7630308749
actual_array: [ 18.1 -119. ]
varray: [-65.266, -66.285, -66.492, -66.725]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -66.192
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [59260.0, 59796.0, 60332.0, 60868.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.1142
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.1687
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 23 - Sample 23 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 23
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.15225
varray: [2.159, 2.141, 2.164, 2.145]
inst: mean_all
operation: mean_all(residuals)
tarray: [61454.0, 61991.0, 62526.0, 63062.0]
source_col: residuals
SAMP_residuals_STD | 0.0109962114688
varray: [2.159, 2.141, 2.164, 2.145]
inst: std_all
operation: std_all(residuals)
tarray: [61454.0, 61991.0, 62526.0, 63062.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.43225
varray: [-1.475, -1.376, -1.454, -1.424]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [61454.0, 61991.0, 62526.0, 63062.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0429447319237
varray: [-1.475, -1.376, -1.454, -1.424]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [61454.0, 61991.0, 62526.0, 63062.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 54.1235
varray: [54.072, 53.908, 54.311, 54.203]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [61454.0, 61991.0, 62526.0, 63062.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.1737517386
varray: [54.072, 53.908, 54.311, 54.203]
inst: std_all
operation: std_all(baseline_shift)
tarray: [61454.0, 61991.0, 62526.0, 63062.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.39625
varray: [26.606, 26.52, 26.435, 26.024]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [61454.0, 61991.0, 62526.0, 63062.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.2577988557
varray: [26.606, 26.52, 26.435, 26.024]
inst: std_all
operation: std_all(slope_shift)
tarray: [61454.0, 61991.0, 62526.0, 63062.0]
source_col: slope_shift
SAMP_d18O_STD | 0.107871451274
varray: [-3.455, -3.296, -3.212, -3.244]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61454.0, 61991.0, 62526.0, 63062.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.595681892176
varray: [-56.143, -55.427, -54.733, -55.128]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61454.0, 61991.0, 62526.0, 63062.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.09825
varray: [61.785, 61.803, 62.499, 62.306]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [61454.0, 61991.0, 62526.0, 63062.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.360119771743
varray: [61.785, 61.803, 62.499, 62.306]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [61454.0, 61991.0, 62526.0, 63062.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.15225
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.43225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 54.1235
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.39625
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0758
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0758
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.40207805031
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.0661
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-3.455, -3.296, -3.212, -3.244]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [61454.0, 61991.0, 62526.0, 63062.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-56.143, -55.427, -54.733, -55.128]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [61454.0, 61991.0, 62526.0, 63062.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.30175
varray: [-3.455, -3.296, -3.212, -3.244]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [61454.0, 61991.0, 62526.0, 63062.0]
source_col: d(18_16)Mean
dD_mean | -55.35775
varray: [-56.143, -55.427, -54.733, -55.128]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [61454.0, 61991.0, 62526.0, 63062.0]
source_col: d(D_H)Mean
d18O_stdev | 0.107871451274
varray: [-3.455, -3.296, -3.212, -3.244]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61454.0, 61991.0, 62526.0, 63062.0]
source_col: d(18_16)Mean
dD_stdev | 0.595681892176
varray: [-56.143, -55.427, -54.733, -55.128]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61454.0, 61991.0, 62526.0, 63062.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.30175
varray: [-3.455, -3.296, -3.212, -3.244]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [61454.0, 61991.0, 62526.0, 63062.0]
source_col: d(18_16)Mean
measured_dD_mean | -55.35775
varray: [-56.143, -55.427, -54.733, -55.128]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [61454.0, 61991.0, 62526.0, 63062.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.107871451274
varray: [-3.455, -3.296, -3.212, -3.244]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61454.0, 61991.0, 62526.0, 63062.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.595681892176
varray: [-56.143, -55.427, -54.733, -55.128]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61454.0, 61991.0, 62526.0, 63062.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.36492429075
actual_array: [ 1.93 -15.84]
varray: [-3.455, -3.296, -3.212, -3.244]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.30175
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [61454.0, 61991.0, 62526.0, 63062.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -60.976056314
actual_array: [ 18.1 -119. ]
varray: [-56.143, -55.427, -54.733, -55.128]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -55.35775
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [61454.0, 61991.0, 62526.0, 63062.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.0892
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -9.11795
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 24 - Sample 24 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 24
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.16375
varray: [2.161, 2.161, 2.183, 2.15]
inst: mean_all
operation: mean_all(residuals)
tarray: [63649.0, 64185.0, 64721.0, 65258.0]
source_col: residuals
SAMP_residuals_STD | 0.0138413631313
varray: [2.161, 2.161, 2.183, 2.15]
inst: std_all
operation: std_all(residuals)
tarray: [63649.0, 64185.0, 64721.0, 65258.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.26775
varray: [-1.502, -1.27, -1.135, -1.164]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [63649.0, 64185.0, 64721.0, 65258.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.166598069217
varray: [-1.502, -1.27, -1.135, -1.164]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [63649.0, 64185.0, 64721.0, 65258.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 53.8025
varray: [54.075, 53.626, 54.004, 53.505]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [63649.0, 64185.0, 64721.0, 65258.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.279593156807
varray: [54.075, 53.626, 54.004, 53.505]
inst: std_all
operation: std_all(baseline_shift)
tarray: [63649.0, 64185.0, 64721.0, 65258.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.09175
varray: [28.504, 27.898, 27.925, 28.04]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [63649.0, 64185.0, 64721.0, 65258.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.281645610653
varray: [28.504, 27.898, 27.925, 28.04]
inst: std_all
operation: std_all(slope_shift)
tarray: [63649.0, 64185.0, 64721.0, 65258.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0723342933884
varray: [-1.585, -1.478, -1.438, -1.426]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63649.0, 64185.0, 64721.0, 65258.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.28339768877
varray: [-61.114, -61.639, -61.612, -61.752]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63649.0, 64185.0, 64721.0, 65258.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.0145
varray: [61.731, 62.866, 64.34, 63.121]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [63649.0, 64185.0, 64721.0, 65258.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.07046983454
varray: [61.731, 62.866, 64.34, 63.121]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [63649.0, 64185.0, 64721.0, 65258.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.16375
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.26775
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 53.8025
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.09175
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0643
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0643
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.88593164426
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.14985
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-1.585, -1.478, -1.438, -1.426]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [63649.0, 64185.0, 64721.0, 65258.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-61.114, -61.639, -61.612, -61.752]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [63649.0, 64185.0, 64721.0, 65258.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.48175
varray: [-1.585, -1.478, -1.438, -1.426]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63649.0, 64185.0, 64721.0, 65258.0]
source_col: d(18_16)Mean
dD_mean | -61.52925
varray: [-61.114, -61.639, -61.612, -61.752]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63649.0, 64185.0, 64721.0, 65258.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0723342933884
varray: [-1.585, -1.478, -1.438, -1.426]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63649.0, 64185.0, 64721.0, 65258.0]
source_col: d(18_16)Mean
dD_stdev | 0.28339768877
varray: [-61.114, -61.639, -61.612, -61.752]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63649.0, 64185.0, 64721.0, 65258.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.48175
varray: [-1.585, -1.478, -1.438, -1.426]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63649.0, 64185.0, 64721.0, 65258.0]
source_col: d(18_16)Mean
measured_dD_mean | -61.52925
varray: [-61.114, -61.639, -61.612, -61.752]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63649.0, 64185.0, 64721.0, 65258.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0723342933884
varray: [-1.585, -1.478, -1.438, -1.426]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63649.0, 64185.0, 64721.0, 65258.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.28339768877
varray: [-61.114, -61.639, -61.612, -61.752]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63649.0, 64185.0, 64721.0, 65258.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.48936466283
actual_array: [ 1.93 -15.84]
varray: [-1.585, -1.478, -1.438, -1.426]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -1.48175
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [63649.0, 64185.0, 64721.0, 65258.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -67.6902555897
actual_array: [ 18.1 -119. ]
varray: [-61.114, -61.639, -61.612, -61.752]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -61.52925
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [63649.0, 64185.0, 64721.0, 65258.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.0777
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -9.43895
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 25 - Sample 25 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 25
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.11225
varray: [2.092, 2.117, 2.128, 2.112]
inst: mean_all
operation: mean_all(residuals)
tarray: [65845.0, 66380.0, 66917.0, 67453.0]
source_col: residuals
SAMP_residuals_STD | 0.0150637534057
varray: [2.092, 2.117, 2.128, 2.112]
inst: std_all
operation: std_all(residuals)
tarray: [65845.0, 66380.0, 66917.0, 67453.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.367
varray: [-1.488, -1.363, -1.315, -1.302]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [65845.0, 66380.0, 66917.0, 67453.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0848253106881
varray: [-1.488, -1.363, -1.315, -1.302]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [65845.0, 66380.0, 66917.0, 67453.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 357.04775
varray: [353.262, 359.865, 357.967, 357.097]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [65845.0, 66380.0, 66917.0, 67453.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.77586303396
varray: [353.262, 359.865, 357.967, 357.097]
inst: std_all
operation: std_all(baseline_shift)
tarray: [65845.0, 66380.0, 66917.0, 67453.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 17.1865
varray: [17.145, 17.423, 17.423, 16.755]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [65845.0, 66380.0, 66917.0, 67453.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.316111267331
varray: [17.145, 17.423, 17.423, 16.755]
inst: std_all
operation: std_all(slope_shift)
tarray: [65845.0, 66380.0, 66917.0, 67453.0]
source_col: slope_shift
SAMP_d18O_STD | 0.160389110187
varray: [-4.721, -5.031, -5.014, -5.07]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65845.0, 66380.0, 66917.0, 67453.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.235232331961
varray: [-60.91, -61.468, -61.274, -61.303]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65845.0, 66380.0, 66917.0, 67453.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 55.586
varray: [55.625, 55.542, 55.195, 55.982]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [65845.0, 66380.0, 66917.0, 67453.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.323086159819
varray: [55.625, 55.542, 55.195, 55.982]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [65845.0, 66380.0, 66917.0, 67453.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.11225
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.367
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 357.04775
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 17.1865
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1158
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1158
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.19736989744
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -9.57835
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-4.721, -5.031, -5.014, -5.07]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [65845.0, 66380.0, 66917.0, 67453.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-60.91, -61.468, -61.274, -61.303]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [65845.0, 66380.0, 66917.0, 67453.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.959
varray: [-4.721, -5.031, -5.014, -5.07]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65845.0, 66380.0, 66917.0, 67453.0]
source_col: d(18_16)Mean
dD_mean | -61.23875
varray: [-60.91, -61.468, -61.274, -61.303]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65845.0, 66380.0, 66917.0, 67453.0]
source_col: d(D_H)Mean
d18O_stdev | 0.160389110187
varray: [-4.721, -5.031, -5.014, -5.07]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65845.0, 66380.0, 66917.0, 67453.0]
source_col: d(18_16)Mean
dD_stdev | 0.235232331961
varray: [-60.91, -61.468, -61.274, -61.303]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65845.0, 66380.0, 66917.0, 67453.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.959
varray: [-4.721, -5.031, -5.014, -5.07]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65845.0, 66380.0, 66917.0, 67453.0]
source_col: d(18_16)Mean
measured_dD_mean | -61.23875
varray: [-60.91, -61.468, -61.274, -61.303]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65845.0, 66380.0, 66917.0, 67453.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.160389110187
varray: [-4.721, -5.031, -5.014, -5.07]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65845.0, 66380.0, 66917.0, 67453.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.235232331961
varray: [-60.91, -61.468, -61.274, -61.303]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65845.0, 66380.0, 66917.0, 67453.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.07276560579
actual_array: [ 1.93 -15.84]
varray: [-4.721, -5.031, -5.014, -5.07]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -4.959
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [65845.0, 66380.0, 66917.0, 67453.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -67.3742100757
actual_array: [ 18.1 -119. ]
varray: [-60.91, -61.468, -61.274, -61.303]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -61.23875
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [65845.0, 66380.0, 66917.0, 67453.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.1292
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 293.8063
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 26 - Sample 26 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 26
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.10075
varray: [2.153, 2.079, 2.08, 2.091]
inst: mean_all
operation: mean_all(residuals)
tarray: [68040.0, 68576.0, 69111.0, 69648.0]
source_col: residuals
SAMP_residuals_STD | 0.0352550232827
varray: [2.153, 2.079, 2.08, 2.091]
inst: std_all
operation: std_all(residuals)
tarray: [68040.0, 68576.0, 69111.0, 69648.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0415
varray: [-1.114, -1.231, -0.82, -1.001]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [68040.0, 68576.0, 69111.0, 69648.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.174994285621
varray: [-1.114, -1.231, -0.82, -1.001]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [68040.0, 68576.0, 69111.0, 69648.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 54.50475
varray: [57.025, 54.164, 53.376, 53.454]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [68040.0, 68576.0, 69111.0, 69648.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.7171607137
varray: [57.025, 54.164, 53.376, 53.454]
inst: std_all
operation: std_all(baseline_shift)
tarray: [68040.0, 68576.0, 69111.0, 69648.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.523
varray: [28.113, 27.535, 27.097, 27.347]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [68040.0, 68576.0, 69111.0, 69648.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.432317784352
varray: [28.113, 27.535, 27.097, 27.347]
inst: std_all
operation: std_all(slope_shift)
tarray: [68040.0, 68576.0, 69111.0, 69648.0]
source_col: slope_shift
SAMP_d18O_STD | 0.021156165374
varray: [-5.244, -5.223, -5.2, -5.2]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [68040.0, 68576.0, 69111.0, 69648.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.705198021835
varray: [-51.089, -50.093, -49.822, -49.439]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [68040.0, 68576.0, 69111.0, 69648.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.11525
varray: [63.559, 60.413, 62.734, 61.755]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [68040.0, 68576.0, 69111.0, 69648.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.35335395099
varray: [63.559, 60.413, 62.734, 61.755]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [68040.0, 68576.0, 69111.0, 69648.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.10075
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0415
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 54.50475
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.523
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1273
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1273
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.71351630349
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.0491
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-5.244, -5.223, -5.2, -5.2]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [68040.0, 68576.0, 69111.0, 69648.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-51.089, -50.093, -49.822, -49.439]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [68040.0, 68576.0, 69111.0, 69648.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.21675
varray: [-5.244, -5.223, -5.2, -5.2]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [68040.0, 68576.0, 69111.0, 69648.0]
source_col: d(18_16)Mean
dD_mean | -50.11075
varray: [-51.089, -50.093, -49.822, -49.439]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [68040.0, 68576.0, 69111.0, 69648.0]
source_col: d(D_H)Mean
d18O_stdev | 0.021156165374
varray: [-5.244, -5.223, -5.2, -5.2]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [68040.0, 68576.0, 69111.0, 69648.0]
source_col: d(18_16)Mean
dD_stdev | 0.705198021835
varray: [-51.089, -50.093, -49.822, -49.439]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [68040.0, 68576.0, 69111.0, 69648.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.21675
varray: [-5.244, -5.223, -5.2, -5.2]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [68040.0, 68576.0, 69111.0, 69648.0]
source_col: d(18_16)Mean
measured_dD_mean | -50.11075
varray: [-51.089, -50.093, -49.822, -49.439]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [68040.0, 68576.0, 69111.0, 69648.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.021156165374
varray: [-5.244, -5.223, -5.2, -5.2]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [68040.0, 68576.0, 69111.0, 69648.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.705198021835
varray: [-51.089, -50.093, -49.822, -49.439]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [68040.0, 68576.0, 69111.0, 69648.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.33838400914
actual_array: [ 1.93 -15.84]
varray: [-5.244, -5.223, -5.2, -5.2]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.21675
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [68040.0, 68576.0, 69111.0, 69648.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -55.2676542076
actual_array: [ 18.1 -119. ]
varray: [-51.089, -50.093, -49.822, -49.439]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -50.11075
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [68040.0, 68576.0, 69111.0, 69648.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.1407
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -8.7367
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 27 - Sample 27 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 27
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.10975
varray: [2.119, 2.126, 2.109, 2.085]
inst: mean_all
operation: mean_all(residuals)
tarray: [70234.0, 70770.0, 71306.0, 71842.0]
source_col: residuals
SAMP_residuals_STD | 0.0179141471097
varray: [2.119, 2.126, 2.109, 2.085]
inst: std_all
operation: std_all(residuals)
tarray: [70234.0, 70770.0, 71306.0, 71842.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9115
varray: [-0.976, -0.757, -0.912, -1.001]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [70234.0, 70770.0, 71306.0, 71842.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.109606873264
varray: [-0.976, -0.757, -0.912, -1.001]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [70234.0, 70770.0, 71306.0, 71842.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 52.74
varray: [52.849, 52.818, 52.826, 52.467]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [70234.0, 70770.0, 71306.0, 71842.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.182473742403
varray: [52.849, 52.818, 52.826, 52.467]
inst: std_all
operation: std_all(baseline_shift)
tarray: [70234.0, 70770.0, 71306.0, 71842.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.39375
varray: [27.702, 27.327, 27.614, 26.932]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [70234.0, 70770.0, 71306.0, 71842.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.346985470397
varray: [27.702, 27.327, 27.614, 26.932]
inst: std_all
operation: std_all(slope_shift)
tarray: [70234.0, 70770.0, 71306.0, 71842.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0426692707539
varray: [-3.142, -3.121, -3.047, -3.078]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70234.0, 70770.0, 71306.0, 71842.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.209171978684
varray: [-47.208, -47.125, -46.791, -46.829]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70234.0, 70770.0, 71306.0, 71842.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.9875
varray: [62.414, 64.063, 63.196, 62.277]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [70234.0, 70770.0, 71306.0, 71842.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.823384276427
varray: [62.414, 64.063, 63.196, 62.277]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [70234.0, 70770.0, 71306.0, 71842.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.10975
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9115
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 52.74
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.39375
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1183
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1183
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.30957563789
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.17685
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-3.142, -3.121, -3.047, -3.078]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [70234.0, 70770.0, 71306.0, 71842.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-47.208, -47.125, -46.791, -46.829]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [70234.0, 70770.0, 71306.0, 71842.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.097
varray: [-3.142, -3.121, -3.047, -3.078]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [70234.0, 70770.0, 71306.0, 71842.0]
source_col: d(18_16)Mean
dD_mean | -46.98825
varray: [-47.208, -47.125, -46.791, -46.829]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [70234.0, 70770.0, 71306.0, 71842.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0426692707539
varray: [-3.142, -3.121, -3.047, -3.078]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70234.0, 70770.0, 71306.0, 71842.0]
source_col: d(18_16)Mean
dD_stdev | 0.209171978684
varray: [-47.208, -47.125, -46.791, -46.829]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70234.0, 70770.0, 71306.0, 71842.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.097
varray: [-3.142, -3.121, -3.047, -3.078]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [70234.0, 70770.0, 71306.0, 71842.0]
source_col: d(18_16)Mean
measured_dD_mean | -46.98825
varray: [-47.208, -47.125, -46.791, -46.829]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [70234.0, 70770.0, 71306.0, 71842.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0426692707539
varray: [-3.142, -3.121, -3.047, -3.078]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70234.0, 70770.0, 71306.0, 71842.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.209171978684
varray: [-47.208, -47.125, -46.791, -46.829]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70234.0, 70770.0, 71306.0, 71842.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.15392383261
actual_array: [ 1.93 -15.84]
varray: [-3.142, -3.121, -3.047, -3.078]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.097
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [70234.0, 70770.0, 71306.0, 71842.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -51.8705729084
actual_array: [ 18.1 -119. ]
varray: [-47.208, -47.125, -46.791, -46.829]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -46.98825
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [70234.0, 70770.0, 71306.0, 71842.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.1317
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.50145
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 28 - Sample 28 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 28
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.21275
varray: [2.217, 2.194, 2.208, 2.232]
inst: mean_all
operation: mean_all(residuals)
tarray: [72429.0, 72966.0, 73502.0, 74038.0]
source_col: residuals
SAMP_residuals_STD | 0.0159452187191
varray: [2.217, 2.194, 2.208, 2.232]
inst: std_all
operation: std_all(residuals)
tarray: [72429.0, 72966.0, 73502.0, 74038.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.26
varray: [-2.346, -2.141, -2.306, -2.247]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [72429.0, 72966.0, 73502.0, 74038.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0891478173223
varray: [-2.346, -2.141, -2.306, -2.247]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [72429.0, 72966.0, 73502.0, 74038.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 193.39975
varray: [192.856, 191.34, 193.739, 195.664]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [72429.0, 72966.0, 73502.0, 74038.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.80556110854
varray: [192.856, 191.34, 193.739, 195.664]
inst: std_all
operation: std_all(baseline_shift)
tarray: [72429.0, 72966.0, 73502.0, 74038.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 18.8965
varray: [19.352, 18.354, 18.974, 18.906]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [72429.0, 72966.0, 73502.0, 74038.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.411453926137
varray: [19.352, 18.354, 18.974, 18.906]
inst: std_all
operation: std_all(slope_shift)
tarray: [72429.0, 72966.0, 73502.0, 74038.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0600555298592
varray: [-2.742, -2.828, -2.846, -2.884]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72429.0, 72966.0, 73502.0, 74038.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.658258814348
varray: [-54.752, -55.596, -56.053, -56.223]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72429.0, 72966.0, 73502.0, 74038.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 57.60375
varray: [57.361, 57.84, 57.403, 57.811]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [72429.0, 72966.0, 73502.0, 74038.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.256901245618
varray: [57.361, 57.84, 57.403, 57.811]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [72429.0, 72966.0, 73502.0, 74038.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.21275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.26
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 193.39975
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 18.8965
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0153
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0153
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.686699131528
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -7.5606
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-2.742, -2.828, -2.846, -2.884]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [72429.0, 72966.0, 73502.0, 74038.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-54.752, -55.596, -56.053, -56.223]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [72429.0, 72966.0, 73502.0, 74038.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.825
varray: [-2.742, -2.828, -2.846, -2.884]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72429.0, 72966.0, 73502.0, 74038.0]
source_col: d(18_16)Mean
dD_mean | -55.656
varray: [-54.752, -55.596, -56.053, -56.223]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72429.0, 72966.0, 73502.0, 74038.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0600555298592
varray: [-2.742, -2.828, -2.846, -2.884]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72429.0, 72966.0, 73502.0, 74038.0]
source_col: d(18_16)Mean
dD_stdev | 0.658258814348
varray: [-54.752, -55.596, -56.053, -56.223]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72429.0, 72966.0, 73502.0, 74038.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.825
varray: [-2.742, -2.828, -2.846, -2.884]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72429.0, 72966.0, 73502.0, 74038.0]
source_col: d(18_16)Mean
measured_dD_mean | -55.656
varray: [-54.752, -55.596, -56.053, -56.223]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72429.0, 72966.0, 73502.0, 74038.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0600555298592
varray: [-2.742, -2.828, -2.846, -2.884]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72429.0, 72966.0, 73502.0, 74038.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.658258814348
varray: [-54.752, -55.596, -56.053, -56.223]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72429.0, 72966.0, 73502.0, 74038.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.87362041569
actual_array: [ 1.93 -15.84]
varray: [-2.742, -2.828, -2.846, -2.884]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [72429.0, 72966.0, 73502.0, 74038.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -61.3005333348
actual_array: [ 18.1 -119. ]
varray: [-54.752, -55.596, -56.053, -56.223]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -55.656
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [72429.0, 72966.0, 73502.0, 74038.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.0287
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 130.1583
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 29 - Sample 29 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 29
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.10125
varray: [2.105, 2.026, 2.12, 2.154]
inst: mean_all
operation: mean_all(residuals)
tarray: [74622.0, 75158.0, 75693.0, 76228.0]
source_col: residuals
SAMP_residuals_STD | 0.0541933267725
varray: [2.105, 2.026, 2.12, 2.154]
inst: std_all
operation: std_all(residuals)
tarray: [74622.0, 75158.0, 75693.0, 76228.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.97075
varray: [-1.05, -1.091, -0.973, -0.769]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [74622.0, 75158.0, 75693.0, 76228.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.143118540611
varray: [-1.05, -1.091, -0.973, -0.769]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [74622.0, 75158.0, 75693.0, 76228.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 53.7695
varray: [54.868, 53.01, 53.703, 53.497]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [74622.0, 75158.0, 75693.0, 76228.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.787870759283
varray: [54.868, 53.01, 53.703, 53.497]
inst: std_all
operation: std_all(baseline_shift)
tarray: [74622.0, 75158.0, 75693.0, 76228.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.625
varray: [27.829, 26.646, 27.909, 28.116]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [74622.0, 75158.0, 75693.0, 76228.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.663775564479
varray: [27.829, 26.646, 27.909, 28.116]
inst: std_all
operation: std_all(slope_shift)
tarray: [74622.0, 75158.0, 75693.0, 76228.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0474789427009
varray: [-5.827, -5.828, -5.816, -5.918]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74622.0, 75158.0, 75693.0, 76228.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.274664037204
varray: [-55.832, -55.204, -55.309, -55.445]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74622.0, 75158.0, 75693.0, 76228.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.4115
varray: [62.143, 60.269, 62.773, 64.461]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [74622.0, 75158.0, 75693.0, 76228.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.73142821586
varray: [62.143, 60.269, 62.773, 64.461]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [74622.0, 75158.0, 75693.0, 76228.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.10125
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.97075
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 53.7695
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.625
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1268
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1268
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.6910751554
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.75285
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-5.827, -5.828, -5.816, -5.918]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [74622.0, 75158.0, 75693.0, 76228.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-55.832, -55.204, -55.309, -55.445]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [74622.0, 75158.0, 75693.0, 76228.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.84725
varray: [-5.827, -5.828, -5.816, -5.918]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [74622.0, 75158.0, 75693.0, 76228.0]
source_col: d(18_16)Mean
dD_mean | -55.4475
varray: [-55.832, -55.204, -55.309, -55.445]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [74622.0, 75158.0, 75693.0, 76228.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0474789427009
varray: [-5.827, -5.828, -5.816, -5.918]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74622.0, 75158.0, 75693.0, 76228.0]
source_col: d(18_16)Mean
dD_stdev | 0.274664037204
varray: [-55.832, -55.204, -55.309, -55.445]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74622.0, 75158.0, 75693.0, 76228.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.84725
varray: [-5.827, -5.828, -5.816, -5.918]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [74622.0, 75158.0, 75693.0, 76228.0]
source_col: d(18_16)Mean
measured_dD_mean | -55.4475
varray: [-55.832, -55.204, -55.309, -55.445]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [74622.0, 75158.0, 75693.0, 76228.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0474789427009
varray: [-5.827, -5.828, -5.816, -5.918]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74622.0, 75158.0, 75693.0, 76228.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.274664037204
varray: [-55.832, -55.204, -55.309, -55.445]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74622.0, 75158.0, 75693.0, 76228.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.98813145167
actual_array: [ 1.93 -15.84]
varray: [-5.827, -5.828, -5.816, -5.918]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.84725
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [74622.0, 75158.0, 75693.0, 76228.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -61.0736986027
actual_array: [ 18.1 -119. ]
varray: [-55.832, -55.204, -55.309, -55.445]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -55.4475
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [74622.0, 75158.0, 75693.0, 76228.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.1402
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -9.47195
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 30 - Sample 30 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 30
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.09625
varray: [2.055, 2.108, 2.073, 2.149]
inst: mean_all
operation: mean_all(residuals)
tarray: [76812.0, 77348.0, 77883.0, 78418.0]
source_col: residuals
SAMP_residuals_STD | 0.0414839326326
varray: [2.055, 2.108, 2.073, 2.149]
inst: std_all
operation: std_all(residuals)
tarray: [76812.0, 77348.0, 77883.0, 78418.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.173
varray: [-1.112, -1.299, -1.142, -1.139]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [76812.0, 77348.0, 77883.0, 78418.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0850764362206
varray: [-1.112, -1.299, -1.142, -1.139]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [76812.0, 77348.0, 77883.0, 78418.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 54.3905
varray: [53.836, 54.642, 54.117, 54.967]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [76812.0, 77348.0, 77883.0, 78418.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.509207554278
varray: [53.836, 54.642, 54.117, 54.967]
inst: std_all
operation: std_all(baseline_shift)
tarray: [76812.0, 77348.0, 77883.0, 78418.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.93875
varray: [26.319, 27.215, 26.476, 27.745]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [76812.0, 77348.0, 77883.0, 78418.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.664474918012
varray: [26.319, 27.215, 26.476, 27.745]
inst: std_all
operation: std_all(slope_shift)
tarray: [76812.0, 77348.0, 77883.0, 78418.0]
source_col: slope_shift
SAMP_d18O_STD | 0.111661691431
varray: [-2.102, -1.949, -1.84, -1.903]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76812.0, 77348.0, 77883.0, 78418.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.272824730062
varray: [-56.55, -56.768, -56.174, -56.736]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76812.0, 77348.0, 77883.0, 78418.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.879
varray: [61.127, 61.648, 61.225, 63.516]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [76812.0, 77348.0, 77883.0, 78418.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.11450287872
varray: [61.127, 61.648, 61.225, 63.516]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [76812.0, 77348.0, 77883.0, 78418.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.09625
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.173
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 54.3905
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.93875
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1318
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1318
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.9154866363
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.28535
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-2.102, -1.949, -1.84, -1.903]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [76812.0, 77348.0, 77883.0, 78418.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-56.55, -56.768, -56.174, -56.736]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [76812.0, 77348.0, 77883.0, 78418.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.9485
varray: [-2.102, -1.949, -1.84, -1.903]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [76812.0, 77348.0, 77883.0, 78418.0]
source_col: d(18_16)Mean
dD_mean | -56.557
varray: [-56.55, -56.768, -56.174, -56.736]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [76812.0, 77348.0, 77883.0, 78418.0]
source_col: d(D_H)Mean
d18O_stdev | 0.111661691431
varray: [-2.102, -1.949, -1.84, -1.903]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76812.0, 77348.0, 77883.0, 78418.0]
source_col: d(18_16)Mean
dD_stdev | 0.272824730062
varray: [-56.55, -56.768, -56.174, -56.736]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76812.0, 77348.0, 77883.0, 78418.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.9485
varray: [-2.102, -1.949, -1.84, -1.903]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [76812.0, 77348.0, 77883.0, 78418.0]
source_col: d(18_16)Mean
measured_dD_mean | -56.557
varray: [-56.55, -56.768, -56.174, -56.736]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [76812.0, 77348.0, 77883.0, 78418.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.111661691431
varray: [-2.102, -1.949, -1.84, -1.903]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76812.0, 77348.0, 77883.0, 78418.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.272824730062
varray: [-56.55, -56.768, -56.174, -56.736]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76812.0, 77348.0, 77883.0, 78418.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.97036326521
actual_array: [ 1.93 -15.84]
varray: [-2.102, -1.949, -1.84, -1.903]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -1.9485
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [76812.0, 77348.0, 77883.0, 78418.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -62.2807639995
actual_array: [ 18.1 -119. ]
varray: [-56.55, -56.768, -56.174, -56.736]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -56.557
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [76812.0, 77348.0, 77883.0, 78418.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.1452
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -8.85095
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 31 - Sample 31 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 31
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.11175
varray: [2.146, 2.119, 2.09, 2.092]
inst: mean_all
operation: mean_all(residuals)
tarray: [79004.0, 79540.0, 80076.0, 80612.0]
source_col: residuals
SAMP_residuals_STD | 0.0263865496039
varray: [2.146, 2.119, 2.09, 2.092]
inst: std_all
operation: std_all(residuals)
tarray: [79004.0, 79540.0, 80076.0, 80612.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.969
varray: [-1.067, -0.861, -0.994, -0.954]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [79004.0, 79540.0, 80076.0, 80612.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.08586423392
varray: [-1.067, -0.861, -0.994, -0.954]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [79004.0, 79540.0, 80076.0, 80612.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 52.921
varray: [53.419, 52.891, 52.736, 52.638]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [79004.0, 79540.0, 80076.0, 80612.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.347954977931
varray: [53.419, 52.891, 52.736, 52.638]
inst: std_all
operation: std_all(baseline_shift)
tarray: [79004.0, 79540.0, 80076.0, 80612.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.51075
varray: [29.131, 28.173, 28.515, 28.224]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [79004.0, 79540.0, 80076.0, 80612.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.440086639197
varray: [29.131, 28.173, 28.515, 28.224]
inst: std_all
operation: std_all(slope_shift)
tarray: [79004.0, 79540.0, 80076.0, 80612.0]
source_col: slope_shift
SAMP_d18O_STD | 0.149762033462
varray: [-8.618, -8.832, -8.947, -8.923]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [79004.0, 79540.0, 80076.0, 80612.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.51988652608
varray: [-65.528, -66.347, -66.773, -66.332]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [79004.0, 79540.0, 80076.0, 80612.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.98775
varray: [63.59, 64.038, 61.557, 62.766]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [79004.0, 79540.0, 80076.0, 80612.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.08964042234
varray: [63.59, 64.038, 61.557, 62.766]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [79004.0, 79540.0, 80076.0, 80612.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.11175
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.969
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 52.921
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.51075
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1163
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1163
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.21981104553
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.1766
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-8.618, -8.832, -8.947, -8.923]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [79004.0, 79540.0, 80076.0, 80612.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-65.528, -66.347, -66.773, -66.332]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [79004.0, 79540.0, 80076.0, 80612.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.83
varray: [-8.618, -8.832, -8.947, -8.923]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [79004.0, 79540.0, 80076.0, 80612.0]
source_col: d(18_16)Mean
dD_mean | -66.245
varray: [-65.528, -66.347, -66.773, -66.332]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [79004.0, 79540.0, 80076.0, 80612.0]
source_col: d(D_H)Mean
d18O_stdev | 0.149762033462
varray: [-8.618, -8.832, -8.947, -8.923]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [79004.0, 79540.0, 80076.0, 80612.0]
source_col: d(18_16)Mean
dD_stdev | 0.51988652608
varray: [-65.528, -66.347, -66.773, -66.332]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [79004.0, 79540.0, 80076.0, 80612.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.83
varray: [-8.618, -8.832, -8.947, -8.923]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [79004.0, 79540.0, 80076.0, 80612.0]
source_col: d(18_16)Mean
measured_dD_mean | -66.245
varray: [-65.528, -66.347, -66.773, -66.332]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [79004.0, 79540.0, 80076.0, 80612.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.149762033462
varray: [-8.618, -8.832, -8.947, -8.923]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [79004.0, 79540.0, 80076.0, 80612.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.51988652608
varray: [-65.528, -66.347, -66.773, -66.332]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [79004.0, 79540.0, 80076.0, 80612.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.0619366606
actual_array: [ 1.93 -15.84]
varray: [-8.618, -8.832, -8.947, -8.923]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -8.83
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [79004.0, 79540.0, 80076.0, 80612.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -72.8206915022
actual_array: [ 18.1 -119. ]
varray: [-65.528, -66.347, -66.773, -66.332]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -66.245
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [79004.0, 79540.0, 80076.0, 80612.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.1297
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.32045
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 32 - Sample 32 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 32
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.05575
varray: [1.991, 2.068, 2.07, 2.094]
inst: mean_all
operation: mean_all(residuals)
tarray: [81198.0, 81733.0, 82269.0, 82804.0]
source_col: residuals
SAMP_residuals_STD | 0.0447539569945
varray: [1.991, 2.068, 2.07, 2.094]
inst: std_all
operation: std_all(residuals)
tarray: [81198.0, 81733.0, 82269.0, 82804.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.111
varray: [-0.946, -1.213, -1.143, -1.142]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [81198.0, 81733.0, 82269.0, 82804.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.114911560196
varray: [-0.946, -1.213, -1.143, -1.142]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [81198.0, 81733.0, 82269.0, 82804.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 52.4915
varray: [51.534, 52.883, 52.627, 52.922]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [81198.0, 81733.0, 82269.0, 82804.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.65160545322
varray: [51.534, 52.883, 52.627, 52.922]
inst: std_all
operation: std_all(baseline_shift)
tarray: [81198.0, 81733.0, 82269.0, 82804.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.6295
varray: [25.658, 26.913, 26.908, 27.039]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [81198.0, 81733.0, 82269.0, 82804.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.650496477059
varray: [25.658, 26.913, 26.908, 27.039]
inst: std_all
operation: std_all(slope_shift)
tarray: [81198.0, 81733.0, 82269.0, 82804.0]
source_col: slope_shift
SAMP_d18O_STD | 0.139373598648
varray: [-3.907, -3.661, -3.625, -3.609]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [81198.0, 81733.0, 82269.0, 82804.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.669442242965
varray: [-56.875, -55.672, -55.581, -55.409]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [81198.0, 81733.0, 82269.0, 82804.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 60.93525
varray: [59.221, 61.171, 61.208, 62.141]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [81198.0, 81733.0, 82269.0, 82804.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.22779704485
varray: [59.221, 61.171, 61.208, 62.141]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [81198.0, 81733.0, 82269.0, 82804.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.05575
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.111
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 52.4915
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.6295
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1723
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1723
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -7.73321963152
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.2291
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-3.907, -3.661, -3.625, -3.609]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [81198.0, 81733.0, 82269.0, 82804.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-56.875, -55.672, -55.581, -55.409]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [81198.0, 81733.0, 82269.0, 82804.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.7005
varray: [-3.907, -3.661, -3.625, -3.609]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [81198.0, 81733.0, 82269.0, 82804.0]
source_col: d(18_16)Mean
dD_mean | -55.88425
varray: [-56.875, -55.672, -55.581, -55.409]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [81198.0, 81733.0, 82269.0, 82804.0]
source_col: d(D_H)Mean
d18O_stdev | 0.139373598648
varray: [-3.907, -3.661, -3.625, -3.609]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [81198.0, 81733.0, 82269.0, 82804.0]
source_col: d(18_16)Mean
dD_stdev | 0.669442242965
varray: [-56.875, -55.672, -55.581, -55.409]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [81198.0, 81733.0, 82269.0, 82804.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.7005
varray: [-3.907, -3.661, -3.625, -3.609]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [81198.0, 81733.0, 82269.0, 82804.0]
source_col: d(18_16)Mean
measured_dD_mean | -55.88425
varray: [-56.875, -55.672, -55.581, -55.409]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [81198.0, 81733.0, 82269.0, 82804.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.139373598648
varray: [-3.907, -3.661, -3.625, -3.609]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [81198.0, 81733.0, 82269.0, 82804.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.669442242965
varray: [-56.875, -55.672, -55.581, -55.409]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [81198.0, 81733.0, 82269.0, 82804.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.7758470389
actual_array: [ 1.93 -15.84]
varray: [-3.907, -3.661, -3.625, -3.609]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.7005
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [81198.0, 81733.0, 82269.0, 82804.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -61.5488548101
actual_array: [ 18.1 -119. ]
varray: [-56.875, -55.672, -55.581, -55.409]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -55.88425
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [81198.0, 81733.0, 82269.0, 82804.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.1857
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.74995
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 33 - Sample 33 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 33
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.97825
varray: [1.791, 2.012, 2.054, 2.056]
inst: mean_all
operation: mean_all(residuals)
tarray: [83390.0, 83925.0, 84460.0, 84996.0]
source_col: residuals
SAMP_residuals_STD | 0.126471011171
varray: [1.791, 2.012, 2.054, 2.056]
inst: std_all
operation: std_all(residuals)
tarray: [83390.0, 83925.0, 84460.0, 84996.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.04775
varray: [-1.048, -0.823, -0.957, -1.363]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [83390.0, 83925.0, 84460.0, 84996.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.22958712943
varray: [-1.048, -0.823, -0.957, -1.363]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [83390.0, 83925.0, 84460.0, 84996.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 51.08
varray: [49.006, 51.283, 51.785, 52.246]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [83390.0, 83925.0, 84460.0, 84996.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.43750547825
varray: [49.006, 51.283, 51.785, 52.246]
inst: std_all
operation: std_all(baseline_shift)
tarray: [83390.0, 83925.0, 84460.0, 84996.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 25.18275
varray: [22.267, 25.191, 26.067, 27.206]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [83390.0, 83925.0, 84460.0, 84996.0]
source_col: slope_shift
SAMP_slope_shift_STD | 2.11164270573
varray: [22.267, 25.191, 26.067, 27.206]
inst: std_all
operation: std_all(slope_shift)
tarray: [83390.0, 83925.0, 84460.0, 84996.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0940673517575
varray: [-0.803, -0.618, -0.617, -0.61]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83390.0, 83925.0, 84460.0, 84996.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.746142524366
varray: [-45.759, -44.528, -44.204, -44.173]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83390.0, 83925.0, 84460.0, 84996.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 59.49275
varray: [54.365, 61.243, 62.327, 60.036]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [83390.0, 83925.0, 84460.0, 84996.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.54425773846
varray: [54.365, 61.243, 62.327, 60.036]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [83390.0, 83925.0, 84460.0, 84996.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.97825
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.04775
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 51.08
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 25.18275
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.2498
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.2498
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -11.2115975853
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.6716
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-0.803, -0.618, -0.617, -0.61]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [83390.0, 83925.0, 84460.0, 84996.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-45.759, -44.528, -44.204, -44.173]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [83390.0, 83925.0, 84460.0, 84996.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.662
varray: [-0.803, -0.618, -0.617, -0.61]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [83390.0, 83925.0, 84460.0, 84996.0]
source_col: d(18_16)Mean
dD_mean | -44.666
varray: [-45.759, -44.528, -44.204, -44.173]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [83390.0, 83925.0, 84460.0, 84996.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0940673517575
varray: [-0.803, -0.618, -0.617, -0.61]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83390.0, 83925.0, 84460.0, 84996.0]
source_col: d(18_16)Mean
dD_stdev | 0.746142524366
varray: [-45.759, -44.528, -44.204, -44.173]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83390.0, 83925.0, 84460.0, 84996.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.662
varray: [-0.803, -0.618, -0.617, -0.61]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [83390.0, 83925.0, 84460.0, 84996.0]
source_col: d(18_16)Mean
measured_dD_mean | -44.666
varray: [-45.759, -44.528, -44.204, -44.173]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [83390.0, 83925.0, 84460.0, 84996.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0940673517575
varray: [-0.803, -0.618, -0.617, -0.61]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83390.0, 83925.0, 84460.0, 84996.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.746142524366
varray: [-45.759, -44.528, -44.204, -44.173]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83390.0, 83925.0, 84460.0, 84996.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.64458993482
actual_array: [ 1.93 -15.84]
varray: [-0.803, -0.618, -0.617, -0.61]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -0.662
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [83390.0, 83925.0, 84460.0, 84996.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -49.3441126851
actual_array: [ 18.1 -119. ]
varray: [-45.759, -44.528, -44.204, -44.173]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -44.666
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [83390.0, 83925.0, 84460.0, 84996.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.2632
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -12.16145
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 34 - Sample 34 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 34
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0685
varray: [2.082, 2.086, 2.046, 2.06]
inst: mean_all
operation: mean_all(residuals)
tarray: [85583.0, 86119.0, 86656.0, 87192.0]
source_col: residuals
SAMP_residuals_STD | 0.0188591268798
varray: [2.082, 2.086, 2.046, 2.06]
inst: std_all
operation: std_all(residuals)
tarray: [85583.0, 86119.0, 86656.0, 87192.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0835
varray: [-1.225, -0.888, -0.935, -1.286]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [85583.0, 86119.0, 86656.0, 87192.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.201081243945
varray: [-1.225, -0.888, -0.935, -1.286]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [85583.0, 86119.0, 86656.0, 87192.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 52.42225
varray: [52.666, 52.447, 52.217, 52.359]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [85583.0, 86119.0, 86656.0, 87192.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.188108789446
varray: [52.666, 52.447, 52.217, 52.359]
inst: std_all
operation: std_all(baseline_shift)
tarray: [85583.0, 86119.0, 86656.0, 87192.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.0395
varray: [27.668, 26.627, 26.355, 27.508]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [85583.0, 86119.0, 86656.0, 87192.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.646323190569
varray: [27.668, 26.627, 26.355, 27.508]
inst: std_all
operation: std_all(slope_shift)
tarray: [85583.0, 86119.0, 86656.0, 87192.0]
source_col: slope_shift
SAMP_d18O_STD | 0.134156624883
varray: [-4.408, -4.667, -4.682, -4.679]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85583.0, 86119.0, 86656.0, 87192.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.998233606594
varray: [-57.993, -59.737, -60.015, -60.133]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85583.0, 86119.0, 86656.0, 87192.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.67425
varray: [61.381, 62.921, 62.161, 60.234]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [85583.0, 86119.0, 86656.0, 87192.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.14769722343
varray: [61.381, 62.921, 62.161, 60.234]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [85583.0, 86119.0, 86656.0, 87192.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0685
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0835
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 52.42225
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.0395
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.15955
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.15955
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -7.16097035524
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.4901
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-4.408, -4.667, -4.682, -4.679]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [85583.0, 86119.0, 86656.0, 87192.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-57.993, -59.737, -60.015, -60.133]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [85583.0, 86119.0, 86656.0, 87192.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.609
varray: [-4.408, -4.667, -4.682, -4.679]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [85583.0, 86119.0, 86656.0, 87192.0]
source_col: d(18_16)Mean
dD_mean | -59.4695
varray: [-57.993, -59.737, -60.015, -60.133]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [85583.0, 86119.0, 86656.0, 87192.0]
source_col: d(D_H)Mean
d18O_stdev | 0.134156624883
varray: [-4.408, -4.667, -4.682, -4.679]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85583.0, 86119.0, 86656.0, 87192.0]
source_col: d(18_16)Mean
dD_stdev | 0.998233606594
varray: [-57.993, -59.737, -60.015, -60.133]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85583.0, 86119.0, 86656.0, 87192.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.609
varray: [-4.408, -4.667, -4.682, -4.679]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [85583.0, 86119.0, 86656.0, 87192.0]
source_col: d(18_16)Mean
measured_dD_mean | -59.4695
varray: [-57.993, -59.737, -60.015, -60.133]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [85583.0, 86119.0, 86656.0, 87192.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.134156624883
varray: [-4.408, -4.667, -4.682, -4.679]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85583.0, 86119.0, 86656.0, 87192.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.998233606594
varray: [-57.993, -59.737, -60.015, -60.133]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85583.0, 86119.0, 86656.0, 87192.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.71208106196
actual_array: [ 1.93 -15.84]
varray: [-4.408, -4.667, -4.682, -4.679]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -4.609
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [85583.0, 86119.0, 86656.0, 87192.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -65.4493786622
actual_array: [ 18.1 -119. ]
varray: [-57.993, -59.737, -60.015, -60.133]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -59.4695
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [85583.0, 86119.0, 86656.0, 87192.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.17295
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.8192
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 35 - Sample 35 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 35
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.091
varray: [2.062, 2.122, 2.111, 2.069]
inst: mean_all
operation: mean_all(residuals)
tarray: [87777.0, 88312.0, 88848.0, 89384.0]
source_col: residuals
SAMP_residuals_STD | 0.0299221211369
varray: [2.062, 2.122, 2.111, 2.069]
inst: std_all
operation: std_all(residuals)
tarray: [87777.0, 88312.0, 88848.0, 89384.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.06425
varray: [-1.281, -0.936, -1.05, -0.99]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [87777.0, 88312.0, 88848.0, 89384.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.151816501079
varray: [-1.281, -0.936, -1.05, -0.99]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [87777.0, 88312.0, 88848.0, 89384.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 52.5155
varray: [52.379, 52.773, 52.649, 52.261]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [87777.0, 88312.0, 88848.0, 89384.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.236312646015
varray: [52.379, 52.773, 52.649, 52.261]
inst: std_all
operation: std_all(baseline_shift)
tarray: [87777.0, 88312.0, 88848.0, 89384.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.53575
varray: [27.539, 28.119, 27.585, 26.9]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [87777.0, 88312.0, 88848.0, 89384.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.498930439373
varray: [27.539, 28.119, 27.585, 26.9]
inst: std_all
operation: std_all(slope_shift)
tarray: [87777.0, 88312.0, 88848.0, 89384.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0328164084974
varray: [-4.742, -4.678, -4.721, -4.752]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87777.0, 88312.0, 88848.0, 89384.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.631816692826
varray: [-49.42, -48.205, -48.0, -48.381]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87777.0, 88312.0, 88848.0, 89384.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.76225
varray: [59.66, 62.844, 62.188, 62.357]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [87777.0, 88312.0, 88848.0, 89384.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.42882571249
varray: [59.66, 62.844, 62.188, 62.357]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [87777.0, 88312.0, 88848.0, 89384.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.091
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.06425
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 52.5155
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.53575
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.13705
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.13705
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.15111869123
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.4021
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-4.742, -4.678, -4.721, -4.752]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [87777.0, 88312.0, 88848.0, 89384.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-49.42, -48.205, -48.0, -48.381]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [87777.0, 88312.0, 88848.0, 89384.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.72325
varray: [-4.742, -4.678, -4.721, -4.752]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [87777.0, 88312.0, 88848.0, 89384.0]
source_col: d(18_16)Mean
dD_mean | -48.5015
varray: [-49.42, -48.205, -48.0, -48.381]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [87777.0, 88312.0, 88848.0, 89384.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0328164084974
varray: [-4.742, -4.678, -4.721, -4.752]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87777.0, 88312.0, 88848.0, 89384.0]
source_col: d(18_16)Mean
dD_stdev | 0.631816692826
varray: [-49.42, -48.205, -48.0, -48.381]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87777.0, 88312.0, 88848.0, 89384.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.72325
varray: [-4.742, -4.678, -4.721, -4.752]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [87777.0, 88312.0, 88848.0, 89384.0]
source_col: d(18_16)Mean
measured_dD_mean | -48.5015
varray: [-49.42, -48.205, -48.0, -48.381]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [87777.0, 88312.0, 88848.0, 89384.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0328164084974
varray: [-4.742, -4.678, -4.721, -4.752]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87777.0, 88312.0, 88848.0, 89384.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.631816692826
varray: [-49.42, -48.205, -48.0, -48.381]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87777.0, 88312.0, 88848.0, 89384.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.82981880234
actual_array: [ 1.93 -15.84]
varray: [-4.742, -4.678, -4.721, -4.752]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -4.72325
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [87777.0, 88312.0, 88848.0, 89384.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -53.5168926125
actual_array: [ 18.1 -119. ]
varray: [-49.42, -48.205, -48.0, -48.381]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -48.5015
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [87777.0, 88312.0, 88848.0, 89384.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.15045
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.72595
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 36 - Sample 36 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 36
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0605
varray: [2.124, 2.069, 2.05, 1.999]
inst: mean_all
operation: mean_all(residuals)
tarray: [89967.0, 90503.0, 91038.0, 91574.0]
source_col: residuals
SAMP_residuals_STD | 0.0516300945832
varray: [2.124, 2.069, 2.05, 1.999]
inst: std_all
operation: std_all(residuals)
tarray: [89967.0, 90503.0, 91038.0, 91574.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.15625
varray: [-1.217, -1.014, -1.222, -1.172]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [89967.0, 90503.0, 91038.0, 91574.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0974623859069
varray: [-1.217, -1.014, -1.222, -1.172]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [89967.0, 90503.0, 91038.0, 91574.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 52.4405
varray: [53.005, 52.476, 52.489, 51.792]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [89967.0, 90503.0, 91038.0, 91574.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.497602585738
varray: [53.005, 52.476, 52.489, 51.792]
inst: std_all
operation: std_all(baseline_shift)
tarray: [89967.0, 90503.0, 91038.0, 91574.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.23925
varray: [27.111, 26.171, 26.366, 25.309]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [89967.0, 90503.0, 91038.0, 91574.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.740728639022
varray: [27.111, 26.171, 26.366, 25.309]
inst: std_all
operation: std_all(slope_shift)
tarray: [89967.0, 90503.0, 91038.0, 91574.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0353506246998
varray: [-2.466, -2.44, -2.387, -2.405]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89967.0, 90503.0, 91038.0, 91574.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.247945390493
varray: [-44.616, -44.852, -44.473, -44.262]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89967.0, 90503.0, 91038.0, 91574.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.0035
varray: [62.501, 61.541, 60.386, 59.586]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [89967.0, 90503.0, 91038.0, 91574.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.2808883636
varray: [62.501, 61.541, 60.386, 59.586]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [89967.0, 90503.0, 91038.0, 91574.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0605
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.15625
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 52.4405
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.23925
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.16755
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.16755
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -7.52002872467
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.16085
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-2.466, -2.44, -2.387, -2.405]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [89967.0, 90503.0, 91038.0, 91574.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-44.616, -44.852, -44.473, -44.262]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [89967.0, 90503.0, 91038.0, 91574.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.4245
varray: [-2.466, -2.44, -2.387, -2.405]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [89967.0, 90503.0, 91038.0, 91574.0]
source_col: d(18_16)Mean
dD_mean | -44.55075
varray: [-44.616, -44.852, -44.473, -44.262]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [89967.0, 90503.0, 91038.0, 91574.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0353506246998
varray: [-2.466, -2.44, -2.387, -2.405]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89967.0, 90503.0, 91038.0, 91574.0]
source_col: d(18_16)Mean
dD_stdev | 0.247945390493
varray: [-44.616, -44.852, -44.473, -44.262]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89967.0, 90503.0, 91038.0, 91574.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.4245
varray: [-2.466, -2.44, -2.387, -2.405]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [89967.0, 90503.0, 91038.0, 91574.0]
source_col: d(18_16)Mean
measured_dD_mean | -44.55075
varray: [-44.616, -44.852, -44.473, -44.262]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [89967.0, 90503.0, 91038.0, 91574.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0353506246998
varray: [-2.466, -2.44, -2.387, -2.405]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89967.0, 90503.0, 91038.0, 91574.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.247945390493
varray: [-44.616, -44.852, -44.473, -44.262]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89967.0, 90503.0, 91038.0, 91574.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.46089424482
actual_array: [ 1.93 -15.84]
varray: [-2.466, -2.44, -2.387, -2.405]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.4245
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [89967.0, 90503.0, 91038.0, 91574.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -49.2187280191
actual_array: [ 18.1 -119. ]
varray: [-44.616, -44.852, -44.473, -44.262]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -44.55075
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [89967.0, 90503.0, 91038.0, 91574.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.18095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.80095
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 37 - Sample 37 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 37
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.33325
varray: [2.267, 2.362, 2.355, 2.349]
inst: mean_all
operation: mean_all(residuals)
tarray: [92158.0, 92693.0, 93229.0, 93765.0]
source_col: residuals
SAMP_residuals_STD | 0.044485016204
varray: [2.267, 2.362, 2.355, 2.349]
inst: std_all
operation: std_all(residuals)
tarray: [92158.0, 92693.0, 93229.0, 93765.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.15225
varray: [-2.194, -2.09, -2.191, -2.134]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [92158.0, 92693.0, 93229.0, 93765.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.049842251153
varray: [-2.194, -2.09, -2.191, -2.134]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [92158.0, 92693.0, 93229.0, 93765.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 63.268
varray: [62.273, 63.466, 63.753, 63.58]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [92158.0, 92693.0, 93229.0, 93765.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.673745253539
varray: [62.273, 63.466, 63.753, 63.58]
inst: std_all
operation: std_all(baseline_shift)
tarray: [92158.0, 92693.0, 93229.0, 93765.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 24.84225
varray: [24.063, 24.886, 25.074, 25.346]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [92158.0, 92693.0, 93229.0, 93765.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.552755747746
varray: [24.063, 24.886, 25.074, 25.346]
inst: std_all
operation: std_all(slope_shift)
tarray: [92158.0, 92693.0, 93229.0, 93765.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0663645236553
varray: [-2.761, -2.877, -2.901, -2.898]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [92158.0, 92693.0, 93229.0, 93765.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.10341964667
varray: [-60.688, -62.313, -62.904, -63.134]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [92158.0, 92693.0, 93229.0, 93765.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.6645
varray: [62.098, 64.899, 64.121, 63.54]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [92158.0, 92693.0, 93229.0, 93765.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.18347017988
varray: [62.098, 64.899, 64.121, 63.54]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [92158.0, 92693.0, 93229.0, 93765.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.33325
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.15225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 63.268
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 24.84225
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.1052
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1052
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 4.72161755795
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.49985
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-2.761, -2.877, -2.901, -2.898]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [92158.0, 92693.0, 93229.0, 93765.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-60.688, -62.313, -62.904, -63.134]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [92158.0, 92693.0, 93229.0, 93765.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.85925
varray: [-2.761, -2.877, -2.901, -2.898]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [92158.0, 92693.0, 93229.0, 93765.0]
source_col: d(18_16)Mean
dD_mean | -62.25975
varray: [-60.688, -62.313, -62.904, -63.134]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [92158.0, 92693.0, 93229.0, 93765.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0663645236553
varray: [-2.761, -2.877, -2.901, -2.898]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [92158.0, 92693.0, 93229.0, 93765.0]
source_col: d(18_16)Mean
dD_stdev | 1.10341964667
varray: [-60.688, -62.313, -62.904, -63.134]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [92158.0, 92693.0, 93229.0, 93765.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.85925
varray: [-2.761, -2.877, -2.901, -2.898]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [92158.0, 92693.0, 93229.0, 93765.0]
source_col: d(18_16)Mean
measured_dD_mean | -62.25975
varray: [-60.688, -62.313, -62.904, -63.134]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [92158.0, 92693.0, 93229.0, 93765.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0663645236553
varray: [-2.761, -2.877, -2.901, -2.898]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [92158.0, 92693.0, 93229.0, 93765.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.10341964667
varray: [-60.688, -62.313, -62.904, -63.134]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [92158.0, 92693.0, 93229.0, 93765.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.90891597462
actual_array: [ 1.93 -15.84]
varray: [-2.761, -2.877, -2.901, -2.898]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.85925
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [92158.0, 92693.0, 93229.0, 93765.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -68.4849931042
actual_array: [ 18.1 -119. ]
varray: [-60.688, -62.313, -62.904, -63.134]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -62.25975
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [92158.0, 92693.0, 93229.0, 93765.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 60.9082
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.02655
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 38 - Sample 38 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 38
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.08525
varray: [2.052, 2.112, 2.073, 2.104]
inst: mean_all
operation: mean_all(residuals)
tarray: [94348.0, 94884.0, 95420.0, 95956.0]
source_col: residuals
SAMP_residuals_STD | 0.027825348156
varray: [2.052, 2.112, 2.073, 2.104]
inst: std_all
operation: std_all(residuals)
tarray: [94348.0, 94884.0, 95420.0, 95956.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.03725
varray: [-0.87, -1.066, -1.212, -1.001]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [94348.0, 94884.0, 95420.0, 95956.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.142186204207
varray: [-0.87, -1.066, -1.212, -1.001]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [94348.0, 94884.0, 95420.0, 95956.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 52.5
varray: [52.056, 52.969, 52.415, 52.56]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [94348.0, 94884.0, 95420.0, 95956.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.377677993358
varray: [52.056, 52.969, 52.415, 52.56]
inst: std_all
operation: std_all(baseline_shift)
tarray: [94348.0, 94884.0, 95420.0, 95956.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.37125
varray: [26.205, 28.223, 27.488, 27.569]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [94348.0, 94884.0, 95420.0, 95956.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.844265509185
varray: [26.205, 28.223, 27.488, 27.569]
inst: std_all
operation: std_all(slope_shift)
tarray: [94348.0, 94884.0, 95420.0, 95956.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0443499342352
varray: [-5.278, -5.353, -5.369, -5.373]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94348.0, 94884.0, 95420.0, 95956.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.245858190834
varray: [-59.153, -59.096, -58.655, -58.759]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94348.0, 94884.0, 95420.0, 95956.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.948
varray: [62.152, 62.569, 60.423, 62.648]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [94348.0, 94884.0, 95420.0, 95956.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.03969258277
varray: [62.152, 62.569, 60.423, 62.648]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [94348.0, 94884.0, 95420.0, 95956.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.08525
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.03725
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 52.5
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.37125
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1428
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1428
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.40919189426
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.21635
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-5.278, -5.353, -5.369, -5.373]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [94348.0, 94884.0, 95420.0, 95956.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-59.153, -59.096, -58.655, -58.759]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [94348.0, 94884.0, 95420.0, 95956.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.34325
varray: [-5.278, -5.353, -5.369, -5.373]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [94348.0, 94884.0, 95420.0, 95956.0]
source_col: d(18_16)Mean
dD_mean | -58.91575
varray: [-59.153, -59.096, -58.655, -58.759]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [94348.0, 94884.0, 95420.0, 95956.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0443499342352
varray: [-5.278, -5.353, -5.369, -5.373]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94348.0, 94884.0, 95420.0, 95956.0]
source_col: d(18_16)Mean
dD_stdev | 0.245858190834
varray: [-59.153, -59.096, -58.655, -58.759]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94348.0, 94884.0, 95420.0, 95956.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.34325
varray: [-5.278, -5.353, -5.369, -5.373]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [94348.0, 94884.0, 95420.0, 95956.0]
source_col: d(18_16)Mean
measured_dD_mean | -58.91575
varray: [-59.153, -59.096, -58.655, -58.759]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [94348.0, 94884.0, 95420.0, 95956.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0443499342352
varray: [-5.278, -5.353, -5.369, -5.373]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94348.0, 94884.0, 95420.0, 95956.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.245858190834
varray: [-59.153, -59.096, -58.655, -58.759]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94348.0, 94884.0, 95420.0, 95956.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.46874570855
actual_array: [ 1.93 -15.84]
varray: [-5.278, -5.353, -5.369, -5.373]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.34325
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [94348.0, 94884.0, 95420.0, 95956.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -64.8469339001
actual_array: [ 18.1 -119. ]
varray: [-59.153, -59.096, -58.655, -58.759]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -58.91575
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [94348.0, 94884.0, 95420.0, 95956.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.1562
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.74145
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 39 - Sample 39 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 39
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.98825
varray: [1.972, 1.988, 1.992, 2.001]
inst: mean_all
operation: mean_all(residuals)
tarray: [96540.0, 97076.0, 97611.0, 98147.0]
source_col: residuals
SAMP_residuals_STD | 0.0121209185571
varray: [1.972, 1.988, 1.992, 2.001]
inst: std_all
operation: std_all(residuals)
tarray: [96540.0, 97076.0, 97611.0, 98147.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.05875
varray: [-1.043, -1.028, -1.061, -1.103]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [96540.0, 97076.0, 97611.0, 98147.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0324384031666
varray: [-1.043, -1.028, -1.061, -1.103]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [96540.0, 97076.0, 97611.0, 98147.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 51.4065
varray: [51.273, 51.409, 51.353, 51.591]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [96540.0, 97076.0, 97611.0, 98147.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.135069118109
varray: [51.273, 51.409, 51.353, 51.591]
inst: std_all
operation: std_all(baseline_shift)
tarray: [96540.0, 97076.0, 97611.0, 98147.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.053
varray: [25.899, 25.896, 26.232, 26.185]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [96540.0, 97076.0, 97611.0, 98147.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.180582391168
varray: [25.899, 25.896, 26.232, 26.185]
inst: std_all
operation: std_all(slope_shift)
tarray: [96540.0, 97076.0, 97611.0, 98147.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0421772055341
varray: [-4.107, -4.084, -4.037, -4.015]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96540.0, 97076.0, 97611.0, 98147.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.102379604089
varray: [-57.613, -57.826, -57.819, -57.699]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96540.0, 97076.0, 97611.0, 98147.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 59.167
varray: [58.753, 58.558, 59.378, 59.979]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [96540.0, 97076.0, 97611.0, 98147.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.644500323248
varray: [58.753, 58.558, 59.378, 59.979]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [96540.0, 97076.0, 97611.0, 98147.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.98825
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.05875
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 51.4065
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.053
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.2398
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.2398
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -10.7627746235
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.99735
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-4.107, -4.084, -4.037, -4.015]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [96540.0, 97076.0, 97611.0, 98147.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-57.613, -57.826, -57.819, -57.699]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [96540.0, 97076.0, 97611.0, 98147.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.06075
varray: [-4.107, -4.084, -4.037, -4.015]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [96540.0, 97076.0, 97611.0, 98147.0]
source_col: d(18_16)Mean
dD_mean | -57.73925
varray: [-57.613, -57.826, -57.819, -57.699]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [96540.0, 97076.0, 97611.0, 98147.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0421772055341
varray: [-4.107, -4.084, -4.037, -4.015]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96540.0, 97076.0, 97611.0, 98147.0]
source_col: d(18_16)Mean
dD_stdev | 0.102379604089
varray: [-57.613, -57.826, -57.819, -57.699]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96540.0, 97076.0, 97611.0, 98147.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.06075
varray: [-4.107, -4.084, -4.037, -4.015]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [96540.0, 97076.0, 97611.0, 98147.0]
source_col: d(18_16)Mean
measured_dD_mean | -57.73925
varray: [-57.613, -57.826, -57.819, -57.699]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [96540.0, 97076.0, 97611.0, 98147.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0421772055341
varray: [-4.107, -4.084, -4.037, -4.015]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96540.0, 97076.0, 97611.0, 98147.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.102379604089
varray: [-57.613, -57.826, -57.819, -57.699]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96540.0, 97076.0, 97611.0, 98147.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.14709448723
actual_array: [ 1.93 -15.84]
varray: [-4.107, -4.084, -4.037, -4.015]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -4.06075
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [96540.0, 97076.0, 97611.0, 98147.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -63.5669767669
actual_array: [ 18.1 -119. ]
varray: [-57.613, -57.826, -57.819, -57.699]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -57.73925
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [96540.0, 97076.0, 97611.0, 98147.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.2532
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -11.83495
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 40 - Sample 40 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 40
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.09725
varray: [2.146, 2.084, 2.081, 2.078]
inst: mean_all
operation: mean_all(residuals)
tarray: [98731.0, 99266.0, 99801.0, 100337.0]
source_col: residuals
SAMP_residuals_STD | 0.0325921769755
varray: [2.146, 2.084, 2.081, 2.078]
inst: std_all
operation: std_all(residuals)
tarray: [98731.0, 99266.0, 99801.0, 100337.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.05725
varray: [-1.123, -0.941, -1.054, -1.111]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [98731.0, 99266.0, 99801.0, 100337.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.083139942266
varray: [-1.123, -0.941, -1.054, -1.111]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [98731.0, 99266.0, 99801.0, 100337.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 52.7005
varray: [53.379, 52.343, 52.619, 52.461]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [98731.0, 99266.0, 99801.0, 100337.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.466251362822
varray: [53.379, 52.343, 52.619, 52.461]
inst: std_all
operation: std_all(baseline_shift)
tarray: [98731.0, 99266.0, 99801.0, 100337.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.11675
varray: [29.04, 27.877, 27.715, 27.835]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [98731.0, 99266.0, 99801.0, 100337.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.61931595329
varray: [29.04, 27.877, 27.715, 27.835]
inst: std_all
operation: std_all(slope_shift)
tarray: [98731.0, 99266.0, 99801.0, 100337.0]
source_col: slope_shift
SAMP_d18O_STD | 0.147326168755
varray: [-8.67, -8.916, -8.922, -9.014]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [98731.0, 99266.0, 99801.0, 100337.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.50334638504
varray: [-65.503, -66.067, -66.326, -66.701]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [98731.0, 99266.0, 99801.0, 100337.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.57125
varray: [63.173, 62.799, 62.718, 61.595]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [98731.0, 99266.0, 99801.0, 100337.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.680336375136
varray: [63.173, 62.799, 62.718, 61.595]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [98731.0, 99266.0, 99801.0, 100337.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.09725
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.05725
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 52.7005
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.11675
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1308
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1308
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.87060434012
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.5931
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-8.67, -8.916, -8.922, -9.014]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [98731.0, 99266.0, 99801.0, 100337.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-65.503, -66.067, -66.326, -66.701]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [98731.0, 99266.0, 99801.0, 100337.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.8805
varray: [-8.67, -8.916, -8.922, -9.014]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [98731.0, 99266.0, 99801.0, 100337.0]
source_col: d(18_16)Mean
dD_mean | -66.14925
varray: [-65.503, -66.067, -66.326, -66.701]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [98731.0, 99266.0, 99801.0, 100337.0]
source_col: d(D_H)Mean
d18O_stdev | 0.147326168755
varray: [-8.67, -8.916, -8.922, -9.014]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [98731.0, 99266.0, 99801.0, 100337.0]
source_col: d(18_16)Mean
dD_stdev | 0.50334638504
varray: [-65.503, -66.067, -66.326, -66.701]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [98731.0, 99266.0, 99801.0, 100337.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.8805
varray: [-8.67, -8.916, -8.922, -9.014]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [98731.0, 99266.0, 99801.0, 100337.0]
source_col: d(18_16)Mean
measured_dD_mean | -66.14925
varray: [-65.503, -66.067, -66.326, -66.701]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [98731.0, 99266.0, 99801.0, 100337.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.147326168755
varray: [-8.67, -8.916, -8.922, -9.014]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [98731.0, 99266.0, 99801.0, 100337.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.50334638504
varray: [-65.503, -66.067, -66.326, -66.701]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [98731.0, 99266.0, 99801.0, 100337.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.1139782876
actual_array: [ 1.93 -15.84]
varray: [-8.67, -8.916, -8.922, -9.014]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -8.8805
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [98731.0, 99266.0, 99801.0, 100337.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -72.7165215953
actual_array: [ 18.1 -119. ]
varray: [-65.503, -66.067, -66.326, -66.701]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -66.14925
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [98731.0, 99266.0, 99801.0, 100337.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.1442
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.54095
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 41 - Sample 41 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 41
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.96425
varray: [1.946, 1.954, 1.96, 1.997]
inst: mean_all
operation: mean_all(residuals)
tarray: [100922.0, 101458.0, 101995.0, 102530.0]
source_col: residuals
SAMP_residuals_STD | 0.0225739525412
varray: [1.946, 1.954, 1.96, 1.997]
inst: std_all
operation: std_all(residuals)
tarray: [100922.0, 101458.0, 101995.0, 102530.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.91375
varray: [-0.782, -0.888, -1.092, -0.893]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [100922.0, 101458.0, 101995.0, 102530.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.12938927055
varray: [-0.782, -0.888, -1.092, -0.893]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [100922.0, 101458.0, 101995.0, 102530.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 50.951
varray: [50.663, 50.943, 50.999, 51.199]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [100922.0, 101458.0, 101995.0, 102530.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.221220854954
varray: [50.663, 50.943, 50.999, 51.199]
inst: std_all
operation: std_all(baseline_shift)
tarray: [100922.0, 101458.0, 101995.0, 102530.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 25.30425
varray: [24.807, 24.889, 25.599, 25.922]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [100922.0, 101458.0, 101995.0, 102530.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.544114800999
varray: [24.807, 24.889, 25.599, 25.922]
inst: std_all
operation: std_all(slope_shift)
tarray: [100922.0, 101458.0, 101995.0, 102530.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0993210283206
varray: [-5.863, -5.69, -5.694, -5.633]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100922.0, 101458.0, 101995.0, 102530.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.981598823349
varray: [-52.308, -50.715, -50.587, -50.019]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100922.0, 101458.0, 101995.0, 102530.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 59.39
varray: [58.811, 60.124, 58.546, 60.079]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [100922.0, 101458.0, 101995.0, 102530.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.828865489691
varray: [58.811, 60.124, 58.546, 60.079]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [100922.0, 101458.0, 101995.0, 102530.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.96425
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.91375
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 50.951
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 25.30425
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.2638
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.2638
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -11.8399497318
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.77435
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-5.863, -5.69, -5.694, -5.633]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [100922.0, 101458.0, 101995.0, 102530.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-52.308, -50.715, -50.587, -50.019]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [100922.0, 101458.0, 101995.0, 102530.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.72
varray: [-5.863, -5.69, -5.694, -5.633]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [100922.0, 101458.0, 101995.0, 102530.0]
source_col: d(18_16)Mean
dD_mean | -50.90725
varray: [-52.308, -50.715, -50.587, -50.019]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [100922.0, 101458.0, 101995.0, 102530.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0993210283206
varray: [-5.863, -5.69, -5.694, -5.633]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100922.0, 101458.0, 101995.0, 102530.0]
source_col: d(18_16)Mean
dD_stdev | 0.981598823349
varray: [-52.308, -50.715, -50.587, -50.019]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100922.0, 101458.0, 101995.0, 102530.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.72
varray: [-5.863, -5.69, -5.694, -5.633]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [100922.0, 101458.0, 101995.0, 102530.0]
source_col: d(18_16)Mean
measured_dD_mean | -50.90725
varray: [-52.308, -50.715, -50.587, -50.019]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [100922.0, 101458.0, 101995.0, 102530.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0993210283206
varray: [-5.863, -5.69, -5.694, -5.633]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100922.0, 101458.0, 101995.0, 102530.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.981598823349
varray: [-52.308, -50.715, -50.587, -50.019]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100922.0, 101458.0, 101995.0, 102530.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.8569968568
actual_array: [ 1.93 -15.84]
varray: [-5.863, -5.69, -5.694, -5.633]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.72
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [100922.0, 101458.0, 101995.0, 102530.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -56.1341955222
actual_array: [ 18.1 -119. ]
varray: [-52.308, -50.715, -50.587, -50.019]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -50.90725
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [100922.0, 101458.0, 101995.0, 102530.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.2772
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -12.29045
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 42 - Sample 42 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 42
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0605
varray: [2.032, 2.143, 2.02, 2.047]
inst: mean_all
operation: mean_all(residuals)
tarray: [103114.0, 103650.0, 104186.0, 104722.0]
source_col: residuals
SAMP_residuals_STD | 0.056098128311
varray: [2.032, 2.143, 2.02, 2.047]
inst: std_all
operation: std_all(residuals)
tarray: [103114.0, 103650.0, 104186.0, 104722.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.02625
varray: [-1.159, -0.853, -0.983, -1.11]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [103114.0, 103650.0, 104186.0, 104722.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.137262218643
varray: [-1.159, -0.853, -0.983, -1.11]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [103114.0, 103650.0, 104186.0, 104722.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 52.38375
varray: [52.181, 53.213, 51.902, 52.239]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [103114.0, 103650.0, 104186.0, 104722.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.572071892335
varray: [52.181, 53.213, 51.902, 52.239]
inst: std_all
operation: std_all(baseline_shift)
tarray: [103114.0, 103650.0, 104186.0, 104722.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.7
varray: [26.248, 28.144, 25.716, 26.692]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [103114.0, 103650.0, 104186.0, 104722.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.04207485336
varray: [26.248, 28.144, 25.716, 26.692]
inst: std_all
operation: std_all(slope_shift)
tarray: [103114.0, 103650.0, 104186.0, 104722.0]
source_col: slope_shift
SAMP_d18O_STD | 0.111219003172
varray: [-3.6, -3.337, -3.487, -3.53]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [103114.0, 103650.0, 104186.0, 104722.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.43893763794
varray: [-54.468, -54.536, -55.247, -55.274]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [103114.0, 103650.0, 104186.0, 104722.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.24825
varray: [60.392, 63.665, 60.595, 60.341]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [103114.0, 103650.0, 104186.0, 104722.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.61489759737
varray: [60.392, 63.665, 60.595, 60.341]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [103114.0, 103650.0, 104186.0, 104722.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0605
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.02625
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 52.38375
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.7
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.16755
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.16755
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -7.52002872467
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.9161
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-3.6, -3.337, -3.487, -3.53]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [103114.0, 103650.0, 104186.0, 104722.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-54.468, -54.536, -55.247, -55.274]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [103114.0, 103650.0, 104186.0, 104722.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.4885
varray: [-3.6, -3.337, -3.487, -3.53]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [103114.0, 103650.0, 104186.0, 104722.0]
source_col: d(18_16)Mean
dD_mean | -54.88125
varray: [-54.468, -54.536, -55.247, -55.274]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [103114.0, 103650.0, 104186.0, 104722.0]
source_col: d(D_H)Mean
d18O_stdev | 0.111219003172
varray: [-3.6, -3.337, -3.487, -3.53]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [103114.0, 103650.0, 104186.0, 104722.0]
source_col: d(18_16)Mean
dD_stdev | 0.43893763794
varray: [-54.468, -54.536, -55.247, -55.274]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [103114.0, 103650.0, 104186.0, 104722.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.4885
varray: [-3.6, -3.337, -3.487, -3.53]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [103114.0, 103650.0, 104186.0, 104722.0]
source_col: d(18_16)Mean
measured_dD_mean | -54.88125
varray: [-54.468, -54.536, -55.247, -55.274]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [103114.0, 103650.0, 104186.0, 104722.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.111219003172
varray: [-3.6, -3.337, -3.487, -3.53]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [103114.0, 103650.0, 104186.0, 104722.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.43893763794
varray: [-54.468, -54.536, -55.247, -55.274]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [103114.0, 103650.0, 104186.0, 104722.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.55737525807
actual_array: [ 1.93 -15.84]
varray: [-3.6, -3.337, -3.487, -3.53]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.4885
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [103114.0, 103650.0, 104186.0, 104722.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -60.4576546361
actual_array: [ 18.1 -119. ]
varray: [-54.468, -54.536, -55.247, -55.274]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -54.88125
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [103114.0, 103650.0, 104186.0, 104722.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.18095
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.8577
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 43 - Sample 43 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 43
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.08825
varray: [2.054, 2.119, 2.07, 2.11]
inst: mean_all
operation: mean_all(residuals)
tarray: [105306.0, 105842.0, 106378.0, 106914.0]
source_col: residuals
SAMP_residuals_STD | 0.0312236555622
varray: [2.054, 2.119, 2.07, 2.11]
inst: std_all
operation: std_all(residuals)
tarray: [105306.0, 105842.0, 106378.0, 106914.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.03525
varray: [-0.906, -1.119, -1.107, -1.009]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [105306.0, 105842.0, 106378.0, 106914.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0992585008954
varray: [-0.906, -1.119, -1.107, -1.009]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [105306.0, 105842.0, 106378.0, 106914.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 52.33175
varray: [51.923, 52.8, 52.198, 52.406]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [105306.0, 105842.0, 106378.0, 106914.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.369565848891
varray: [51.923, 52.8, 52.198, 52.406]
inst: std_all
operation: std_all(baseline_shift)
tarray: [105306.0, 105842.0, 106378.0, 106914.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.376
varray: [26.593, 28.201, 27.282, 27.428]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [105306.0, 105842.0, 106378.0, 106914.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.659609480425
varray: [26.593, 28.201, 27.282, 27.428]
inst: std_all
operation: std_all(slope_shift)
tarray: [105306.0, 105842.0, 106378.0, 106914.0]
source_col: slope_shift
SAMP_d18O_STD | 0.00939858145325
varray: [-4.149, -4.162, -4.141, -4.158]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [105306.0, 105842.0, 106378.0, 106914.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.342395482836
varray: [-49.546, -48.912, -48.896, -48.798]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [105306.0, 105842.0, 106378.0, 106914.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.74925
varray: [61.794, 62.111, 60.894, 62.198]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [105306.0, 105842.0, 106378.0, 106914.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.596013073123
varray: [61.794, 62.111, 60.894, 62.198]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [105306.0, 105842.0, 106378.0, 106914.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.08825
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.03525
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 52.33175
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.376
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1398
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1398
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.27454500572
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.4151
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-4.149, -4.162, -4.141, -4.158]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [105306.0, 105842.0, 106378.0, 106914.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-49.546, -48.912, -48.896, -48.798]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [105306.0, 105842.0, 106378.0, 106914.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.1525
varray: [-4.149, -4.162, -4.141, -4.158]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [105306.0, 105842.0, 106378.0, 106914.0]
source_col: d(18_16)Mean
dD_mean | -49.038
varray: [-49.546, -48.912, -48.896, -48.798]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [105306.0, 105842.0, 106378.0, 106914.0]
source_col: d(D_H)Mean
d18O_stdev | 0.00939858145325
varray: [-4.149, -4.162, -4.141, -4.158]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [105306.0, 105842.0, 106378.0, 106914.0]
source_col: d(18_16)Mean
dD_stdev | 0.342395482836
varray: [-49.546, -48.912, -48.896, -48.798]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [105306.0, 105842.0, 106378.0, 106914.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.1525
varray: [-4.149, -4.162, -4.141, -4.158]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [105306.0, 105842.0, 106378.0, 106914.0]
source_col: d(18_16)Mean
measured_dD_mean | -49.038
varray: [-49.546, -48.912, -48.896, -48.798]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [105306.0, 105842.0, 106378.0, 106914.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.00939858145325
varray: [-4.149, -4.162, -4.141, -4.158]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [105306.0, 105842.0, 106378.0, 106914.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.342395482836
varray: [-49.546, -48.912, -48.896, -48.798]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [105306.0, 105842.0, 106378.0, 106914.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.24164536408
actual_array: [ 1.93 -15.84]
varray: [-4.149, -4.162, -4.141, -4.158]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -4.1525
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [105306.0, 105842.0, 106378.0, 106914.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -54.1005704722
actual_array: [ 18.1 -119. ]
varray: [-49.546, -48.912, -48.896, -48.798]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -49.038
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [105306.0, 105842.0, 106378.0, 106914.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.1532
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.9097
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 44 - Sample 44 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 44
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0165
varray: [2.018, 2.024, 2.005, 2.019]
inst: mean_all
operation: mean_all(residuals)
tarray: [107498.0, 108034.0, 108569.0, 109104.0]
source_col: residuals
SAMP_residuals_STD | 0.00810349718743
varray: [2.018, 2.024, 2.005, 2.019]
inst: std_all
operation: std_all(residuals)
tarray: [107498.0, 108034.0, 108569.0, 109104.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.032
varray: [-0.831, -1.142, -1.047, -1.108]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [107498.0, 108034.0, 108569.0, 109104.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.139644787467
varray: [-0.831, -1.142, -1.047, -1.108]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [107498.0, 108034.0, 108569.0, 109104.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 51.362
varray: [51.185, 51.654, 51.251, 51.358]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [107498.0, 108034.0, 108569.0, 109104.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.207308144236
varray: [51.185, 51.654, 51.251, 51.358]
inst: std_all
operation: std_all(baseline_shift)
tarray: [107498.0, 108034.0, 108569.0, 109104.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.19125
varray: [26.027, 26.416, 26.054, 26.268]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [107498.0, 108034.0, 108569.0, 109104.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.184588506323
varray: [26.027, 26.416, 26.054, 26.268]
inst: std_all
operation: std_all(slope_shift)
tarray: [107498.0, 108034.0, 108569.0, 109104.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0341906419946
varray: [-2.907, -2.913, -2.849, -2.853]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [107498.0, 108034.0, 108569.0, 109104.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.20598786372
varray: [-45.2, -45.263, -44.853, -44.906]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [107498.0, 108034.0, 108569.0, 109104.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 60.19925
varray: [61.044, 60.436, 59.789, 59.528]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [107498.0, 108034.0, 108569.0, 109104.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.680327065952
varray: [61.044, 60.436, 59.789, 59.528]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [107498.0, 108034.0, 108569.0, 109104.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0165
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.032
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 51.362
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.19125
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.21155
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.21155
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -9.49484975651
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.9651
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-2.907, -2.913, -2.849, -2.853]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [107498.0, 108034.0, 108569.0, 109104.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-45.2, -45.263, -44.853, -44.906]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [107498.0, 108034.0, 108569.0, 109104.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.8805
varray: [-2.907, -2.913, -2.849, -2.853]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [107498.0, 108034.0, 108569.0, 109104.0]
source_col: d(18_16)Mean
dD_mean | -45.0555
varray: [-45.2, -45.263, -44.853, -44.906]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [107498.0, 108034.0, 108569.0, 109104.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0341906419946
varray: [-2.907, -2.913, -2.849, -2.853]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [107498.0, 108034.0, 108569.0, 109104.0]
source_col: d(18_16)Mean
dD_stdev | 0.20598786372
varray: [-45.2, -45.263, -44.853, -44.906]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [107498.0, 108034.0, 108569.0, 109104.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.8805
varray: [-2.907, -2.913, -2.849, -2.853]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [107498.0, 108034.0, 108569.0, 109104.0]
source_col: d(18_16)Mean
measured_dD_mean | -45.0555
varray: [-45.2, -45.263, -44.853, -44.906]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [107498.0, 108034.0, 108569.0, 109104.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0341906419946
varray: [-2.907, -2.913, -2.849, -2.853]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [107498.0, 108034.0, 108569.0, 109104.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.20598786372
varray: [-45.2, -45.263, -44.853, -44.906]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [107498.0, 108034.0, 108569.0, 109104.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.93081467907
actual_array: [ 1.93 -15.84]
varray: [-2.907, -2.913, -2.849, -2.853]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.8805
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [107498.0, 108034.0, 108569.0, 109104.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -49.7678638992
actual_array: [ 18.1 -119. ]
varray: [-45.2, -45.263, -44.853, -44.906]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -45.0555
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [107498.0, 108034.0, 108569.0, 109104.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.22495
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -11.87945
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 45 - Sample 45 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 45
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.042
varray: [1.989, 2.054, 2.043, 2.082]
inst: mean_all
operation: mean_all(residuals)
tarray: [109688.0, 110223.0, 110759.0, 111294.0]
source_col: residuals
SAMP_residuals_STD | 0.0389615194776
varray: [1.989, 2.054, 2.043, 2.082]
inst: std_all
operation: std_all(residuals)
tarray: [109688.0, 110223.0, 110759.0, 111294.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.07175
varray: [-1.123, -1.085, -1.118, -0.961]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [109688.0, 110223.0, 110759.0, 111294.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0757336340956
varray: [-1.123, -1.085, -1.118, -0.961]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [109688.0, 110223.0, 110759.0, 111294.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 51.6735
varray: [51.182, 51.863, 51.878, 51.771]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [109688.0, 110223.0, 110759.0, 111294.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.33106343803
varray: [51.182, 51.863, 51.878, 51.771]
inst: std_all
operation: std_all(baseline_shift)
tarray: [109688.0, 110223.0, 110759.0, 111294.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.46975
varray: [25.47, 26.874, 26.704, 26.831]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [109688.0, 110223.0, 110759.0, 111294.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.670396089388
varray: [25.47, 26.874, 26.704, 26.831]
inst: std_all
operation: std_all(slope_shift)
tarray: [109688.0, 110223.0, 110759.0, 111294.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0522653486228
varray: [-2.362, -2.318, -2.346, -2.244]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [109688.0, 110223.0, 110759.0, 111294.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.158049570283
varray: [-45.058, -44.958, -44.844, -44.69]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [109688.0, 110223.0, 110759.0, 111294.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 60.68475
varray: [59.638, 60.805, 59.996, 62.3]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [109688.0, 110223.0, 110759.0, 111294.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.18230773631
varray: [59.638, 60.805, 59.996, 62.3]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [109688.0, 110223.0, 110759.0, 111294.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.042
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.07175
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 51.6735
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.46975
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.18605
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.18605
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -8.35035120397
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.4796
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-2.362, -2.318, -2.346, -2.244]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [109688.0, 110223.0, 110759.0, 111294.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-45.058, -44.958, -44.844, -44.69]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [109688.0, 110223.0, 110759.0, 111294.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.3175
varray: [-2.362, -2.318, -2.346, -2.244]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [109688.0, 110223.0, 110759.0, 111294.0]
source_col: d(18_16)Mean
dD_mean | -44.8875
varray: [-45.058, -44.958, -44.844, -44.69]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [109688.0, 110223.0, 110759.0, 111294.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0522653486228
varray: [-2.362, -2.318, -2.346, -2.244]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [109688.0, 110223.0, 110759.0, 111294.0]
source_col: d(18_16)Mean
dD_stdev | 0.158049570283
varray: [-45.058, -44.958, -44.844, -44.69]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [109688.0, 110223.0, 110759.0, 111294.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.3175
varray: [-2.362, -2.318, -2.346, -2.244]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [109688.0, 110223.0, 110759.0, 111294.0]
source_col: d(18_16)Mean
measured_dD_mean | -44.8875
varray: [-45.058, -44.958, -44.844, -44.69]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [109688.0, 110223.0, 110759.0, 111294.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0522653486228
varray: [-2.362, -2.318, -2.346, -2.244]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [109688.0, 110223.0, 110759.0, 111294.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.158049570283
varray: [-45.058, -44.958, -44.844, -44.69]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [109688.0, 110223.0, 110759.0, 111294.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.35062782714
actual_array: [ 1.93 -15.84]
varray: [-2.362, -2.318, -2.346, -2.244]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.3175
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [109688.0, 110223.0, 110759.0, 111294.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -49.5850905899
actual_array: [ 18.1 -119. ]
varray: [-45.058, -44.958, -44.844, -44.69]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -44.8875
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [109688.0, 110223.0, 110759.0, 111294.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.19945
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -11.56795
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 46 - Sample 46 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 46
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.05225
varray: [2.084, 2.047, 2.036, 2.042]
inst: mean_all
operation: mean_all(residuals)
tarray: [111880.0, 112416.0, 112950.0, 113486.0]
source_col: residuals
SAMP_residuals_STD | 0.0216390850084
varray: [2.084, 2.047, 2.036, 2.042]
inst: std_all
operation: std_all(residuals)
tarray: [111880.0, 112416.0, 112950.0, 113486.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.11075
varray: [-1.108, -1.135, -1.141, -1.059]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [111880.0, 112416.0, 112950.0, 113486.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0373664287831
varray: [-1.108, -1.135, -1.141, -1.059]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [111880.0, 112416.0, 112950.0, 113486.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 52.37375
varray: [52.829, 52.311, 52.176, 52.179]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [111880.0, 112416.0, 112950.0, 113486.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.309958464959
varray: [52.829, 52.311, 52.176, 52.179]
inst: std_all
operation: std_all(baseline_shift)
tarray: [111880.0, 112416.0, 112950.0, 113486.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.33025
varray: [26.586, 26.105, 26.088, 26.542]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [111880.0, 112416.0, 112950.0, 113486.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.270597333074
varray: [26.586, 26.105, 26.088, 26.542]
inst: std_all
operation: std_all(slope_shift)
tarray: [111880.0, 112416.0, 112950.0, 113486.0]
source_col: slope_shift
SAMP_d18O_STD | 0.085429795739
varray: [-0.49, -0.336, -0.36, -0.291]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [111880.0, 112416.0, 112950.0, 113486.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.297362853542
varray: [-43.088, -42.66, -42.799, -42.373]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [111880.0, 112416.0, 112950.0, 113486.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 60.646
varray: [61.966, 61.288, 59.152, 60.178]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [111880.0, 112416.0, 112950.0, 113486.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.23903510846
varray: [61.966, 61.288, 59.152, 60.178]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [111880.0, 112416.0, 112950.0, 113486.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.05225
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.11075
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 52.37375
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.33025
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1758
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1758
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -7.89030766814
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.51835
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-0.49, -0.336, -0.36, -0.291]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [111880.0, 112416.0, 112950.0, 113486.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-43.088, -42.66, -42.799, -42.373]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [111880.0, 112416.0, 112950.0, 113486.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.36925
varray: [-0.49, -0.336, -0.36, -0.291]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [111880.0, 112416.0, 112950.0, 113486.0]
source_col: d(18_16)Mean
dD_mean | -42.73
varray: [-43.088, -42.66, -42.799, -42.373]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [111880.0, 112416.0, 112950.0, 113486.0]
source_col: d(D_H)Mean
d18O_stdev | 0.085429795739
varray: [-0.49, -0.336, -0.36, -0.291]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [111880.0, 112416.0, 112950.0, 113486.0]
source_col: d(18_16)Mean
dD_stdev | 0.297362853542
varray: [-43.088, -42.66, -42.799, -42.373]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [111880.0, 112416.0, 112950.0, 113486.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.36925
varray: [-0.49, -0.336, -0.36, -0.291]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [111880.0, 112416.0, 112950.0, 113486.0]
source_col: d(18_16)Mean
measured_dD_mean | -42.73
varray: [-43.088, -42.66, -42.799, -42.373]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [111880.0, 112416.0, 112950.0, 113486.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.085429795739
varray: [-0.49, -0.336, -0.36, -0.291]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [111880.0, 112416.0, 112950.0, 113486.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.297362853542
varray: [-43.088, -42.66, -42.799, -42.373]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [111880.0, 112416.0, 112950.0, 113486.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.34290307708
actual_array: [ 1.93 -15.84]
varray: [-0.49, -0.336, -0.36, -0.291]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -0.36925
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [111880.0, 112416.0, 112950.0, 113486.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -47.2378678828
actual_array: [ 18.1 -119. ]
varray: [-43.088, -42.66, -42.799, -42.373]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -42.73
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [111880.0, 112416.0, 112950.0, 113486.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.1892
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.8677
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 47 - Sample 47 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 47
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.034
varray: [2.055, 2.051, 2.01, 2.02]
inst: mean_all
operation: mean_all(residuals)
tarray: [114070.0, 114606.0, 115142.0, 115677.0]
source_col: residuals
SAMP_residuals_STD | 0.0223755819291
varray: [2.055, 2.051, 2.01, 2.02]
inst: std_all
operation: std_all(residuals)
tarray: [114070.0, 114606.0, 115142.0, 115677.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.98925
varray: [-0.874, -1.089, -1.053, -0.941]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [114070.0, 114606.0, 115142.0, 115677.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0993726152754
varray: [-0.874, -1.089, -1.053, -0.941]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [114070.0, 114606.0, 115142.0, 115677.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 51.99675
varray: [52.127, 52.341, 51.662, 51.857]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [114070.0, 114606.0, 115142.0, 115677.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.29836261495
varray: [52.127, 52.341, 51.662, 51.857]
inst: std_all
operation: std_all(baseline_shift)
tarray: [114070.0, 114606.0, 115142.0, 115677.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.08025
varray: [26.49, 26.387, 25.826, 25.618]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [114070.0, 114606.0, 115142.0, 115677.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.424385339835
varray: [26.49, 26.387, 25.826, 25.618]
inst: std_all
operation: std_all(slope_shift)
tarray: [114070.0, 114606.0, 115142.0, 115677.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0988820846598
varray: [-3.453, -3.636, -3.631, -3.674]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [114070.0, 114606.0, 115142.0, 115677.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.480359847753
varray: [-49.154, -49.829, -50.242, -50.086]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [114070.0, 114606.0, 115142.0, 115677.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 60.83325
varray: [62.037, 60.938, 59.564, 60.794]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [114070.0, 114606.0, 115142.0, 115677.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.01201494554
varray: [62.037, 60.938, 59.564, 60.794]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [114070.0, 114606.0, 115142.0, 115677.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.034
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.98925
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 51.99675
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.08025
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.19405
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.19405
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -8.70940957339
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.3311
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-3.453, -3.636, -3.631, -3.674]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [114070.0, 114606.0, 115142.0, 115677.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-49.154, -49.829, -50.242, -50.086]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [114070.0, 114606.0, 115142.0, 115677.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.5985
varray: [-3.453, -3.636, -3.631, -3.674]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [114070.0, 114606.0, 115142.0, 115677.0]
source_col: d(18_16)Mean
dD_mean | -49.82775
varray: [-49.154, -49.829, -50.242, -50.086]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [114070.0, 114606.0, 115142.0, 115677.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0988820846598
varray: [-3.453, -3.636, -3.631, -3.674]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [114070.0, 114606.0, 115142.0, 115677.0]
source_col: d(18_16)Mean
dD_stdev | 0.480359847753
varray: [-49.154, -49.829, -50.242, -50.086]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [114070.0, 114606.0, 115142.0, 115677.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.5985
varray: [-3.453, -3.636, -3.631, -3.674]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [114070.0, 114606.0, 115142.0, 115677.0]
source_col: d(18_16)Mean
measured_dD_mean | -49.82775
varray: [-49.154, -49.829, -50.242, -50.086]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [114070.0, 114606.0, 115142.0, 115677.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0988820846598
varray: [-3.453, -3.636, -3.631, -3.674]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [114070.0, 114606.0, 115142.0, 115677.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.480359847753
varray: [-49.154, -49.829, -50.242, -50.086]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [114070.0, 114606.0, 115142.0, 115677.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.67073325756
actual_array: [ 1.93 -15.84]
varray: [-3.453, -3.636, -3.631, -3.674]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.5985
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [114070.0, 114606.0, 115142.0, 115677.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -54.9597682164
actual_array: [ 18.1 -119. ]
varray: [-49.154, -49.829, -50.242, -50.086]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -49.82775
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [114070.0, 114606.0, 115142.0, 115677.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.20745
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -11.2447
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 48 - Sample 48 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 48
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.101
varray: [2.038, 2.194, 2.04, 2.132]
inst: mean_all
operation: mean_all(residuals)
tarray: [116262.0, 116798.0, 117332.0, 117867.0]
source_col: residuals
SAMP_residuals_STD | 0.075938571666
varray: [2.038, 2.194, 2.04, 2.132]
inst: std_all
operation: std_all(residuals)
tarray: [116262.0, 116798.0, 117332.0, 117867.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.09675
varray: [-1.017, -0.983, -1.158, -1.229]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [116262.0, 116798.0, 117332.0, 117867.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.11624793905
varray: [-1.017, -0.983, -1.158, -1.229]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [116262.0, 116798.0, 117332.0, 117867.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 52.8005
varray: [52.117, 53.678, 52.073, 53.334]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [116262.0, 116798.0, 117332.0, 117867.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.826852868814
varray: [52.117, 53.678, 52.073, 53.334]
inst: std_all
operation: std_all(baseline_shift)
tarray: [116262.0, 116798.0, 117332.0, 117867.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.2915
varray: [26.167, 28.734, 26.326, 27.939]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [116262.0, 116798.0, 117332.0, 117867.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.25123312509
varray: [26.167, 28.734, 26.326, 27.939]
inst: std_all
operation: std_all(slope_shift)
tarray: [116262.0, 116798.0, 117332.0, 117867.0]
source_col: slope_shift
SAMP_d18O_STD | 0.11883707334
varray: [-2.397, -2.133, -2.315, -2.37]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [116262.0, 116798.0, 117332.0, 117867.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.670216382969
varray: [-42.692, -41.206, -41.345, -41.729]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [116262.0, 116798.0, 117332.0, 117867.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.4
varray: [61.229, 64.858, 60.285, 63.228]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [116262.0, 116798.0, 117332.0, 117867.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.04709664322
varray: [61.229, 64.858, 60.285, 63.228]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [116262.0, 116798.0, 117332.0, 117867.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.101
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.09675
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 52.8005
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.2915
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.12705
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.12705
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.70229572945
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.76435
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-2.397, -2.133, -2.315, -2.37]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [116262.0, 116798.0, 117332.0, 117867.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-42.692, -41.206, -41.345, -41.729]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [116262.0, 116798.0, 117332.0, 117867.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.30375
varray: [-2.397, -2.133, -2.315, -2.37]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [116262.0, 116798.0, 117332.0, 117867.0]
source_col: d(18_16)Mean
dD_mean | -41.743
varray: [-42.692, -41.206, -41.345, -41.729]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [116262.0, 116798.0, 117332.0, 117867.0]
source_col: d(D_H)Mean
d18O_stdev | 0.11883707334
varray: [-2.397, -2.133, -2.315, -2.37]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [116262.0, 116798.0, 117332.0, 117867.0]
source_col: d(18_16)Mean
dD_stdev | 0.670216382969
varray: [-42.692, -41.206, -41.345, -41.729]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [116262.0, 116798.0, 117332.0, 117867.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.30375
varray: [-2.397, -2.133, -2.315, -2.37]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [116262.0, 116798.0, 117332.0, 117867.0]
source_col: d(18_16)Mean
measured_dD_mean | -41.743
varray: [-42.692, -41.206, -41.345, -41.729]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [116262.0, 116798.0, 117332.0, 117867.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.11883707334
varray: [-2.397, -2.133, -2.315, -2.37]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [116262.0, 116798.0, 117332.0, 117867.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.670216382969
varray: [-42.692, -41.206, -41.345, -41.729]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [116262.0, 116798.0, 117332.0, 117867.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.3364580772
actual_array: [ 1.93 -15.84]
varray: [-2.397, -2.133, -2.315, -2.37]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.30375
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [116262.0, 116798.0, 117332.0, 117867.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -46.1640746907
actual_array: [ 18.1 -119. ]
varray: [-42.692, -41.206, -41.345, -41.729]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -41.743
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [116262.0, 116798.0, 117332.0, 117867.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.14045
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.44095
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 49 - Sample 49 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Sample Variables
Sample 49
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0935
varray: [2.149, 2.076, 2.078, 2.071]
inst: mean_all
operation: mean_all(residuals)
tarray: [118453.0, 118988.0, 119524.0, 120059.0]
source_col: residuals
SAMP_residuals_STD | 0.0371169323445
varray: [2.149, 2.076, 2.078, 2.071]
inst: std_all
operation: std_all(residuals)
tarray: [118453.0, 118988.0, 119524.0, 120059.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9865
varray: [-0.903, -1.018, -1.029, -0.996]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [118453.0, 118988.0, 119524.0, 120059.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0573323643329
varray: [-0.903, -1.018, -1.029, -0.996]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [118453.0, 118988.0, 119524.0, 120059.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 52.57925
varray: [53.059, 52.413, 52.529, 52.316]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [118453.0, 118988.0, 119524.0, 120059.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.331473855178
varray: [53.059, 52.413, 52.529, 52.316]
inst: std_all
operation: std_all(baseline_shift)
tarray: [118453.0, 118988.0, 119524.0, 120059.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.03325
varray: [28.958, 27.509, 27.65, 28.016]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [118453.0, 118988.0, 119524.0, 120059.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.652477011089
varray: [28.958, 27.509, 27.65, 28.016]
inst: std_all
operation: std_all(slope_shift)
tarray: [118453.0, 118988.0, 119524.0, 120059.0]
source_col: slope_shift
SAMP_d18O_STD | 0.162618930837
varray: [-8.57, -8.842, -8.919, -8.904]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [118453.0, 118988.0, 119524.0, 120059.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.26102164269
varray: [-63.388, -65.386, -65.768, -66.263]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [118453.0, 118988.0, 119524.0, 120059.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.49575
varray: [64.019, 62.191, 62.038, 61.735]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [118453.0, 118988.0, 119524.0, 120059.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.03302771018
varray: [64.019, 62.191, 62.038, 61.735]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [118453.0, 118988.0, 119524.0, 120059.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0935
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9865
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 52.57925
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.03325
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.13455
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.13455
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.03891295079
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.6686
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4369
varray: [-8.57, -8.842, -8.919, -8.904]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [118453.0, 118988.0, 119524.0, 120059.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6917
varray: [-63.388, -65.386, -65.768, -66.263]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [118453.0, 118988.0, 119524.0, 120059.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.80875
varray: [-8.57, -8.842, -8.919, -8.904]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [118453.0, 118988.0, 119524.0, 120059.0]
source_col: d(18_16)Mean
dD_mean | -65.20125
varray: [-63.388, -65.386, -65.768, -66.263]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [118453.0, 118988.0, 119524.0, 120059.0]
source_col: d(D_H)Mean
d18O_stdev | 0.162618930837
varray: [-8.57, -8.842, -8.919, -8.904]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [118453.0, 118988.0, 119524.0, 120059.0]
source_col: d(18_16)Mean
dD_stdev | 1.26102164269
varray: [-63.388, -65.386, -65.768, -66.263]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [118453.0, 118988.0, 119524.0, 120059.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.80875
varray: [-8.57, -8.842, -8.919, -8.904]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [118453.0, 118988.0, 119524.0, 120059.0]
source_col: d(18_16)Mean
measured_dD_mean | -65.20125
varray: [-63.388, -65.386, -65.768, -66.263]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [118453.0, 118988.0, 119524.0, 120059.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.162618930837
varray: [-8.57, -8.842, -8.919, -8.904]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [118453.0, 118988.0, 119524.0, 120059.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.26102164269
varray: [-63.388, -65.386, -65.768, -66.263]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [118453.0, 118988.0, 119524.0, 120059.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.0400379561
actual_array: [ 1.93 -15.84]
varray: [-8.57, -8.842, -8.919, -8.904]
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -8.80875
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: [118453.0, 118988.0, 119524.0, 120059.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -71.6851579214
actual_array: [ 18.1 -119. ]
varray: [-63.388, -65.386, -65.768, -66.263]
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -65.20125
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: [118453.0, 118988.0, 119524.0, 120059.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 61.14795
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.6622
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 1 - Sample 1 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 1 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.875
operation: col(d(18_16)Mean)
inst: col
dD_actual | -75.572
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.1388376557
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -9.875
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.9678739771
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -75.572
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 1 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.474
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.771
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.7255962212
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -9.474
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.8326278543
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -71.771
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 1 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.258
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.755
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5030023313
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -9.258
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.6393481428
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -69.755
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 1 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.183
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.948
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4257127862
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -9.183
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.7613834964
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -68.948
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 1 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.176
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.891
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4184990953
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -9.176
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.6993711236
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -68.891
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 1 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.107
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.321
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3473927138
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -9.107
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.0792473956
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -68.321
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 1 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.126
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.064
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3669727319
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -9.126
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.7996477499
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -68.064
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 1 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.074
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.978
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.313385314
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -9.074
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.7060852225
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -67.978
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 1 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.043
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.716
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2814389687
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -9.043
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.4210458949
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -67.716
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 1 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.082
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.968
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3216295321
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -9.082
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.6952058588
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -67.968
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 2 - Sample 2 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 2 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.974
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.713
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8067
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3267
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.239
operation: col(d(18_16)Mean)
inst: col
dD_actual | 8.309
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.34496966991
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: 2.239
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 8.28931624271
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: 8.309
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 2 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.974
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.713
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8067
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3267
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.687
operation: col(d(18_16)Mean)
inst: col
dD_actual | 15.198
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.80664588601
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: 2.687
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 15.7841098601
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: 15.198
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 2 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.974
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.713
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8067
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3267
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.769
operation: col(d(18_16)Mean)
inst: col
dD_actual | 16.918
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.89114912199
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: 2.769
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 17.6553604077
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: 16.918
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 2 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.974
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.713
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8067
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3267
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.883
operation: col(d(18_16)Mean)
inst: col
dD_actual | 18.143
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.00862923056
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: 2.883
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 18.9880824546
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: 18.143
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 2 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.974
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.713
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8067
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3267
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.846
operation: col(d(18_16)Mean)
inst: col
dD_actual | 18.353
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.97049972164
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: 2.846
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.2165490912
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: 18.353
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 2 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.974
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.713
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8067
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3267
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.873
operation: col(d(18_16)Mean)
inst: col
dD_actual | 18.732
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.99832395787
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: 2.873
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.6288769735
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: 18.732
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 2 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.974
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.713
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8067
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3267
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.926
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.289
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.05294190308
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: 2.926
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.2348575287
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: 19.289
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 2 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.974
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.713
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8067
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3267
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.923
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.124
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.04985032128
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: 2.923
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.0553480285
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: 19.124
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 2 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.974
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.713
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8067
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3267
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.947
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.488
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.07458297571
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: 2.947
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.4513568653
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: 19.488
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 2 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.974
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.713
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8067
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3267
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.974
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.713
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.10240721195
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: 2.974
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.6961425474
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: 19.713
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 3 - Sample 3 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 3 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.561
operation: col(d(18_16)Mean)
inst: col
dD_actual | -95.448
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.9373611659
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -13.561
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -104.591697165
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -95.448
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 3 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.282
operation: col(d(18_16)Mean)
inst: col
dD_actual | -105.739
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.6803713262
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -14.282
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -115.787650296
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -105.739
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 3 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.497
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.726
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9019346888
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -14.497
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -119.037316217
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -108.726
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 3 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.482
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.713
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.8864767798
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -14.482
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.111109409
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -109.713
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 3 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.55
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.198
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.956552634
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -14.55
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.638758546
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -110.198
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 3 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.588
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.9
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9957126702
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -14.588
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.402489874
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -110.9
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 3 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.593
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.092
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0008653065
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -14.593
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.611373657
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -111.092
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 3 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.611
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.537
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0194147974
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -14.611
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.095505339
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -111.537
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 3 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.604
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.624
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0122011065
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -14.604
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.190155803
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -111.624
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 3 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.601
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.94
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0091095247
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -14.601
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.533943694
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -111.94
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 4 - Sample 4 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 4 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.327
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.374
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5741087128
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -9.327
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.4886534823
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -72.374
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 4 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.156
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.081
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.39788855
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -9.156
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.9060790329
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -69.081
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 4 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.052
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.062
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2907137141
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -9.052
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.7974718771
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -68.062
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 4 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.074
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.733
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.313385314
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -9.074
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.4395408131
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -67.733
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 4 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.101
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.034
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3412095502
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -9.101
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.7670096589
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -68.034
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 4 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.048
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.499
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.286591605
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -9.048
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.1849637037
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -67.499
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 4 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.08
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.604
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3195684776
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -9.08
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.299197022
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -67.604
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 4 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.035
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.439
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2731947505
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -9.035
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.1196875218
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -67.439
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 4 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.012
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.402
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2494926234
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -9.012
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.0794338763
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -67.402
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 4 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.015
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.24
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2525842052
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -9.015
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.9031881852
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -67.24
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 5 - Sample 5 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 5 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.591
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.307
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.72405883922
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.591
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.9208435435
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -55.307
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 5 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.456
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.775
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.58493765803
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.456
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -59.2541250325
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -53.775
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 5 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.322
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.214
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.44684700411
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.322
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.6437927319
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -53.214
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 5 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.422
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.47
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.54989973091
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.422
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.9223044413
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -53.47
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 6 - Sample 6 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 6 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.962380883342
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.750347012817
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 6 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.962380883342
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.750347012817
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 6 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.962380883342
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.750347012817
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 6 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -0.962380883342
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -0.750347012817
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 7 - Sample 7 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 7 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.239
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.017
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.33078597277
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -4.239
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.6932783625
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -56.017
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 7 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.083
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.354
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.17002371895
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -4.083
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.9719765526
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -55.354
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 7 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.043
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.548
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.12880262822
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -4.043
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.1830362074
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -55.548
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 7 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.111
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.697
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.19887848245
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -4.111
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.3451387258
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -55.697
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 8 - Sample 8 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 8 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.127
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.044
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.27642145492
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -6.127
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.4588435498
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -53.044
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 8 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.186
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.282
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.33722256373
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -6.186
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.6298360398
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -52.282
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 8 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.245
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.402
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.39802367255
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -6.245
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.7603884036
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -52.402
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 8 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.258
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.378
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.41142052704
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -6.258
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.7342779308
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -52.378
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 9 - Sample 9 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 9 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.245
operation: col(d(18_16)Mean)
inst: col
dD_actual | -46.704
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.30644186829
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.245
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.5613269967
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -46.704
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 9 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.169
operation: col(d(18_16)Mean)
inst: col
dD_actual | -46.343
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.22812179591
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.169
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.168581969
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -46.343
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 9 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.172
operation: col(d(18_16)Mean)
inst: col
dD_actual | -46.194
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.23121337772
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.172
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.0064794506
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -46.194
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 9 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.135
operation: col(d(18_16)Mean)
inst: col
dD_actual | -46.294
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.1930838688
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.135
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.1152730871
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -46.294
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 10 - Sample 10 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 10 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.441
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.51
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.6000070171
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -6.441
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.1416946255
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -55.51
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 10 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.587
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.316
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.75046399824
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -6.587
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.0185713356
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -56.316
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 10 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.632
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.394
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.7968377253
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -6.632
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.1034303721
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -56.394
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 10 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.632
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.684
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.7968377253
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -6.632
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.4189319179
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -56.684
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 11 - Sample 11 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 11 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.788
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.644
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.83549090677
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.788
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.0236690039
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -52.644
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 11 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.559
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.147
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.59950016238
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.559
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.4829646306
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -52.147
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 11 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.605
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.525
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.64690441671
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.605
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.8942045765
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -52.525
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 11 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.574
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.372
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.6149580714
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.574
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.7277503127
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -52.372
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 12 - Sample 12 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 12 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.588
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.622
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.72096725742
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.588
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.911798039
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -51.622
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 12 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.692
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.495
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.8281420933
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.692
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.7736301207
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -51.495
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 12 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.741
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.364
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.87863792943
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.741
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.6311104569
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -51.364
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 12 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.738
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.408
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.87554634763
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.738
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.6789796569
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -51.408
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 13 - Sample 13 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 13 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.869
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.125
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.102127224
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -8.869
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.6022027736
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -65.125
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 13 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.941
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.111
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1763251873
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -8.941
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.6749080293
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -66.111
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 13 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.982
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.619
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2185768053
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -8.982
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.2275797026
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -66.619
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 13 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.944
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.315
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1794167691
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -8.944
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.8968470477
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -66.315
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 14 - Sample 14 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 14 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.597
operation: col(d(18_16)Mean)
inst: col
dD_actual | -54.193
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.73024200283
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.597
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -59.708882433
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -54.193
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 14 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.497
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.342
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.62718927602
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.497
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.7830485866
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -53.342
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 14 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.459
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.187
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.58802923983
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.459
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.61441845
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -53.187
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 14 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.379
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.523
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.50558705839
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.379
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.8920287037
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -52.523
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 15 - Sample 15 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 15 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.341
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.228
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.419556183164
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: 1.341
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.5710874761
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -52.228
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 15 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.406
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.915
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.486540455589
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: 1.406
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.3184997588
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -52.915
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 15 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.407
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.169
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.487570982857
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: 1.407
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.5948355954
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -53.169
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 15 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.395
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.133
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.47520465564
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: 1.395
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.5556698863
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -53.133
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 16 - Sample 16 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 16 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.241
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.852
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.30231975922
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.241
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.1620234029
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -51.852
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 16 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.166
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.439
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.22503021411
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.166
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.6247693194
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -50.439
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 16 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.145
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.552
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.20338914148
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.145
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.7477061286
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -50.552
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 16 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.132
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.325
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.18999228699
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.132
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.5007445738
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -50.325
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 17 - Sample 17 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 17 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.725
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.535
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.80109495697
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.725
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -63.3447657644
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -57.535
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 17 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.8
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.425
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.87838450208
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.8
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.3130291291
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -58.425
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 17 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.847
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.676
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.92681928368
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.847
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.5861011567
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -58.676
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 17 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.786
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.623
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.86395712032
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.786
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.5284405293
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -58.623
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 18 - Sample 18 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 18 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.46
operation: col(d(18_16)Mean)
inst: col
dD_actual | -54.253
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.49747796284
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.46
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -59.7741586149
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -54.253
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 18 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.343
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.593
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.37690627247
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.343
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -59.0561206141
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -53.593
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 18 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.296
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.214
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.32847149087
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.296
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.6437927319
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -53.214
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 18 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.366
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.529
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.40060839964
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.366
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.9864926868
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -53.529
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 19 - Sample 19 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 19 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.292
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.28
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.41593118606
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.292
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.5949601804
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -63.28
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 19 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.438
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.535
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.5663881672
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.438
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.9603203183
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -64.535
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 19 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.434
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.694
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.56226605813
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.434
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.1333022003
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -64.694
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 19 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.455
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.758
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.58390713076
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.455
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.2029301277
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -64.758
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 20 - Sample 20 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 20 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.562
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.128
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.63311901227
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.562
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.7261029342
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -55.128
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 20 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.522
operation: col(d(18_16)Mean)
inst: col
dD_actual | -54.363
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.59189792155
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.522
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -59.8938316151
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -54.363
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 20 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.527
operation: col(d(18_16)Mean)
inst: col
dD_actual | -54.059
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.59705055789
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.527
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -59.5630989602
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -54.059
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 20 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.507
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.887
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.57644001253
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.507
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -59.3759739054
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -53.887
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 21 - Sample 21 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 21 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.75
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.353
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.82685813867
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.75
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.9708886163
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -55.353
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 21 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.726
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.565
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.80212548424
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.726
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.2015311256
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -55.565
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 21 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.77
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.688
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.84746868403
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.77
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.3353472985
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -55.688
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 21 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.724
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.834
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.8000644297
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.724
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.4941860078
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -55.834
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 22 - Sample 22 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 22 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.7
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.266
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.92796811571
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -8.7
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.755601801
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -65.266
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 22 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.893
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.285
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1268598784
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -8.893
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.8642089568
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -66.285
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 22 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.916
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.492
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1505620056
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -8.916
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.0894117843
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -66.492
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 22 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.986
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.725
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2226989144
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -8.986
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.3429009573
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -66.725
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 23 - Sample 23 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 23 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.455
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.143
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.52285259459
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.455
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.8303583445
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -56.143
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 23 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.296
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.427
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.35899875896
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.296
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.0513959073
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -55.427
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 23 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.212
operation: col(d(18_16)Mean)
inst: col
dD_actual | -54.733
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.27243446844
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.212
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.2963680701
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -54.733
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 23 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.244
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.128
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.30541134102
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.244
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.7261029342
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -55.128
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 24 - Sample 24 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 24 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.585
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.114
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.59576660326
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -1.585
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.2384900142
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -61.114
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 24 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.478
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.639
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.48550018558
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -1.478
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.8096566057
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -61.639
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 24 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.438
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.612
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.44427909485
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -1.438
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.7802823239
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -61.612
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 24 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.426
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.752
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.43191276764
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -1.426
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.9325934149
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -61.752
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 25 - Sample 25 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 25 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.721
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.91
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.82750011599
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -4.721
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.0165509957
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -60.91
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 25 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.031
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.468
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.14696356909
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.031
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.6236194873
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -61.468
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 25 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.014
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.274
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.12944460553
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.014
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.4125598325
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -61.274
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 25 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.07
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.303
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.18715413255
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.07
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.4441099871
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -61.303
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 26 - Sample 26 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 26 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.244
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.089
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.3664658772
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.244
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.3319279566
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -51.089
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 26 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.223
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.093
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.34482480457
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.223
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.2483433372
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -50.093
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 26 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.2
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.822
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.3211226774
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.2
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.9535125823
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -49.822
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 26 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.2
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.439
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.3211226774
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.2
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.5368329546
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -49.439
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 27 - Sample 27 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 27 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.142
operation: col(d(18_16)Mean)
inst: col
dD_actual | -47.208
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.20029755967
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.142
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.1096469246
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -47.208
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 27 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.121
operation: col(d(18_16)Mean)
inst: col
dD_actual | -47.125
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.17865648704
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.121
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.0193482063
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -47.125
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 27 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.047
operation: col(d(18_16)Mean)
inst: col
dD_actual | -46.791
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.10239746921
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.047
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.6559774604
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -46.791
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 27 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.078
operation: col(d(18_16)Mean)
inst: col
dD_actual | -46.829
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.13434381452
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.078
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.6973190423
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -46.829
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 28 - Sample 28 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 28 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.742
operation: col(d(18_16)Mean)
inst: col
dD_actual | -54.752
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.78808665244
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.742
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.317038861
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -54.752
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 28 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.828
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.596
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.87671199749
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.828
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.2352571529
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -55.596
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 28 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.846
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.053
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.89526148832
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.846
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.7324440717
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -56.053
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 28 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.884
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.223
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.93442152451
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.884
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.9173932537
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -56.223
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 29 - Sample 29 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 29 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.827
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.832
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.96726327449
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.827
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.492010135
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -55.832
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 29 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.828
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.204
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.96829380176
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.828
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.8087860979
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -55.204
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 29 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.816
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.309
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.95592747454
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.816
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.9230194162
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -55.309
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 29 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.918
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.445
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.06104125589
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.918
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.0709787618
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -55.445
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 30 - Sample 30 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 30 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.102
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.55
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.12854920086
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.102
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.273148445
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -56.55
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 30 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.949
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.768
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.97087852885
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -1.949
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.5103185725
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -56.768
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 30 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.84
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.174
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.85855105662
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -1.84
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.8640843718
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -56.174
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 30 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.903
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.736
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.92347427451
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -1.903
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.4755046089
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -56.736
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 31 - Sample 31 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 31 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.618
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.528
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.84346487972
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -8.618
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.0406411286
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -65.528
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 31 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.832
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.347
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0639977151
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -8.832
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.9316610114
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -66.347
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 31 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.947
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.773
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1825083509
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -8.947
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.3951219028
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -66.773
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 31 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.923
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.332
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1577756965
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -8.923
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.9153419659
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -66.332
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 32 - Sample 32 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 32 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.907
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.875
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.98865091976
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.907
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.6267277636
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -56.875
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 32 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.661
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.672
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.73514121181
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.661
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.3179403167
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -55.672
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 32 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.625
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.581
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.69804223016
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.625
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.2189381075
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -55.581
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 32 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.609
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.409
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.68155379387
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.609
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.0318130527
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -55.409
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 33 - Sample 33 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 33 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.803
operation: col(d(18_16)Mean)
inst: col
dD_actual | -45.759
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.78989427962
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -0.803
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -50.5332271319
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -45.759
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 33 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.618
operation: col(d(18_16)Mean)
inst: col
dD_actual | -44.528
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.59924673502
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -0.618
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -49.1939774668
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -44.528
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 33 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.617
operation: col(d(18_16)Mean)
inst: col
dD_actual | -44.204
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.59821620775
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -0.617
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -48.8414860846
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -44.204
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 33 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.61
operation: col(d(18_16)Mean)
inst: col
dD_actual | -44.173
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.59100251688
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -0.61
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -48.8077600573
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -44.173
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 34 - Sample 34 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 34 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.408
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.993
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.50494508107
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -4.408
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -63.8430406195
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -57.993
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 34 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.667
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.737
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.77185164351
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -4.667
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -65.7404016398
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -59.737
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 34 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.682
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.015
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.78730955253
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -4.682
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.0428479492
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -60.015
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 34 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.679
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.133
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.78421797073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -4.679
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.1712244402
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -60.133
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 35 - Sample 35 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 35 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.742
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.42
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.84914118861
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -4.742
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.5161621636
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -49.42
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 35 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.678
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.205
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.78318744346
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -4.678
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.1943194803
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -48.205
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 35 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.721
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.0
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.82750011599
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -4.721
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.9712925255
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -48.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 35 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.752
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.381
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.8594464613
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -4.752
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.3857962805
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -48.381
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 36 - Sample 36 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 36 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.466
operation: col(d(18_16)Mean)
inst: col
dD_actual | -44.616
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.50366112645
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.466
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -49.2897158669
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -44.616
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 36 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.44
operation: col(d(18_16)Mean)
inst: col
dD_actual | -44.852
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.47686741748
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.44
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -49.546468849
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -44.852
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 36 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.387
operation: col(d(18_16)Mean)
inst: col
dD_actual | -44.473
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.42224947227
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.387
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -49.1341409667
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -44.473
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 36 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.405
operation: col(d(18_16)Mean)
inst: col
dD_actual | -44.262
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.44079896309
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.405
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -48.9045863937
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -44.262
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 37 - Sample 37 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 37 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.761
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.688
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.80766667053
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.761
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.7750291227
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -60.688
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 37 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.877
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.313
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.92720783363
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.877
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -68.5429257156
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -62.313
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 37 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.901
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.904
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.95194048807
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.901
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.1858961072
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -62.904
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 37 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.898
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.134
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.94884890626
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.898
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.4361214712
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -63.134
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 38 - Sample 38 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 38 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.278
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.153
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.40150380431
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.278
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -65.1050468027
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -59.153
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 38 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.353
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.096
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.47879334942
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.353
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -65.0430344299
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -59.096
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 38 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.369
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.655
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.49528178571
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.369
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.563254493
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -58.655
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 38 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.373
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.759
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.49940389478
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.373
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.6763998749
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -58.759
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 39 - Sample 39 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 39 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.107
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.613
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.19475637338
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -4.107
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -63.4296248008
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -57.613
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 39 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.084
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.826
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.17105424621
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -4.084
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -63.6613552465
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -57.826
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 39 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.037
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.819
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.12261946461
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -4.037
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -63.653739692
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -57.819
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 39 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.015
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.699
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.09994786472
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -4.015
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -63.5231873282
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -57.699
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 40 - Sample 40 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 40 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.67
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.503
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.89705229766
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -8.67
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.0134427195
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -65.503
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 40 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.916
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.067
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1505620056
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -8.916
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.6270388293
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -66.067
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 40 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.922
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.326
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1567451692
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -8.922
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.9088143477
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -66.326
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 40 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.014
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.701
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2515536779
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -9.014
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.3167904846
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -66.701
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 41 - Sample 41 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 41 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.863
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.308
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.00436225614
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.863
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.6581223853
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -52.308
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 41 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.69
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.715
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.82608103876
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.69
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.9250397561
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -50.715
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 41 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.694
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.587
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.83020314783
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.694
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.7857839014
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -50.587
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 41 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.633
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.019
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.76734098448
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -5.633
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.1678360462
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -50.019
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 42 - Sample 42 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 42 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.6
operation: col(d(18_16)Mean)
inst: col
dD_actual | -54.468
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.67227904846
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.6
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.0080649334
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -54.468
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 42 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.337
operation: col(d(18_16)Mean)
inst: col
dD_actual | -54.536
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.40125037695
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.337
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.0820446062
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -54.536
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 42 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.487
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.247
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.55582946716
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.487
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.8555673616
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -55.247
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 42 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.53
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.274
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.60014213969
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.53
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.8849416434
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -55.274
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 43 - Sample 43 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 43 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.149
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.546
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.23803851864
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -4.149
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.6532421456
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -49.546
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 43 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.162
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.912
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.25143537312
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -4.162
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.9634904903
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -48.912
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 43 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.141
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.896
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.22979430049
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -4.141
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.9460835084
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -48.896
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 43 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.158
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.798
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.24731326405
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -4.158
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.8394657447
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -48.798
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 44 - Sample 44 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 44 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.907
operation: col(d(18_16)Mean)
inst: col
dD_actual | -45.2
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.95812365167
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.907
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -49.925070704
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -45.2
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 44 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.913
operation: col(d(18_16)Mean)
inst: col
dD_actual | -45.263
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.96430681528
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.913
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -49.9936106949
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -45.263
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 44 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.849
operation: col(d(18_16)Mean)
inst: col
dD_actual | -44.853
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.89835307012
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.849
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -49.5475567854
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -44.853
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 44 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.853
operation: col(d(18_16)Mean)
inst: col
dD_actual | -44.906
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.9024751792
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.853
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -49.6052174127
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -44.906
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 45 - Sample 45 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 45 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.362
operation: col(d(18_16)Mean)
inst: col
dD_actual | -45.058
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.39648629057
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.362
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -49.7705837402
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -45.058
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 45 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.318
operation: col(d(18_16)Mean)
inst: col
dD_actual | -44.958
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.35114309077
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.318
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -49.6617901037
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -44.958
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 45 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.346
operation: col(d(18_16)Mean)
inst: col
dD_actual | -44.844
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.37999785428
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.346
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -49.5377653581
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -44.844
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 45 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.244
operation: col(d(18_16)Mean)
inst: col
dD_actual | -44.69
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.27488407293
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.244
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -49.3702231579
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -44.69
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 46 - Sample 46 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 46 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.49
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.088
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.46733924471
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -0.49
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -47.6273491014
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -43.088
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 46 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.336
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.66
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.30863804542
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -0.336
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -47.1617123372
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -42.66
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 46 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.36
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.799
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.33337069985
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -0.36
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -47.312935492
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -42.799
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 46 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.291
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.373
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.26226431836
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -0.291
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -46.8494746005
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -42.373
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 47 - Sample 47 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 47 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.453
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.154
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.52079154005
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.453
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.2267710906
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -49.154
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 47 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.636
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.829
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.70937803011
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.636
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.9611281368
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -49.829
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 47 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.631
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.242
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.70422539377
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.631
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.4104458555
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -50.242
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 47 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.674
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.086
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.7485380663
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -3.674
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.2407277826
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -50.086
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 48 - Sample 48 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 48 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.397
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.692
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.43255474495
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.397
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -47.1965263009
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -42.692
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 48 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.133
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.206
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.16049554617
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.133
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -45.5798528628
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -41.206
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 48 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.315
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.345
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.34805150897
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.315
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -45.7310760175
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -41.345
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 48 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.37
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.729
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.40473050871
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -2.37
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -46.1488435816
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -41.729
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean

## 49 - Sample 49 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:05:32
Detail Variables
Detail 49 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.57
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.388
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.79399957086
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -8.57
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.7124573078
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -63.388
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 49 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.842
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.386
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0743029878
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -8.842
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.8861541648
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -65.386
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 49 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.919
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.768
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1536535874
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -8.919
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.3017458562
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -65.768
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean
Detail 49 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8067
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3267
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.601
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.94
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4369
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6917
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.904
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.263
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1381956784
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.933872312887
measure_array: [ 2.8067 -14.4369]
inst: fit_type1
actual_mean: -8.904
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.970377039955
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.8402743568
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.689697520058
measure_array: [ 17.3267 -108.6917]
inst: fit_type1
actual_mean: -66.263
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.919171407732
tarray: []
source_col: d(D_H)Mean