## Summary
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
221213_22-309_22-283_Precip_HIDS2052
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 | * | EV | -9.7755 | -71.2768 |  |  |  | 2.4761 | 20.4084
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.27536 | 3.995
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.3814 | 5.694
4 |  | EV | -10.3829 | -74.8295 |  |  |  | 0.086927 | 1.6789
5 |  | 22-0684 | -9.9635 | -70.0889 |  |  |  | 0.041987 | 0.18007
6 |  | 22-0685 | -9.0531 | -65.1799 |  |  |  | 0.037986 | 0.13216
7 |  | 22-0679 | -8.9021 | -66.4077 |  |  |  | 0.031623 | 0.20342
8 |  | 22-0693 | -10.7323 | -74.1018 |  |  |  | 0.1011 | 0.90501
9 |  | 22-0678 | -9.317 | -63.4031 |  |  |  | 0.10666 | 0.99868
10 |  | 22-0191 | -5.3256 | -42.6319 |  |  |  | 0.14263 | 1.2177
11 |  | 22-0691 | -9.4387 | -66.8237 |  |  |  | 0.10785 | 1.047
12 |  | 22-0185 | -6.8818 | -55.7721 |  |  |  | 0.2955 | 2.0191
13 |  | EV | -10.1684 | -72.6089 |  |  |  | 0.13409 | 1.0697
14 | * | 22-0201 | 8.458 | 26.9625 |  |  |  | 7.4737 | 31.9029
15 |  | 22-0688 | -9.1059 | -66.0571 |  |  |  | 0.60057 | 6.27
16 |  | 22-0184 | -6.9745 | -64.3331 |  |  |  | 0.16871 | 0.63774
17 |  | 22-0173 | -6.0262 | -47.9163 |  |  |  | 0.0415 | 0.99676
18 |  | 22-0187 | -5.137 | -47.646 |  |  |  | 0.10763 | 0.24023
19 |  | 22-0694 | -9.7449 | -71.7402 |  |  |  | 0.13352 | 1.3088
20 |  | 22-0692 | -9.4272 | -63.2253 |  |  |  | 0.097841 | 0.80711
21 |  | 22-0190 | 0.083589 | -26.1025 |  |  |  | 0.29639 | 1.9952
22 |  | EV | -10.0141 | -71.2357 |  |  |  | 0.27316 | 2.5155
23 |  | 22-0200 | 9.3727 | 26.0964 |  |  |  | 0.56787 | 5.7627
24 | * | 22-0175 | -4.4955 | -33.0518 |  |  |  | 1.1641 | 10.808
25 | * | 22-0180 | -3.5475 | -38.7727 |  |  |  | 0.5184 | 14.7929
26 |  | 22-0198 | 3.2291 | -6.8237 |  |  |  | 0.35041 | 2.8198
27 |  | 22-0681 | -9.2361 | -67.7365 |  |  |  | 0.38168 | 3.945
28 |  | 22-0686 | -8.7939 | -63.6918 |  |  |  | 0.050153 | 0.34941
29 |  | 22-0683 | -9.0153 | -64.9124 |  |  |  | 0.11554 | 0.46803
30 |  | 22-0177 | -4.807 | -42.3024 |  |  |  | 0.089876 | 1.0806
31 |  | EV | -10.2357 | -72.2419 |  |  |  | 0.16049 | 1.8519
32 |  | 22-0179 | -5.2407 | -43.4109 |  |  |  | 0.21804 | 1.8222
33 |  | 22-0197 | 2.9251 | -7.0569 |  |  |  | 0.20242 | 2.0156
34 |  | 22-0178 | -4.8511 | -39.2981 |  |  |  | 0.24405 | 2.0393
35 |  | 22-0172 | -5.875 | -44.7299 |  |  |  | 0.069168 | 0.30893
36 |  | 22-0690 | -9.799 | -72.3269 |  |  |  | 0.13961 | 1.6581
37 | * | 22-0181 | 1.7605 | 1.634 |  |  |  | 5.2233 | 24.0974
38 |  | 22-0202 | -3.8214 | -33.5338 |  |  |  | 0.32497 | 3.4399
39 |  | 22-0193 | -5.5026 | -49.8975 |  |  |  | 0.092174 | 1.254
40 |  | EV | -10.2159 | -72.477 |  |  |  | 0.17223 | 1.4322
41 | * | 22-0186 | -4.8801 | -39.9681 |  |  |  | 1.4061 | 15.6408
42 |  | 22-0682 | -9.5661 | -70.3644 |  |  |  | 0.12953 | 1.3423
43 |  | 22-0174 | -5.0814 | -36.6576 |  |  |  | 0.1981 | 2.3132
44 |  | 22-0195 | -10.4223 | -79.1229 |  |  |  | 0.20949 | 2.659
45 |  | 22-0188 | -6.4441 | -54.0742 |  |  |  | 0.12793 | 1.6809
46 |  | 22-0695 | -10.0134 | -73.7923 |  |  |  | 0.12544 | 1.1166
47 |  | 22-0192 | 2.3244 | -4.1211 |  |  |  | 0.39552 | 4.666
48 |  | 22-0689 | -8.7661 | -61.9271 |  |  |  | 0.28305 | 3.3385
49 |  | EV | -10.2492 | -72.894 |  |  |  | 0.049295 | 0.6091
50 | * | 22-0199 | 2.3277 | -5.9369 |  |  |  | 3.3945 | 8.5405
51 | * | 22-0189 | -6.0628 | -46.7203 |  |  |  | 2.1915 | 20.4315
52 |  | 22-0203 | 10.4983 | 29.666 |  |  |  | 0.55239 | 5.3629
53 |  | 22-0680 | -9.3843 | -67.1196 |  |  |  | 0.60075 | 5.8684
54 |  | 22-0196 | -7.1342 | -57.178 |  |  |  | 0.054485 | 0.65617
55 |  | 22-0194 | -9.6455 | -72.5585 |  |  |  | 0.1129 | 1.2408
56 |  | SCBI4-B-3 | -6.4036 | -52.6883 |  | 636.6517 |  | 0.089463 | 1.1772 | -23.4175 | -21.971
57 | * | SCBI4-B-2 | -4.8979 | -43.1925 |  |  |  | 1.4171 | 17.9312 | -12.5335
58 |  | EV | -10.1959 | -72.473 |  |  |  | 0.10632 | 1.2961
59 |  | SLCP-22-020 | -22.2743 | -165.009 |  |  |  | 0.43209 | 5.478
60 |  | EV | -10.6511 | -78.3261 |  |  |  | 0.32554 | 5.3809

## Detail
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
1 |  | EV |  |  | -2.7702 | -17.9541 |  |  | 2022/12/13 11:19:11
1 |  | EV | -8.54 | -73.804 | -11.324 | -87.2282 | 0.11 | 1405.818 | 2022/12/13 11:29:13
1 |  | EV | -7.996 | -66.476 | -10.7791 | -80.35 | 0.524 | 1398.85 | 2022/12/13 11:38:10
1 |  | EV | -7.81 | -63.642 | -10.5928 | -77.6899 | 0.365 | 1330.216 | 2022/12/13 11:47:07
1 |  | EV | -7.663 | -61.793 | -10.4455 | -75.9544 | 0.443 | 1308.256 | 2022/12/13 11:56:04
1 |  | EV | -7.596 | -60.773 | -10.3784 | -74.997 | 0.722 | 1305.883 | 2022/12/13 12:05:01
1 |  | EV | -7.629 | -60.621 | -10.4115 | -74.8543 | 0.379 | 1305.126 | 2022/12/13 12:13:58
1 |  | EV | -7.592 | -60.485 | -10.3744 | -74.7267 | 0.48 | 1304.33 | 2022/12/13 12:22:54
1 |  | EV | -7.572 | -60.328 | -10.3544 | -74.5793 | 0.366 | 1304.046 | 2022/12/13 12:31:51
1 |  | EV | -7.542 | -60.173 | -10.3244 | -74.4338 | 0.094 | 1304.06 | 2022/12/13 12:40:49
Sample 1 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 | * | EV | -9.7755 | -71.2768 |  |  |  | 2.4761 | 20.4084
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
2 |  | PZ | 4.002 | 27.931 | 1.2382 | 8.2626 | 0.252 | 1301.761 | 2022/12/13 12:50:38
2 |  | PZ | 4.502 | 35.95 | 1.739 | 15.7894 | 0.503 | 1304.158 | 2022/12/13 12:59:34
2 |  | PZ | 4.634 | 38.323 | 1.8712 | 18.0167 | 0.156 | 1303.736 | 2022/12/13 13:08:31
2 |  | PZ | 4.702 | 39.038 | 1.9393 | 18.6879 | 0.455 | 1301.942 | 2022/12/13 13:17:28
2 |  | PZ | 4.734 | 39.387 | 1.9714 | 19.0154 | 0.36 | 1301.732 | 2022/12/13 13:26:26
2 |  | PZ | 4.845 | 40.433 | 2.0825 | 19.9972 | 0.471 | 1304.965 | 2022/12/13 13:35:23
2 |  | PZ | 4.825 | 40.454 | 2.0625 | 20.0169 | 0.434 | 1302.532 | 2022/12/13 13:44:20
2 |  | PZ | 4.881 | 40.7 | 2.1186 | 20.2479 | 0.395 | 1303.933 | 2022/12/13 13:53:17
2 |  | PZ | 4.881 | 40.95 | 2.1186 | 20.4825 | 0.594 | 1303.41 | 2022/12/13 14:02:15
2 |  | PZ | 4.921 | 40.951 | 2.1587 | 20.4834 | 0.527 | 1303.426 | 2022/12/13 14:11:12
Sample 2 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.27536 | 3.995
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
3 |  | UT2 | -12.045 | -92.364 | -14.8346 | -104.649 | 0.66 | 1303.861 | 2022/12/13 14:21:01
3 |  | UT2 | -12.85 | -104.802 | -15.6409 | -116.3236 | 0.633 | 1304.325 | 2022/12/13 14:29:58
3 |  | UT2 | -13.025 | -107.803 | -15.8162 | -119.1404 | 0.698 | 1303.662 | 2022/12/13 14:38:55
3 |  | UT2 | -13.068 | -108.829 | -15.8592 | -120.1034 | 0.524 | 1305.269 | 2022/12/13 14:47:52
3 |  | UT2 | -13.188 | -109.722 | -15.9794 | -120.9416 | 0.721 | 1302.529 | 2022/12/13 14:56:49
3 |  | UT2 | -13.205 | -110.03 | -15.9965 | -121.2307 | 0.703 | 1304.861 | 2022/12/13 15:05:46
3 |  | UT2 | -13.267 | -110.336 | -16.0586 | -121.5179 | 0.775 | 1303.338 | 2022/12/13 15:14:43
3 |  | UT2 | -13.27 | -110.728 | -16.0616 | -121.8859 | 0.247 | 1303.014 | 2022/12/13 15:23:41
3 |  | UT2 | -13.217 | -110.81 | -16.0085 | -121.9629 | 0.579 | 1302.753 | 2022/12/13 15:32:38
3 |  | UT2 | -13.353 | -111.11 | -16.1447 | -122.2444 | 0.616 | 1303.078 | 2022/12/13 15:41:35
Sample 3 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.3814 | 5.694
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
4 |  | EV | -7.83 | -65.169 | -10.6128 | -79.1232 | 0.209 | 1306.229 | 2022/12/13 15:51:25
4 |  | EV | -7.649 | -61.279 | -10.4315 | -75.4719 | 0.39 | 1303.556 | 2022/12/13 16:00:23
4 |  | EV | -7.59 | -60.391 | -10.3724 | -74.6384 | 0.598 | 1304.477 | 2022/12/13 16:09:20
4 |  | EV | -7.575 | -60.146 | -10.3574 | -74.4085 | 0.41 | 1304.719 | 2022/12/13 16:18:17
4 |  | EV | -7.582 | -59.969 | -10.3644 | -74.2423 | 0.954 | 1302.87 | 2022/12/13 16:27:14
4 |  | EV | -7.549 | -59.972 | -10.3314 | -74.2452 | 0.428 | 1304.427 | 2022/12/13 16:36:12
4 |  | EV | -7.567 | -59.801 | -10.3494 | -74.0847 | 0.498 | 1304.429 | 2022/12/13 16:45:09
4 |  | EV | -7.588 | -60.029 | -10.3704 | -74.2987 | 0.673 | 1302.921 | 2022/12/13 16:54:07
4 |  | EV | -7.54 | -59.642 | -10.3224 | -73.9354 | 0.61 | 1304.213 | 2022/12/13 17:03:04
4 |  | EV | -7.535 | -59.547 | -10.3173 | -73.8462 | 0.605 | 1303.175 | 2022/12/13 17:12:02
Sample 4 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
4 |  | EV | -10.3829 | -74.8295 |  |  |  | 0.086927 | 1.6789
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
5 |  | 22-0684 | -7.132 | -55.458 | -9.9137 | -70.0082 | 0.094 | 1307.487 | 2022/12/13 17:21:52
5 |  | 22-0684 | -7.205 | -55.637 | -9.9868 | -70.1762 | 0.508 | 1301.111 | 2022/12/13 17:30:49
5 |  | 22-0684 | -7.164 | -55.339 | -9.9457 | -69.8965 | 0.237 | 1300.273 | 2022/12/13 17:39:46
5 |  | 22-0684 | -7.226 | -55.742 | -10.0078 | -70.2748 | 0.518 | 1300.058 | 2022/12/13 17:48:42
Sample 5 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
5 |  | 22-0684 | -9.9635 | -70.0889 |  |  |  | 0.041987 | 0.18007
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
6 |  | 22-0685 | -6.24 | -50.322 | -9.0203 | -65.1875 | 0.1 | 1308.241 | 2022/12/13 17:58:33
6 |  | 22-0685 | -6.323 | -50.493 | -9.1034 | -65.348 | 0.436 | 1303.661 | 2022/12/13 18:07:30
6 |  | 22-0685 | -6.247 | -50.183 | -9.0273 | -65.057 | 0.412 | 1303.846 | 2022/12/13 18:16:28
6 |  | 22-0685 | -6.281 | -50.258 | -9.0613 | -65.1274 | 0.402 | 1302.874 | 2022/12/13 18:25:25
Sample 6 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
6 |  | 22-0685 | -9.0531 | -65.1799 |  |  |  | 0.037986 | 0.13216
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
7 |  | 22-0679 | -6.102 | -51.369 | -8.882 | -66.1702 | 0.397 | 1298.391 | 2022/12/13 18:35:15
7 |  | 22-0679 | -6.162 | -51.867 | -8.9421 | -66.6376 | 0.415 | 1296.662 | 2022/12/13 18:44:12
7 |  | 22-0679 | -6.132 | -51.62 | -8.9121 | -66.4058 | 0.456 | 1297.715 | 2022/12/13 18:53:09
7 |  | 22-0679 | -6.092 | -51.632 | -8.872 | -66.417 | 0.325 | 1297.967 | 2022/12/13 19:02:06
Sample 7 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
7 |  | 22-0679 | -8.9021 | -66.4077 |  |  |  | 0.031623 | 0.20342
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
8 |  | 22-0693 | -7.8 | -58.579 | -10.5828 | -72.9377 | 0.249 | 1305.632 | 2022/12/13 19:11:53
8 |  | 22-0693 | -7.98 | -59.743 | -10.7631 | -74.0302 | 0.143 | 1302.875 | 2022/12/13 19:20:50
8 |  | 22-0693 | -7.994 | -60.325 | -10.7771 | -74.5765 | 0.299 | 1300.36 | 2022/12/13 19:29:47
8 |  | 22-0693 | -8.023 | -60.63 | -10.8061 | -74.8628 | 0.823 | 1300.823 | 2022/12/13 19:38:44
Sample 8 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
8 |  | 22-0693 | -10.7323 | -74.1018 |  |  |  | 0.1011 | 0.90501
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
9 |  | 22-0678 | -6.684 | -49.748 | -9.465 | -64.6487 | 0.582 | 1284.684 | 2022/12/13 19:48:32
9 |  | 22-0678 | -6.496 | -48.403 | -9.2767 | -63.3862 | 0.521 | 1287.236 | 2022/12/13 19:57:29
9 |  | 22-0678 | -6.532 | -48.199 | -9.3127 | -63.1948 | 0.41 | 1288.285 | 2022/12/13 20:06:27
9 |  | 22-0678 | -6.433 | -47.334 | -9.2136 | -62.3828 | 0.27 | 1293.205 | 2022/12/13 20:15:24
Sample 9 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
9 |  | 22-0678 | -9.317 | -63.4031 |  |  |  | 0.10666 | 0.99868
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
10 |  | 22-0191 | -2.761 | -28.095 | -5.5357 | -44.3247 | 0.411 | 1298.315 | 2022/12/13 20:25:13
10 |  | 22-0191 | -2.442 | -25.908 | -5.2162 | -42.2719 | 0.483 | 1298.5 | 2022/12/13 20:34:09
10 |  | 22-0191 | -2.506 | -25.724 | -5.2803 | -42.0992 | 0.725 | 1295.596 | 2022/12/13 20:43:06
10 |  | 22-0191 | -2.496 | -25.439 | -5.2702 | -41.8317 | 0.242 | 1296.947 | 2022/12/13 20:52:04
Sample 10 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
10 |  | 22-0191 | -5.3256 | -42.6319 |  |  |  | 0.14263 | 1.2177
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
11 |  | 22-0691 | -6.612 | -50.686 | -9.3929 | -65.5291 | 0.569 | 1287.144 | 2022/12/13 21:01:53
11 |  | 22-0691 | -6.527 | -51.862 | -9.3077 | -66.6329 | 0.406 | 1296.915 | 2022/12/13 21:10:50
11 |  | 22-0691 | -6.738 | -53.072 | -9.5191 | -67.7687 | 0.244 | 1291.969 | 2022/12/13 21:19:47
11 |  | 22-0691 | -6.754 | -52.641 | -9.5351 | -67.3641 | 0.504 | 1297.051 | 2022/12/13 21:28:44
Sample 11 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
11 |  | 22-0691 | -9.4387 | -66.8237 |  |  |  | 0.10785 | 1.047
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
12 |  | 22-0185 | -4.542 | -43.252 | -7.3195 | -58.5514 | 0.138 | 1270.065 | 2022/12/13 21:38:34
12 |  | 22-0185 | -4.002 | -39.884 | -6.7787 | -55.3901 | 0.383 | 1288.323 | 2022/12/13 21:47:31
12 |  | 22-0185 | -3.986 | -39.147 | -6.7626 | -54.6983 | 0.453 | 1288.43 | 2022/12/13 21:56:29
12 |  | 22-0185 | -3.89 | -38.881 | -6.6665 | -54.4487 | 0.218 | 1294.489 | 2022/12/13 22:05:25
Sample 12 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
12 |  | 22-0185 | -6.8818 | -55.7721 |  |  |  | 0.2955 | 2.0191
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
13 |  | EV | -7.201 | -56.688 | -9.9828 | -71.1627 | -0.075 | 1301.214 | 2022/12/13 22:15:16
13 |  | EV | -7.38 | -58.454 | -10.1621 | -72.8203 | 0.507 | 1300.128 | 2022/12/13 22:24:12
13 |  | EV | -7.457 | -58.619 | -10.2392 | -72.9752 | 0.412 | 1299.488 | 2022/12/13 22:33:09
13 |  | EV | -7.507 | -59.154 | -10.2893 | -73.4774 | 0.566 | 1299.675 | 2022/12/13 22:42:05
Sample 13 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
13 |  | EV | -10.1684 | -72.6089 |  |  |  | 0.13409 | 1.0697
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
14 |  | 22-0201 |  |  | -2.7702 | -17.9541 |  |  | 2022/12/13 22:51:55
14 |  | 22-0201 | 14.907 | 63.694 | 12.1607 | 41.8305 | 0.36 | 1284.884 | 2022/12/13 23:00:53
14 |  | 22-0201 | 15.027 | 64.027 | 12.2809 | 42.1431 | 0.931 | 1294.674 | 2022/12/13 23:09:50
14 |  | 22-0201 | 14.907 | 63.694 | 12.1607 | 41.8305 | 0.36 | 1284.884 | 2022/12/13 23:18:47
Sample 14 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
14 | * | 22-0201 | 8.458 | 26.9625 |  |  |  | 7.4737 | 31.9029
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
15 |  | 22-0688 | -5.504 | -42.282 | -8.2831 | -57.6409 | 0.196 | 1286.694 | 2022/12/13 23:28:34
15 |  | 22-0688 | -6.254 | -51.614 | -9.0343 | -66.4002 | 0.46 | 1299.596 | 2022/12/13 23:37:30
15 |  | 22-0688 | -6.731 | -55.091 | -9.5121 | -69.6637 | 0.572 | 1287.764 | 2022/12/13 23:46:27
15 |  | 22-0688 | -6.813 | -56.007 | -9.5942 | -70.5235 | 0.134 | 1286.495 | 2022/12/13 23:55:24
Sample 15 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
15 |  | 22-0688 | -9.1059 | -66.0571 |  |  |  | 0.60057 | 6.27
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
16 |  | 22-0184 | -4.358 | -50.174 | -7.1352 | -65.0485 | 0.633 | 1286.438 | 2022/12/14 00:05:12
16 |  | 22-0184 | -3.969 | -48.663 | -6.7456 | -63.6303 | 0.717 | 1294.389 | 2022/12/14 00:14:09
16 |  | 22-0184 | -4.181 | -49.207 | -6.958 | -64.1409 | 0.332 | 1288.167 | 2022/12/14 00:23:06
16 |  | 22-0184 | -4.282 | -49.603 | -7.0591 | -64.5126 | 0.321 | 1286.331 | 2022/12/14 00:32:03
Sample 16 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
16 |  | 22-0184 | -6.9745 | -64.3331 |  |  |  | 0.16871 | 0.63774
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
17 |  | 22-0173 | -3.312 | -33.383 | -6.0876 | -49.2881 | 0.459 | 1289.447 | 2022/12/14 00:41:51
17 |  | 22-0173 | -3.24 | -31.682 | -6.0154 | -47.6915 | 0.277 | 1291.149 | 2022/12/14 00:50:47
17 |  | 22-0173 | -3.222 | -31.168 | -5.9974 | -47.2091 | 0.397 | 1291.149 | 2022/12/14 00:59:44
17 |  | 22-0173 | -3.229 | -31.453 | -6.0044 | -47.4766 | 0.118 | 1292.124 | 2022/12/14 01:08:40
Sample 17 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
17 |  | 22-0173 | -6.0262 | -47.9163 |  |  |  | 0.0415 | 0.99676
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
18 |  | 22-0187 | -2.454 | -31.79 | -5.2282 | -47.7929 | 0.113 | 1296.393 | 2022/12/14 01:18:29
18 |  | 22-0187 | -2.207 | -31.349 | -4.9808 | -47.379 | 0.506 | 1301.035 | 2022/12/14 01:27:26
18 |  | 22-0187 | -2.399 | -31.87 | -5.1731 | -47.868 | 0.437 | 1298.004 | 2022/12/14 01:36:23
18 |  | 22-0187 | -2.392 | -31.525 | -5.1661 | -47.5442 | 0.393 | 1298.996 | 2022/12/14 01:45:20
Sample 18 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
18 |  | 22-0187 | -5.137 | -47.646 |  |  |  | 0.10763 | 0.24023
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
19 |  | 22-0694 | -6.831 | -55.566 | -9.6122 | -70.1096 | 0.173 | 1282.159 | 2022/12/14 01:55:10
19 |  | 22-0694 | -6.867 | -57.028 | -9.6483 | -71.4819 | 0.509 | 1288.465 | 2022/12/14 02:04:07
19 |  | 22-0694 | -7.064 | -58.21 | -9.8456 | -72.5913 | 0.174 | 1282.574 | 2022/12/14 02:13:04
19 |  | 22-0694 | -7.092 | -58.409 | -9.8736 | -72.7781 | 0.59 | 1283.408 | 2022/12/14 02:22:01
Sample 19 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
19 |  | 22-0694 | -9.7449 | -71.7402 |  |  |  | 0.13352 | 1.3088
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
20 |  | 22-0692 | -6.721 | -49.391 | -9.502 | -64.3136 | 0.47 | 1283.051 | 2022/12/14 02:31:50
20 |  | 22-0692 | -6.503 | -47.574 | -9.2837 | -62.6081 | 0.28 | 1287.724 | 2022/12/14 02:40:47
20 |  | 22-0692 | -6.669 | -48.141 | -9.45 | -63.1403 | 0.575 | 1282.124 | 2022/12/14 02:49:45
20 |  | 22-0692 | -6.692 | -47.82 | -9.473 | -62.839 | 0.436 | 1282.727 | 2022/12/14 02:58:42
Sample 20 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
20 |  | 22-0692 | -9.4272 | -63.2253 |  |  |  | 0.097841 | 0.80711
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
21 |  | 22-0190 | 2.412 | -11.651 | -0.35436 | -28.89 | 0.416 | 1292.937 | 2022/12/14 03:08:31
21 |  | 22-0190 | 3.042 | -7.642 | 0.27665 | -25.127 | 0.532 | 1293.019 | 2022/12/14 03:17:28
21 |  | 22-0190 | 2.92 | -8.016 | 0.15445 | -25.4781 | 0.448 | 1285.178 | 2022/12/14 03:26:25
21 |  | 22-0190 | 3.023 | -7.416 | 0.25762 | -24.9149 | 0.606 | 1285.304 | 2022/12/14 03:35:22
Sample 21 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
21 |  | 22-0190 | 0.083589 | -26.1025 |  |  |  | 0.29639 | 1.9952
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
22 |  | EV | -6.835 | -53.081 | -9.6162 | -67.7771 | 0.319 | 1299.422 | 2022/12/14 03:45:08
22 |  | EV | -7.273 | -57.245 | -10.0549 | -71.6855 | 0.335 | 1299.64 | 2022/12/14 03:54:05
22 |  | EV | -7.39 | -58.231 | -10.1721 | -72.611 | 0.399 | 1299.894 | 2022/12/14 04:03:02
22 |  | EV | -7.431 | -58.506 | -10.2132 | -72.8691 | 0.499 | 1298.977 | 2022/12/14 04:11:59
Sample 22 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
22 |  | EV | -10.0141 | -71.2357 |  |  |  | 0.27316 | 2.5155
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
23 |  | 22-0200 | 11.277 | 38.442 | 8.5249 | 18.1284 | 0.507 | 1292.844 | 2022/12/14 04:21:46
23 |  | 22-0200 | 12.322 | 48.366 | 9.5716 | 27.4433 | 0.32 | 1291.68 | 2022/12/14 04:30:43
23 |  | 22-0200 | 12.42 | 49.9 | 9.6697 | 28.8832 | 0.243 | 1289.822 | 2022/12/14 04:39:40
23 |  | 22-0200 | 12.475 | 51.016 | 9.7248 | 29.9307 | 0.35 | 1290.979 | 2022/12/14 04:48:37
Sample 23 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
23 |  | 22-0200 | 9.3727 | 26.0964 |  |  |  | 0.56787 | 5.7627
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
24 |  | 22-0175 |  |  | -2.7702 | -17.9541 |  |  | 2022/12/14 04:58:24
24 |  | 22-0175 | -2.027 | -19.615 | -4.8005 | -36.3652 | 0.235 | 1303.154 | 2022/12/14 05:07:21
24 |  | 22-0175 | -2.414 | -21.897 | -5.1881 | -38.5071 | 0.509 | 1290.541 | 2022/12/14 05:16:18
24 |  | 22-0175 | -2.449 | -22.828 | -5.2232 | -39.381 | 0.579 | 1291.056 | 2022/12/14 05:25:14
Sample 24 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
24 | * | 22-0175 | -4.4955 | -33.0518 |  |  |  | 1.1641 | 10.808
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
25 |  | 22-0180 | -1.063 | -29.011 | -3.8349 | -45.1845 | 0.376 | 1285.187 | 2022/12/14 05:35:02
25 |  | 22-0180 |  |  | -2.7702 | -17.9541 |  |  | 2022/12/14 05:43:59
25 |  | 22-0180 | -0.988 | -29.654 | -3.7598 | -45.788 | 0.174 | 1282.227 | 2022/12/14 05:52:55
25 |  | 22-0180 | -1.053 | -30.055 | -3.8249 | -46.1644 | 0.593 | 1285.321 | 2022/12/14 06:01:52
Sample 25 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
25 | * | 22-0180 | -3.5475 | -38.7727 |  |  |  | 0.5184 | 14.7929
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
26 |  | 22-0198 | 5.553 | 7.71 | 2.7917 | -10.7173 | 0.478 | 1288.976 | 2022/12/14 06:11:40
26 |  | 22-0198 | 6.411 | 14.018 | 3.6511 | -4.7965 | 0.857 | 1303.183 | 2022/12/14 06:20:36
26 |  | 22-0198 | 6.002 | 12.809 | 3.2414 | -5.9313 | 0.442 | 1293.583 | 2022/12/14 06:29:33
26 |  | 22-0198 | 5.993 | 12.896 | 3.2324 | -5.8496 | 0.501 | 1292.871 | 2022/12/14 06:38:30
Sample 26 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
26 |  | 22-0198 | 3.2291 | -6.8237 |  |  |  | 0.35041 | 2.8198
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
27 |  | 22-0681 | -5.955 | -47.445 | -8.7348 | -62.487 | 0.262 | 1289.823 | 2022/12/14 06:48:19
27 |  | 22-0681 | -6.365 | -53.137 | -9.1455 | -67.8297 | 0.594 | 1295.25 | 2022/12/14 06:57:16
27 |  | 22-0681 | -6.708 | -55.39 | -9.489 | -69.9444 | 0.28 | 1282.7 | 2022/12/14 07:06:12
27 |  | 22-0681 | -6.794 | -56.179 | -9.5752 | -70.685 | 0.29 | 1282.644 | 2022/12/14 07:15:10
Sample 27 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
27 |  | 22-0681 | -9.2361 | -67.7365 |  |  |  | 0.38168 | 3.945
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
28 |  | 22-0686 | -6.022 | -48.814 | -8.8019 | -63.772 | 0.513 | 1289.458 | 2022/12/14 07:24:59
28 |  | 22-0686 | -5.941 | -48.251 | -8.7208 | -63.2436 | 0.289 | 1285.774 | 2022/12/14 07:33:56
28 |  | 22-0686 | -6.051 | -48.76 | -8.831 | -63.7213 | 0.276 | 1279.26 | 2022/12/14 07:42:54
28 |  | 22-0686 | -6.042 | -49.089 | -8.8219 | -64.0301 | 0.243 | 1279.649 | 2022/12/14 07:51:51
Sample 28 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
28 |  | 22-0686 | -8.7939 | -63.6918 |  |  |  | 0.050153 | 0.34941
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
29 |  | 22-0683 | -6.286 | -50.408 | -9.0663 | -65.2682 | 0.125 | 1278.46 | 2022/12/14 08:01:35
29 |  | 22-0683 | -6.067 | -49.357 | -8.847 | -64.2817 | 0.341 | 1286.367 | 2022/12/14 08:10:32
29 |  | 22-0683 | -6.328 | -50.273 | -9.1084 | -65.1415 | 0.201 | 1280.403 | 2022/12/14 08:19:30
29 |  | 22-0683 | -6.259 | -50.078 | -9.0393 | -64.9584 | 0.589 | 1280.547 | 2022/12/14 08:28:26
Sample 29 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
29 |  | 22-0683 | -9.0153 | -64.9124 |  |  |  | 0.11554 | 0.46803
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
30 |  | 22-0177 | -2.166 | -27.532 | -4.9397 | -43.7962 | 0.336 | 1296.059 | 2022/12/14 08:38:14
30 |  | 22-0177 | -1.999 | -25.691 | -4.7724 | -42.0682 | 0.294 | 1295.693 | 2022/12/14 08:47:10
30 |  | 22-0177 | -2.003 | -25.33 | -4.7765 | -41.7294 | 0.245 | 1294.09 | 2022/12/14 08:56:06
30 |  | 22-0177 | -1.966 | -25.209 | -4.7394 | -41.6158 | 0.324 | 1294.009 | 2022/12/14 09:05:03
Sample 30 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
30 |  | 22-0177 | -4.807 | -42.3024 |  |  |  | 0.089876 | 1.0806
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
31 |  | EV | -7.217 | -55.14 | -9.9988 | -69.7097 | 0.633 | 1294.735 | 2022/12/14 09:14:49
31 |  | EV | -7.491 | -58.174 | -10.2733 | -72.5575 | 0.226 | 1295.856 | 2022/12/14 09:23:46
31 |  | EV | -7.545 | -58.785 | -10.3274 | -73.131 | 0.231 | 1295.337 | 2022/12/14 09:32:42
31 |  | EV | -7.561 | -59.252 | -10.3434 | -73.5694 | 0.522 | 1296.322 | 2022/12/14 09:41:39
Sample 31 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
31 |  | EV | -10.2357 | -72.2419 |  |  |  | 0.16049 | 1.8519
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
32 |  | 22-0179 | -2.758 | -29.838 | -5.5327 | -45.9607 | 0.271 | 1279.71 | 2022/12/14 09:51:25
32 |  | 22-0179 | -2.251 | -26.156 | -5.0249 | -42.5047 | 0.378 | 1281.518 | 2022/12/14 10:00:23
32 |  | 22-0179 | -2.493 | -26.487 | -5.2672 | -42.8154 | 0.124 | 1271.765 | 2022/12/14 10:09:19
32 |  | 22-0179 | -2.364 | -26.005 | -5.138 | -42.363 | 0.166 | 1275.938 | 2022/12/14 10:18:17
Sample 32 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
32 |  | 22-0179 | -5.2407 | -43.4109 |  |  |  | 0.21804 | 1.8222
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
33 |  | 22-0197 | 5.384 | 8.649 | 2.6224 | -9.8359 | 0.049 | 1289.293 | 2022/12/14 10:28:04
33 |  | 22-0197 | 5.77 | 12.038 | 3.009 | -6.655 | 0.226 | 1290.925 | 2022/12/14 10:37:00
33 |  | 22-0197 | 5.777 | 12.751 | 3.016 | -5.9857 | 0.35 | 1291.273 | 2022/12/14 10:45:57
33 |  | 22-0197 | 5.814 | 13.001 | 3.0531 | -5.7511 | 0.58 | 1290.964 | 2022/12/14 10:54:54
Sample 33 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
33 |  | 22-0197 | 2.9251 | -7.0569 |  |  |  | 0.20242 | 2.0156
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
34 |  | 22-0178 | -1.735 | -19.798 | -4.508 | -36.5369 | 0.35 | 1287.558 | 2022/12/14 11:04:41
34 |  | 22-0178 | -2.073 | -22.93 | -4.8466 | -39.4767 | 0.441 | 1287.695 | 2022/12/14 11:13:39
34 |  | 22-0178 | -2.227 | -24.09 | -5.0008 | -40.5655 | 0.327 | 1286.867 | 2022/12/14 11:22:35
34 |  | 22-0178 | -2.275 | -24.141 | -5.0489 | -40.6134 | 0.299 | 1288.022 | 2022/12/14 11:31:31
Sample 34 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
34 |  | 22-0178 | -4.8511 | -39.2981 |  |  |  | 0.24405 | 2.0393
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
35 |  | 22-0172 | -3 | -28.089 | -5.7751 | -44.3191 | 0.365 | 1294.337 | 2022/12/14 11:41:19
35 |  | 22-0172 | -3.159 | -28.666 | -5.9343 | -44.8606 | 0.696 | 1294.385 | 2022/12/14 11:50:17
35 |  | 22-0172 | -3.126 | -28.552 | -5.9013 | -44.7536 | 0.438 | 1293.698 | 2022/12/14 11:59:13
35 |  | 22-0172 | -3.114 | -28.8 | -5.8892 | -44.9864 | 0.187 | 1292.502 | 2022/12/14 12:08:11
Sample 35 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
35 |  | 22-0172 | -5.875 | -44.7299 |  |  |  | 0.069168 | 0.30893
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
36 |  | 22-0690 | -6.892 | -55.672 | -9.6733 | -70.2091 | 0.183 | 1281.724 | 2022/12/14 12:17:55
36 |  | 22-0690 | -6.902 | -57.699 | -9.6833 | -72.1117 | 0.239 | 1286.479 | 2022/12/14 12:26:51
36 |  | 22-0690 | -7.125 | -59.082 | -9.9067 | -73.4098 | 0.45 | 1278.922 | 2022/12/14 12:35:48
36 |  | 22-0690 | -7.151 | -59.26 | -9.9327 | -73.5769 | 0.424 | 1278.006 | 2022/12/14 12:44:46
Sample 36 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
36 |  | 22-0690 | -9.799 | -72.3269 |  |  |  | 0.13961 | 1.6581
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
37 |  | 22-0181 |  |  | -2.7702 | -17.9541 |  |  | 2022/12/14 12:54:30
37 |  | 22-0181 |  |  | -2.7702 | -17.9541 |  |  | 2022/12/14 13:03:26
37 |  | 22-0181 | 9.047 | 41.738 | 6.2913 | 21.2221 | 0.449 | 1274.773 | 2022/12/14 13:12:23
37 |  | 22-0181 | 9.047 | 41.738 | 6.2913 | 21.2221 | 0.449 | 1274.773 | 2022/12/14 13:21:20
Sample 37 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
37 | * | 22-0181 | 1.7605 | 1.634 |  |  |  | 5.2233 | 24.0974
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
38 |  | 22-0202 | -0.596 | -11.692 | -3.3672 | -28.9285 | 0.108 | 1271.705 | 2022/12/14 13:31:05
38 |  | 22-0202 | -1.043 | -16.765 | -3.8149 | -33.6901 | 0.302 | 1279.858 | 2022/12/14 13:40:02
38 |  | 22-0202 | -1.228 | -18.683 | -4.0002 | -35.4904 | 0.056 | 1273.778 | 2022/12/14 13:48:57
38 |  | 22-0202 | -1.331 | -19.254 | -4.1034 | -36.0263 | 0.074 | 1274.614 | 2022/12/14 13:57:54
Sample 38 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
38 |  | 22-0202 | -3.8214 | -33.5338 |  |  |  | 0.32497 | 3.4399
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
39 |  | 22-0193 | -2.598 | -32.291 | -5.3724 | -48.2632 | 0.609 | 1298.455 | 2022/12/14 14:07:40
39 |  | 22-0193 | -2.73 | -33.945 | -5.5046 | -49.8156 | 0.429 | 1298.053 | 2022/12/14 14:16:38
39 |  | 22-0193 | -2.778 | -34.89 | -5.5527 | -50.7026 | 0.186 | 1296.022 | 2022/12/14 14:25:35
39 |  | 22-0193 | -2.806 | -35.003 | -5.5807 | -50.8087 | 0.447 | 1294.972 | 2022/12/14 14:34:32
Sample 39 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
39 |  | 22-0193 | -5.5026 | -49.8975 |  |  |  | 0.092174 | 1.254
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
40 |  | EV | -7.182 | -56.013 | -9.9638 | -70.5292 | 0.257 | 1294.462 | 2022/12/14 14:44:17
40 |  | EV | -7.485 | -58.389 | -10.2673 | -72.7593 | 0.342 | 1293.412 | 2022/12/14 14:53:13
40 |  | EV | -7.496 | -58.674 | -10.2783 | -73.0268 | 0.241 | 1293.753 | 2022/12/14 15:02:10
40 |  | EV | -7.572 | -59.277 | -10.3544 | -73.5928 | 0.309 | 1293.819 | 2022/12/14 15:11:07
Sample 40 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
40 |  | EV | -10.2159 | -72.477 |  |  |  | 0.17223 | 1.4322
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
41 |  | 22-0186 | -2.759 | -30.987 | -5.5337 | -47.0392 | 0.321 | 1285.309 | 2022/12/14 15:20:53
41 |  | 22-0186 |  |  | -2.7702 | -17.9541 |  |  | 2022/12/14 15:29:50
41 |  | 22-0186 | -2.759 | -30.987 | -5.5337 | -47.0392 | 0.321 | 1285.309 | 2022/12/14 15:38:46
41 |  | 22-0186 | -2.908 | -31.84 | -5.6829 | -47.8398 | 0.619 | 1277.297 | 2022/12/14 15:47:44
Sample 41 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
41 | * | 22-0186 | -4.8801 | -39.9681 |  |  |  | 1.4061 | 15.6408
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
42 |  | 22-0682 | -6.615 | -53.961 | -9.3959 | -68.6031 | 0.443 | 1287.175 | 2022/12/14 15:57:29
42 |  | 22-0682 | -6.758 | -55.784 | -9.5391 | -70.3142 | 0.41 | 1287.448 | 2022/12/14 16:06:26
42 |  | 22-0682 | -6.858 | -56.705 | -9.6393 | -71.1787 | 0.426 | 1287.978 | 2022/12/14 16:15:24
42 |  | 22-0682 | -6.909 | -56.9 | -9.6903 | -71.3617 | 0.411 | 1287.251 | 2022/12/14 16:24:20
Sample 42 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
42 |  | 22-0682 | -9.5661 | -70.3644 |  |  |  | 0.12953 | 1.3423
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
43 |  | 22-0174 | -2.593 | -23.386 | -5.3674 | -39.9047 | 0.528 | 1286.097 | 2022/12/14 16:34:05
43 |  | 22-0174 | -2.142 | -18.81 | -4.9157 | -35.6096 | 0.324 | 1292.65 | 2022/12/14 16:43:02
43 |  | 22-0174 | -2.219 | -18.539 | -4.9928 | -35.3552 | 0.39 | 1291.556 | 2022/12/14 16:51:59
43 |  | 22-0174 | -2.276 | -18.971 | -5.0499 | -35.7607 | 0.21 | 1290.693 | 2022/12/14 17:00:55
Sample 43 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
43 |  | 22-0174 | -5.0814 | -36.6576 |  |  |  | 0.1981 | 2.3132
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
44 |  | 22-0195 | -7.469 | -61.536 | -10.2512 | -75.7132 | 0.267 | 1287.4 | 2022/12/14 17:10:41
44 |  | 22-0195 | -7.448 | -64.846 | -10.2302 | -78.82 | 0.6 | 1294.356 | 2022/12/14 17:19:39
44 |  | 22-0195 | -7.817 | -66.939 | -10.5998 | -80.7845 | 0.079 | 1284.472 | 2022/12/14 17:28:36
44 |  | 22-0195 | -7.825 | -67.354 | -10.6078 | -81.1741 | 0.237 | 1286.236 | 2022/12/14 17:37:32
Sample 44 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
44 |  | 22-0195 | -10.4223 | -79.1229 |  |  |  | 0.20949 | 2.659
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
45 |  | 22-0188 | -3.853 | -40.975 | -6.6294 | -56.4141 | 0.616 | 1297.622 | 2022/12/15 11:30:58
45 |  | 22-0188 | -3.577 | -37.955 | -6.353 | -53.5795 | 0.212 | 1299.157 | 2022/12/15 11:39:54
45 |  | 22-0188 | -3.654 | -37.659 | -6.4301 | -53.3017 | 0.3 | 1297.112 | 2022/12/15 11:48:51
45 |  | 22-0188 | -3.588 | -37.339 | -6.364 | -53.0013 | 0.537 | 1296.195 | 2022/12/15 11:57:47
Sample 45 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
45 |  | 22-0188 | -6.4441 | -54.0742 |  |  |  | 0.12793 | 1.6809
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
46 |  | 22-0695 | -7.111 | -58.013 | -9.8927 | -72.4064 | 0.37 | 1293.523 | 2022/12/15 12:07:33
46 |  | 22-0695 | -7.154 | -59.354 | -9.9357 | -73.6651 | 0.526 | 1298.084 | 2022/12/15 12:16:29
46 |  | 22-0695 | -7.271 | -59.947 | -10.0529 | -74.2217 | 0.683 | 1296.615 | 2022/12/15 12:25:26
46 |  | 22-0695 | -7.39 | -60.644 | -10.1721 | -74.8759 | 0.425 | 1295.833 | 2022/12/15 12:34:22
Sample 46 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
46 |  | 22-0695 | -10.0134 | -73.7923 |  |  |  | 0.12544 | 1.1166
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
47 |  | 22-0192 | 4.495 | 7.861 | 1.732 | -10.5756 | 0.343 | 1289.661 | 2022/12/15 12:44:08
47 |  | 22-0192 | 5.263 | 15.828 | 2.5012 | -3.0976 | 0.78 | 1292.488 | 2022/12/15 12:53:03
47 |  | 22-0192 | 5.261 | 17.406 | 2.4992 | -1.6164 | 0.289 | 1294.346 | 2022/12/15 13:01:59
47 |  | 22-0192 | 5.327 | 17.855 | 2.5653 | -1.195 | 0.193 | 1294.09 | 2022/12/15 13:10:55
Sample 47 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
47 |  | 22-0192 | 2.3244 | -4.1211 |  |  |  | 0.39552 | 4.666
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
48 |  | 22-0689 | -5.594 | -41.996 | -8.3732 | -57.3725 | 0.338 | 1283.553 | 2022/12/15 13:20:42
48 |  | 22-0689 | -5.964 | -47.317 | -8.7438 | -62.3669 | 0.634 | 1290.524 | 2022/12/15 13:29:39
48 |  | 22-0689 | -6.187 | -48.866 | -8.9672 | -63.8208 | 0.079 | 1293.863 | 2022/12/15 13:38:35
48 |  | 22-0689 | -6.2 | -49.215 | -8.9802 | -64.1484 | 0.124 | 1295.904 | 2022/12/15 13:47:32
Sample 48 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
48 |  | 22-0689 | -8.7661 | -61.9271 |  |  |  | 0.28305 | 3.3385
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
49 |  | EV | -7.403 | -57.745 | -10.1851 | -72.1548 | 0.279 | 1297.031 | 2022/12/15 13:57:18
49 |  | EV | -7.454 | -58.502 | -10.2362 | -72.8654 | 0.472 | 1299.435 | 2022/12/15 14:06:14
49 |  | EV | -7.5 | -58.661 | -10.2823 | -73.0146 | 0.545 | 1297.353 | 2022/12/15 14:15:11
49 |  | EV | -7.511 | -59.222 | -10.2933 | -73.5412 | 0.106 | 1298.184 | 2022/12/15 14:24:07
Sample 49 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
49 |  | EV | -10.2492 | -72.894 |  |  |  | 0.049295 | 0.6091
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
50 |  | 22-0199 |  |  | -2.7702 | -17.9541 |  |  | 2022/12/15 14:33:51
50 |  | 22-0199 | 6.919 | 16.727 | 4.1599 | -2.2537 | 0.364 | 1301.242 | 2022/12/15 14:42:47
50 |  | 22-0199 | 6.724 | 17.034 | 3.9646 | -1.9656 | 0.383 | 1293.108 | 2022/12/15 14:51:42
50 |  | 22-0199 | 6.716 | 17.451 | 3.9566 | -1.5742 | 0.375 | 1293.445 | 2022/12/15 15:00:38
Sample 50 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
50 | * | 22-0199 | 2.3277 | -5.9369 |  |  |  | 3.3945 | 8.5405
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
51 |  | 22-0189 |  |  | -2.7702 | -17.9541 |  |  | 2022/12/15 15:10:24
51 |  | 22-0189 | -4.383 | -40.863 | -7.1603 | -56.309 | 0.205 | 1280.037 | 2022/12/15 15:19:20
51 |  | 22-0189 | -4.383 | -40.863 | -7.1603 | -56.309 | 0.205 | 1280.037 | 2022/12/15 15:28:15
51 |  | 22-0189 | -4.383 | -40.863 | -7.1603 | -56.309 | 0.205 | 1280.037 | 2022/12/15 15:37:12
Sample 51 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
51 | * | 22-0189 | -6.0628 | -46.7203 |  |  |  | 2.1915 | 20.4315
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
52 |  | 22-0203 | 12.429 | 42.888 | 9.6787 | 22.3016 | 0.557 | 1291.079 | 2022/12/15 15:46:58
52 |  | 22-0203 | 13.461 | 51.812 | 10.7124 | 30.6778 | 0.367 | 1293.164 | 2022/12/15 15:55:53
52 |  | 22-0203 | 13.456 | 53.544 | 10.7074 | 32.3035 | 0.269 | 1292.156 | 2022/12/15 16:04:49
52 |  | 22-0203 | 13.643 | 54.692 | 10.8947 | 33.3811 | 0.671 | 1292.631 | 2022/12/15 16:13:45
Sample 52 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
52 |  | 22-0203 | 10.4983 | 29.666 |  |  |  | 0.55239 | 5.3629
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
53 |  | 22-0680 | -5.75 | -43.865 | -8.5295 | -59.1268 | 0.172 | 1289.896 | 2022/12/15 16:23:31
53 |  | 22-0680 | -6.629 | -53.188 | -9.4099 | -67.8775 | 0.338 | 1293.585 | 2022/12/15 16:32:28
53 |  | 22-0680 | -6.944 | -55.764 | -9.7254 | -70.2954 | 0.476 | 1291.709 | 2022/12/15 16:41:23
53 |  | 22-0680 | -7.091 | -56.705 | -9.8726 | -71.1787 | 0.411 | 1290.72 | 2022/12/15 16:50:19
Sample 53 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
53 |  | 22-0680 | -9.3843 | -67.1196 |  |  |  | 0.60075 | 5.8684
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
54 |  | 22-0196 | -4.437 | -42.758 | -7.2144 | -58.0877 | 0.472 | 1293.239 | 2022/12/15 17:00:06
54 |  | 22-0196 | -4.321 | -41.596 | -7.0982 | -56.997 | 0.481 | 1292.311 | 2022/12/15 17:09:02
54 |  | 22-0196 | -4.324 | -41.318 | -7.1012 | -56.7361 | 0.514 | 1293.637 | 2022/12/15 17:17:57
54 |  | 22-0196 | -4.346 | -41.483 | -7.1232 | -56.891 | 0.564 | 1293.673 | 2022/12/15 17:26:53
Sample 54 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
54 |  | 22-0196 | -7.1342 | -57.178 |  |  |  | 0.054485 | 0.65617
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
55 |  | 22-0194 | -6.695 | -56.318 | -9.476 | -70.8154 | 0.348 | 1291.326 | 2022/12/15 17:36:41
55 |  | 22-0194 | -6.926 | -58.759 | -9.7074 | -73.1066 | 0.375 | 1291.274 | 2022/12/15 17:45:37
55 |  | 22-0194 | -6.918 | -58.908 | -9.6994 | -73.2465 | 0.329 | 1291.821 | 2022/12/15 17:54:33
55 |  | 22-0194 | -6.918 | -58.715 | -9.6994 | -73.0653 | 0.461 | 1292.634 | 2022/12/15 18:03:29
Sample 55 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
55 |  | 22-0194 | -9.6455 | -72.5585 |  |  |  | 0.1129 | 1.2408
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
56 |  | SCBI4-B-3 | -3.761 | -38.704 | -6.5373 | -54.2825 | -0.922 | 1924.243 | 2022/12/15 18:13:16
56 |  | SCBI4-B-3 | -3.593 | -36.757 | -6.369 | -52.455 | -1.205 | 1948.377 | 2022/12/15 18:22:12
56 |  | SCBI4-B-3 | -3.585 | -36.563 | -6.361 | -52.2729 | -0.591 | 1949.033 | 2022/12/15 18:31:08
56 |  | SCBI4-B-3 | -3.571 | -35.998 | -6.347 | -51.7426 | -1.025 | 1938.611 | 2022/12/15 18:40:04
Sample 56 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
56 |  | SCBI4-B-3 | -6.4036 | -52.6883 |  | 636.6517 |  | 0.089463 | 1.1772 | -23.4175 | -21.971
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
57 |  | SCBI4-B-2 |  |  | -2.7702 | -17.9541 |  |  | 2022/12/15 18:49:50
57 |  | SCBI4-B-2 | -2.846 | -35.663 | -5.6208 | -51.4282 | -0.807 | 1728.92 | 2022/12/15 18:58:46
57 |  | SCBI4-B-2 | -2.889 | -36.458 | -5.6639 | -52.1744 | -0.935 | 1697.275 | 2022/12/15 19:07:42
57 |  | SCBI4-B-2 | -2.762 | -35.434 | -5.5367 | -51.2132 | -0.517 | 1754.782 | 2022/12/15 19:16:37
Sample 57 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
57 | * | SCBI4-B-2 | -4.8979 | -43.1925 |  |  |  | 1.4171 | 17.9312 | -12.5335
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
58 |  | EV | -7.264 | -56.223 | -10.0459 | -70.7263 | 0.51 | 1303.56 | 2022/12/15 19:26:23
58 |  | EV | -7.422 | -58.18 | -10.2042 | -72.5631 | 0.479 | 1298.865 | 2022/12/15 19:35:19
58 |  | EV | -7.511 | -58.898 | -10.2933 | -73.2371 | 0.105 | 1297.078 | 2022/12/15 19:44:14
58 |  | EV | -7.458 | -59.035 | -10.2402 | -73.3657 | 0.662 | 1296.546 | 2022/12/15 19:53:10
Sample 58 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
58 |  | EV | -10.1959 | -72.473 |  |  |  | 0.10632 | 1.2961
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
59 |  | SLCP-22-020 | -19.131 | -149.208 | -21.932 | -158.0041 | 0.405 | 1289.631 | 2022/12/15 20:02:57
59 |  | SLCP-22-020 | -19.071 | -155.98 | -21.8719 | -164.3604 | 0.527 | 1308.455 | 2022/12/15 20:11:53
59 |  | SLCP-22-020 | -19.792 | -160.197 | -22.594 | -168.3186 | 0.213 | 1294.765 | 2022/12/15 20:20:48
59 |  | SLCP-22-020 | -19.897 | -161.299 | -22.6992 | -169.353 | 0.567 | 1295.312 | 2022/12/15 20:29:44
Sample 59 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
59 |  | SLCP-22-020 | -22.2743 | -165.009 |  |  |  | 0.43209 | 5.478
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
60 |  | EV | -8.348 | -72.198 | -11.1316 | -85.7208 | 0.457 | 1296.07 | 2022/12/15 20:39:31
60 |  | EV | -7.794 | -63.258 | -10.5768 | -77.3295 | 0.423 | 1295.658 | 2022/12/15 20:48:27
60 |  | EV | -7.67 | -61.378 | -10.4526 | -75.5649 | 0.54 | 1296.861 | 2022/12/15 20:57:23
60 |  | EV | -7.661 | -60.445 | -10.4435 | -74.6891 | 0.304 | 1296.593 | 2022/12/15 21:06:19
Sample 60 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
60 |  | EV | -10.6511 | -78.3261 |  |  |  | 0.32554 | 5.3809

## Source
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Line | Analysis | Time Code | Port | Inj Nr | d(18_16)Mean | d(D_H)Mean | H2O_Mean | Ignore | Good | Identifier 1 | Identifier 2 | Gas Configuration | Timestamp Mean | d(18_16)_SD | d(D_H)_SD | H2O_SD | d(18_16)_Sl | d(D_H)_Sl | H2O_Sl | baseline_shift | slope_shift | residuals | baseline_curvature | interval | ch4_ppm | h16od_adjust | h16od_shift | n2_flag | DAS Temp | Tray | Sample | Job | Method | Error Code
1 | P-2217 | 2022/12/13 11:19:11 | 2-01 | 1 | [5 empty] | EV | standard | H2O | [17 empty] | 1 | 1 | 1 | GJB_Rinse | 0
2 | P-2217 | 2022/12/13 11:29:13 | 2-01 | 2 | -8.54 | -73.804 | 21590 | -1 | 1 | EV | standard | H2O | 1670931260 | 0.326 | 0.789 | 83.12 | 0 | -0.001 | 1.956 | 1405.818 | -1.22 | 3.532 | 132.728 | 0.883 | 0.11 | 0 | 0 | 1 | 42.625 | 1 | 1 | 1 | GJB_Rinse | 0
3 | P-2217 | 2022/12/13 11:38:10 | 2-01 | 3 | -7.996 | -66.476 | 21729 | -1 | 1 | EV | standard | H2O | 1670931796 | 0.357 | 0.737 | 87.769 | 0.001 | 0 | 2.055 | 1398.85 | -1.28 | 3.55 | 133.667 | 0.883 | 0.524 | 0 | 0 | 1 | 42.312 | 1 | 1 | 1 | GJB_Rinse | 0
4 | P-2217 | 2022/12/13 11:47:07 | 2-01 | 4 | -7.81 | -63.642 | 19307 | 0 | 1 | EV | standard | H2O | 1670932334 | 0.293 | 0.754 | 82.289 | 0.001 | -0.001 | 1.96 | 1330.216 | 0.795 | 3.078 | 119.009 | 0.874 | 0.365 | 0 | 0 | 1 | 42.375 | 1 | 1 | 1 | GJB_Rinse | 0
5 | P-2217 | 2022/12/13 11:56:04 | 2-01 | 5 | -7.663 | -61.793 | 18891 | 0 | 1 | EV | standard | H2O | 1670932871 | 0.328 | 0.698 | 81.878 | 0 | 0.001 | 1.948 | 1308.256 | 1.38 | 3.06 | 117.547 | 0.871 | 0.443 | 0 | 0 | 1 | 42.25 | 1 | 1 | 1 | GJB_Rinse | 0
6 | P-2217 | 2022/12/13 12:05:01 | 2-01 | 6 | -7.596 | -60.773 | 19011 | 0 | 1 | EV | standard | H2O | 1670933407 | 0.313 | 0.741 | 82.62 | 0.001 | 0 | 1.95 | 1305.883 | 2.272 | 3.009 | 117.417 | 0.87 | 0.722 | 0 | 0 | 1 | 42.188 | 1 | 1 | 1 | GJB_Rinse | 0
7 | P-2217 | 2022/12/13 12:13:58 | 2-01 | 7 | -7.629 | -60.621 | 19004 | 0 | 1 | EV | standard | H2O | 1670933945 | 0.325 | 0.701 | 84.27 | 0.001 | 0 | 1.997 | 1305.126 | 2.256 | 3.021 | 116.029 | 0.872 | 0.379 | 0 | 0 | 1 | 42.312 | 1 | 1 | 1 | GJB_Rinse | 0
8 | P-2217 | 2022/12/13 12:22:54 | 2-01 | 8 | -7.592 | -60.485 | 18978 | 0 | 1 | EV | standard | H2O | 1670934481 | 0.3 | 0.76 | 80.359 | 0.001 | 0 | 1.906 | 1304.33 | 1.336 | 3.094 | 117.913 | 0.873 | 0.48 | 0 | 0 | 1 | 42.562 | 1 | 1 | 1 | GJB_Rinse | 0
9 | P-2217 | 2022/12/13 12:31:51 | 2-01 | 9 | -7.572 | -60.328 | 18977 | 0 | 1 | EV | standard | H2O | 1670935018 | 0.33 | 0.696 | 78.879 | 0.001 | 0 | 1.878 | 1304.046 | 1.622 | 3.035 | 116.688 | 0.869 | 0.366 | 0 | 0 | 1 | 42.375 | 1 | 1 | 1 | GJB_Rinse | 0
10 | P-2217 | 2022/12/13 12:40:49 | 2-01 | 10 | -7.542 | -60.173 | 19007 | 0 | 1 | EV | standard | H2O | 1670935555 | 0.302 | 0.702 | 78.333 | 0.001 | -0.002 | 1.868 | 1304.06 | 1.592 | 2.996 | 116.143 | 0.868 | 0.094 | 0 | 0 | 1 | 42.375 | 1 | 1 | 1 | GJB_Rinse | 0
11 | P-2218 | 2022/12/13 12:50:38 | 2-02 | 1 | 4.002 | 27.931 | 18863 | -1 | 1 | PZ | standard | H2O | 1670936145 | 0.287 | 0.776 | 82.464 | 0 | -0.002 | 1.967 | 1301.761 | 1.476 | 3.052 | 116.181 | 0.868 | 0.252 | 0 | 0 | 1 | 42.75 | 1 | 2 | 1 | GJB_Rinse | 0
12 | P-2218 | 2022/12/13 12:59:34 | 2-02 | 2 | 4.502 | 35.95 | 19108 | -1 | 1 | PZ | standard | H2O | 1670936681 | 0.317 | 0.773 | 80.94 | 0 | -0.008 | 1.925 | 1304.158 | 1.199 | 3.145 | 119.363 | 0.876 | 0.503 | 0 | 0 | 1 | 42.562 | 1 | 2 | 1 | GJB_Rinse | 0
13 | P-2218 | 2022/12/13 13:08:31 | 2-02 | 3 | 4.634 | 38.323 | 19028 | -1 | 1 | PZ | standard | H2O | 1670937217 | 0.329 | 0.675 | 81.904 | -0.001 | -0.004 | 1.946 | 1303.736 | 1.349 | 3.057 | 116.554 | 0.876 | 0.156 | 0 | 0 | 1 | 42.438 | 1 | 2 | 1 | GJB_Rinse | 0
14 | P-2218 | 2022/12/13 13:17:28 | 2-02 | 4 | 4.702 | 39.038 | 18834 | 0 | 1 | PZ | standard | H2O | 1670937755 | 0.334 | 0.801 | 82.505 | 0 | -0.004 | 1.954 | 1301.942 | 1.544 | 3.066 | 116.577 | 0.873 | 0.455 | 0 | 0 | 1 | 42.5 | 1 | 2 | 1 | GJB_Rinse | 0
15 | P-2218 | 2022/12/13 13:26:26 | 2-02 | 5 | 4.734 | 39.387 | 18819 | 0 | 1 | PZ | standard | H2O | 1670938292 | 0.313 | 0.81 | 81.568 | 0.001 | -0.007 | 1.934 | 1301.732 | 1.344 | 3.06 | 116.549 | 0.868 | 0.36 | 0 | 0 | 1 | 42.562 | 1 | 2 | 1 | GJB_Rinse | 0
16 | P-2218 | 2022/12/13 13:35:23 | 2-02 | 6 | 4.845 | 40.433 | 19240 | 0 | 1 | PZ | standard | H2O | 1670938830 | 0.327 | 0.744 | 80.069 | 0 | -0.003 | 1.889 | 1304.965 | 1.197 | 3.155 | 119.057 | 0.866 | 0.471 | 0 | 0 | 1 | 42.625 | 1 | 2 | 1 | GJB_Rinse | 0
17 | P-2218 | 2022/12/13 13:44:20 | 2-02 | 7 | 4.825 | 40.454 | 18917 | 0 | 1 | PZ | standard | H2O | 1670939367 | 0.305 | 0.784 | 81.884 | 0.001 | -0.003 | 1.938 | 1302.532 | 1.293 | 3.082 | 116.992 | 0.869 | 0.434 | 0 | 0 | 1 | 42.5 | 1 | 2 | 1 | GJB_Rinse | 0
18 | P-2218 | 2022/12/13 13:53:17 | 2-02 | 8 | 4.881 | 40.7 | 19076 | 0 | 1 | PZ | standard | H2O | 1670939904 | 0.326 | 0.749 | 80.588 | 0.001 | -0.004 | 1.915 | 1303.933 | 0.989 | 3.125 | 118.679 | 0.871 | 0.395 | 0 | 0 | 1 | 42.438 | 1 | 2 | 1 | GJB_Rinse | 0
19 | P-2218 | 2022/12/13 14:02:15 | 2-02 | 9 | 4.881 | 40.95 | 19039 | 0 | 1 | PZ | standard | H2O | 1670940441 | 0.304 | 0.728 | 79.46 | 0 | -0.001 | 1.896 | 1303.41 | 0.85 | 3.161 | 119.746 | 0.872 | 0.594 | 0 | 0 | 1 | 42.5 | 1 | 2 | 1 | GJB_Rinse | 0
20 | P-2218 | 2022/12/13 14:11:12 | 2-02 | 10 | 4.921 | 40.951 | 19033 | 0 | 1 | PZ | standard | H2O | 1670940979 | 0.302 | 0.718 | 79.552 | 0 | -0.003 | 1.883 | 1303.426 | 1.425 | 3.128 | 118.195 | 0.869 | 0.527 | 0 | 0 | 1 | 42.375 | 1 | 2 | 1 | GJB_Rinse | 0
21 | P-2219 | 2022/12/13 14:21:01 | 2-03 | 1 | -12.045 | -92.364 | 19163 | -1 | 1 | UT2 | standard | H2O | 1670941568 | 0.312 | 0.814 | 81.329 | 0.003 | 0.011 | 1.925 | 1303.861 | 2.283 | 3.086 | 117.977 | 0.869 | 0.66 | 0 | 0 | 1 | 42.25 | 1 | 3 | 1 | GJB_Rinse | 0
22 | P-2219 | 2022/12/13 14:29:58 | 2-03 | 2 | -12.85 | -104.802 | 19167 | -1 | 1 | UT2 | standard | H2O | 1670942104 | 0.321 | 0.796 | 80.605 | 0.001 | 0.01 | 1.929 | 1304.325 | 2.012 | 3.052 | 118.31 | 0.869 | 0.633 | 0 | 0 | 1 | 42.312 | 1 | 3 | 1 | GJB_Rinse | 0
23 | P-2219 | 2022/12/13 14:38:55 | 2-03 | 3 | -13.025 | -107.803 | 19098 | -1 | 1 | UT2 | standard | H2O | 1670942642 | 0.306 | 0.811 | 77.592 | 0.001 | 0.007 | 1.844 | 1303.662 | 2.25 | 3.082 | 118.568 | 0.881 | 0.698 | 0 | 0 | 1 | 42.25 | 1 | 3 | 1 | GJB_Rinse | 0
24 | P-2219 | 2022/12/13 14:47:52 | 2-03 | 4 | -13.068 | -108.829 | 19257 | 0 | 1 | UT2 | standard | H2O | 1670943179 | 0.338 | 0.729 | 80.086 | 0.002 | 0.005 | 1.889 | 1305.269 | 2 | 3.111 | 119.107 | 0.876 | 0.524 | 0 | 0 | 1 | 42.312 | 1 | 3 | 1 | GJB_Rinse | 0
25 | P-2219 | 2022/12/13 14:56:49 | 2-03 | 5 | -13.188 | -109.722 | 18938 | 0 | 1 | UT2 | standard | H2O | 1670943716 | 0.324 | 0.738 | 78.769 | 0.001 | 0.006 | 1.866 | 1302.529 | 2.002 | 3.115 | 118.617 | 0.873 | 0.721 | 0 | 0 | 1 | 42.312 | 1 | 3 | 1 | GJB_Rinse | 0
26 | P-2219 | 2022/12/13 15:05:46 | 2-03 | 6 | -13.205 | -110.03 | 19229 | 0 | 1 | UT2 | standard | H2O | 1670944253 | 0.356 | 0.7 | 81.45 | 0.002 | 0.007 | 1.929 | 1304.861 | 1.933 | 3.153 | 120.097 | 0.868 | 0.703 | 0 | 0 | 1 | 42.375 | 1 | 3 | 1 | GJB_Rinse | 0
27 | P-2219 | 2022/12/13 15:14:43 | 2-03 | 7 | -13.267 | -110.336 | 19057 | 0 | 1 | UT2 | standard | H2O | 1670944790 | 0.29 | 0.727 | 80.8 | 0.001 | 0.006 | 1.917 | 1303.338 | 2.639 | 3.13 | 118.244 | 0.867 | 0.775 | 0 | 0 | 1 | 42.375 | 1 | 3 | 1 | GJB_Rinse | 0
28 | P-2219 | 2022/12/13 15:23:41 | 2-03 | 8 | -13.27 | -110.728 | 18925 | 0 | 1 | UT2 | standard | H2O | 1670945327 | 0.303 | 0.774 | 80.515 | 0.001 | 0.004 | 1.903 | 1303.014 | 2.266 | 3.008 | 115.027 | 0.87 | 0.247 | 0 | 0 | 1 | 42.25 | 1 | 3 | 1 | GJB_Rinse | 0
29 | P-2219 | 2022/12/13 15:32:38 | 2-03 | 9 | -13.217 | -110.81 | 18975 | 0 | 1 | UT2 | standard | H2O | 1670945864 | 0.303 | 0.744 | 81.261 | 0.002 | 0.003 | 1.909 | 1302.753 | 2.038 | 3.085 | 118.261 | 0.87 | 0.579 | 0 | 0 | 1 | 42.188 | 1 | 3 | 1 | GJB_Rinse | 0
30 | P-2219 | 2022/12/13 15:41:35 | 2-03 | 10 | -13.353 | -111.11 | 18925 | 0 | 1 | UT2 | standard | H2O | 1670946402 | 0.307 | 0.704 | 80 | 0 | 0.005 | 1.892 | 1303.078 | 1.684 | 3.079 | 117.724 | 0.87 | 0.616 | 0 | 0 | 1 | 42.25 | 1 | 3 | 1 | GJB_Rinse | 0
31 | P-2220 | 2022/12/13 15:51:25 | 2-04 | 1 | -7.83 | -65.169 | 19368 | -1 | 1 | EV | standard | H2O | 1670946992 | 0.296 | 0.664 | 81.854 | 0 | 0 | 1.953 | 1306.229 | 1.628 | 3.103 | 119.331 | 0.877 | 0.209 | 0 | 0 | 1 | 42.438 | 1 | 4 | 1 | GJB_Rinse | 0
32 | P-2220 | 2022/12/13 16:00:23 | 2-04 | 2 | -7.649 | -61.279 | 19039 | -1 | 1 | EV | standard | H2O | 1670947529 | 0.284 | 0.753 | 81.744 | 0 | 0 | 1.929 | 1303.556 | 1.564 | 3.09 | 117.572 | 0.872 | 0.39 | 0 | 0 | 1 | 42.375 | 1 | 4 | 1 | GJB_Rinse | 0
33 | P-2220 | 2022/12/13 16:09:20 | 2-04 | 3 | -7.59 | -60.391 | 19104 | -1 | 1 | EV | standard | H2O | 1670948067 | 0.335 | 0.75 | 79.422 | 0.002 | 0 | 1.887 | 1304.477 | 1.238 | 3.117 | 119.738 | 0.872 | 0.598 | 0 | 0 | 1 | 42.375 | 1 | 4 | 1 | GJB_Rinse | 0
34 | P-2220 | 2022/12/13 16:18:17 | 2-04 | 4 | -7.575 | -60.146 | 19105 | 0 | 1 | EV | standard | H2O | 1670948604 | 0.304 | 0.767 | 82.989 | 0.001 | -0.001 | 1.953 | 1304.719 | 1.651 | 3.054 | 117.38 | 0.873 | 0.41 | 0 | 0 | 1 | 42.438 | 1 | 4 | 1 | GJB_Rinse | 0
35 | P-2220 | 2022/12/13 16:27:14 | 2-04 | 5 | -7.582 | -59.969 | 18998 | 0 | 1 | EV | standard | H2O | 1670949141 | 0.3 | 0.756 | 80.588 | 0.001 | 0.002 | 1.917 | 1302.87 | 2.479 | 3.075 | 118.796 | 0.875 | 0.954 | 0 | 0 | 1 | 42.375 | 1 | 4 | 1 | GJB_Rinse | 0
36 | P-2220 | 2022/12/13 16:36:12 | 2-04 | 6 | -7.549 | -59.972 | 19114 | 0 | 1 | EV | standard | H2O | 1670949679 | 0.312 | 0.701 | 79.336 | 0.001 | 0.003 | 1.881 | 1304.427 | 2.411 | 3.045 | 116.695 | 0.867 | 0.428 | 0 | 0 | 1 | 42.375 | 1 | 4 | 1 | GJB_Rinse | 0
37 | P-2220 | 2022/12/13 16:45:09 | 2-04 | 7 | -7.567 | -59.801 | 19076 | 0 | 1 | EV | standard | H2O | 1670950216 | 0.313 | 0.689 | 81.748 | 0.002 | 0.002 | 1.944 | 1304.429 | 1.548 | 3.077 | 118.083 | 0.876 | 0.498 | 0 | 0 | 1 | 42.375 | 1 | 4 | 1 | GJB_Rinse | 0
38 | P-2220 | 2022/12/13 16:54:07 | 2-04 | 8 | -7.588 | -60.029 | 18952 | 0 | 1 | EV | standard | H2O | 1670950754 | 0.3 | 0.759 | 80.103 | 0.001 | 0.004 | 1.894 | 1302.921 | 2.065 | 3.067 | 117.936 | 0.871 | 0.673 | 0 | 0 | 1 | 42.312 | 1 | 4 | 1 | GJB_Rinse | 0
39 | P-2220 | 2022/12/13 17:03:04 | 2-04 | 9 | -7.54 | -59.642 | 19095 | 0 | 1 | EV | standard | H2O | 1670951290 | 0.304 | 0.732 | 82.345 | 0.001 | 0.003 | 1.947 | 1304.213 | 1.814 | 3.097 | 118.735 | 0.875 | 0.61 | 0 | 0 | 1 | 42.25 | 1 | 4 | 1 | GJB_Rinse | 0
40 | P-2220 | 2022/12/13 17:12:02 | 2-04 | 10 | -7.535 | -59.547 | 18957 | 0 | 1 | EV | standard | H2O | 1670951829 | 0.304 | 0.706 | 81.859 | 0.001 | 0.001 | 1.939 | 1303.175 | 1.982 | 3.093 | 117.61 | 0.872 | 0.605 | 0 | 0 | 1 | 42.312 | 1 | 4 | 1 | GJB_Rinse | 0
41 | P-2221 | 2022/12/13 17:21:52 | 2-05 | 1 | -7.132 | -55.458 | 19500 | -1 | 1 | 22-0684 | 22-309 | H2O | 1670952419 | 0.322 | 0.682 | 81.859 | 0.001 | 0.001 | 1.937 | 1307.487 | 1.645 | 3.111 | 119.329 | 0.879 | 0.094 | 0 | 0 | 1 | 42.5 | 1 | 5 | 1 | GJB_Rinse | 0
42 | P-2221 | 2022/12/13 17:30:49 | 2-05 | 2 | -7.205 | -55.637 | 18640 | -1 | 1 | 22-0684 | 22-309 | H2O | 1670952956 | 0.291 | 0.727 | 79.014 | 0 | 0.001 | 1.886 | 1301.111 | 2.227 | 2.991 | 115.001 | 0.872 | 0.508 | 0 | 0 | 1 | 42.5 | 1 | 5 | 1 | GJB_Rinse | 0
43 | P-2221 | 2022/12/13 17:39:46 | 2-05 | 3 | -7.164 | -55.339 | 18595 | -1 | 1 | 22-0684 | 22-309 | H2O | 1670953493 | 0.327 | 0.648 | 80.458 | 0.001 | 0.001 | 1.912 | 1300.273 | 1.884 | 2.93 | 113.598 | 0.871 | 0.237 | 0 | 0 | 1 | 42.375 | 1 | 5 | 1 | GJB_Rinse | 0
44 | P-2221 | 2022/12/13 17:48:42 | 2-05 | 4 | -7.226 | -55.742 | 18554 | 0 | 1 | 22-0684 | 22-309 | H2O | 1670954029 | 0.282 | 0.739 | 79.816 | 0.001 | 0.003 | 1.893 | 1300.058 | 2.177 | 3.027 | 113.966 | 0.867 | 0.518 | 0 | 0 | 1 | 42.5 | 1 | 5 | 1 | GJB_Rinse | 0
45 | P-2222 | 2022/12/13 17:58:33 | 2-06 | 1 | -6.24 | -50.322 | 19622 | -1 | 1 | 22-0685 | 22-309 | H2O | 1670954620 | 0.312 | 0.707 | 79.049 | 0.001 | 0.002 | 1.863 | 1308.241 | 1.632 | 3.111 | 119.468 | 0.882 | 0.1 | 0 | 0 | 1 | 42.5 | 1 | 6 | 1 | GJB_Rinse | 0
46 | P-2222 | 2022/12/13 18:07:30 | 2-06 | 2 | -6.323 | -50.493 | 19068 | -1 | 1 | 22-0685 | 22-309 | H2O | 1670955157 | 0.333 | 0.756 | 78.729 | 0.001 | 0.001 | 1.862 | 1303.661 | 2.218 | 3.057 | 117.021 | 0.865 | 0.436 | 0 | 0 | 1 | 42.5 | 1 | 6 | 1 | GJB_Rinse | 0
47 | P-2222 | 2022/12/13 18:16:28 | 2-06 | 3 | -6.247 | -50.183 | 19081 | -1 | 1 | 22-0685 | 22-309 | H2O | 1670955694 | 0.347 | 0.746 | 84.07 | 0.002 | 0.004 | 1.983 | 1303.846 | 1.748 | 3.096 | 118.081 | 0.872 | 0.412 | 0 | 0 | 1 | 42.375 | 1 | 6 | 1 | GJB_Rinse | 0
48 | P-2222 | 2022/12/13 18:25:25 | 2-06 | 4 | -6.281 | -50.258 | 18927 | 0 | 1 | 22-0685 | 22-309 | H2O | 1670956231 | 0.327 | 0.723 | 80.238 | 0.003 | 0 | 1.905 | 1302.874 | 1.879 | 3.006 | 115.901 | 0.871 | 0.402 | 0 | 0 | 1 | 42.438 | 1 | 6 | 1 | GJB_Rinse | 0
49 | P-2223 | 2022/12/13 18:35:15 | 2-07 | 1 | -6.102 | -51.369 | 18362 | -1 | 1 | 22-0679 | 22-309 | H2O | 1670956822 | 0.309 | 0.696 | 80.12 | 0.002 | 0.002 | 1.897 | 1298.391 | 1.539 | 2.926 | 113.843 | 0.869 | 0.397 | 0 | 0 | 1 | 42.625 | 1 | 7 | 1 | GJB_Rinse | 0
50 | P-2223 | 2022/12/13 18:44:12 | 2-07 | 2 | -6.162 | -51.867 | 18117 | -1 | 1 | 22-0679 | 22-309 | H2O | 1670957359 | 0.314 | 0.8 | 78.682 | 0.001 | 0.003 | 1.88 | 1296.662 | 1.545 | 2.928 | 112.275 | 0.866 | 0.415 | 0 | 0 | 1 | 42.5 | 1 | 7 | 1 | GJB_Rinse | 0
51 | P-2223 | 2022/12/13 18:53:09 | 2-07 | 3 | -6.132 | -51.62 | 18286 | -1 | 1 | 22-0679 | 22-309 | H2O | 1670957896 | 0.334 | 0.756 | 78.355 | 0.001 | 0.002 | 1.86 | 1297.715 | 1.883 | 2.969 | 112.913 | 0.866 | 0.456 | 0 | 0 | 1 | 42.375 | 1 | 7 | 1 | GJB_Rinse | 0
52 | P-2223 | 2022/12/13 19:02:06 | 2-07 | 4 | -6.092 | -51.632 | 18300 | 0 | 1 | 22-0679 | 22-309 | H2O | 1670958432 | 0.299 | 0.75 | 77.402 | 0.001 | 0.002 | 1.836 | 1297.967 | 1.454 | 2.877 | 113.62 | 0.871 | 0.325 | 0 | 0 | 1 | 42.312 | 1 | 7 | 1 | GJB_Rinse | 0
53 | P-2224 | 2022/12/13 19:11:53 | 2-08 | 1 | -7.8 | -58.579 | 19379 | -1 | 1 | 22-0693 | 22-309 | H2O | 1670959019 | 0.31 | 0.747 | 80.079 | 0.001 | 0.003 | 1.889 | 1305.632 | 1.858 | 3.061 | 118.812 | 0.877 | 0.249 | 0 | 0 | 1 | 42.5 | 1 | 8 | 1 | GJB_Rinse | 0
54 | P-2224 | 2022/12/13 19:20:50 | 2-08 | 2 | -7.98 | -59.743 | 18965 | -1 | 1 | 22-0693 | 22-309 | H2O | 1670959556 | 0.312 | 0.699 | 78.192 | 0.001 | 0 | 1.849 | 1302.875 | 1.899 | 2.996 | 115.828 | 0.874 | 0.143 | 0 | 0 | 1 | 42.562 | 1 | 8 | 1 | GJB_Rinse | 0
55 | P-2224 | 2022/12/13 19:29:47 | 2-08 | 3 | -7.994 | -60.325 | 18701 | -1 | 1 | 22-0693 | 22-309 | H2O | 1670960093 | 0.334 | 0.729 | 76.921 | 0.001 | 0.003 | 1.828 | 1300.36 | 1.762 | 3.011 | 115.626 | 0.871 | 0.299 | 0 | 0 | 1 | 42.562 | 1 | 8 | 1 | GJB_Rinse | 0
56 | P-2224 | 2022/12/13 19:38:44 | 2-08 | 4 | -8.023 | -60.63 | 18782 | 0 | 1 | 22-0693 | 22-309 | H2O | 1670960630 | 0.317 | 0.648 | 79.29 | 0 | 0.003 | 1.879 | 1300.823 | 2.335 | 3.04 | 116.095 | 0.868 | 0.823 | 0 | 0 | 1 | 42.312 | 1 | 8 | 1 | GJB_Rinse | 0
57 | P-2225 | 2022/12/13 19:48:32 | 2-09 | 1 | -6.684 | -49.748 | 16558 | -1 | 1 | 22-0678 | 22-309 | H2O | 1670961220 | 0.318 | 0.699 | 76.569 | 0 | 0.002 | 1.841 | 1284.684 | 1.655 | 2.707 | 103.516 | 0.866 | 0.582 | 0 | 0 | 1 | 42.438 | 1 | 9 | 1 | GJB_Rinse | 0
58 | P-2225 | 2022/12/13 19:57:29 | 2-09 | 2 | -6.496 | -48.403 | 16896 | -1 | 1 | 22-0678 | 22-309 | H2O | 1670961756 | 0.331 | 0.79 | 75.639 | 0.001 | 0.002 | 1.822 | 1287.236 | 1.688 | 2.731 | 104.669 | 0.864 | 0.521 | 0 | 0 | 1 | 42.438 | 1 | 9 | 1 | GJB_Rinse | 0
59 | P-2225 | 2022/12/13 20:06:27 | 2-09 | 3 | -6.532 | -48.199 | 17035 | -1 | 1 | 22-0678 | 22-309 | H2O | 1670962293 | 0.302 | 0.78 | 76.998 | 0 | -0.001 | 1.844 | 1288.285 | 1.748 | 2.695 | 105.035 | 0.865 | 0.41 | 0 | 0 | 1 | 42.375 | 1 | 9 | 1 | GJB_Rinse | 0
60 | P-2225 | 2022/12/13 20:15:24 | 2-09 | 4 | -6.433 | -47.334 | 17723 | 0 | 1 | 22-0678 | 22-309 | H2O | 1670962830 | 0.291 | 0.687 | 79.574 | 0 | 0.003 | 1.895 | 1293.205 | 1.785 | 2.826 | 109.512 | 0.869 | 0.27 | 0 | 0 | 1 | 42.5 | 1 | 9 | 1 | GJB_Rinse | 0
61 | P-2226 | 2022/12/13 20:25:13 | 2-10 | 1 | -2.761 | -28.095 | 18415 | -1 | 1 | 22-0191 | 22-309 | H2O | 1670963420 | 0.294 | 0.673 | 77.368 | 0.001 | -0.001 | 1.842 | 1298.315 | 1.237 | 2.969 | 114.221 | 0.869 | 0.411 | 0 | 0 | 1 | 42.562 | 1 | 10 | 1 | GJB_Rinse | 0
62 | P-2226 | 2022/12/13 20:34:09 | 2-10 | 2 | -2.442 | -25.908 | 18502 | -1 | 1 | 22-0191 | 22-309 | H2O | 1670963956 | 0.306 | 0.834 | 77.156 | 0.001 | 0 | 1.836 | 1298.5 | 1.545 | 3.028 | 115.103 | 0.864 | 0.483 | 0 | 0 | 1 | 42.5 | 1 | 10 | 1 | GJB_Rinse | 0
63 | P-2226 | 2022/12/13 20:43:06 | 2-10 | 3 | -2.506 | -25.724 | 18115 | -1 | 1 | 22-0191 | 22-309 | H2O | 1670964493 | 0.329 | 0.729 | 78.427 | 0.001 | 0.001 | 1.86 | 1295.596 | 1.582 | 2.932 | 113.418 | 0.868 | 0.725 | 0 | 0 | 1 | 42.438 | 1 | 10 | 1 | GJB_Rinse | 0
64 | P-2226 | 2022/12/13 20:52:04 | 2-10 | 4 | -2.496 | -25.439 | 18262 | 0 | 1 | 22-0191 | 22-309 | H2O | 1670965030 | 0.359 | 0.695 | 79.361 | 0.001 | 0.002 | 1.891 | 1296.947 | 1.806 | 2.914 | 111.931 | 0.873 | 0.242 | 0 | 0 | 1 | 42.875 | 1 | 10 | 1 | GJB_Rinse | 0
65 | P-2227 | 2022/12/13 21:01:53 | 2-11 | 1 | -6.612 | -50.686 | 16949 | -1 | 1 | 22-0691 | 22-309 | H2O | 1670965619 | 0.287 | 0.752 | 76.682 | 0.002 | 0.004 | 1.836 | 1287.144 | 1.738 | 2.724 | 105.957 | 0.869 | 0.569 | 0 | 0 | 1 | 42.562 | 1 | 11 | 1 | GJB_Rinse | 0
66 | P-2227 | 2022/12/13 21:10:50 | 2-11 | 2 | -6.527 | -51.862 | 18300 | -1 | 1 | 22-0691 | 22-309 | H2O | 1670966156 | 0.309 | 0.694 | 78.46 | 0.001 | 0.001 | 1.877 | 1296.915 | 1.857 | 2.936 | 113.056 | 0.87 | 0.406 | 0 | 0 | 1 | 42.625 | 1 | 11 | 1 | GJB_Rinse | 0
67 | P-2227 | 2022/12/13 21:19:47 | 2-11 | 3 | -6.738 | -53.072 | 17579 | -1 | 1 | 22-0691 | 22-309 | H2O | 1670966694 | 0.28 | 0.705 | 75.84 | 0 | 0.003 | 1.807 | 1291.969 | 1.639 | 2.765 | 108.293 | 0.874 | 0.244 | 0 | 0 | 1 | 42.625 | 1 | 11 | 1 | GJB_Rinse | 0
68 | P-2227 | 2022/12/13 21:28:44 | 2-11 | 4 | -6.754 | -52.641 | 18357 | 0 | 1 | 22-0691 | 22-309 | H2O | 1670967230 | 0.353 | 0.727 | 81.769 | 0.001 | 0.004 | 1.939 | 1297.051 | 2.645 | 2.942 | 112.529 | 0.872 | 0.504 | 0 | 0 | 1 | 42.562 | 1 | 11 | 1 | GJB_Rinse | 0
69 | P-2228 | 2022/12/13 21:38:34 | 2-12 | 1 | -4.542 | -43.252 | 14491 | -1 | 1 | 22-0185 | 22-309 | H2O | 1670967821 | 0.333 | 0.715 | 73.164 | 0 | -0.001 | 1.78 | 1270.065 | 1.812 | 2.217 | 88.797 | 0.862 | 0.138 | 0 | 0 | 1 | 42.5 | 1 | 12 | 1 | GJB_Rinse | 0
70 | P-2228 | 2022/12/13 21:47:31 | 2-12 | 2 | -4.002 | -39.884 | 17086 | -1 | 1 | 22-0185 | 22-309 | H2O | 1670968358 | 0.317 | 0.769 | 74.185 | 0.001 | 0.002 | 1.768 | 1288.323 | 1.22 | 2.747 | 106.598 | 0.868 | 0.383 | 0 | 0 | 1 | 42.438 | 1 | 12 | 1 | GJB_Rinse | 0
71 | P-2228 | 2022/12/13 21:56:29 | 2-12 | 3 | -3.986 | -39.147 | 17119 | -1 | 1 | 22-0185 | 22-309 | H2O | 1670968895 | 0.337 | 0.832 | 77.916 | 0.002 | 0.004 | 1.873 | 1288.43 | 1.899 | 2.715 | 105.431 | 0.861 | 0.453 | 0 | 0 | 1 | 42.5 | 1 | 12 | 1 | GJB_Rinse | 0
72 | P-2228 | 2022/12/13 22:05:25 | 2-12 | 4 | -3.89 | -38.881 | 17940 | 0 | 1 | 22-0185 | 22-309 | H2O | 1670969432 | 0.285 | 0.747 | 78.623 | 0 | 0 | 1.875 | 1294.489 | 1.523 | 2.838 | 110.093 | 0.864 | 0.218 | 0 | 0 | 1 | 42.438 | 1 | 12 | 1 | GJB_Rinse | 0
73 | P-2229 | 2022/12/13 22:15:16 | 2-13 | 1 | -7.201 | -56.688 | 18872 | -1 | 1 | EV | standard | H2O | 1670970022 | 0.307 | 0.769 | 78.856 | 0.001 | 0.003 | 1.865 | 1301.214 | 2.161 | 2.974 | 114.237 | 0.869 | -0.075 | 0 | 0 | 1 | 42.438 | 1 | 13 | 1 | GJB_Rinse | 0
74 | P-2229 | 2022/12/13 22:24:12 | 2-13 | 2 | -7.38 | -58.454 | 18741 | -1 | 1 | EV | standard | H2O | 1670970559 | 0.32 | 0.739 | 78.383 | 0.001 | 0.003 | 1.855 | 1300.128 | 1.251 | 3.051 | 117.109 | 0.868 | 0.507 | 0 | 0 | 1 | 42.375 | 1 | 13 | 1 | GJB_Rinse | 0
75 | P-2229 | 2022/12/13 22:33:09 | 2-13 | 3 | -7.457 | -58.619 | 18676 | -1 | 1 | EV | standard | H2O | 1670971096 | 0.335 | 0.755 | 82.749 | 0.001 | 0.003 | 1.968 | 1299.488 | 2 | 2.994 | 115.026 | 0.87 | 0.412 | 0 | 0 | 1 | 42.25 | 1 | 13 | 1 | GJB_Rinse | 0
76 | P-2229 | 2022/12/13 22:42:05 | 2-13 | 4 | -7.507 | -59.154 | 18672 | 0 | 1 | EV | standard | H2O | 1670971632 | 0.317 | 0.693 | 80.781 | 0.001 | 0.004 | 1.915 | 1299.675 | 1.283 | 3.035 | 115.888 | 0.867 | 0.566 | 0 | 0 | 1 | 42.312 | 1 | 13 | 1 | GJB_Rinse | 0
77 | P-2230 | 2022/12/13 22:51:55 | 2-14 | 1 | [5 empty] | 22-0201 | 22-309 | [18 empty] | 1 | 14 | 1 | GJB_Rinse | 0
78 | P-2230 | 2022/12/13 23:00:53 | 2-14 | 2 | 14.907 | 63.694 | 16622 | 0 | 1 | 22-0201 | 22-309 | H2O | 1670973834 | 0.318 | 0.784 | 77.65 | 0 | -0.003 | 1.862 | 1284.884 | 1.01 | 2.708 | 103.178 | 0.863 | 0.36 | 0 | 0 | 1 | 42.312 | 1 | 14 | 1 | GJB_Rinse | 0
79 | P-2230 | 2022/12/13 23:09:50 | 2-14 | 3 | 15.027 | 64.027 | 18060 | -1 | 1 | 22-0201 | 22-309 | H2O | 1670973297 | 0.301 | 0.78 | 82.159 | 0 | -0.002 | 1.956 | 1294.674 | 1.156 | 3.025 | 114.757 | 0.869 | 0.931 | 0 | 0 | 1 | 42.375 | 1 | 14 | 1 | GJB_Rinse | 0
80 | P-2230 | 2022/12/13 23:18:47 | 2-14 | 4 | 14.907 | 63.694 | 16622 | 0 | 1 | 22-0201 | 22-309 | H2O | 1670973834 | 0.318 | 0.784 | 77.65 | 0 | -0.003 | 1.862 | 1284.884 | 1.01 | 2.708 | 103.178 | 0.863 | 0.36 | 0 | 0 | 1 | 42.312 | 1 | 14 | 1 | GJB_Rinse | 0
81 | P-2231 | 2022/12/13 23:28:34 | 2-15 | 1 | -5.504 | -42.282 | 16877 | -1 | 1 | 22-0688 | 22-309 | H2O | 1670974420 | 0.324 | 0.736 | 74.251 | 0.001 | 0.004 | 1.78 | 1286.694 | 1.362 | 2.678 | 105.002 | 0.869 | 0.196 | 0 | 0 | 1 | 42.375 | 1 | 15 | 1 | GJB_Rinse | 0
82 | P-2231 | 2022/12/13 23:37:30 | 2-15 | 2 | -6.254 | -51.614 | 18678 | -1 | 1 | 22-0688 | 22-309 | H2O | 1670974957 | 0.32 | 0.748 | 77.42 | 0.002 | 0.007 | 1.84 | 1299.596 | 1.379 | 2.994 | 116.397 | 0.871 | 0.46 | 0 | 0 | 1 | 42.5 | 1 | 15 | 1 | GJB_Rinse | 0
83 | P-2231 | 2022/12/13 23:46:27 | 2-15 | 3 | -6.731 | -55.091 | 17075 | -1 | 1 | 22-0688 | 22-309 | H2O | 1670975493 | 0.322 | 0.764 | 75.283 | 0.002 | 0.005 | 1.795 | 1287.764 | 2.329 | 2.684 | 104.671 | 0.864 | 0.572 | 0 | 0 | 1 | 42.625 | 1 | 15 | 1 | GJB_Rinse | 0
84 | P-2231 | 2022/12/13 23:55:24 | 2-15 | 4 | -6.813 | -56.007 | 16851 | 0 | 1 | 22-0688 | 22-309 | H2O | 1670976031 | 0.3 | 0.808 | 77.326 | 0 | 0.004 | 1.843 | 1286.495 | 2.041 | 2.594 | 102.91 | 0.868 | 0.134 | 0 | 0 | 1 | 42.625 | 1 | 15 | 1 | GJB_Rinse | 0
85 | P-2232 | 2022/12/14 00:05:12 | 2-16 | 1 | -4.358 | -50.174 | 16903 | -1 | 1 | 22-0184 | 22-309 | H2O | 1670976619 | 0.336 | 0.815 | 76.144 | 0 | 0.004 | 1.815 | 1286.438 | 1.857 | 2.724 | 105.861 | 0.862 | 0.633 | 0 | 0 | 1 | 42.688 | 1 | 16 | 1 | GJB_Rinse | 0
86 | P-2232 | 2022/12/14 00:14:09 | 2-16 | 2 | -3.969 | -48.663 | 17972 | -1 | 1 | 22-0184 | 22-309 | H2O | 1670977155 | 0.299 | 0.738 | 78.962 | 0.001 | 0 | 1.881 | 1294.389 | 1.23 | 2.925 | 112.821 | 0.868 | 0.717 | 0 | 0 | 1 | 42.5 | 1 | 16 | 1 | GJB_Rinse | 0
87 | P-2232 | 2022/12/14 00:23:06 | 2-16 | 3 | -4.181 | -49.207 | 17125 | -1 | 1 | 22-0184 | 22-309 | H2O | 1670977692 | 0.281 | 0.845 | 79.329 | 0.001 | 0.006 | 1.9 | 1288.167 | 2.041 | 2.701 | 105.994 | 0.869 | 0.332 | 0 | 0 | 1 | 42.312 | 1 | 16 | 1 | GJB_Rinse | 0
88 | P-2232 | 2022/12/14 00:32:03 | 2-16 | 4 | -4.282 | -49.603 | 16861 | 0 | 1 | 22-0184 | 22-309 | H2O | 1670978230 | 0.304 | 0.86 | 78.59 | 0.002 | 0.006 | 1.873 | 1286.331 | 2.011 | 2.662 | 103.803 | 0.868 | 0.321 | 0 | 0 | 1 | 42.438 | 1 | 16 | 1 | GJB_Rinse | 0
89 | P-2233 | 2022/12/14 00:41:51 | 2-17 | 1 | -3.312 | -33.383 | 17295 | -1 | 1 | 22-0173 | 22-309 | H2O | 1670978818 | 0.313 | 0.711 | 77.77 | 0.002 | 0 | 1.867 | 1289.447 | 1.75 | 2.799 | 107.17 | 0.868 | 0.459 | 0 | 0 | 1 | 42.562 | 1 | 17 | 1 | GJB_Rinse | 0
90 | P-2233 | 2022/12/14 00:50:47 | 2-17 | 2 | -3.24 | -31.682 | 17509 | -1 | 1 | 22-0173 | 22-309 | H2O | 1670979354 | 0.311 | 0.787 | 77.565 | 0 | 0.002 | 1.86 | 1291.149 | 1.479 | 2.8 | 107.594 | 0.863 | 0.277 | 0 | 0 | 1 | 42.562 | 1 | 17 | 1 | GJB_Rinse | 0
91 | P-2233 | 2022/12/14 00:59:44 | 2-17 | 3 | -3.222 | -31.168 | 17542 | -1 | 1 | 22-0173 | 22-309 | H2O | 1670979890 | 0.333 | 0.735 | 80.62 | 0.002 | 0.004 | 1.913 | 1291.149 | 1.897 | 2.803 | 108.016 | 0.866 | 0.397 | 0 | 0 | 1 | 42.312 | 1 | 17 | 1 | GJB_Rinse | 0
92 | P-2233 | 2022/12/14 01:08:40 | 2-17 | 4 | -3.229 | -31.453 | 17635 | 0 | 1 | 22-0173 | 22-309 | H2O | 1670980426 | 0.309 | 0.744 | 78.103 | 0.001 | 0.002 | 1.869 | 1292.124 | 1.59 | 2.798 | 108.512 | 0.871 | 0.118 | 0 | 0 | 1 | 42.375 | 1 | 17 | 1 | GJB_Rinse | 0
93 | P-2234 | 2022/12/14 01:18:29 | 2-18 | 1 | -2.454 | -31.79 | 18236 | -1 | 1 | 22-0187 | 22-309 | H2O | 1670981016 | 0.323 | 0.73 | 78.573 | 0.002 | 0.002 | 1.859 | 1296.393 | 2.203 | 2.871 | 110.441 | 0.864 | 0.113 | 0 | 0 | 1 | 42.375 | 1 | 18 | 1 | GJB_Rinse | 0
94 | P-2234 | 2022/12/14 01:27:26 | 2-18 | 2 | -2.207 | -31.349 | 18865 | -1 | 1 | 22-0187 | 22-309 | H2O | 1670981553 | 0.318 | 0.712 | 79.146 | -0.001 | 0.001 | 1.887 | 1301.035 | 0.998 | 3.054 | 117.612 | 0.867 | 0.506 | 0 | 0 | 1 | 42.438 | 1 | 18 | 1 | GJB_Rinse | 0
95 | P-2234 | 2022/12/14 01:36:23 | 2-18 | 3 | -2.399 | -31.87 | 18484 | -1 | 1 | 22-0187 | 22-309 | H2O | 1670982090 | 0.298 | 0.694 | 76.673 | 0.001 | 0.004 | 1.813 | 1298.004 | 1.7 | 2.953 | 114.672 | 0.864 | 0.437 | 0 | 0 | 1 | 42.438 | 1 | 18 | 1 | GJB_Rinse | 0
96 | P-2234 | 2022/12/14 01:45:20 | 2-18 | 4 | -2.392 | -31.525 | 18601 | 0 | 1 | 22-0187 | 22-309 | H2O | 1670982627 | 0.328 | 0.718 | 79.117 | 0.002 | 0.002 | 1.869 | 1298.996 | 1.376 | 2.965 | 115.451 | 0.876 | 0.393 | 0 | 0 | 1 | 42.562 | 1 | 18 | 1 | GJB_Rinse | 0
97 | P-2235 | 2022/12/14 01:55:10 | 2-19 | 1 | -6.831 | -55.566 | 16242 | -1 | 1 | 22-0694 | 22-309 | H2O | 1670983216 | 0.336 | 0.793 | 76.156 | 0.003 | 0.004 | 1.831 | 1282.159 | 1.762 | 2.566 | 99.713 | 0.866 | 0.173 | 0 | 0 | 1 | 42.625 | 1 | 19 | 1 | GJB_Rinse | 0
98 | P-2235 | 2022/12/14 02:04:07 | 2-19 | 2 | -6.867 | -57.028 | 17182 | -1 | 1 | 22-0694 | 22-309 | H2O | 1670983754 | 0.327 | 0.728 | 75.645 | 0.001 | 0.004 | 1.811 | 1288.465 | 1.742 | 2.763 | 107.522 | 0.869 | 0.509 | 0 | 0 | 1 | 42.562 | 1 | 19 | 1 | GJB_Rinse | 0
99 | P-2235 | 2022/12/14 02:13:04 | 2-19 | 3 | -7.064 | -58.21 | 16298 | -1 | 1 | 22-0694 | 22-309 | H2O | 1670984291 | 0.279 | 0.712 | 74.675 | 0 | 0.003 | 1.803 | 1282.574 | 1.637 | 2.582 | 100.47 | 0.868 | 0.174 | 0 | 0 | 1 | 42.5 | 1 | 19 | 1 | GJB_Rinse | 0
100 | P-2235 | 2022/12/14 02:22:01 | 2-19 | 4 | -7.092 | -58.409 | 16455 | 0 | 1 | 22-0694 | 22-309 | H2O | 1670984828 | 0.293 | 0.761 | 74.691 | 0 | 0.004 | 1.793 | 1283.408 | 1.556 | 2.662 | 103.442 | 0.868 | 0.59 | 0 | 0 | 1 | 42.562 | 1 | 19 | 1 | GJB_Rinse | 0
101 | P-2236 | 2022/12/14 02:31:50 | 2-20 | 1 | -6.721 | -49.391 | 16406 | -1 | 1 | 22-0692 | 22-309 | H2O | 1670985417 | 0.312 | 0.75 | 75.575 | 0.001 | 0 | 1.819 | 1283.051 | 1.783 | 2.636 | 102.513 | 0.87 | 0.47 | 0 | 0 | 1 | 42.625 | 1 | 20 | 1 | GJB_Rinse | 0
102 | P-2236 | 2022/12/14 02:40:47 | 2-20 | 2 | -6.503 | -47.574 | 17041 | -1 | 1 | 22-0692 | 22-309 | H2O | 1670985954 | 0.317 | 0.773 | 74.847 | 0.002 | 0 | 1.791 | 1287.724 | 1.801 | 2.713 | 104.958 | 0.865 | 0.28 | 0 | 0 | 1 | 42.688 | 1 | 20 | 1 | GJB_Rinse | 0
103 | P-2236 | 2022/12/14 02:49:45 | 2-20 | 3 | -6.669 | -48.141 | 16234 | -1 | 1 | 22-0692 | 22-309 | H2O | 1670986492 | 0.279 | 0.802 | 72.837 | 0.002 | -0.001 | 1.752 | 1282.124 | 1.318 | 2.594 | 101.146 | 0.865 | 0.575 | 0 | 0 | 1 | 42.562 | 1 | 20 | 1 | GJB_Rinse | 0
104 | P-2236 | 2022/12/14 02:58:42 | 2-20 | 4 | -6.692 | -47.82 | 16345 | 0 | 1 | 22-0692 | 22-309 | H2O | 1670987028 | 0.323 | 0.769 | 72.736 | 0.001 | 0 | 1.748 | 1282.727 | 2.173 | 2.617 | 100.562 | 0.861 | 0.436 | 0 | 0 | 1 | 42.375 | 1 | 20 | 1 | GJB_Rinse | 0
105 | P-2237 | 2022/12/14 03:08:31 | 2-21 | 1 | 2.412 | -11.651 | 17756 | -1 | 1 | 22-0190 | 22-309 | H2O | 1670987618 | 0.327 | 0.708 | 76.435 | 0.001 | -0.001 | 1.81 | 1292.937 | 1.678 | 2.856 | 109.563 | 0.869 | 0.416 | 0 | 0 | 1 | 42.5 | 1 | 21 | 1 | GJB_Rinse | 0
106 | P-2237 | 2022/12/14 03:17:28 | 2-21 | 2 | 3.042 | -7.642 | 17783 | -1 | 1 | 22-0190 | 22-309 | H2O | 1670988155 | 0.307 | 0.751 | 77.136 | 0.001 | 0 | 1.841 | 1293.019 | 1.607 | 2.892 | 110.204 | 0.867 | 0.532 | 0 | 0 | 1 | 42.375 | 1 | 21 | 1 | GJB_Rinse | 0
107 | P-2237 | 2022/12/14 03:26:25 | 2-21 | 3 | 2.92 | -8.016 | 16659 | -1 | 1 | 22-0190 | 22-309 | H2O | 1670988692 | 0.305 | 0.762 | 73.527 | 0.001 | 0.002 | 1.771 | 1285.178 | 1.215 | 2.687 | 103.531 | 0.86 | 0.448 | 0 | 0 | 1 | 42.5 | 1 | 21 | 1 | GJB_Rinse | 0
108 | P-2237 | 2022/12/14 03:35:22 | 2-21 | 4 | 3.023 | -7.416 | 16718 | 0 | 1 | 22-0190 | 22-309 | H2O | 1670989229 | 0.328 | 0.763 | 75.136 | 0.001 | 0.002 | 1.799 | 1285.304 | 1.485 | 2.741 | 104.336 | 0.865 | 0.606 | 0 | 0 | 1 | 42.438 | 1 | 21 | 1 | GJB_Rinse | 0
109 | P-2238 | 2022/12/14 03:45:08 | 2-22 | 1 | -6.835 | -53.081 | 18641 | -1 | 1 | EV | standard | H2O | 1670989815 | 0.316 | 0.74 | 76.974 | 0.001 | 0.003 | 1.827 | 1299.422 | 1.806 | 2.946 | 114.036 | 0.867 | 0.319 | 0 | 0 | 1 | 42.438 | 1 | 22 | 1 | GJB_Rinse | 0
110 | P-2238 | 2022/12/14 03:54:05 | 2-22 | 2 | -7.273 | -57.245 | 18668 | -1 | 1 | EV | standard | H2O | 1670990352 | 0.348 | 0.736 | 79.466 | 0 | 0.002 | 1.889 | 1299.64 | 1.909 | 2.993 | 113.954 | 0.866 | 0.335 | 0 | 0 | 1 | 42.312 | 1 | 22 | 1 | GJB_Rinse | 0
111 | P-2238 | 2022/12/14 04:03:02 | 2-22 | 3 | -7.39 | -58.231 | 18713 | -1 | 1 | EV | standard | H2O | 1670990889 | 0.299 | 0.712 | 81.202 | 0 | 0.006 | 1.929 | 1299.894 | 1.575 | 3.025 | 115.65 | 0.866 | 0.399 | 0 | 0 | 1 | 42.375 | 1 | 22 | 1 | GJB_Rinse | 0
112 | P-2238 | 2022/12/14 04:11:59 | 2-22 | 4 | -7.431 | -58.506 | 18636 | 0 | 1 | EV | standard | H2O | 1670991426 | 0.341 | 0.733 | 80.986 | 0.002 | 0.002 | 1.923 | 1298.977 | 2.281 | 2.98 | 114.927 | 0.871 | 0.499 | 0 | 0 | 1 | 42.375 | 1 | 22 | 1 | GJB_Rinse | 0
113 | P-2239 | 2022/12/14 04:21:46 | 2-23 | 1 | 11.277 | 38.442 | 17754 | -1 | 1 | 22-0200 | 22-309 | H2O | 1670992013 | 0.292 | 0.766 | 79.838 | -0.001 | -0.004 | 1.896 | 1292.844 | 1.218 | 2.887 | 111.217 | 0.871 | 0.507 | 0 | 0 | 1 | 42.375 | 1 | 23 | 1 | GJB_Rinse | 0
114 | P-2239 | 2022/12/14 04:30:43 | 2-23 | 2 | 12.322 | 48.366 | 17576 | -1 | 1 | 22-0200 | 22-309 | H2O | 1670992550 | 0.325 | 0.773 | 75.911 | 0 | -0.006 | 1.815 | 1291.68 | 0.964 | 2.82 | 109.469 | 0.876 | 0.32 | 0 | 0 | 1 | 42.438 | 1 | 23 | 1 | GJB_Rinse | 0
115 | P-2239 | 2022/12/14 04:39:40 | 2-23 | 3 | 12.42 | 49.9 | 17285 | -1 | 1 | 22-0200 | 22-309 | H2O | 1670993087 | 0.316 | 0.786 | 77.928 | 0 | -0.003 | 1.866 | 1289.822 | 0.879 | 2.822 | 107.534 | 0.868 | 0.243 | 0 | 0 | 1 | 42.562 | 1 | 23 | 1 | GJB_Rinse | 0
116 | P-2239 | 2022/12/14 04:48:37 | 2-23 | 4 | 12.475 | 51.016 | 17474 | 0 | 1 | 22-0200 | 22-309 | H2O | 1670993623 | 0.306 | 0.794 | 79.362 | 0 | -0.003 | 1.88 | 1290.979 | 0.866 | 2.834 | 108.938 | 0.869 | 0.35 | 0 | 0 | 1 | 42.438 | 1 | 23 | 1 | GJB_Rinse | 0
117 | P-2240 | 2022/12/14 04:58:24 | 2-24 | 1 | [5 empty] | 22-0175 | 22-309 | [18 empty] | 1 | 24 | 1 | GJB_Rinse | 0
118 | P-2240 | 2022/12/14 05:07:21 | 2-24 | 2 | -2.027 | -19.615 | 19125 | -1 | 1 | 22-0175 | 22-309 | H2O | 1670994748 | 0.321 | 0.746 | 78.886 | 0.001 | 0.003 | 1.868 | 1303.154 | 1.012 | 3.025 | 117.747 | 0.873 | 0.235 | 0 | 0 | 1 | 42.562 | 1 | 24 | 1 | GJB_Rinse | 0
119 | P-2240 | 2022/12/14 05:16:18 | 2-24 | 3 | -2.414 | -21.897 | 17439 | -1 | 1 | 22-0175 | 22-309 | H2O | 1670995284 | 0.352 | 0.714 | 77.248 | 0.002 | 0.002 | 1.838 | 1290.541 | 1.387 | 2.822 | 109.444 | 0.874 | 0.509 | 0 | 0 | 1 | 42.625 | 1 | 24 | 1 | GJB_Rinse | 0
120 | P-2240 | 2022/12/14 05:25:14 | 2-24 | 4 | -2.449 | -22.828 | 17511 | 0 | 1 | 22-0175 | 22-309 | H2O | 1670995821 | 0.318 | 0.787 | 77.267 | 0.001 | 0.001 | 1.847 | 1291.056 | 1.281 | 2.852 | 109.882 | 0.871 | 0.579 | 0 | 0 | 1 | 42.75 | 1 | 24 | 1 | GJB_Rinse | 0
121 | P-2241 | 2022/12/14 05:35:02 | 2-25 | 1 | -1.063 | -29.011 | 16648 | -1 | 1 | 22-0180 | 22-309 | H2O | 1670996409 | 0.337 | 0.702 | 75.184 | 0 | 0.002 | 1.79 | 1285.187 | 1.09 | 2.685 | 103.813 | 0.864 | 0.376 | 0 | 0 | 1 | 42.5 | 1 | 25 | 1 | GJB_Rinse | 0
122 | P-2241 | 2022/12/14 05:43:59 | 2-25 | 2 | [5 empty] | 22-0180 | 22-309 | [18 empty] | 1 | 25 | 1 | GJB_Rinse | 0
123 | P-2241 | 2022/12/14 05:52:55 | 2-25 | 3 | -0.988 | -29.654 | 16248 | -1 | 1 | 22-0180 | 22-309 | H2O | 1670997482 | 0.342 | 0.766 | 77.098 | 0.002 | 0.005 | 1.84 | 1282.227 | 1.804 | 2.522 | 99.676 | 0.872 | 0.174 | 0 | 0 | 1 | 42.375 | 1 | 25 | 1 | GJB_Rinse | 0
124 | P-2241 | 2022/12/14 06:01:52 | 2-25 | 4 | -1.053 | -30.055 | 16721 | 0 | 1 | 22-0180 | 22-309 | H2O | 1670998019 | 0.316 | 0.807 | 73.573 | 0.001 | 0.003 | 1.766 | 1285.321 | 1.291 | 2.695 | 105.525 | 0.872 | 0.593 | 0 | 0 | 1 | 42.438 | 1 | 25 | 1 | GJB_Rinse | 0
125 | P-2242 | 2022/12/14 06:11:40 | 2-26 | 1 | 5.553 | 7.71 | 17218 | -1 | 1 | 22-0198 | 22-309 | H2O | 1670998607 | 0.335 | 0.815 | 75.109 | 0 | -0.004 | 1.796 | 1288.976 | 1.428 | 2.755 | 107.069 | 0.865 | 0.478 | 0 | 0 | 1 | 42.375 | 1 | 26 | 1 | GJB_Rinse | 0
126 | P-2242 | 2022/12/14 06:20:36 | 2-26 | 2 | 6.411 | 14.018 | 19230 | -1 | 1 | 22-0198 | 22-309 | H2O | 1670999143 | 0.316 | 0.725 | 80.565 | 0.002 | 0 | 1.894 | 1303.183 | 1.216 | 3.188 | 121.276 | 0.865 | 0.857 | 0 | 0 | 1 | 42.25 | 1 | 26 | 1 | GJB_Rinse | 0
127 | P-2242 | 2022/12/14 06:29:33 | 2-26 | 3 | 6.002 | 12.809 | 17847 | -1 | 1 | 22-0198 | 22-309 | H2O | 1670999680 | 0.294 | 0.737 | 77.192 | 0.001 | 0 | 1.84 | 1293.583 | 1.003 | 2.932 | 111.653 | 0.874 | 0.442 | 0 | 0 | 1 | 42.312 | 1 | 26 | 1 | GJB_Rinse | 0
128 | P-2242 | 2022/12/14 06:38:30 | 2-26 | 4 | 5.993 | 12.896 | 17765 | 0 | 1 | 22-0198 | 22-309 | H2O | 1671000217 | 0.3 | 0.725 | 77.121 | 0.001 | 0.002 | 1.842 | 1292.871 | 1.384 | 2.846 | 110.731 | 0.867 | 0.501 | 0 | 0 | 1 | 42.375 | 1 | 26 | 1 | GJB_Rinse | 0
129 | P-2243 | 2022/12/14 06:48:19 | 2-27 | 1 | -5.955 | -47.445 | 17317 | -1 | 1 | 22-0681 | 22-309 | H2O | 1671000806 | 0.311 | 0.832 | 78.982 | 0.001 | 0.007 | 1.888 | 1289.823 | 1.275 | 2.743 | 107.239 | 0.866 | 0.262 | 0 | 0 | 1 | 42.562 | 1 | 27 | 1 | GJB_Rinse | 0
130 | P-2243 | 2022/12/14 06:57:16 | 2-27 | 2 | -6.365 | -53.137 | 18127 | -1 | 1 | 22-0681 | 22-309 | H2O | 1671001342 | 0.332 | 0.736 | 79.651 | 0.002 | 0.007 | 1.9 | 1295.25 | 1.724 | 2.943 | 113.194 | 0.867 | 0.594 | 0 | 0 | 1 | 42.625 | 1 | 27 | 1 | GJB_Rinse | 0
131 | P-2243 | 2022/12/14 07:06:12 | 2-27 | 3 | -6.708 | -55.39 | 16348 | -1 | 1 | 22-0681 | 22-309 | H2O | 1671001879 | 0.298 | 0.853 | 74.795 | 0.001 | 0.007 | 1.808 | 1282.7 | 1.927 | 2.603 | 101.008 | 0.867 | 0.28 | 0 | 0 | 1 | 42.625 | 1 | 27 | 1 | GJB_Rinse | 0
132 | P-2243 | 2022/12/14 07:15:10 | 2-27 | 4 | -6.794 | -56.179 | 16339 | 0 | 1 | 22-0681 | 22-309 | H2O | 1671002417 | 0.3 | 0.765 | 75.212 | 0.001 | 0.004 | 1.806 | 1282.644 | 1.958 | 2.577 | 100.906 | 0.867 | 0.29 | 0 | 0 | 1 | 42.625 | 1 | 27 | 1 | GJB_Rinse | 0
133 | P-2244 | 2022/12/14 07:24:59 | 2-28 | 1 | -6.022 | -48.814 | 17320 | -1 | 1 | 22-0686 | 22-309 | H2O | 1671003006 | 0.305 | 0.777 | 77.538 | 0.001 | 0.005 | 1.85 | 1289.458 | 1.87 | 2.767 | 107.557 | 0.872 | 0.513 | 0 | 0 | 1 | 42.438 | 1 | 28 | 1 | GJB_Rinse | 0
134 | P-2244 | 2022/12/14 07:33:56 | 2-28 | 2 | -5.941 | -48.251 | 16771 | -1 | 1 | 22-0686 | 22-309 | H2O | 1671003543 | 0.301 | 0.775 | 74.516 | 0.002 | 0.006 | 1.779 | 1285.774 | 1.831 | 2.625 | 103.088 | 0.867 | 0.289 | 0 | 0 | 1 | 42.25 | 1 | 28 | 1 | GJB_Rinse | 0
135 | P-2244 | 2022/12/14 07:42:54 | 2-28 | 3 | -6.051 | -48.76 | 15869 | -1 | 1 | 22-0686 | 22-309 | H2O | 1671004080 | 0.287 | 0.715 | 73.717 | 0 | 0.002 | 1.766 | 1279.26 | 2.309 | 2.472 | 96.779 | 0.864 | 0.276 | 0 | 0 | 1 | 42.25 | 1 | 28 | 1 | GJB_Rinse | 0
136 | P-2244 | 2022/12/14 07:51:51 | 2-28 | 4 | -6.042 | -49.089 | 15902 | 0 | 1 | 22-0686 | 22-309 | H2O | 1671004618 | 0.316 | 0.673 | 72.383 | 0.001 | 0.002 | 1.737 | 1279.649 | 1.896 | 2.486 | 97.512 | 0.863 | 0.243 | 0 | 0 | 1 | 42.312 | 1 | 28 | 1 | GJB_Rinse | 0
137 | P-2245 | 2022/12/14 08:01:35 | 2-29 | 1 | -6.286 | -50.408 | 15750 | -1 | 1 | 22-0683 | 22-309 | H2O | 1671005202 | 0.31 | 0.793 | 75.783 | 0.001 | 0.003 | 1.838 | 1278.46 | 2.295 | 2.489 | 96.299 | 0.864 | 0.125 | 0 | 0 | 1 | 42.562 | 1 | 29 | 1 | GJB_Rinse | 0
138 | P-2245 | 2022/12/14 08:10:32 | 2-29 | 2 | -6.067 | -49.357 | 16853 | -1 | 1 | 22-0683 | 22-309 | H2O | 1671005739 | 0.325 | 0.736 | 75.233 | 0.001 | 0.004 | 1.796 | 1286.367 | 1.541 | 2.726 | 104.583 | 0.866 | 0.341 | 0 | 0 | 1 | 43.188 | 1 | 29 | 1 | GJB_Rinse | 0
139 | P-2245 | 2022/12/14 08:19:30 | 2-29 | 3 | -6.328 | -50.273 | 15985 | -1 | 1 | 22-0683 | 22-309 | H2O | 1671006277 | 0.267 | 0.843 | 70.49 | 0.001 | 0.001 | 1.698 | 1280.403 | 1.344 | 2.505 | 98.884 | 0.865 | 0.201 | 0 | 0 | 1 | 45.25 | 1 | 29 | 1 | GJB_Rinse | 0
140 | P-2245 | 2022/12/14 08:28:26 | 2-29 | 4 | -6.259 | -50.078 | 16073 | 0 | 1 | 22-0683 | 22-309 | H2O | 1671006813 | 0.298 | 0.768 | 73.96 | 0 | 0.002 | 1.776 | 1280.547 | 1.691 | 2.568 | 100.904 | 0.867 | 0.589 | 0 | 0 | 1 | 45.75 | 1 | 29 | 1 | GJB_Rinse | 0
141 | P-2246 | 2022/12/14 08:38:14 | 2-30 | 1 | -2.166 | -27.532 | 18196 | -1 | 1 | 22-0177 | 22-309 | H2O | 1671007400 | 0.305 | 0.722 | 78.37 | 0.001 | 0.001 | 1.867 | 1296.059 | 1.182 | 2.879 | 113.528 | 0.874 | 0.336 | 0 | 0 | 1 | 45.75 | 1 | 30 | 1 | GJB_Rinse | 0
142 | P-2246 | 2022/12/14 08:47:10 | 2-30 | 2 | -1.999 | -25.691 | 18146 | -1 | 1 | 22-0177 | 22-309 | H2O | 1671007937 | 0.305 | 0.718 | 77.819 | 0 | 0 | 1.859 | 1295.693 | 1.877 | 2.895 | 111.788 | 0.872 | 0.294 | 0 | 0 | 1 | 46 | 1 | 30 | 1 | GJB_Rinse | 0
143 | P-2246 | 2022/12/14 08:56:06 | 2-30 | 3 | -2.003 | -25.33 | 17900 | -1 | 1 | 22-0177 | 22-309 | H2O | 1671008473 | 0.351 | 0.717 | 78.046 | 0.001 | 0.002 | 1.857 | 1294.09 | 1.317 | 2.842 | 110.855 | 0.875 | 0.245 | 0 | 0 | 1 | 46.125 | 1 | 30 | 1 | GJB_Rinse | 0
144 | P-2246 | 2022/12/14 09:05:03 | 2-30 | 4 | -1.966 | -25.209 | 17923 | 0 | 1 | 22-0177 | 22-309 | H2O | 1671009010 | 0.328 | 0.72 | 78.303 | 0.001 | 0.002 | 1.85 | 1294.009 | 1.908 | 2.887 | 110.085 | 0.865 | 0.324 | 0 | 0 | 1 | 44.938 | 1 | 30 | 1 | GJB_Rinse | 0
145 | P-2247 | 2022/12/14 09:14:49 | 2-31 | 1 | -7.217 | -55.14 | 18071 | -1 | 1 | EV | standard | H2O | 1671009596 | 0.331 | 0.808 | 79.926 | 0.002 | 0.004 | 1.905 | 1294.735 | 1.962 | 2.941 | 112.416 | 0.873 | 0.633 | 0 | 0 | 1 | 43.125 | 1 | 31 | 1 | GJB_Rinse | 0
146 | P-2247 | 2022/12/14 09:23:46 | 2-31 | 2 | -7.491 | -58.174 | 18190 | -1 | 1 | EV | standard | H2O | 1671010133 | 0.312 | 0.729 | 78.671 | 0.001 | 0.005 | 1.866 | 1295.856 | 2.204 | 2.856 | 110.845 | 0.878 | 0.226 | 0 | 0 | 1 | 43 | 1 | 31 | 1 | GJB_Rinse | 0
147 | P-2247 | 2022/12/14 09:32:42 | 2-31 | 3 | -7.545 | -58.785 | 18096 | -1 | 1 | EV | standard | H2O | 1671010669 | 0.328 | 0.85 | 80.038 | 0.001 | 0.004 | 1.899 | 1295.337 | 1.797 | 2.878 | 110.894 | 0.876 | 0.231 | 0 | 0 | 1 | 42.812 | 1 | 31 | 1 | GJB_Rinse | 0
148 | P-2247 | 2022/12/14 09:41:39 | 2-31 | 4 | -7.561 | -59.252 | 18304 | 0 | 1 | EV | standard | H2O | 1671011205 | 0.296 | 0.711 | 78.365 | 0.002 | 0.002 | 1.859 | 1296.322 | 1.946 | 2.968 | 114.457 | 0.873 | 0.522 | 0 | 0 | 1 | 42.938 | 1 | 31 | 1 | GJB_Rinse | 0
149 | P-2248 | 2022/12/14 09:51:25 | 2-32 | 1 | -2.758 | -29.838 | 15930 | -1 | 1 | 22-0179 | 22-309 | H2O | 1671011792 | 0.316 | 0.731 | 73.19 | 0.001 | 0.002 | 1.756 | 1279.71 | 1.763 | 2.507 | 98.651 | 0.867 | 0.271 | 0 | 0 | 1 | 42.875 | 1 | 32 | 1 | GJB_Rinse | 0
150 | P-2248 | 2022/12/14 10:00:23 | 2-32 | 2 | -2.251 | -26.156 | 16180 | -1 | 1 | 22-0179 | 22-309 | H2O | 1671012329 | 0.335 | 0.807 | 75.603 | 0.001 | 0.002 | 1.811 | 1281.518 | 1.466 | 2.579 | 100.913 | 0.869 | 0.378 | 0 | 0 | 1 | 42.625 | 1 | 32 | 1 | GJB_Rinse | 0
151 | P-2248 | 2022/12/14 10:09:19 | 2-32 | 3 | -2.493 | -26.487 | 14758 | -1 | 1 | 22-0179 | 22-309 | H2O | 1671012866 | 0.299 | 0.816 | 71.041 | 0 | 0.001 | 1.738 | 1271.765 | 1.579 | 2.32 | 90.386 | 0.862 | 0.124 | 0 | 0 | 1 | 42.625 | 1 | 32 | 1 | GJB_Rinse | 0
152 | P-2248 | 2022/12/14 10:18:17 | 2-32 | 4 | -2.364 | -26.005 | 15367 | 0 | 1 | 22-0179 | 22-309 | H2O | 1671013403 | 0.298 | 0.755 | 70.818 | 0.002 | -0.001 | 1.71 | 1275.938 | 1.858 | 2.409 | 94.038 | 0.862 | 0.166 | 0 | 0 | 1 | 42.5 | 1 | 32 | 1 | GJB_Rinse | 0
153 | P-2249 | 2022/12/14 10:28:04 | 2-33 | 1 | 5.384 | 8.649 | 17244 | -1 | 1 | 22-0197 | 22-309 | H2O | 1671013990 | 0.296 | 0.858 | 75.404 | 0.001 | -0.002 | 1.798 | 1289.293 | 1.662 | 2.724 | 105.399 | 0.868 | 0.049 | 0 | 0 | 1 | 42.5 | 1 | 33 | 1 | GJB_Rinse | 0
154 | P-2249 | 2022/12/14 10:37:00 | 2-33 | 2 | 5.77 | 12.038 | 17488 | -1 | 1 | 22-0197 | 22-309 | H2O | 1671014527 | 0.311 | 0.731 | 78.895 | 0.001 | 0.002 | 1.883 | 1290.925 | 1.744 | 2.761 | 107.524 | 0.867 | 0.226 | 0 | 0 | 1 | 42.438 | 1 | 33 | 1 | GJB_Rinse | 0
155 | P-2249 | 2022/12/14 10:45:57 | 2-33 | 3 | 5.777 | 12.751 | 17510 | -1 | 1 | 22-0197 | 22-309 | H2O | 1671015063 | 0.307 | 0.735 | 78.085 | 0 | -0.002 | 1.858 | 1291.273 | 1.413 | 2.848 | 107.773 | 0.864 | 0.35 | 0 | 0 | 1 | 42.375 | 1 | 33 | 1 | GJB_Rinse | 0
156 | P-2249 | 2022/12/14 10:54:54 | 2-33 | 4 | 5.814 | 13.001 | 17525 | 0 | 1 | 22-0197 | 22-309 | H2O | 1671015601 | 0.316 | 0.736 | 76.853 | 0.001 | -0.001 | 1.828 | 1290.964 | 1.584 | 2.871 | 109.642 | 0.865 | 0.58 | 0 | 0 | 1 | 42.25 | 1 | 33 | 1 | GJB_Rinse | 0
157 | P-2250 | 2022/12/14 11:04:41 | 2-34 | 1 | -1.735 | -19.798 | 17019 | -1 | 1 | 22-0178 | 22-309 | H2O | 1671016188 | 0.338 | 0.677 | 77.494 | 0 | 0.005 | 1.848 | 1287.558 | 1.681 | 2.692 | 104.907 | 0.863 | 0.35 | 0 | 0 | 1 | 42.188 | 1 | 34 | 1 | GJB_Rinse | 0
158 | P-2250 | 2022/12/14 11:13:39 | 2-34 | 2 | -2.073 | -22.93 | 17058 | -1 | 1 | 22-0178 | 22-309 | H2O | 1671016725 | 0.309 | 0.74 | 72.177 | 0.001 | 0.002 | 1.735 | 1287.695 | 1.764 | 2.754 | 105.72 | 0.865 | 0.441 | 0 | 0 | 1 | 42.25 | 1 | 34 | 1 | GJB_Rinse | 0
159 | P-2250 | 2022/12/14 11:22:35 | 2-34 | 3 | -2.227 | -24.09 | 16941 | -1 | 1 | 22-0178 | 22-309 | H2O | 1671017262 | 0.296 | 0.8 | 76.08 | 0.001 | 0.004 | 1.821 | 1286.867 | 1.836 | 2.683 | 104.823 | 0.865 | 0.327 | 0 | 0 | 1 | 42.25 | 1 | 34 | 1 | GJB_Rinse | 0
160 | P-2250 | 2022/12/14 11:31:31 | 2-34 | 4 | -2.275 | -24.141 | 17074 | 0 | 1 | 22-0178 | 22-309 | H2O | 1671017798 | 0.306 | 0.773 | 75.682 | 0.001 | 0 | 1.821 | 1288.022 | 1.638 | 2.726 | 105.529 | 0.86 | 0.299 | 0 | 0 | 1 | 42.188 | 1 | 34 | 1 | GJB_Rinse | 0
161 | P-2251 | 2022/12/14 11:41:19 | 2-35 | 1 | -3 | -28.089 | 17959 | -1 | 1 | 22-0172 | 22-309 | H2O | 1671018386 | 0.295 | 0.668 | 76.337 | 0 | 0 | 1.81 | 1294.337 | 1.497 | 2.899 | 111.345 | 0.867 | 0.365 | 0 | 0 | 1 | 42.25 | 1 | 35 | 1 | GJB_Rinse | 0
162 | P-2251 | 2022/12/14 11:50:17 | 2-35 | 2 | -3.159 | -28.666 | 18028 | -1 | 1 | 22-0172 | 22-309 | H2O | 1671018923 | 0.3 | 0.725 | 78.235 | 0 | 0.001 | 1.867 | 1294.385 | 1.93 | 2.946 | 112.803 | 0.867 | 0.696 | 0 | 0 | 1 | 42.312 | 1 | 35 | 1 | GJB_Rinse | 0
163 | P-2251 | 2022/12/14 11:59:13 | 2-35 | 3 | -3.126 | -28.552 | 17915 | -1 | 1 | 22-0172 | 22-309 | H2O | 1671019460 | 0.311 | 0.701 | 80.323 | 0 | 0 | 1.918 | 1293.698 | 2.151 | 2.882 | 110.718 | 0.872 | 0.438 | 0 | 0 | 1 | 42.312 | 1 | 35 | 1 | GJB_Rinse | 0
164 | P-2251 | 2022/12/14 12:08:11 | 2-35 | 4 | -3.114 | -28.8 | 17713 | 0 | 1 | 22-0172 | 22-309 | H2O | 1671019998 | 0.288 | 0.777 | 77.665 | 0.001 | 0.002 | 1.852 | 1292.502 | 1.657 | 2.82 | 109.298 | 0.868 | 0.187 | 0 | 0 | 1 | 42.375 | 1 | 35 | 1 | GJB_Rinse | 0
165 | P-2252 | 2022/12/14 12:17:55 | 2-36 | 1 | -6.892 | -55.672 | 16204 | -1 | 1 | 22-0690 | 22-309 | H2O | 1671020581 | 0.284 | 0.776 | 74.613 | 0.001 | 0.005 | 1.789 | 1281.724 | 2.106 | 2.524 | 98.633 | 0.863 | 0.183 | 0 | 0 | 1 | 42.375 | 1 | 36 | 1 | GJB_Rinse | 0
166 | P-2252 | 2022/12/14 12:26:51 | 2-36 | 2 | -6.902 | -57.699 | 16905 | -1 | 1 | 22-0690 | 22-309 | H2O | 1671021118 | 0.298 | 0.766 | 75.635 | 0.001 | 0.006 | 1.81 | 1286.479 | 2.463 | 2.68 | 103.08 | 0.865 | 0.239 | 0 | 0 | 1 | 42.438 | 1 | 36 | 1 | GJB_Rinse | 0
167 | P-2252 | 2022/12/14 12:35:48 | 2-36 | 3 | -7.125 | -59.082 | 15816 | -1 | 1 | 22-0690 | 22-309 | H2O | 1671021655 | 0.3 | 0.761 | 72.656 | 0.001 | 0.002 | 1.755 | 1278.922 | 1.721 | 2.505 | 98.029 | 0.862 | 0.45 | 0 | 0 | 1 | 42.312 | 1 | 36 | 1 | GJB_Rinse | 0
168 | P-2252 | 2022/12/14 12:44:46 | 2-36 | 4 | -7.151 | -59.26 | 15687 | 0 | 1 | 22-0690 | 22-309 | H2O | 1671022193 | 0.318 | 0.804 | 72.613 | 0.001 | 0.006 | 1.756 | 1278.006 | 1.779 | 2.483 | 97.458 | 0.861 | 0.424 | 0 | 0 | 1 | 42.375 | 1 | 36 | 1 | GJB_Rinse | 0
169 | P-2253 | 2022/12/14 12:54:30 | 2-37 | 1 | [5 empty] | 22-0181 | 22-309 | [18 empty] | 1 | 37 | 1 | GJB_Rinse | 0
170 | P-2253 | 2022/12/14 13:03:26 | 2-37 | 2 | [5 empty] | 22-0181 | 22-309 | [18 empty] | 1 | 37 | 1 | GJB_Rinse | 0
171 | P-2253 | 2022/12/14 13:12:23 | 2-37 | 3 | 9.047 | 41.738 | 15204 | -1 | 1 | 22-0181 | 22-309 | H2O | 1671023850 | 0.31 | 0.814 | 71.069 | 0.001 | -0.003 | 1.734 | 1274.773 | 0.788 | 2.442 | 96.098 | 0.864 | 0.449 | 0 | 0 | 1 | 42.562 | 1 | 37 | 1 | GJB_Rinse | 0
172 | P-2253 | 2022/12/14 13:21:20 | 2-37 | 4 | 9.047 | 41.738 | 15204 | -1 | 1 | 22-0181 | 22-309 | H2O | 1671023850 | 0.31 | 0.814 | 71.069 | 0.001 | -0.003 | 1.734 | 1274.773 | 0.788 | 2.442 | 96.098 | 0.864 | 0.449 | 0 | 0 | 1 | 42.562 | 1 | 37 | 1 | GJB_Rinse | 0
173 | P-2254 | 2022/12/14 13:31:05 | 2-38 | 1 | -0.596 | -11.692 | 14725 | -1 | 1 | 22-0202 | 22-309 | H2O | 1671024971 | 0.331 | 0.77 | 69.209 | 0.001 | 0.003 | 1.697 | 1271.705 | 1.244 | 2.308 | 91.111 | 0.865 | 0.108 | 0 | 0 | 1 | 42.375 | 1 | 38 | 1 | GJB_Rinse | 0
174 | P-2254 | 2022/12/14 13:40:02 | 2-38 | 2 | -1.043 | -16.765 | 15922 | -1 | 1 | 22-0202 | 22-309 | H2O | 1671025508 | 0.323 | 0.82 | 73.087 | 0.001 | 0.003 | 1.752 | 1279.858 | 1.467 | 2.527 | 99.027 | 0.863 | 0.302 | 0 | 0 | 1 | 42.25 | 1 | 38 | 1 | GJB_Rinse | 0
175 | P-2254 | 2022/12/14 13:48:57 | 2-38 | 3 | -1.228 | -18.683 | 15025 | -1 | 1 | 22-0202 | 22-309 | H2O | 1671026044 | 0.31 | 0.799 | 71.011 | 0.002 | 0.004 | 1.725 | 1273.778 | 1.707 | 2.325 | 91.947 | 0.862 | 0.056 | 0 | 0 | 1 | 42.5 | 1 | 38 | 1 | GJB_Rinse | 0
176 | P-2254 | 2022/12/14 13:57:54 | 2-38 | 4 | -1.331 | -19.254 | 15165 | 0 | 1 | 22-0202 | 22-309 | H2O | 1671026581 | 0.295 | 0.764 | 72.432 | 0.001 | 0 | 1.763 | 1274.614 | 1.841 | 2.336 | 92.641 | 0.871 | 0.074 | 0 | 0 | 1 | 42.5 | 1 | 38 | 1 | GJB_Rinse | 0
177 | P-2255 | 2022/12/14 14:07:40 | 2-39 | 1 | -2.598 | -32.291 | 18559 | -1 | 1 | 22-0193 | 22-309 | H2O | 1671027167 | 0.328 | 0.681 | 76.869 | 0.001 | 0.003 | 1.824 | 1298.455 | 1.976 | 3.024 | 114.777 | 0.867 | 0.609 | 0 | 0 | 1 | 42.438 | 1 | 39 | 1 | GJB_Rinse | 0
178 | P-2255 | 2022/12/14 14:16:38 | 2-39 | 2 | -2.73 | -33.945 | 18504 | -1 | 1 | 22-0193 | 22-309 | H2O | 1671027704 | 0.339 | 0.831 | 78.33 | 0.002 | 0.004 | 1.868 | 1298.053 | 1.799 | 2.96 | 114.546 | 0.863 | 0.429 | 0 | 0 | 1 | 42.375 | 1 | 39 | 1 | GJB_Rinse | 0
179 | P-2255 | 2022/12/14 14:25:35 | 2-39 | 3 | -2.778 | -34.89 | 18181 | -1 | 1 | 22-0193 | 22-309 | H2O | 1671028241 | 0.296 | 0.794 | 77.998 | 0.001 | 0.004 | 1.862 | 1296.022 | 1.525 | 2.918 | 112.24 | 0.877 | 0.186 | 0 | 0 | 1 | 42.375 | 1 | 39 | 1 | GJB_Rinse | 0
180 | P-2255 | 2022/12/14 14:34:32 | 2-39 | 4 | -2.806 | -35.003 | 18071 | 0 | 1 | 22-0193 | 22-309 | H2O | 1671028778 | 0.306 | 0.802 | 78.381 | 0 | 0.006 | 1.864 | 1294.972 | 1.749 | 2.918 | 112.469 | 0.87 | 0.447 | 0 | 0 | 1 | 42.688 | 1 | 39 | 1 | GJB_Rinse | 0
181 | P-2256 | 2022/12/14 14:44:17 | 2-40 | 1 | -7.182 | -56.013 | 18000 | -1 | 1 | EV | standard | H2O | 1671029363 | 0.341 | 0.82 | 80.048 | 0.003 | 0.008 | 1.922 | 1294.462 | 1.957 | 2.878 | 111.012 | 0.873 | 0.257 | 0 | 0 | 1 | 42.562 | 1 | 40 | 1 | GJB_Rinse | 0
182 | P-2256 | 2022/12/14 14:53:13 | 2-40 | 2 | -7.485 | -58.389 | 17863 | -1 | 1 | EV | standard | H2O | 1671029900 | 0.314 | 0.74 | 80.182 | 0.001 | 0.002 | 1.906 | 1293.412 | 2.07 | 2.836 | 110.304 | 0.87 | 0.342 | 0 | 0 | 1 | 42.5 | 1 | 40 | 1 | GJB_Rinse | 0
183 | P-2256 | 2022/12/14 15:02:10 | 2-40 | 3 | -7.496 | -58.674 | 17891 | -1 | 1 | EV | standard | H2O | 1671030437 | 0.313 | 0.754 | 78.329 | 0.001 | 0.004 | 1.862 | 1293.753 | 2.087 | 2.822 | 110.09 | 0.875 | 0.241 | 0 | 0 | 1 | 42.625 | 1 | 40 | 1 | GJB_Rinse | 0
184 | P-2256 | 2022/12/14 15:11:07 | 2-40 | 4 | -7.572 | -59.277 | 17915 | 0 | 1 | EV | standard | H2O | 1671030974 | 0.313 | 0.754 | 77.723 | 0.002 | 0.004 | 1.859 | 1293.819 | 2.13 | 2.851 | 110.137 | 0.87 | 0.309 | 0 | 0 | 1 | 42.5 | 1 | 40 | 1 | GJB_Rinse | 0
185 | P-2257 | 2022/12/14 15:20:53 | 2-41 | 1 | -2.759 | -30.987 | 16677 | -1 | 1 | 22-0186 | 22-309 | H2O | 1671032633 | 0.31 | 0.795 | 72.849 | 0.001 | -0.001 | 1.74 | 1285.309 | 1.218 | 2.696 | 103.932 | 0.876 | 0.321 | 0 | 0 | 1 | 42.625 | 1 | 41 | 1 | GJB_Rinse | 0
186 | P-2257 | 2022/12/14 15:29:50 | 2-41 | 2 | [5 empty] | 22-0186 | 22-309 | [18 empty] | 1 | 41 | 1 | GJB_Rinse | 0
187 | P-2257 | 2022/12/14 15:38:46 | 2-41 | 3 | -2.759 | -30.987 | 16677 | -1 | 1 | 22-0186 | 22-309 | H2O | 1671032633 | 0.31 | 0.795 | 72.849 | 0.001 | -0.001 | 1.74 | 1285.309 | 1.218 | 2.696 | 103.932 | 0.876 | 0.321 | 0 | 0 | 1 | 42.625 | 1 | 41 | 1 | GJB_Rinse | 0
188 | P-2257 | 2022/12/14 15:47:44 | 2-41 | 4 | -2.908 | -31.84 | 15580 | 0 | 1 | 22-0186 | 22-309 | H2O | 1671033170 | 0.308 | 0.767 | 71.988 | 0 | 0 | 1.74 | 1277.297 | 1.216 | 2.536 | 98.374 | 0.866 | 0.619 | 0 | 0 | 1 | 42.688 | 1 | 41 | 1 | GJB_Rinse | 0
189 | P-2258 | 2022/12/14 15:57:29 | 2-42 | 1 | -6.615 | -53.961 | 16969 | -1 | 1 | 22-0682 | 22-309 | H2O | 1671033756 | 0.321 | 0.816 | 75.521 | 0.001 | 0.006 | 1.809 | 1287.175 | 1.542 | 2.747 | 105.567 | 0.864 | 0.443 | 0 | 0 | 1 | 42.5 | 1 | 42 | 1 | GJB_Rinse | 0
190 | P-2258 | 2022/12/14 16:06:26 | 2-42 | 2 | -6.758 | -55.784 | 16997 | -1 | 1 | 22-0682 | 22-309 | H2O | 1671034294 | 0.321 | 0.775 | 75.344 | 0.001 | 0.002 | 1.804 | 1287.448 | 1.742 | 2.746 | 105.144 | 0.864 | 0.41 | 0 | 0 | 1 | 42.438 | 1 | 42 | 1 | GJB_Rinse | 0
191 | P-2258 | 2022/12/14 16:15:24 | 2-42 | 3 | -6.858 | -56.705 | 17098 | -1 | 1 | 22-0682 | 22-309 | H2O | 1671034831 | 0.262 | 0.768 | 76.069 | 0.001 | 0.004 | 1.814 | 1287.978 | 1.957 | 2.706 | 105.904 | 0.868 | 0.426 | 0 | 0 | 1 | 42.5 | 1 | 42 | 1 | GJB_Rinse | 0
192 | P-2258 | 2022/12/14 16:24:20 | 2-42 | 4 | -6.909 | -56.9 | 16972 | 0 | 1 | 22-0682 | 22-309 | H2O | 1671035367 | 0.303 | 0.824 | 77.287 | 0.001 | 0.005 | 1.853 | 1287.251 | 1.273 | 2.713 | 106.114 | 0.869 | 0.411 | 0 | 0 | 1 | 42.5 | 1 | 42 | 1 | GJB_Rinse | 0
193 | P-2259 | 2022/12/14 16:34:05 | 2-43 | 1 | -2.593 | -23.386 | 16790 | -1 | 1 | 22-0174 | 22-309 | H2O | 1671035952 | 0.331 | 0.727 | 74.084 | 0.001 | -0.002 | 1.78 | 1286.097 | 1.176 | 2.664 | 104.761 | 0.866 | 0.528 | 0 | 0 | 1 | 42.438 | 1 | 43 | 1 | GJB_Rinse | 0
194 | P-2259 | 2022/12/14 16:43:02 | 2-43 | 2 | -2.142 | -18.81 | 17727 | -1 | 1 | 22-0174 | 22-309 | H2O | 1671036488 | 0.301 | 0.733 | 77.796 | 0.001 | 0.001 | 1.854 | 1292.65 | 1.525 | 2.883 | 110.324 | 0.873 | 0.324 | 0 | 0 | 1 | 42.375 | 1 | 43 | 1 | GJB_Rinse | 0
195 | P-2259 | 2022/12/14 16:51:59 | 2-43 | 3 | -2.219 | -18.539 | 17581 | -1 | 1 | 22-0174 | 22-309 | H2O | 1671037025 | 0.307 | 0.8 | 75.723 | 0.001 | 0.001 | 1.812 | 1291.556 | 1.551 | 2.832 | 108.877 | 0.875 | 0.39 | 0 | 0 | 1 | 42.5 | 1 | 43 | 1 | GJB_Rinse | 0
196 | P-2259 | 2022/12/14 17:00:55 | 2-43 | 4 | -2.276 | -18.971 | 17449 | 0 | 1 | 22-0174 | 22-309 | H2O | 1671037562 | 0.346 | 0.718 | 74.121 | 0.001 | 0.001 | 1.779 | 1290.693 | 1.576 | 2.788 | 107.749 | 0.867 | 0.21 | 0 | 0 | 1 | 42.5 | 1 | 43 | 1 | GJB_Rinse | 0
197 | P-2260 | 2022/12/14 17:10:41 | 2-44 | 1 | -7.469 | -61.536 | 16994 | -1 | 1 | 22-0195 | 22-309 | H2O | 1671038148 | 0.307 | 0.709 | 77.279 | 0.001 | 0.003 | 1.844 | 1287.4 | 1.857 | 2.668 | 104.163 | 0.863 | 0.267 | 0 | 0 | 1 | 42.5 | 1 | 44 | 1 | GJB_Rinse | 0
198 | P-2260 | 2022/12/14 17:19:39 | 2-44 | 2 | -7.448 | -64.846 | 17978 | -1 | 1 | 22-0195 | 22-309 | H2O | 1671038686 | 0.289 | 0.736 | 77.447 | 0 | 0.004 | 1.84 | 1294.356 | 1.439 | 2.941 | 112.386 | 0.871 | 0.6 | 0 | 0 | 1 | 42.5 | 1 | 44 | 1 | GJB_Rinse | 0
199 | P-2260 | 2022/12/14 17:28:36 | 2-44 | 3 | -7.817 | -66.939 | 16596 | -1 | 1 | 22-0195 | 22-309 | H2O | 1671039222 | 0.316 | 0.763 | 73.643 | 0 | 0.004 | 1.764 | 1284.472 | 2.24 | 2.586 | 101.09 | 0.869 | 0.079 | 0 | 0 | 1 | 42.625 | 1 | 44 | 1 | GJB_Rinse | 0
200 | P-2260 | 2022/12/14 17:37:32 | 2-44 | 4 | -7.825 | -67.354 | 16828 | 0 | 1 | 22-0195 | 22-309 | H2O | 1671039759 | 0.294 | 0.751 | 75.541 | 0 | 0.004 | 1.813 | 1286.236 | 1.83 | 2.646 | 103.5 | 0.868 | 0.237 | 0 | 0 | 1 | 42.625 | 1 | 44 | 1 | GJB_Rinse | 0
201 | P-2263 | 2022/12/15 11:30:58 | 2-45 | 1 | -3.853 | -40.975 | 18195 | -1 | 1 | 22-0188 | 22-309 | H2O | 1671104165 | 0.266 | 0.734 | 74.557 | 0.001 | 0.001 | 1.771 | 1297.622 | 1.776 | 2.926 | 113.014 | 0.871 | 0.616 | 0 | 0 | 1 | 42.188 | 1 | 45 | 1 | GJB_Rinse | 0
202 | P-2263 | 2022/12/15 11:39:54 | 2-45 | 2 | -3.577 | -37.955 | 18176 | -1 | 1 | 22-0188 | 22-309 | H2O | 1671104701 | 0.33 | 0.752 | 74.926 | 0 | -0.003 | 1.788 | 1299.157 | 2.02 | 2.882 | 110.593 | 0.873 | 0.212 | 0 | 0 | 1 | 42.125 | 1 | 45 | 1 | GJB_Rinse | 0
203 | P-2263 | 2022/12/15 11:48:51 | 2-45 | 3 | -3.654 | -37.659 | 18053 | -1 | 1 | 22-0188 | 22-309 | H2O | 1671105237 | 0.292 | 0.691 | 74.421 | 0.001 | -0.002 | 1.77 | 1297.112 | 1.463 | 2.863 | 110.733 | 0.864 | 0.3 | 0 | 0 | 1 | 42.312 | 1 | 45 | 1 | GJB_Rinse | 0
204 | P-2263 | 2022/12/15 11:57:47 | 2-45 | 4 | -3.588 | -37.339 | 18087 | 0 | 1 | 22-0188 | 22-309 | H2O | 1671105773 | 0.318 | 0.754 | 73.948 | 0.001 | 0.001 | 1.757 | 1296.195 | 1.827 | 2.961 | 112.104 | 0.87 | 0.537 | 0 | 0 | 1 | 42.375 | 1 | 45 | 1 | GJB_Rinse | 0
205 | P-2264 | 2022/12/15 12:07:33 | 2-46 | 1 | -7.111 | -58.013 | 17848 | -1 | 1 | 22-0695 | 22-309 | H2O | 1671106359 | 0.294 | 0.727 | 74.784 | 0.002 | 0.002 | 1.779 | 1293.523 | 2.224 | 2.825 | 109.118 | 0.87 | 0.37 | 0 | 0 | 1 | 42.688 | 1 | 46 | 1 | GJB_Rinse | 0
206 | P-2264 | 2022/12/15 12:16:29 | 2-46 | 2 | -7.154 | -59.354 | 18452 | -1 | 1 | 22-0695 | 22-309 | H2O | 1671106896 | 0.317 | 0.789 | 77.855 | 0.002 | 0.003 | 1.843 | 1298.084 | 2.284 | 2.987 | 113.902 | 0.874 | 0.526 | 0 | 0 | 1 | 42.5 | 1 | 46 | 1 | GJB_Rinse | 0
207 | P-2264 | 2022/12/15 12:25:26 | 2-46 | 3 | -7.271 | -59.947 | 18253 | -1 | 1 | 22-0695 | 22-309 | H2O | 1671107433 | 0.305 | 0.759 | 75.811 | 0.002 | 0.004 | 1.798 | 1296.615 | 1.878 | 2.935 | 113.242 | 0.867 | 0.683 | 0 | 0 | 1 | 42.375 | 1 | 46 | 1 | GJB_Rinse | 0
208 | P-2264 | 2022/12/15 12:34:22 | 2-46 | 4 | -7.39 | -60.644 | 18130 | 0 | 1 | 22-0695 | 22-309 | H2O | 1671107968 | 0.286 | 0.767 | 77.025 | 0 | 0.004 | 1.839 | 1295.833 | 1.584 | 2.902 | 112.19 | 0.872 | 0.425 | 0 | 0 | 1 | 42.25 | 1 | 46 | 1 | GJB_Rinse | 0
209 | P-2265 | 2022/12/15 12:44:08 | 2-47 | 1 | 4.495 | 7.861 | 17324 | -1 | 1 | 22-0192 | 22-309 | H2O | 1671108554 | 0.327 | 0.822 | 74.603 | 0.001 | -0.004 | 1.782 | 1289.661 | 1.628 | 2.772 | 107.286 | 0.87 | 0.343 | 0 | 0 | 1 | 42.25 | 1 | 47 | 1 | GJB_Rinse | 0
210 | P-2265 | 2022/12/15 12:53:03 | 2-47 | 2 | 5.263 | 15.828 | 17743 | -1 | 1 | 22-0192 | 22-309 | H2O | 1671109090 | 0.312 | 0.778 | 76.629 | 0 | -0.001 | 1.829 | 1292.488 | 1.709 | 2.889 | 110.702 | 0.867 | 0.78 | 0 | 0 | 1 | 42.312 | 1 | 47 | 1 | GJB_Rinse | 0
211 | P-2265 | 2022/12/15 13:01:59 | 2-47 | 3 | 5.261 | 17.406 | 17957 | -1 | 1 | 22-0192 | 22-309 | H2O | 1671109625 | 0.326 | 0.756 | 75.624 | 0 | -0.002 | 1.805 | 1294.346 | 1.68 | 2.857 | 110.899 | 0.862 | 0.289 | 0 | 0 | 1 | 42.312 | 1 | 47 | 1 | GJB_Rinse | 0
212 | P-2265 | 2022/12/15 13:10:55 | 2-47 | 4 | 5.327 | 17.855 | 17897 | 0 | 1 | 22-0192 | 22-309 | H2O | 1671110162 | 0.293 | 0.679 | 76.931 | 0.001 | -0.002 | 1.826 | 1294.09 | 1.379 | 2.892 | 109.421 | 0.866 | 0.193 | 0 | 0 | 1 | 42.125 | 1 | 47 | 1 | GJB_Rinse | 0
213 | P-2266 | 2022/12/15 13:20:42 | 2-48 | 1 | -5.594 | -41.996 | 16461 | -1 | 1 | 22-0689 | 22-309 | H2O | 1671110747 | 0.319 | 0.781 | 72.544 | 0.002 | 0.004 | 1.769 | 1283.553 | 1.614 | 2.579 | 101.39 | 0.864 | 0.338 | 0 | 0 | 1 | 42.125 | 1 | 48 | 1 | GJB_Rinse | 0
214 | P-2266 | 2022/12/15 13:29:39 | 2-48 | 2 | -5.964 | -47.317 | 17450 | -1 | 1 | 22-0689 | 22-309 | H2O | 1671111286 | 0.311 | 0.759 | 71.582 | 0.001 | 0.001 | 1.716 | 1290.524 | 1.745 | 2.827 | 108.586 | 0.862 | 0.634 | 0 | 0 | 1 | 43 | 1 | 48 | 1 | GJB_Rinse | 0
215 | P-2266 | 2022/12/15 13:38:35 | 2-48 | 3 | -6.187 | -48.866 | 17888 | -1 | 1 | 22-0689 | 22-309 | H2O | 1671111821 | 0.303 | 0.724 | 75.012 | 0.002 | 0.004 | 1.788 | 1293.863 | 2.243 | 2.791 | 108.273 | 0.869 | 0.079 | 0 | 0 | 1 | 45 | 1 | 48 | 1 | GJB_Rinse | 0
216 | P-2266 | 2022/12/15 13:47:32 | 2-48 | 4 | -6.2 | -49.215 | 18167 | 0 | 1 | 22-0689 | 22-309 | H2O | 1671112358 | 0.294 | 0.809 | 78.949 | 0.001 | 0.005 | 1.869 | 1295.904 | 1.952 | 2.858 | 110.333 | 0.869 | 0.124 | 0 | 0 | 1 | 45.438 | 1 | 48 | 1 | GJB_Rinse | 0
217 | P-2267 | 2022/12/15 13:57:18 | 2-49 | 1 | -7.403 | -57.745 | 18349 | -1 | 1 | EV | standard | H2O | 1671112945 | 0.282 | 0.782 | 78.829 | 0.001 | 0.006 | 1.876 | 1297.031 | 1.863 | 2.917 | 112.536 | 0.874 | 0.279 | 0 | 0 | 1 | 45.438 | 1 | 49 | 1 | GJB_Rinse | 0
218 | P-2267 | 2022/12/15 14:06:14 | 2-49 | 2 | -7.454 | -58.502 | 18677 | -1 | 1 | EV | standard | H2O | 1671113481 | 0.293 | 0.74 | 78.685 | 0.002 | 0.005 | 1.867 | 1299.435 | 1.859 | 3 | 115.316 | 0.872 | 0.472 | 0 | 0 | 1 | 44.125 | 1 | 49 | 1 | GJB_Rinse | 0
219 | P-2267 | 2022/12/15 14:15:11 | 2-49 | 3 | -7.5 | -58.661 | 18397 | -1 | 1 | EV | standard | H2O | 1671114018 | 0.317 | 0.794 | 79.126 | 0.002 | 0.005 | 1.885 | 1297.353 | 1.768 | 2.972 | 114.208 | 0.868 | 0.545 | 0 | 0 | 1 | 44.25 | 1 | 49 | 1 | GJB_Rinse | 0
220 | P-2267 | 2022/12/15 14:24:07 | 2-49 | 4 | -7.511 | -59.222 | 18508 | 0 | 1 | EV | standard | H2O | 1671114553 | 0.326 | 0.722 | 82.268 | 0.001 | 0.004 | 1.965 | 1298.184 | 2.183 | 2.904 | 113.132 | 0.872 | 0.106 | 0 | 0 | 1 | 45 | 1 | 49 | 1 | GJB_Rinse | 0
221 | P-2268 | 2022/12/15 14:33:51 | 2-50 | 1 | [5 empty] | 22-0199 | 22-309 | [18 empty] | 1 | 50 | 1 | GJB_Rinse | 0
222 | P-2268 | 2022/12/15 14:42:47 | 2-50 | 2 | 6.919 | 16.727 | 18915 | -1 | 1 | 22-0199 | 22-309 | H2O | 1671115673 | 0.359 | 0.807 | 80.426 | 0 | -0.001 | 1.912 | 1301.242 | 1.495 | 3.037 | 117.017 | 0.881 | 0.364 | 0 | 0 | 1 | 42.688 | 1 | 50 | 1 | GJB_Rinse | 0
223 | P-2268 | 2022/12/15 14:51:42 | 2-50 | 3 | 6.724 | 17.034 | 17796 | -1 | 1 | 22-0199 | 22-309 | H2O | 1671116209 | 0.316 | 0.762 | 78.55 | 0 | -0.001 | 1.88 | 1293.108 | 1.662 | 2.832 | 110.354 | 0.874 | 0.383 | 0 | 0 | 1 | 42.625 | 1 | 50 | 1 | GJB_Rinse | 0
224 | P-2268 | 2022/12/15 15:00:38 | 2-50 | 4 | 6.716 | 17.451 | 17819 | 0 | 1 | 22-0199 | 22-309 | H2O | 1671116745 | 0.284 | 0.756 | 77.348 | 0 | 0 | 1.846 | 1293.445 | 1.035 | 2.877 | 110.668 | 0.867 | 0.375 | 0 | 0 | 1 | 42.438 | 1 | 50 | 1 | GJB_Rinse | 0
225 | P-2269 | 2022/12/15 15:10:24 | 2-51 | 1 | [5 empty] | 22-0189 | 22-309 | [18 empty] | 1 | 51 | 1 | GJB_Rinse | 0
226 | P-2269 | 2022/12/15 15:19:20 | 2-51 | 2 | -4.383 | -40.863 | 15949 | 0 | 1 | 22-0189 | 22-309 | H2O | 1671118938 | 0.296 | 0.764 | 74.834 | 0 | 0.004 | 1.796 | 1280.037 | 1.554 | 2.489 | 97.862 | 0.863 | 0.205 | 0 | 0 | 1 | 42.438 | 1 | 51 | 1 | GJB_Rinse | 0
227 | P-2269 | 2022/12/15 15:28:15 | 2-51 | 3 | -4.383 | -40.863 | 15949 | 0 | 1 | 22-0189 | 22-309 | H2O | 1671118938 | 0.296 | 0.764 | 74.834 | 0 | 0.004 | 1.796 | 1280.037 | 1.554 | 2.489 | 97.862 | 0.863 | 0.205 | 0 | 0 | 1 | 42.438 | 1 | 51 | 1 | GJB_Rinse | 0
228 | P-2269 | 2022/12/15 15:37:12 | 2-51 | 4 | -4.383 | -40.863 | 15949 | 0 | 1 | 22-0189 | 22-309 | H2O | 1671118938 | 0.296 | 0.764 | 74.834 | 0 | 0.004 | 1.796 | 1280.037 | 1.554 | 2.489 | 97.862 | 0.863 | 0.205 | 0 | 0 | 1 | 42.438 | 1 | 51 | 1 | GJB_Rinse | 0
229 | P-2270 | 2022/12/15 15:46:58 | 2-52 | 1 | 12.429 | 42.888 | 17515 | -1 | 1 | 22-0203 | 22-309 | H2O | 1671119524 | 0.306 | 0.732 | 78.144 | 0 | -0.005 | 1.865 | 1291.079 | 1.001 | 2.857 | 109.262 | 0.867 | 0.557 | 0 | 0 | 1 | 42.312 | 1 | 52 | 1 | GJB_Rinse | 0
230 | P-2270 | 2022/12/15 15:55:53 | 2-52 | 2 | 13.461 | 51.812 | 17781 | -1 | 1 | 22-0203 | 22-309 | H2O | 1671120060 | 0.322 | 0.828 | 77.01 | 0 | -0.006 | 1.836 | 1293.164 | 0.704 | 2.905 | 110.927 | 0.867 | 0.367 | 0 | 0 | 1 | 42.25 | 1 | 52 | 1 | GJB_Rinse | 0
231 | P-2270 | 2022/12/15 16:04:49 | 2-52 | 3 | 13.456 | 53.544 | 17620 | -1 | 1 | 22-0203 | 22-309 | H2O | 1671120596 | 0.358 | 0.705 | 76.034 | 0.001 | -0.002 | 1.818 | 1292.156 | 0.819 | 2.838 | 108.618 | 0.866 | 0.269 | 0 | 0 | 1 | 42.188 | 1 | 52 | 1 | GJB_Rinse | 0
232 | P-2270 | 2022/12/15 16:13:45 | 2-52 | 4 | 13.643 | 54.692 | 17740 | 0 | 1 | 22-0203 | 22-309 | H2O | 1671121131 | 0.315 | 0.774 | 77.612 | 0.001 | -0.002 | 1.845 | 1292.631 | 0.808 | 2.915 | 112 | 0.87 | 0.671 | 0 | 0 | 1 | 42.125 | 1 | 52 | 1 | GJB_Rinse | 0
233 | P-2271 | 2022/12/15 16:23:31 | 2-53 | 1 | -5.75 | -43.865 | 17305 | -1 | 1 | 22-0680 | 22-309 | H2O | 1671121719 | 0.307 | 0.733 | 79.236 | 0.001 | 0.008 | 1.894 | 1289.896 | 1.182 | 2.728 | 106.829 | 0.871 | 0.172 | 0 | 0 | 1 | 42.25 | 1 | 53 | 1 | GJB_Rinse | 0
234 | P-2271 | 2022/12/15 16:32:28 | 2-53 | 2 | -6.629 | -53.188 | 17888 | -1 | 1 | 22-0680 | 22-309 | H2O | 1671122254 | 0.313 | 0.827 | 78.993 | 0.002 | 0.011 | 1.887 | 1293.585 | 2.568 | 2.812 | 108.836 | 0.871 | 0.338 | 0 | 0 | 1 | 42.312 | 1 | 53 | 1 | GJB_Rinse | 0
235 | P-2271 | 2022/12/15 16:41:23 | 2-53 | 3 | -6.944 | -55.764 | 17624 | -1 | 1 | 22-0680 | 22-309 | H2O | 1671122790 | 0.317 | 0.724 | 76.687 | 0.001 | 0.005 | 1.83 | 1291.709 | 1.707 | 2.8 | 109.48 | 0.866 | 0.476 | 0 | 0 | 1 | 42.25 | 1 | 53 | 1 | GJB_Rinse | 0
236 | P-2271 | 2022/12/15 16:50:19 | 2-53 | 4 | -7.091 | -56.705 | 17470 | 0 | 1 | 22-0680 | 22-309 | H2O | 1671123326 | 0.304 | 0.714 | 75.753 | 0.001 | 0.004 | 1.794 | 1290.72 | 1.895 | 2.764 | 107.074 | 0.868 | 0.411 | 0 | 0 | 1 | 42.188 | 1 | 53 | 1 | GJB_Rinse | 0
237 | P-2272 | 2022/12/15 17:00:06 | 2-54 | 1 | -4.437 | -42.758 | 17825 | -1 | 1 | 22-0196 | 22-309 | H2O | 1671123912 | 0.312 | 0.772 | 76.922 | 0.001 | 0.002 | 1.838 | 1293.239 | 1.626 | 2.839 | 111.055 | 0.872 | 0.472 | 0 | 0 | 1 | 42.188 | 1 | 54 | 1 | GJB_Rinse | 0
238 | P-2272 | 2022/12/15 17:09:02 | 2-54 | 2 | -4.321 | -41.596 | 17721 | -1 | 1 | 22-0196 | 22-309 | H2O | 1671124448 | 0.339 | 0.764 | 77.428 | 0.002 | 0.002 | 1.843 | 1292.311 | 2.022 | 2.864 | 109.611 | 0.864 | 0.481 | 0 | 0 | 1 | 42.125 | 1 | 54 | 1 | GJB_Rinse | 0
239 | P-2272 | 2022/12/15 17:17:57 | 2-54 | 3 | -4.324 | -41.318 | 17849 | -1 | 1 | 22-0196 | 22-309 | H2O | 1671124984 | 0.313 | 0.667 | 75.571 | 0.002 | 0.002 | 1.812 | 1293.637 | 1.135 | 2.887 | 111.558 | 0.868 | 0.514 | 0 | 0 | 1 | 42.062 | 1 | 54 | 1 | GJB_Rinse | 0
240 | P-2272 | 2022/12/15 17:26:53 | 2-54 | 4 | -4.346 | -41.483 | 17875 | 0 | 1 | 22-0196 | 22-309 | H2O | 1671125519 | 0.3 | 0.743 | 76.347 | 0.002 | 0 | 1.818 | 1293.673 | 1.128 | 2.908 | 112.069 | 0.869 | 0.564 | 0 | 0 | 1 | 42.188 | 1 | 54 | 1 | GJB_Rinse | 0
241 | P-2273 | 2022/12/15 17:36:41 | 2-55 | 1 | -6.695 | -56.318 | 17538 | -1 | 1 | 22-0194 | 22-309 | H2O | 1671126107 | 0.316 | 0.796 | 81.777 | 0.002 | 0.008 | 1.954 | 1291.326 | 1.809 | 2.811 | 108.011 | 0.867 | 0.348 | 0 | 0 | 1 | 42.25 | 1 | 55 | 1 | GJB_Rinse | 0
242 | P-2273 | 2022/12/15 17:45:37 | 2-55 | 2 | -6.926 | -58.759 | 17569 | -1 | 1 | 22-0194 | 22-309 | H2O | 1671126643 | 0.363 | 0.764 | 75.875 | 0.002 | 0.005 | 1.816 | 1291.274 | 2.208 | 2.78 | 107.721 | 0.87 | 0.375 | 0 | 0 | 1 | 42.25 | 1 | 55 | 1 | GJB_Rinse | 0
243 | P-2273 | 2022/12/15 17:54:33 | 2-55 | 3 | -6.918 | -58.908 | 17608 | -1 | 1 | 22-0194 | 22-309 | H2O | 1671127180 | 0.296 | 0.758 | 77.211 | 0.001 | -0.001 | 1.853 | 1291.821 | 1.606 | 2.784 | 108.254 | 0.873 | 0.329 | 0 | 0 | 1 | 42.188 | 1 | 55 | 1 | GJB_Rinse | 0
244 | P-2273 | 2022/12/15 18:03:29 | 2-55 | 4 | -6.918 | -58.715 | 17759 | 0 | 1 | 22-0194 | 22-309 | H2O | 1671127716 | 0.302 | 0.764 | 76.505 | 0.001 | 0.006 | 1.827 | 1292.634 | 2.048 | 2.826 | 109.405 | 0.867 | 0.461 | 0 | 0 | 1 | 42.188 | 1 | 55 | 1 | GJB_Rinse | 0
245 | P-2274 | 2022/12/15 18:13:16 | 2-56 | 1 | -3.761 | -38.704 | 16105 | -1 | 1 | SCBI4-B-3 | 22-283 | H2O | 1671128302 | 0.407 | 1.178 | 73.486 | 0 | 0.003 | 1.771 | 1924.243 | -22.285 | 2.505 | 93.834 | 0.868 | -0.922 | 0 | 0 | 1 | 42.438 | 1 | 56 | 1 | GJB_Rinse | 0
246 | P-2274 | 2022/12/15 18:22:12 | 2-56 | 2 | -3.593 | -36.757 | 16421 | -1 | 1 | SCBI4-B-3 | 22-283 | H2O | 1671128838 | 0.373 | 1.054 | 73.845 | 0.002 | 0 | 1.784 | 1948.377 | -24.02 | 2.534 | 96.612 | 0.877 | -1.205 | 0 | 0 | 1 | 42.25 | 1 | 56 | 1 | GJB_Rinse | 0
247 | P-2274 | 2022/12/15 18:31:08 | 2-56 | 3 | -3.585 | -36.563 | 16394 | -1 | 1 | SCBI4-B-3 | 22-283 | H2O | 1671129375 | 0.394 | 1.057 | 74.955 | -0.001 | 0.004 | 1.807 | 1949.033 | -23.586 | 2.65 | 97.645 | 0.878 | -0.591 | 0 | 0 | 1 | 42.188 | 1 | 56 | 1 | GJB_Rinse | 0
248 | P-2274 | 2022/12/15 18:40:04 | 2-56 | 4 | -3.571 | -35.998 | 16150 | 0 | 1 | SCBI4-B-3 | 22-283 | H2O | 1671129910 | 0.371 | 1.153 | 74.589 | 0 | 0.003 | 1.797 | 1938.611 | -23.779 | 2.537 | 95.99 | 0.875 | -1.025 | 0 | 0 | 1 | 42.25 | 1 | 56 | 1 | GJB_Rinse | 0
249 | P-2275 | 2022/12/15 18:49:50 | 2-57 | 1 | [5 empty] | SCBI4-B-2 | 22-283 | [18 empty] | 1 | 57 | 1 | GJB_Rinse | 0
250 | P-2275 | 2022/12/15 18:58:46 | 2-57 | 2 | -2.846 | -35.663 | 16485 | -1 | 1 | SCBI4-B-2 | 22-283 | H2O | 1671131032 | 0.336 | 0.949 | 75.542 | 0.001 | 0.005 | 1.814 | 1728.92 | -16.626 | 2.54 | 100.037 | 0.867 | -0.807 | 0 | 0 | 1 | 42.188 | 1 | 57 | 1 | GJB_Rinse | 0
251 | P-2275 | 2022/12/15 19:07:42 | 2-57 | 3 | -2.889 | -36.458 | 15573 | -1 | 1 | SCBI4-B-2 | 22-283 | H2O | 1671131568 | 0.359 | 1.001 | 73.469 | 0.001 | 0.004 | 1.776 | 1697.275 | -15.561 | 2.462 | 94.401 | 0.866 | -0.935 | 0 | 0 | 1 | 42.188 | 1 | 57 | 1 | GJB_Rinse | 0
252 | P-2275 | 2022/12/15 19:16:37 | 2-57 | 4 | -2.762 | -35.434 | 17242 | 0 | 1 | SCBI4-B-2 | 22-283 | H2O | 1671132104 | 0.404 | 0.919 | 75.732 | 0.001 | 0.003 | 1.81 | 1754.782 | -17.947 | 2.788 | 108.01 | 0.876 | -0.517 | 0 | 0 | 1 | 42.125 | 1 | 57 | 1 | GJB_Rinse | 0
253 | P-2276 | 2022/12/15 19:26:23 | 2-58 | 1 | -7.264 | -56.223 | 18017 | -1 | 1 | EV | standard | H2O | 1671132690 | 0.337 | 0.759 | 79.173 | 0.001 | 0.006 | 1.882 | 1303.56 | 1.587 | 2.869 | 110.657 | 0.865 | 0.51 | 0 | 0 | 1 | 42.188 | 1 | 58 | 1 | GJB_Rinse | 0
254 | P-2276 | 2022/12/15 19:35:19 | 2-58 | 2 | -7.422 | -58.18 | 18072 | -1 | 1 | EV | standard | H2O | 1671133226 | 0.311 | 0.792 | 79.197 | 0.002 | 0.005 | 1.891 | 1298.865 | 1.801 | 2.865 | 110.911 | 0.872 | 0.479 | 0 | 0 | 1 | 42.188 | 1 | 58 | 1 | GJB_Rinse | 0
255 | P-2276 | 2022/12/15 19:44:14 | 2-58 | 3 | -7.511 | -58.898 | 17989 | -1 | 1 | EV | standard | H2O | 1671133761 | 0.34 | 0.793 | 81.845 | 0.001 | 0.006 | 1.951 | 1297.078 | 2.015 | 2.826 | 108.629 | 0.866 | 0.105 | 0 | 0 | 1 | 42.25 | 1 | 58 | 1 | GJB_Rinse | 0
256 | P-2276 | 2022/12/15 19:53:10 | 2-58 | 4 | -7.458 | -59.035 | 18049 | 0 | 1 | EV | standard | H2O | 1671134297 | 0.312 | 0.707 | 79.459 | 0 | 0.002 | 1.904 | 1296.546 | 1.353 | 2.952 | 112.907 | 0.873 | 0.662 | 0 | 0 | 1 | 42.188 | 1 | 58 | 1 | GJB_Rinse | 0
257 | P-2277 | 2022/12/15 20:02:57 | 2-59 | 1 | -19.131 | -149.208 | 17220 | -1 | 1 | SLCP-22-020 | Precipitation | H2O | 1671134883 | 0.323 | 0.909 | 80.166 | 0.003 | 0.013 | 1.911 | 1289.631 | 2.787 | 2.72 | 104.946 | 0.869 | 0.405 | 0 | 0 | 1 | 42.312 | 1 | 59 | 1 | GJB_Rinse | 0
258 | P-2277 | 2022/12/15 20:11:53 | 2-59 | 2 | -19.071 | -155.98 | 19765 | -1 | 1 | SLCP-22-020 | Precipitation | H2O | 1671135419 | 0.338 | 0.751 | 81.038 | 0.002 | 0.011 | 1.902 | 1308.455 | 2.386 | 3.144 | 121.682 | 0.876 | 0.527 | 0 | 0 | 1 | 42.25 | 1 | 59 | 1 | GJB_Rinse | 0
259 | P-2277 | 2022/12/15 20:20:48 | 2-59 | 3 | -19.792 | -160.197 | 17902 | -1 | 1 | SLCP-22-020 | Precipitation | H2O | 1671135955 | 0.346 | 0.831 | 80.131 | 0.003 | 0.011 | 1.912 | 1294.765 | 2.628 | 2.832 | 109.023 | 0.872 | 0.213 | 0 | 0 | 1 | 42.188 | 1 | 59 | 1 | GJB_Rinse | 0
260 | P-2277 | 2022/12/15 20:29:44 | 2-59 | 4 | -19.897 | -161.299 | 18005 | 0 | 1 | SLCP-22-020 | Precipitation | H2O | 1671136491 | 0.31 | 0.811 | 80.798 | 0.001 | 0.01 | 1.925 | 1295.312 | 2.512 | 2.859 | 110.566 | 0.869 | 0.567 | 0 | 0 | 1 | 42.25 | 1 | 59 | 1 | GJB_Rinse | 0
261 | P-2278 | 2022/12/15 20:39:31 | 2-60 | 1 | -8.348 | -72.198 | 18148 | -1 | 1 | EV | standard | H2O | 1671137078 | 0.321 | 0.805 | 78.164 | 0.001 | 0.002 | 1.854 | 1296.07 | 1.909 | 2.871 | 111.583 | 0.867 | 0.457 | 0 | 0 | 1 | 42.312 | 1 | 60 | 1 | GJB_Rinse | 0
262 | P-2278 | 2022/12/15 20:48:27 | 2-60 | 2 | -7.794 | -63.258 | 18071 | -1 | 1 | EV | standard | H2O | 1671137614 | 0.335 | 0.763 | 79.094 | 0 | 0.003 | 1.878 | 1295.658 | 2.081 | 2.872 | 110.816 | 0.876 | 0.423 | 0 | 0 | 1 | 42.188 | 1 | 60 | 1 | GJB_Rinse | 0
263 | P-2278 | 2022/12/15 20:57:23 | 2-60 | 3 | -7.67 | -61.378 | 18278 | -1 | 1 | EV | standard | H2O | 1671138150 | 0.288 | 0.756 | 81.096 | 0 | 0 | 1.934 | 1296.861 | 2.031 | 2.92 | 113.528 | 0.877 | 0.54 | 0 | 0 | 1 | 42.188 | 1 | 60 | 1 | GJB_Rinse | 0
264 | P-2278 | 2022/12/15 21:06:19 | 2-60 | 4 | -7.661 | -60.445 | 18217 | 0 | 1 | EV | standard | H2O | 1671138686 | 0.302 | 0.722 | 79.766 | 0.001 | 0.003 | 1.9 | 1296.593 | 1.84 | 2.909 | 111.884 | 0.874 | 0.304 | 0 | 0 | 1 | 42.375 | 1 | 60 | 1 | GJB_Rinse | 0

## Instructions
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Variable | Instruction | Comment
StandardsFile | standards file.csv
OutputBase | ChemCorrect
CorrectionType | M
timecode | Time Code
identifier | Identifier 1
sample | Sample
injection | Inj Nr | Inj Nr
line | Line
job | Job
Variable | Instruction | Comment
baseline_shift_MEAN | mean_standards(baseline_shift) | The mean of the organic baseline of the standards.
baseline_shift_STD | std_standards(baseline_shift) | The standard deviation of the organic baseline of the standards.
ch4_ppm_MEAN | mean_standards(ch4_ppm) | The mean methane concentration of the standards.
ch4_ppm_STD | std_standards(ch4_ppm) | The standard deviation of the methane concentrations of the standards.
residuals_MEAN | mean_standards(residuals) | The mean difference between the measured and modeled spectra of the standards.
residuals_STD | std_standards(residuals) | The standard deviation of the difference between measured and modeled spectra of the standards.
slope_shift_MEAN | mean_standards(slope_shift) | The mean organic slope of the standards.
slope_shift_STD | std_standards(slope_shift) | The standard deviation of the organic slope of the standards.
baseline_curvature_MEAN | mean_standards(baseline_curvature) | The mean of the baseline curvature of the standards.
baseline_curvature_STD | std_standards(baseline_curvature) | The standard deviation of the baseline curvature of the standards.
residuals_MEAN_ABS | abs(residuals_MEAN) | The absolute value of the mean difference between the measured and modeled spectra of the standards.
baseline_shift_MEAN_ABS | abs(baseline_shift_MEAN) | The absolute value of the mean of the organic baseline of the standards.
ch4_ppm_MEAN_ABS | abs(ch4_ppm_MEAN) | The absolute value of the mean methane concentration of the standards.
slope_shift_MEAN_ABS | abs(slope_shift_MEAN) | The absolute value of the mean organic slope of the standards.
residuals_15SIGMA | evaluation(residuals_STD*1.5) | One and half times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_15SIGMA | evaluation(residuals_MEAN+residuals_15SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_2SIGMA | evaluation(residuals_STD*2) | Twice the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_2SIGMA | evaluation(residuals_MEAN+residuals_2SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_3SIGMA | evaluation(residuals_STD*3) | Three times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_3SIGMA | evaluation(residuals_MEAN+residuals_3SIGMA) | The mean difference between measured and modeled spectra of the standards plus three standard deviation.
residuals_10PCT | evaluation(residuals_MEAN*0.1) | 30% of the difference between measured and modeled spectra of the standards.
ch4_ppm_3SIGMA | evaluation(ch4_ppm_STD*3) | Three times the standard deviation of the methanol amplitudes of the standards.
ch4_ppm_PLUS_3SIGMA | evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA) | The mean methanol amplitude of the standards plus three standard deviation.
HALF | float(0.5) | A fixed value of 0.5 or half.
baseline_shift_HALFSIGMA | evaluation(baseline_shift_STD*HALF) | Half the standard deviation of the organic baseline of the standards.
baseline_shift_PLUS_HALFSIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA) | The absolute value of the mean organic baseline of the standards plus one-half standard deviation.
baseline_shift_PLUS_1SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD) | The absolute value of the mean organic baseline of the standards plus one standard deviation.
baseline_shift_3SIGMA | evaluation(baseline_shift_STD*3) | Three times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_3SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA) | The absolute value of the mean organic baseline of the standards plus three standard deviation.
baseline_shift_18SIGMA | evaluation(baseline_shift_STD*18) | Eighteen times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_18SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA) | The absolute value of the mean organic baseline of the standards plus eighteen standard deviation.
slope_shift_5SIGMA | evaluation(slope_shift_STD*5) | Five times the standard deviation of the organic slope of the standards.
slope_shift_PLUS_5SIGMA | evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA) | The absolute value of the mean organic slope of the standards plus five standard deviation.
baseline_curvature_3SIGMA | evaluation(baseline_curvature_STD*3) | Three times the standard deviation of the baseline curvature of the standards.
baseline_curvature_PLUS_3SIGMA | evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
TRUE | evaluation(1>0)
FALSE | evaluation(1<0)
PROC_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the uncalibrated d18O value in order to calibrate it.
PROC_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the uncalibrated dD value in order to calibrate it.
Variable | Instruction | Comment
IS_STANDARD | is_standard_sample() | Standard recognition.
IS_NOT_STANDARD | bool_xor(IS_STANDARD,TRUE)
SAMP_residuals_MEAN | mean_all(residuals) | The mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_residuals_STD | std_all(residuals) | One standard deviation of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN | mean_all(ch4_ppm) | The mean methane concentration of a sample.
SAMP_ch4_ppm_STD | std_all(ch4_ppm) | The standard deviation of the methane concentrations of a sample.
SAMP_baseline_shift_MEAN | mean_all(baseline_shift) | The mean of the organic baseline of a sample.
SAMP_baseline_shift_STD | std_all(baseline_shift) | The standard deviation of the organic baseline of a sample.
SAMP_slope_shift_MEAN | mean_all(slope_shift) | The mean organic slope of a sample.
SAMP_slope_shift_STD | std_all(slope_shift) | The standard deviation of the organic slope of a sample.
SAMP_d18O_STD | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
SAMP_dD_STD | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_baseline_curvature_MEAN | mean_all(baseline_curvature) | The mean of the baseline curvature of the ssamples
SAMP_baseline_curvature_STD | std_all(baseline_curvature) | The standard deviation of the baseline curvature of the samples
SAMP_residuals_MEAN_ABS | abs(SAMP_residuals_MEAN) | Absolute value of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN_ABS | abs(SAMP_ch4_ppm_MEAN) | Absolute value of the mean methane concentration of a sample.
SAMP_baseline_shift_MEAN_ABS | abs(SAMP_baseline_shift_MEAN) | Absolute value of the mean organic baseline of a sample.
SAMP_slope_shift_MEAN_ABS | abs(SAMP_slope_shift_MEAN) | Absolute value of the mean organic slope of a sample.
SAMP_residuals_DIFFERENTIAL | evaluation(SAMP_residuals_MEAN-residuals_MEAN)
SAMP_residuals_DIFFERENTIAL_ABS | abs(SAMP_residuals_DIFFERENTIAL) | Absolute value of the difference between the mean of the sample residual and the mean of all standards residuals
SAMP_residuals_DIFFERENTIAL_PCT | evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
SAMP_baseline_curvature_DIFFERENTIAL | evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
SAMP_STD_DELTA_18_16O_CURR_MEAN | current_standards_mean(d(18_16)Mean) | The d18O mean of the last run standard.
SAMP_STD_DELTA_DH_CURR_MEAN | current_standards_mean(d(D_H)Mean) | The dD mean of the last run standard.
known_dD | known_standard(d(D_H)Mean) | The dD value of the known Standard.
known_d18O | known_standard(d(18_16)Mean) | The d18O value of the known Standard.
d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
measured_d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
measured_dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
measured_d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
measured_dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the mean uncalibrated d18O value in order to calibrate it.
SAMP_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the mean uncalibrated dD value in order to calibrate it.
SAMP_TEST_VALUE | evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
residuals_GT_3SIGMA | evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA) | Threshold of red trigger. If the mean difference between the measured spectra of the sample and the modeled spectra of the standard is greater than that of the mean difference of the standard plus three standard deviations, this will trigger.
residuals_GE_2SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GE_15SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GT_10PCT | evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
ch4_ppm_GE_3SIGMA | evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA) | True if the mean methane concentration of the sample is greater than or equal to that of the standard plus three standard deviations, otherwise false.
baseline_shift_GT_1SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus one standard deviation, otherwise false.
baseline_shift_GT_3SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus three standard deviation, otherwise false. Sets yellow trigger for "other", trigger threshold is approximately sample to sample dD precision specification
baseline_shift_GT_18SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus eighteen standard deviation, otherwise false. Sets red trigger for "other", trigger threshold is approximately 24 hour dD drift specification
slope_shift_GT_5SIGMA | evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA) | True if the mean organic slope of the sample is greater than that of the standard plus five standard deviations, otherwise false.
baseline_curvature_GT_3SIGMA | evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
RED | truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA)) | Red trigger. If the sample is not a standard and the threshold is triggered (measured-modeled > standard+3stdev's), then RED.
NOT_RED | bool_xor(RED,TRUE) | If the sample is determined to be "red" than it can not be "not red."
YELLOW | truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA))) | Truth evaluation used to determine whether or not a given sample is yellow.
NOT_YELLOW | bool_xor(YELLOW,TRUE) | If sample is determined to be "yellow" than it can not be "not yellow."
GREEN | truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW)) | This is the pre-requisite for a sample to be colored green. It must not be a standard, must not be yellow, and must not be red.
NOT_GREEN | bool_xor(GREEN,TRUE) | Pre-requisite for being "not-green."
CYAN | truth_evaluation((IS_STANDARD)) | This is the pre-requisite for a sample to be a standard and therefore cyan colored.
NOT_CYAN | bool_xor(CYAN,TRUE) | Pre-requisite for being not a standard.
METHANE | truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA)) | This is the threshold for amounts of methane being included in the main ChemCorrect spreadsheet. Methane will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured methane concentration must be greater than or equal to that of the standard plus three standard deviations.
OTHER1 | truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA)) | This is the threshold for amounts of "OTHER1" being included in the main ChemCorrect spreadsheet. This column (OTHER1) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic baseline must be greater than that of the standard plus one standard deviation.
OTHER2 | truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA)) | This is the threshold for amounts of "OTHER2" being included in the main ChemCorrect spreadsheet. This column (OTHER2) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic slope must be greater than that of the standard plus three standard deviations.
OTHER3 | truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | truth_evaluation(RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN)) | Organic baseline differential showing the difference between the measured baseline and the standard baseline.
Variable | Instruction | Comment
DTL_TEST | math_evaluation((1 + 2)) | Test evaluation of 1+2 being equal to 3.
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample | FALSE | FALSE | FALSE | FALSE | Testing Sample
B | identifier | TRUE | Name | RED | YELLOW | GREEN | CYAN | Testing Name
C | SAMP_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d<sup>18</sup>O Mean | RED | YELLOW | GREEN | CYAN | Testing18 16 Mean
D | SAMP_DELTA_DH_CALIBRATE | TRUE | Calibrated d<sup>2</sup>H Mean | RED | YELLOW | GREEN | CYAN | This is a test to see that I can use tool tips
E | SAMP_ch4_ppm_MEAN | METHANE | CH<sub>4</sub> | RED | YELLOW | FALSE | FALSE
F | baseline_shift_DIFF | OTHER1 | C<sub>2+</sub> alcohols | RED | YELLOW | FALSE | FALSE
G | SAMP_residuals_DIFFERENTIAL_PCT | OTHER3 | relative %deviation | RED | YELLOW | FALSE | FALSE
H | SAMP_d18O_STD | TRUE | Uncalibrated d<sup>18</sup>O precision | RED | YELLOW | GREEN | CYAN
I | SAMP_dD_STD | TRUE | Uncalibrated d<sup>2</sup>H precision | RED | YELLOW | GREEN | CYAN
J | SAMP_slope_shift_MEAN | OTHER2 | slope | RED | YELLOW | FALSE | FALSE
K | SAMP_baseline_curvature_DIFFERENTIAL | OTHER4 | curvature | RED | YELLOW | FALSE | FALSE
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample
B | identifier | TRUE | Name
C | col(d(18_16)Mean) | TRUE | Measured d(18_16) Mean
D | col(d(D_H)Mean) | TRUE | Measured d(D_H) Mean
E | DTL_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d(18_16) Mean
F | DTL_DELTA_DH_CALIBRATE | TRUE | Calibrated d(D_H) Mean
G | col(ch4_ppm) | METHANE | CH<sub>4</sub>
H | col(baseline_shift) | OTHER1 | Other
I | col(Time Code) | TRUE | Time Code
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | d18O_mean | TRUE | d(18_16) Mean | RED | YELLOW | GREEN | CYAN
B | dD_mean | TRUE | d(D_H) Mean | RED | YELLOW | GREEN | CYAN
C | XY_X(d18O_mean) XY_Y(dD_mean) | TRUE | d18O x dD | gmwl | best_fit

## Process Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Process Variables
Variable | Value | Operations
baseline_shift_MEAN | 1303.41425
varray: [1301.761, 1304.158, 1303.736, 1301.942, 1301.732, 1304.965, 1302.532, 1303.933, 1303.41, 1303.426, 1303.861, 1304.325, 1303.662, 1305.269, 1302.529, 1304.861, 1303.338, 1303.014, 1302.753, 1303.078]
inst: mean_standards
operation: mean_standards(baseline_shift)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0, 10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: baseline_shift
baseline_shift_STD | 1.02504089135
varray: [1301.761, 1304.158, 1303.736, 1301.942, 1301.732, 1304.965, 1302.532, 1303.933, 1303.41, 1303.426, 1303.861, 1304.325, 1303.662, 1305.269, 1302.529, 1304.861, 1303.338, 1303.014, 1302.753, 1303.078]
inst: std_standards
operation: std_standards(baseline_shift)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0, 10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: baseline_shift
ch4_ppm_MEAN | 0.51515
varray: [0.252, 0.503, 0.156, 0.455, 0.36, 0.471, 0.434, 0.395, 0.594, 0.527, 0.66, 0.633, 0.698, 0.524, 0.721, 0.703, 0.775, 0.247, 0.579, 0.616]
inst: mean_standards
operation: mean_standards(ch4_ppm)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0, 10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: ch4_ppm
ch4_ppm_STD | 0.170929432962
varray: [0.252, 0.503, 0.156, 0.455, 0.36, 0.471, 0.434, 0.395, 0.594, 0.527, 0.66, 0.633, 0.698, 0.524, 0.721, 0.703, 0.775, 0.247, 0.579, 0.616]
inst: std_standards
operation: std_standards(ch4_ppm)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0, 10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: ch4_ppm
residuals_MEAN | 3.0966
varray: [3.052, 3.145, 3.057, 3.066, 3.06, 3.155, 3.082, 3.125, 3.161, 3.128, 3.086, 3.052, 3.082, 3.111, 3.115, 3.153, 3.13, 3.008, 3.085, 3.079]
inst: mean_standards
operation: mean_standards(residuals)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0, 10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: residuals
residuals_STD | 0.0417957439284
varray: [3.052, 3.145, 3.057, 3.066, 3.06, 3.155, 3.082, 3.125, 3.161, 3.128, 3.086, 3.052, 3.082, 3.111, 3.115, 3.153, 3.13, 3.008, 3.085, 3.079]
inst: std_standards
operation: std_standards(residuals)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0, 10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: residuals
slope_shift_MEAN | 1.68865
varray: [1.476, 1.199, 1.349, 1.544, 1.344, 1.197, 1.293, 0.989, 0.85, 1.425, 2.283, 2.012, 2.25, 2.0, 2.002, 1.933, 2.639, 2.266, 2.038, 1.684]
inst: mean_standards
operation: mean_standards(slope_shift)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0, 10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: slope_shift
slope_shift_STD | 0.491345122035
varray: [1.476, 1.199, 1.349, 1.544, 1.344, 1.197, 1.293, 0.989, 0.85, 1.425, 2.283, 2.012, 2.25, 2.0, 2.002, 1.933, 2.639, 2.266, 2.038, 1.684]
inst: std_standards
operation: std_standards(slope_shift)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0, 10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: slope_shift
baseline_curvature_MEAN | 117.99125
varray: [116.181, 119.363, 116.554, 116.577, 116.549, 119.057, 116.992, 118.679, 119.746, 118.195, 117.977, 118.31, 118.568, 119.107, 118.617, 120.097, 118.244, 115.027, 118.261, 117.724]
inst: mean_standards
operation: mean_standards(baseline_curvature)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0, 10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: baseline_curvature
baseline_curvature_STD | 1.3085247167
varray: [116.181, 119.363, 116.554, 116.577, 116.549, 119.057, 116.992, 118.679, 119.746, 118.195, 117.977, 118.31, 118.568, 119.107, 118.617, 120.097, 118.244, 115.027, 118.261, 117.724]
inst: std_standards
operation: std_standards(baseline_curvature)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0, 10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: baseline_curvature
residuals_MEAN_ABS | 3.0966
operation: abs(residuals_MEAN)
inst: abs
source_col: residuals_MEAN
baseline_shift_MEAN_ABS | 1303.41425
operation: abs(baseline_shift_MEAN)
inst: abs
source_col: baseline_shift_MEAN
ch4_ppm_MEAN_ABS | 0.51515
operation: abs(ch4_ppm_MEAN)
inst: abs
source_col: ch4_ppm_MEAN
slope_shift_MEAN_ABS | 1.68865
operation: abs(slope_shift_MEAN)
inst: abs
source_col: slope_shift_MEAN
residuals_15SIGMA | 0.0626936158926
operation: evaluation(residuals_STD*1.5)
inst: evaluation
factors: residuals_STD*1.5
residuals_PLUS_15SIGMA | 3.15929361589
operation: evaluation(residuals_MEAN+residuals_15SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_15SIGMA
residuals_2SIGMA | 0.0835914878568
operation: evaluation(residuals_STD*2)
inst: evaluation
factors: residuals_STD*2
residuals_PLUS_2SIGMA | 3.18019148786
operation: evaluation(residuals_MEAN+residuals_2SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_2SIGMA
residuals_3SIGMA | 0.125387231785
operation: evaluation(residuals_STD*3)
inst: evaluation
factors: residuals_STD*3
residuals_PLUS_3SIGMA | 3.22198723179
operation: evaluation(residuals_MEAN+residuals_3SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_3SIGMA
residuals_10PCT | 0.30966
operation: evaluation(residuals_MEAN*0.1)
inst: evaluation
factors: residuals_MEAN*0.1
ch4_ppm_3SIGMA | 0.512788298885
operation: evaluation(ch4_ppm_STD*3)
inst: evaluation
factors: ch4_ppm_STD*3
ch4_ppm_PLUS_3SIGMA | 1.02793829889
operation: evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA)
inst: evaluation
factors: ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA
HALF | 0.5
operation: float(0.5)
inst: float
baseline_shift_HALFSIGMA | 0.512520445677
operation: evaluation(baseline_shift_STD*HALF)
inst: evaluation
factors: baseline_shift_STD*HALF
baseline_shift_PLUS_HALFSIGMA | 1303.92677045
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA
baseline_shift_PLUS_1SIGMA | 1304.43929089
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_STD
baseline_shift_3SIGMA | 3.07512267406
operation: evaluation(baseline_shift_STD*3)
inst: evaluation
factors: baseline_shift_STD*3
baseline_shift_PLUS_3SIGMA | 1306.48937267
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_3SIGMA
baseline_shift_18SIGMA | 18.4507360444
operation: evaluation(baseline_shift_STD*18)
inst: evaluation
factors: baseline_shift_STD*18
baseline_shift_PLUS_18SIGMA | 1321.86498604
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_18SIGMA
slope_shift_5SIGMA | 2.45672561017
operation: evaluation(slope_shift_STD*5)
inst: evaluation
factors: slope_shift_STD*5
slope_shift_PLUS_5SIGMA | 4.14537561017
operation: evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA)
inst: evaluation
factors: slope_shift_MEAN_ABS+slope_shift_5SIGMA
baseline_curvature_3SIGMA | 3.9255741501
operation: evaluation(baseline_curvature_STD*3)
inst: evaluation
factors: baseline_curvature_STD*3
baseline_curvature_PLUS_3SIGMA | 121.91682415
operation: evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
inst: evaluation
factors: baseline_curvature_3SIGMA+baseline_curvature_MEAN
TRUE | True
operation: evaluation(1>0)
inst: evaluation
factors: 1>0
FALSE | False
operation: evaluation(1<0)
inst: evaluation
factors: 1<0
PROC_DELTA_18_16O_CALIBRATE | -6.48352717264
actual_array: [ 1.93 -15.84]
varray: [0.0, -8.54, -7.996, -7.81, -7.663, -7.596, -7.629, -7.592, -7.572, -7.542, 4.002, 4.502, 4.634, 4.702, 4.734, 4.845, 4.825, 4.881, 4.881, 4.921, -12.045, -12.85, -13.025, -13.068, -13.188, -13.205, -13.267, -13.27, -13.217, -13.353, -7.83, -7.649, -7.59, -7.575, -7.582, -7.549, -7.567, -7.588, -7.54, -7.535, -7.132, -7.205, -7.164, -7.226, -6.24, -6.323, -6.247, -6.281, -6.102, -6.162, -6.132, -6.092, -7.8, -7.98, -7.994, -8.023, -6.684, -6.496, -6.532, -6.433, -2.761, -2.442, -2.506, -2.496, -6.612, -6.527, -6.738, -6.754, -4.542, -4.002, -3.986, -3.89, -7.201, -7.38, -7.457, -7.507, 0.0, 14.907, 15.027, 14.907, -5.504, -6.254, -6.731, -6.813, -4.358, -3.969, -4.181, -4.282, -3.312, -3.24, -3.222, -3.229, -2.454, -2.207, -2.399, -2.392, -6.831, -6.867, -7.064, -7.092, -6.721, -6.503, -6.669, -6.692, 2.412, 3.042, 2.92, 3.023, -6.835, -7.273, -7.39, -7.431, 11.277, 12.322, 12.42, 12.475, 0.0, -2.027, -2.414, -2.449, -1.063, 0.0, -0.988, -1.053, 5.553, 6.411, 6.002, 5.993, -5.955, -6.365, -6.708, -6.794, -6.022, -5.941, -6.051, -6.042, -6.286, -6.067, -6.328, -6.259, -2.166, -1.999, -2.003, -1.966, -7.217, -7.491, -7.545, -7.561, -2.758, -2.251, -2.493, -2.364, 5.384, 5.77, 5.777, 5.814, -1.735, -2.073, -2.227, -2.275, -3.0, -3.159, -3.126, -3.114, -6.892, -6.902, -7.125, -7.151, 0.0, 0.0, 9.047, 9.047, -0.596, -1.043, -1.228, -1.331, -2.598, -2.73, -2.778, -2.806, -7.182, -7.485, -7.496, -7.572, -2.759, 0.0, -2.759, -2.908, -6.615, -6.758, -6.858, -6.909, -2.593, -2.142, -2.219, -2.276, -7.469, -7.448, -7.817, -7.825, -3.853, -3.577, -3.654, -3.588, -7.111, -7.154, -7.271, -7.39, 4.495, 5.263, 5.261, 5.327, -5.594, -5.964, -6.187, -6.2, -7.403, -7.454, -7.5, -7.511, 0.0, 6.919, 6.724, 6.716, 0.0, -4.383, -4.383, -4.383, 12.429, 13.461, 13.456, 13.643, -5.75, -6.629, -6.944, -7.091, -4.437, -4.321, -4.324, -4.346, -6.695, -6.926, -6.918, -6.918, -3.761, -3.593, -3.585, -3.571, 0.0, -2.846, -2.889, -2.762, -7.264, -7.422, -7.511, -7.458, -19.131, -19.071, -19.792, -19.897, -8.348, -7.794, -7.67, -7.661]
fit_rsq: 1.0
fit_intercept: 2.76579538548
except: True
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -3.70733333333
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [1.0, 603.0, 1140.0, 1677.0, 2214.0, 2751.0, 3288.0, 3824.0, 4361.0, 4899.0, 5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0, 10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0, 16335.0, 16873.0, 17410.0, 17947.0, 18484.0, 19022.0, 19559.0, 20097.0, 20634.0, 21172.0, 21762.0, 22299.0, 22836.0, 23372.0, 23963.0, 24500.0, 25038.0, 25575.0, 26165.0, 26702.0, 27239.0, 27776.0, 28363.0, 28900.0, 29437.0, 29974.0, 30562.0, 31099.0, 31637.0, 32174.0, 32763.0, 33299.0, 33836.0, 34374.0, 34963.0, 35500.0, 36037.0, 36574.0, 37164.0, 37701.0, 38239.0, 38775.0, 39366.0, 39902.0, 40439.0, 40975.0, 41565.0, 42103.0, 42640.0, 43177.0, 43764.0, 44300.0, 44837.0, 45374.0, 45962.0, 46499.0, 47036.0, 47573.0, 48161.0, 48697.0, 49234.0, 49770.0, 50359.0, 50896.0, 51433.0, 51970.0, 52560.0, 53097.0, 53634.0, 54171.0, 54760.0, 55297.0, 55835.0, 56372.0, 56961.0, 57498.0, 58035.0, 58572.0, 59158.0, 59695.0, 60232.0, 60769.0, 61356.0, 61893.0, 62430.0, 62967.0, 63554.0, 64091.0, 64628.0, 65164.0, 65752.0, 66289.0, 66825.0, 67362.0, 67950.0, 68486.0, 69023.0, 69560.0, 70149.0, 70686.0, 71222.0, 71760.0, 72349.0, 72886.0, 73424.0, 73961.0, 74545.0, 75082.0, 75620.0, 76156.0, 76744.0, 77280.0, 77816.0, 78353.0, 78939.0, 79476.0, 80012.0, 80549.0, 81135.0, 81673.0, 82209.0, 82747.0, 83334.0, 83870.0, 84407.0, 84944.0, 85531.0, 86069.0, 86605.0, 87141.0, 87729.0, 88267.0, 88803.0, 89341.0, 89925.0, 90461.0, 90998.0, 91536.0, 92120.0, 92656.0, 93193.0, 93730.0, 94315.0, 94852.0, 95387.0, 95924.0, 96510.0, 97048.0, 97585.0, 98122.0, 98707.0, 99243.0, 99780.0, 100317.0, 100903.0, 101440.0, 101976.0, 102514.0, 103099.0, 103636.0, 104174.0, 104710.0, 105295.0, 105832.0, 106369.0, 106905.0, 107491.0, 108029.0, 108566.0, 109102.0, 173508.0, 174044.0, 174581.0, 175117.0, 175703.0, 176239.0, 176776.0, 177312.0, 177898.0, 178433.0, 178969.0, 179505.0, 180092.0, 180629.0, 181165.0, 181702.0, 182288.0, 182824.0, 183361.0, 183897.0, 184481.0, 185017.0, 185552.0, 186088.0, 186674.0, 187210.0, 187745.0, 188282.0, 188868.0, 189403.0, 189939.0, 190475.0, 191061.0, 191598.0, 192133.0, 192669.0, 193256.0, 193792.0, 194327.0, 194863.0, 195451.0, 195987.0, 196523.0, 197059.0, 197646.0, 198182.0, 198718.0, 199254.0, 199840.0, 200376.0, 200912.0, 201447.0, 202033.0, 202569.0, 203104.0, 203640.0, 204227.0, 204763.0, 205298.0, 205834.0, 206421.0, 206957.0, 207493.0, 208029.0]
source_col: d(18_16)Mean
PROC_DELTA_DH_CALIBRATE | -52.6487752861
actual_array: [ 18.1 -119. ]
varray: [0.0, -73.804, -66.476, -63.642, -61.793, -60.773, -60.621, -60.485, -60.328, -60.173, 27.931, 35.95, 38.323, 39.038, 39.387, 40.433, 40.454, 40.7, 40.95, 40.951, -92.364, -104.802, -107.803, -108.829, -109.722, -110.03, -110.336, -110.728, -110.81, -111.11, -65.169, -61.279, -60.391, -60.146, -59.969, -59.972, -59.801, -60.029, -59.642, -59.547, -55.458, -55.637, -55.339, -55.742, -50.322, -50.493, -50.183, -50.258, -51.369, -51.867, -51.62, -51.632, -58.579, -59.743, -60.325, -60.63, -49.748, -48.403, -48.199, -47.334, -28.095, -25.908, -25.724, -25.439, -50.686, -51.862, -53.072, -52.641, -43.252, -39.884, -39.147, -38.881, -56.688, -58.454, -58.619, -59.154, 0.0, 63.694, 64.027, 63.694, -42.282, -51.614, -55.091, -56.007, -50.174, -48.663, -49.207, -49.603, -33.383, -31.682, -31.168, -31.453, -31.79, -31.349, -31.87, -31.525, -55.566, -57.028, -58.21, -58.409, -49.391, -47.574, -48.141, -47.82, -11.651, -7.642, -8.016, -7.416, -53.081, -57.245, -58.231, -58.506, 38.442, 48.366, 49.9, 51.016, 0.0, -19.615, -21.897, -22.828, -29.011, 0.0, -29.654, -30.055, 7.71, 14.018, 12.809, 12.896, -47.445, -53.137, -55.39, -56.179, -48.814, -48.251, -48.76, -49.089, -50.408, -49.357, -50.273, -50.078, -27.532, -25.691, -25.33, -25.209, -55.14, -58.174, -58.785, -59.252, -29.838, -26.156, -26.487, -26.005, 8.649, 12.038, 12.751, 13.001, -19.798, -22.93, -24.09, -24.141, -28.089, -28.666, -28.552, -28.8, -55.672, -57.699, -59.082, -59.26, 0.0, 0.0, 41.738, 41.738, -11.692, -16.765, -18.683, -19.254, -32.291, -33.945, -34.89, -35.003, -56.013, -58.389, -58.674, -59.277, -30.987, 0.0, -30.987, -31.84, -53.961, -55.784, -56.705, -56.9, -23.386, -18.81, -18.539, -18.971, -61.536, -64.846, -66.939, -67.354, -40.975, -37.955, -37.659, -37.339, -58.013, -59.354, -59.947, -60.644, 7.861, 15.828, 17.406, 17.855, -41.996, -47.317, -48.866, -49.215, -57.745, -58.502, -58.661, -59.222, 0.0, 16.727, 17.034, 17.451, 0.0, -40.863, -40.863, -40.863, 42.888, 51.812, 53.544, 54.692, -43.865, -53.188, -55.764, -56.705, -42.758, -41.596, -41.318, -41.483, -56.318, -58.759, -58.908, -58.715, -38.704, -36.757, -36.563, -35.998, 0.0, -35.663, -36.458, -35.434, -56.223, -58.18, -58.898, -59.035, -149.208, -155.98, -160.197, -161.299, -72.198, -63.258, -61.378, -60.445]
fit_rsq: 1.0
fit_intercept: 19.1281237053
except: True
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -36.963405303
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [1.0, 603.0, 1140.0, 1677.0, 2214.0, 2751.0, 3288.0, 3824.0, 4361.0, 4899.0, 5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0, 10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0, 16335.0, 16873.0, 17410.0, 17947.0, 18484.0, 19022.0, 19559.0, 20097.0, 20634.0, 21172.0, 21762.0, 22299.0, 22836.0, 23372.0, 23963.0, 24500.0, 25038.0, 25575.0, 26165.0, 26702.0, 27239.0, 27776.0, 28363.0, 28900.0, 29437.0, 29974.0, 30562.0, 31099.0, 31637.0, 32174.0, 32763.0, 33299.0, 33836.0, 34374.0, 34963.0, 35500.0, 36037.0, 36574.0, 37164.0, 37701.0, 38239.0, 38775.0, 39366.0, 39902.0, 40439.0, 40975.0, 41565.0, 42103.0, 42640.0, 43177.0, 43764.0, 44300.0, 44837.0, 45374.0, 45962.0, 46499.0, 47036.0, 47573.0, 48161.0, 48697.0, 49234.0, 49770.0, 50359.0, 50896.0, 51433.0, 51970.0, 52560.0, 53097.0, 53634.0, 54171.0, 54760.0, 55297.0, 55835.0, 56372.0, 56961.0, 57498.0, 58035.0, 58572.0, 59158.0, 59695.0, 60232.0, 60769.0, 61356.0, 61893.0, 62430.0, 62967.0, 63554.0, 64091.0, 64628.0, 65164.0, 65752.0, 66289.0, 66825.0, 67362.0, 67950.0, 68486.0, 69023.0, 69560.0, 70149.0, 70686.0, 71222.0, 71760.0, 72349.0, 72886.0, 73424.0, 73961.0, 74545.0, 75082.0, 75620.0, 76156.0, 76744.0, 77280.0, 77816.0, 78353.0, 78939.0, 79476.0, 80012.0, 80549.0, 81135.0, 81673.0, 82209.0, 82747.0, 83334.0, 83870.0, 84407.0, 84944.0, 85531.0, 86069.0, 86605.0, 87141.0, 87729.0, 88267.0, 88803.0, 89341.0, 89925.0, 90461.0, 90998.0, 91536.0, 92120.0, 92656.0, 93193.0, 93730.0, 94315.0, 94852.0, 95387.0, 95924.0, 96510.0, 97048.0, 97585.0, 98122.0, 98707.0, 99243.0, 99780.0, 100317.0, 100903.0, 101440.0, 101976.0, 102514.0, 103099.0, 103636.0, 104174.0, 104710.0, 105295.0, 105832.0, 106369.0, 106905.0, 107491.0, 108029.0, 108566.0, 109102.0, 173508.0, 174044.0, 174581.0, 175117.0, 175703.0, 176239.0, 176776.0, 177312.0, 177898.0, 178433.0, 178969.0, 179505.0, 180092.0, 180629.0, 181165.0, 181702.0, 182288.0, 182824.0, 183361.0, 183897.0, 184481.0, 185017.0, 185552.0, 186088.0, 186674.0, 187210.0, 187745.0, 188282.0, 188868.0, 189403.0, 189939.0, 190475.0, 191061.0, 191598.0, 192133.0, 192669.0, 193256.0, 193792.0, 194327.0, 194863.0, 195451.0, 195987.0, 196523.0, 197059.0, 197646.0, 198182.0, 198718.0, 199254.0, 199840.0, 200376.0, 200912.0, 201447.0, 202033.0, 202569.0, 203104.0, 203640.0, 204227.0, 204763.0, 205298.0, 205834.0, 206421.0, 206957.0, 207493.0, 208029.0]
source_col: d(D_H)Mean

## 1 - Sample 1 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 1
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8375
varray: [0.0, 3.532, 3.55, 3.078, 3.06, 3.009, 3.021, 3.094, 3.035, 2.996]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [1.0, 603.0, 1140.0, 1677.0, 2214.0, 2751.0, 3288.0, 3824.0, 4361.0, 4899.0]
source_col: residuals
SAMP_residuals_STD | 1.01880891786
varray: [0.0, 3.532, 3.55, 3.078, 3.06, 3.009, 3.021, 3.094, 3.035, 2.996]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [1.0, 603.0, 1140.0, 1677.0, 2214.0, 2751.0, 3288.0, 3824.0, 4361.0, 4899.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3483
varray: [0.0, 0.11, 0.524, 0.365, 0.443, 0.722, 0.379, 0.48, 0.366, 0.094]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [1.0, 603.0, 1140.0, 1677.0, 2214.0, 2751.0, 3288.0, 3824.0, 4361.0, 4899.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.221394093066
varray: [0.0, 0.11, 0.524, 0.365, 0.443, 0.722, 0.379, 0.48, 0.366, 0.094]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [1.0, 603.0, 1140.0, 1677.0, 2214.0, 2751.0, 3288.0, 3824.0, 4361.0, 4899.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1196.6585
varray: [0.0, 1405.818, 1398.85, 1330.216, 1308.256, 1305.883, 1305.126, 1304.33, 1304.046, 1304.06]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [1.0, 603.0, 1140.0, 1677.0, 2214.0, 2751.0, 3288.0, 3824.0, 4361.0, 4899.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 422.330540871
varray: [0.0, 1405.818, 1398.85, 1330.216, 1308.256, 1305.883, 1305.126, 1304.33, 1304.046, 1304.06]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [1.0, 603.0, 1140.0, 1677.0, 2214.0, 2751.0, 3288.0, 3824.0, 4361.0, 4899.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.8753
varray: [0.0, -1.22, -1.28, 0.795, 1.38, 2.272, 2.256, 1.336, 1.622, 1.592]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [1.0, 603.0, 1140.0, 1677.0, 2214.0, 2751.0, 3288.0, 3824.0, 4361.0, 4899.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.29972168388
varray: [0.0, -1.22, -1.28, 0.795, 1.38, 2.272, 2.256, 1.336, 1.622, 1.592]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [1.0, 603.0, 1140.0, 1677.0, 2214.0, 2751.0, 3288.0, 3824.0, 4361.0, 4899.0]
source_col: slope_shift
SAMP_d18O_STD | 2.47613484644
varray: [0.0, -8.54, -7.996, -7.81, -7.663, -7.596, -7.629, -7.592, -7.572, -7.542]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 603.0, 1140.0, 1677.0, 2214.0, 2751.0, 3288.0, 3824.0, 4361.0, 4899.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 20.4084160486
varray: [0.0, -73.804, -66.476, -63.642, -61.793, -60.773, -60.621, -60.485, -60.328, -60.173]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 603.0, 1140.0, 1677.0, 2214.0, 2751.0, 3288.0, 3824.0, 4361.0, 4899.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.7141
varray: [0.0, 132.728, 133.667, 119.009, 117.547, 117.417, 116.029, 117.913, 116.688, 116.143]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [1.0, 603.0, 1140.0, 1677.0, 2214.0, 2751.0, 3288.0, 3824.0, 4361.0, 4899.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 38.7795978245
varray: [0.0, 132.728, 133.667, 119.009, 117.547, 117.417, 116.029, 117.913, 116.688, 116.143]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [1.0, 603.0, 1140.0, 1677.0, 2214.0, 2751.0, 3288.0, 3824.0, 4361.0, 4899.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8375
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3483
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1196.6585
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.8753
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.2591
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.2591
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -8.36724149067
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -9.27715
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 0
varray: [0.0, -8.54, -7.996, -7.81, -7.663, -7.596, -7.629, -7.592, -7.572, -7.542]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [1.0, 603.0, 1140.0, 1677.0, 2214.0, 2751.0, 3288.0, 3824.0, 4361.0, 4899.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 0
varray: [0.0, -73.804, -66.476, -63.642, -61.793, -60.773, -60.621, -60.485, -60.328, -60.173]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [1.0, 603.0, 1140.0, 1677.0, 2214.0, 2751.0, 3288.0, 3824.0, 4361.0, 4899.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.994
varray: [0.0, -8.54, -7.996, -7.81, -7.663, -7.596, -7.629, -7.592, -7.572, -7.542]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 603.0, 1140.0, 1677.0, 2214.0, 2751.0, 3288.0, 3824.0, 4361.0, 4899.0]
source_col: d(18_16)Mean
dD_mean | -56.8095
varray: [0.0, -73.804, -66.476, -63.642, -61.793, -60.773, -60.621, -60.485, -60.328, -60.173]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 603.0, 1140.0, 1677.0, 2214.0, 2751.0, 3288.0, 3824.0, 4361.0, 4899.0]
source_col: d(D_H)Mean
d18O_stdev | 2.47613484644
varray: [0.0, -8.54, -7.996, -7.81, -7.663, -7.596, -7.629, -7.592, -7.572, -7.542]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 603.0, 1140.0, 1677.0, 2214.0, 2751.0, 3288.0, 3824.0, 4361.0, 4899.0]
source_col: d(18_16)Mean
dD_stdev | 20.4084160486
varray: [0.0, -73.804, -66.476, -63.642, -61.793, -60.773, -60.621, -60.485, -60.328, -60.173]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 603.0, 1140.0, 1677.0, 2214.0, 2751.0, 3288.0, 3824.0, 4361.0, 4899.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.994
varray: [0.0, -8.54, -7.996, -7.81, -7.663, -7.596, -7.629, -7.592, -7.572, -7.542]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 603.0, 1140.0, 1677.0, 2214.0, 2751.0, 3288.0, 3824.0, 4361.0, 4899.0]
source_col: d(18_16)Mean
measured_dD_mean | -56.8095
varray: [0.0, -73.804, -66.476, -63.642, -61.793, -60.773, -60.621, -60.485, -60.328, -60.173]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 603.0, 1140.0, 1677.0, 2214.0, 2751.0, 3288.0, 3824.0, 4361.0, 4899.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 2.47613484644
varray: [0.0, -8.54, -7.996, -7.81, -7.663, -7.596, -7.629, -7.592, -7.572, -7.542]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 603.0, 1140.0, 1677.0, 2214.0, 2751.0, 3288.0, 3824.0, 4361.0, 4899.0]
source_col: d(18_16)Mean
measured_dD_stdev | 20.4084160486
varray: [0.0, -73.804, -66.476, -63.642, -61.793, -60.773, -60.621, -60.485, -60.328, -60.173]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 603.0, 1140.0, 1677.0, 2214.0, 2751.0, 3288.0, 3824.0, 4361.0, 4899.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.77547355071
actual_array: [ 1.93 -15.84]
varray: [0.0, -8.54, -7.996, -7.81, -7.663, -7.596, -7.629, -7.592, -7.572, -7.542]
fit_rsq: 1.0
fit_intercept: 2.76579538548
except: True
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.994
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [1.0, 603.0, 1140.0, 1677.0, 2214.0, 2751.0, 3288.0, 3824.0, 4361.0, 4899.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -71.2767677563
actual_array: [ 18.1 -119. ]
varray: [0.0, -73.804, -66.476, -63.642, -61.793, -60.773, -60.621, -60.485, -60.328, -60.173]
fit_rsq: 1.0
fit_intercept: 19.1281237053
except: True
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -56.8095
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [1.0, 603.0, 1140.0, 1677.0, 2214.0, 2751.0, 3288.0, 3824.0, 4361.0, 4899.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.57675
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -106.75575
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 2 - Sample 2 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 2
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.1031
varray: [3.052, 3.145, 3.057, 3.066, 3.06, 3.155, 3.082, 3.125, 3.161, 3.128]
inst: mean_all
operation: mean_all(residuals)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0]
source_col: residuals
SAMP_residuals_STD | 0.043864564286
varray: [3.052, 3.145, 3.057, 3.066, 3.06, 3.155, 3.082, 3.125, 3.161, 3.128]
inst: std_all
operation: std_all(residuals)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4147
varray: [0.252, 0.503, 0.156, 0.455, 0.36, 0.471, 0.434, 0.395, 0.594, 0.527]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.131021669115
varray: [0.252, 0.503, 0.156, 0.455, 0.36, 0.471, 0.434, 0.395, 0.594, 0.527]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1303.1595
varray: [1301.761, 1304.158, 1303.736, 1301.942, 1301.732, 1304.965, 1302.532, 1303.933, 1303.41, 1303.426]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.11509244759
varray: [1301.761, 1304.158, 1303.736, 1301.942, 1301.732, 1304.965, 1302.532, 1303.933, 1303.41, 1303.426]
inst: std_all
operation: std_all(baseline_shift)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.2666
varray: [1.476, 1.199, 1.349, 1.544, 1.344, 1.197, 1.293, 0.989, 0.85, 1.425]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.215869919679
varray: [1.476, 1.199, 1.349, 1.544, 1.344, 1.197, 1.293, 0.989, 0.85, 1.425]
inst: std_all
operation: std_all(slope_shift)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0]
source_col: slope_shift
SAMP_d18O_STD | 0.275359179578
varray: [4.002, 4.502, 4.634, 4.702, 4.734, 4.845, 4.825, 4.881, 4.881, 4.921]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.99500160061
varray: [27.931, 35.95, 38.323, 39.038, 39.387, 40.433, 40.454, 40.7, 40.95, 40.951]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 117.7893
varray: [116.181, 119.363, 116.554, 116.577, 116.549, 119.057, 116.992, 118.679, 119.746, 118.195]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.35911433989
varray: [116.181, 119.363, 116.554, 116.577, 116.549, 119.057, 116.992, 118.679, 119.746, 118.195]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.1031
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4147
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1303.1595
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.2666
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0065
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0065
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.209907640638
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.20195
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 4.6927
varray: [4.002, 4.502, 4.634, 4.702, 4.734, 4.845, 4.825, 4.881, 4.881, 4.921]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 38.4117
varray: [27.931, 35.95, 38.323, 39.038, 39.387, 40.433, 40.454, 40.7, 40.95, 40.951]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0]
source_col: d(D_H)Mean
known_dD | 18.1
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 1.93
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 4.6927
varray: [4.002, 4.502, 4.634, 4.702, 4.734, 4.845, 4.825, 4.881, 4.881, 4.921]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0]
source_col: d(18_16)Mean
dD_mean | 38.4117
varray: [27.931, 35.95, 38.323, 39.038, 39.387, 40.433, 40.454, 40.7, 40.95, 40.951]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0]
source_col: d(D_H)Mean
d18O_stdev | 0.275359179578
varray: [4.002, 4.502, 4.634, 4.702, 4.734, 4.845, 4.825, 4.881, 4.881, 4.921]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0]
source_col: d(18_16)Mean
dD_stdev | 3.99500160061
varray: [27.931, 35.95, 38.323, 39.038, 39.387, 40.433, 40.454, 40.7, 40.95, 40.951]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0]
source_col: d(D_H)Mean
measured_d18O_mean | 4.6927
varray: [4.002, 4.502, 4.634, 4.702, 4.734, 4.845, 4.825, 4.881, 4.881, 4.921]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0]
source_col: d(18_16)Mean
measured_dD_mean | 38.4117
varray: [27.931, 35.95, 38.323, 39.038, 39.387, 40.433, 40.454, 40.7, 40.95, 40.951]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.275359179578
varray: [4.002, 4.502, 4.634, 4.702, 4.734, 4.845, 4.825, 4.881, 4.881, 4.921]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.99500160061
varray: [27.931, 35.95, 38.323, 39.038, 39.387, 40.433, 40.454, 40.7, 40.95, 40.951]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.93
actual_array: [ 1.93 -15.84]
varray: [4.002, 4.502, 4.634, 4.702, 4.734, 4.845, 4.825, 4.881, 4.881, 4.921]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 4.6927
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 18.1
actual_array: [ 18.1 -119. ]
varray: [27.931, 35.95, 38.323, 39.038, 39.387, 40.433, 40.454, 40.7, 40.95, 40.951]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 38.4117
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [5488.0, 6024.0, 6561.0, 7098.0, 7636.0, 8173.0, 8710.0, 9247.0, 9785.0, 10322.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.31115
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.25475
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 3 - Sample 3 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 3
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.0901
varray: [3.086, 3.052, 3.082, 3.111, 3.115, 3.153, 3.13, 3.008, 3.085, 3.079]
inst: mean_all
operation: mean_all(residuals)
tarray: [10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: residuals
SAMP_residuals_STD | 0.0408641380403
varray: [3.086, 3.052, 3.082, 3.111, 3.115, 3.153, 3.13, 3.008, 3.085, 3.079]
inst: std_all
operation: std_all(residuals)
tarray: [10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.6156
varray: [0.66, 0.633, 0.698, 0.524, 0.721, 0.703, 0.775, 0.247, 0.579, 0.616]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.148629442275
varray: [0.66, 0.633, 0.698, 0.524, 0.721, 0.703, 0.775, 0.247, 0.579, 0.616]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1303.669
varray: [1303.861, 1304.325, 1303.662, 1305.269, 1302.529, 1304.861, 1303.338, 1303.014, 1302.753, 1303.078]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.911326018015
varray: [1303.861, 1304.325, 1303.662, 1305.269, 1302.529, 1304.861, 1303.338, 1303.014, 1302.753, 1303.078]
inst: std_all
operation: std_all(baseline_shift)
tarray: [10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.1107
varray: [2.283, 2.012, 2.25, 2.0, 2.002, 1.933, 2.639, 2.266, 2.038, 1.684]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.259284935501
varray: [2.283, 2.012, 2.25, 2.0, 2.002, 1.933, 2.639, 2.266, 2.038, 1.684]
inst: std_all
operation: std_all(slope_shift)
tarray: [10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: slope_shift
SAMP_d18O_STD | 0.381403198728
varray: [-12.045, -12.85, -13.025, -13.068, -13.188, -13.205, -13.267, -13.27, -13.217, -13.353]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 5.69395096572
varray: [-92.364, -104.802, -107.803, -108.829, -109.722, -110.03, -110.336, -110.728, -110.81, -111.11]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 118.1932
varray: [117.977, 118.31, 118.568, 119.107, 118.617, 120.097, 118.244, 115.027, 118.261, 117.724]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.29495171939
varray: [117.977, 118.31, 118.568, 119.107, 118.617, 120.097, 118.244, 115.027, 118.261, 117.724]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.0901
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.6156
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1303.669
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.1107
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0065
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0065
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.209907640638
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.20195
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-12.045, -12.85, -13.025, -13.068, -13.188, -13.205, -13.267, -13.27, -13.217, -13.353]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-92.364, -104.802, -107.803, -108.829, -109.722, -110.03, -110.336, -110.728, -110.81, -111.11]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: d(D_H)Mean
known_dD | -119.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | -15.84
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -13.0488
varray: [-12.045, -12.85, -13.025, -13.068, -13.188, -13.205, -13.267, -13.27, -13.217, -13.353]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: d(18_16)Mean
dD_mean | -107.6534
varray: [-92.364, -104.802, -107.803, -108.829, -109.722, -110.03, -110.336, -110.728, -110.81, -111.11]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: d(D_H)Mean
d18O_stdev | 0.381403198728
varray: [-12.045, -12.85, -13.025, -13.068, -13.188, -13.205, -13.267, -13.27, -13.217, -13.353]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: d(18_16)Mean
dD_stdev | 5.69395096572
varray: [-92.364, -104.802, -107.803, -108.829, -109.722, -110.03, -110.336, -110.728, -110.81, -111.11]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: d(D_H)Mean
measured_d18O_mean | -13.0488
varray: [-12.045, -12.85, -13.025, -13.068, -13.188, -13.205, -13.267, -13.27, -13.217, -13.353]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: d(18_16)Mean
measured_dD_mean | -107.6534
varray: [-92.364, -104.802, -107.803, -108.829, -109.722, -110.03, -110.336, -110.728, -110.81, -111.11]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.381403198728
varray: [-12.045, -12.85, -13.025, -13.068, -13.188, -13.205, -13.267, -13.27, -13.217, -13.353]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: d(18_16)Mean
measured_dD_stdev | 5.69395096572
varray: [-92.364, -104.802, -107.803, -108.829, -109.722, -110.03, -110.336, -110.728, -110.81, -111.11]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -15.84
actual_array: [ 1.93 -15.84]
varray: [-12.045, -12.85, -13.025, -13.068, -13.188, -13.205, -13.267, -13.27, -13.217, -13.353]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -13.0488
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -119.0
actual_array: [ 18.1 -119. ]
varray: [-92.364, -104.802, -107.803, -108.829, -109.722, -110.03, -110.336, -110.728, -110.81, -111.11]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -107.6534
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [10911.0, 11448.0, 11985.0, 12522.0, 13059.0, 13596.0, 14133.0, 14671.0, 15208.0, 15745.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.32415
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.25475
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 4 - Sample 4 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 4
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.0818
varray: [3.103, 3.09, 3.117, 3.054, 3.075, 3.045, 3.077, 3.067, 3.097, 3.093]
inst: mean_all
operation: mean_all(residuals)
tarray: [16335.0, 16873.0, 17410.0, 17947.0, 18484.0, 19022.0, 19559.0, 20097.0, 20634.0, 21172.0]
source_col: residuals
SAMP_residuals_STD | 0.0224291477026
varray: [3.103, 3.09, 3.117, 3.054, 3.075, 3.045, 3.077, 3.067, 3.097, 3.093]
inst: std_all
operation: std_all(residuals)
tarray: [16335.0, 16873.0, 17410.0, 17947.0, 18484.0, 19022.0, 19559.0, 20097.0, 20634.0, 21172.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.5375
varray: [0.209, 0.39, 0.598, 0.41, 0.954, 0.428, 0.498, 0.673, 0.61, 0.605]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [16335.0, 16873.0, 17410.0, 17947.0, 18484.0, 19022.0, 19559.0, 20097.0, 20634.0, 21172.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.201323537289
varray: [0.209, 0.39, 0.598, 0.41, 0.954, 0.428, 0.498, 0.673, 0.61, 0.605]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [16335.0, 16873.0, 17410.0, 17947.0, 18484.0, 19022.0, 19559.0, 20097.0, 20634.0, 21172.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1304.1016
varray: [1306.229, 1303.556, 1304.477, 1304.719, 1302.87, 1304.427, 1304.429, 1302.921, 1304.213, 1303.175]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [16335.0, 16873.0, 17410.0, 17947.0, 18484.0, 19022.0, 19559.0, 20097.0, 20634.0, 21172.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.01873267674
varray: [1306.229, 1303.556, 1304.477, 1304.719, 1302.87, 1304.427, 1304.429, 1302.921, 1304.213, 1303.175]
inst: std_all
operation: std_all(baseline_shift)
tarray: [16335.0, 16873.0, 17410.0, 17947.0, 18484.0, 19022.0, 19559.0, 20097.0, 20634.0, 21172.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.838
varray: [1.628, 1.564, 1.238, 1.651, 2.479, 2.411, 1.548, 2.065, 1.814, 1.982]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [16335.0, 16873.0, 17410.0, 17947.0, 18484.0, 19022.0, 19559.0, 20097.0, 20634.0, 21172.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.395852946882
varray: [1.628, 1.564, 1.238, 1.651, 2.479, 2.411, 1.548, 2.065, 1.814, 1.982]
inst: std_all
operation: std_all(slope_shift)
tarray: [16335.0, 16873.0, 17410.0, 17947.0, 18484.0, 19022.0, 19559.0, 20097.0, 20634.0, 21172.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0869268530305
varray: [-7.83, -7.649, -7.59, -7.575, -7.582, -7.549, -7.567, -7.588, -7.54, -7.535]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16335.0, 16873.0, 17410.0, 17947.0, 18484.0, 19022.0, 19559.0, 20097.0, 20634.0, 21172.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.67894995293
varray: [-65.169, -61.279, -60.391, -60.146, -59.969, -59.972, -59.801, -60.029, -59.642, -59.547]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16335.0, 16873.0, 17410.0, 17947.0, 18484.0, 19022.0, 19559.0, 20097.0, 20634.0, 21172.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 118.1876
varray: [119.331, 117.572, 119.738, 117.38, 118.796, 116.695, 118.083, 117.936, 118.735, 117.61]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [16335.0, 16873.0, 17410.0, 17947.0, 18484.0, 19022.0, 19559.0, 20097.0, 20634.0, 21172.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.945621629758
varray: [119.331, 117.572, 119.738, 117.38, 118.796, 116.695, 118.083, 117.936, 118.735, 117.61]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [16335.0, 16873.0, 17410.0, 17947.0, 18484.0, 19022.0, 19559.0, 20097.0, 20634.0, 21172.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.0818
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.5375
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1304.1016
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.838
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0148
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0148
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.477943550991
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.19635
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-7.83, -7.649, -7.59, -7.575, -7.582, -7.549, -7.567, -7.588, -7.54, -7.535]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [16335.0, 16873.0, 17410.0, 17947.0, 18484.0, 19022.0, 19559.0, 20097.0, 20634.0, 21172.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-65.169, -61.279, -60.391, -60.146, -59.969, -59.972, -59.801, -60.029, -59.642, -59.547]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [16335.0, 16873.0, 17410.0, 17947.0, 18484.0, 19022.0, 19559.0, 20097.0, 20634.0, 21172.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.6005
varray: [-7.83, -7.649, -7.59, -7.575, -7.582, -7.549, -7.567, -7.588, -7.54, -7.535]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16335.0, 16873.0, 17410.0, 17947.0, 18484.0, 19022.0, 19559.0, 20097.0, 20634.0, 21172.0]
source_col: d(18_16)Mean
dD_mean | -60.5945
varray: [-65.169, -61.279, -60.391, -60.146, -59.969, -59.972, -59.801, -60.029, -59.642, -59.547]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16335.0, 16873.0, 17410.0, 17947.0, 18484.0, 19022.0, 19559.0, 20097.0, 20634.0, 21172.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0869268530305
varray: [-7.83, -7.649, -7.59, -7.575, -7.582, -7.549, -7.567, -7.588, -7.54, -7.535]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16335.0, 16873.0, 17410.0, 17947.0, 18484.0, 19022.0, 19559.0, 20097.0, 20634.0, 21172.0]
source_col: d(18_16)Mean
dD_stdev | 1.67894995293
varray: [-65.169, -61.279, -60.391, -60.146, -59.969, -59.972, -59.801, -60.029, -59.642, -59.547]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16335.0, 16873.0, 17410.0, 17947.0, 18484.0, 19022.0, 19559.0, 20097.0, 20634.0, 21172.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.6005
varray: [-7.83, -7.649, -7.59, -7.575, -7.582, -7.549, -7.567, -7.588, -7.54, -7.535]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16335.0, 16873.0, 17410.0, 17947.0, 18484.0, 19022.0, 19559.0, 20097.0, 20634.0, 21172.0]
source_col: d(18_16)Mean
measured_dD_mean | -60.5945
varray: [-65.169, -61.279, -60.391, -60.146, -59.969, -59.972, -59.801, -60.029, -59.642, -59.547]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16335.0, 16873.0, 17410.0, 17947.0, 18484.0, 19022.0, 19559.0, 20097.0, 20634.0, 21172.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0869268530305
varray: [-7.83, -7.649, -7.59, -7.575, -7.582, -7.549, -7.567, -7.588, -7.54, -7.535]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16335.0, 16873.0, 17410.0, 17947.0, 18484.0, 19022.0, 19559.0, 20097.0, 20634.0, 21172.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.67894995293
varray: [-65.169, -61.279, -60.391, -60.146, -59.969, -59.972, -59.801, -60.029, -59.642, -59.547]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16335.0, 16873.0, 17410.0, 17947.0, 18484.0, 19022.0, 19559.0, 20097.0, 20634.0, 21172.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.3829478342
actual_array: [ 1.93 -15.84]
varray: [-7.83, -7.649, -7.59, -7.575, -7.582, -7.549, -7.567, -7.588, -7.54, -7.535]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.6005
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [16335.0, 16873.0, 17410.0, 17947.0, 18484.0, 19022.0, 19559.0, 20097.0, 20634.0, 21172.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -74.8294541954
actual_array: [ 18.1 -119. ]
varray: [-65.169, -61.279, -60.391, -60.146, -59.969, -59.972, -59.801, -60.029, -59.642, -59.547]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -60.5945
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [16335.0, 16873.0, 17410.0, 17947.0, 18484.0, 19022.0, 19559.0, 20097.0, 20634.0, 21172.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.33245
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.68735
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 5 - Sample 5 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 5
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.01475
varray: [3.111, 2.991, 2.93, 3.027]
inst: mean_all
operation: mean_all(residuals)
tarray: [21762.0, 22299.0, 22836.0, 23372.0]
source_col: residuals
SAMP_residuals_STD | 0.0756323343551
varray: [3.111, 2.991, 2.93, 3.027]
inst: std_all
operation: std_all(residuals)
tarray: [21762.0, 22299.0, 22836.0, 23372.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.33925
varray: [0.094, 0.508, 0.237, 0.518]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [21762.0, 22299.0, 22836.0, 23372.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.208990231032
varray: [0.094, 0.508, 0.237, 0.518]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [21762.0, 22299.0, 22836.0, 23372.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1302.23225
varray: [1307.487, 1301.111, 1300.273, 1300.058]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [21762.0, 22299.0, 22836.0, 23372.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.53249782968
varray: [1307.487, 1301.111, 1300.273, 1300.058]
inst: std_all
operation: std_all(baseline_shift)
tarray: [21762.0, 22299.0, 22836.0, 23372.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.98325
varray: [1.645, 2.227, 1.884, 2.177]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [21762.0, 22299.0, 22836.0, 23372.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.271549105443
varray: [1.645, 2.227, 1.884, 2.177]
inst: std_all
operation: std_all(slope_shift)
tarray: [21762.0, 22299.0, 22836.0, 23372.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0419871011939
varray: [-7.132, -7.205, -7.164, -7.226]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21762.0, 22299.0, 22836.0, 23372.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.180068505482
varray: [-55.458, -55.637, -55.339, -55.742]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21762.0, 22299.0, 22836.0, 23372.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 115.4735
varray: [119.329, 115.001, 113.598, 113.966]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [21762.0, 22299.0, 22836.0, 23372.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.63806703225
varray: [119.329, 115.001, 113.598, 113.966]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [21762.0, 22299.0, 22836.0, 23372.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.01475
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.33925
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1302.23225
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.98325
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.08185
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.08185
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.64322159788
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.51775
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-7.132, -7.205, -7.164, -7.226]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [21762.0, 22299.0, 22836.0, 23372.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-55.458, -55.637, -55.339, -55.742]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [21762.0, 22299.0, 22836.0, 23372.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.18175
varray: [-7.132, -7.205, -7.164, -7.226]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21762.0, 22299.0, 22836.0, 23372.0]
source_col: d(18_16)Mean
dD_mean | -55.544
varray: [-55.458, -55.637, -55.339, -55.742]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21762.0, 22299.0, 22836.0, 23372.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0419871011939
varray: [-7.132, -7.205, -7.164, -7.226]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21762.0, 22299.0, 22836.0, 23372.0]
source_col: d(18_16)Mean
dD_stdev | 0.180068505482
varray: [-55.458, -55.637, -55.339, -55.742]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21762.0, 22299.0, 22836.0, 23372.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.18175
varray: [-7.132, -7.205, -7.164, -7.226]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21762.0, 22299.0, 22836.0, 23372.0]
source_col: d(18_16)Mean
measured_dD_mean | -55.544
varray: [-55.458, -55.637, -55.339, -55.742]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21762.0, 22299.0, 22836.0, 23372.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0419871011939
varray: [-7.132, -7.205, -7.164, -7.226]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21762.0, 22299.0, 22836.0, 23372.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.180068505482
varray: [-55.458, -55.637, -55.339, -55.742]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21762.0, 22299.0, 22836.0, 23372.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.96352515289
actual_array: [ 1.93 -15.84]
varray: [-7.132, -7.205, -7.164, -7.226]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.18175
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [21762.0, 22299.0, 22836.0, 23372.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -70.0889408901
actual_array: [ 18.1 -119. ]
varray: [-55.458, -55.637, -55.339, -55.742]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -55.544
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [21762.0, 22299.0, 22836.0, 23372.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.3995
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.182
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 6 - Sample 6 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 6
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.0675
varray: [3.111, 3.057, 3.096, 3.006]
inst: mean_all
operation: mean_all(residuals)
tarray: [23963.0, 24500.0, 25038.0, 25575.0]
source_col: residuals
SAMP_residuals_STD | 0.0468934963508
varray: [3.111, 3.057, 3.096, 3.006]
inst: std_all
operation: std_all(residuals)
tarray: [23963.0, 24500.0, 25038.0, 25575.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3375
varray: [0.1, 0.436, 0.412, 0.402]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [23963.0, 24500.0, 25038.0, 25575.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.158974840777
varray: [0.1, 0.436, 0.412, 0.402]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [23963.0, 24500.0, 25038.0, 25575.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1304.6555
varray: [1308.241, 1303.661, 1303.846, 1302.874]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [23963.0, 24500.0, 25038.0, 25575.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.42719817897
varray: [1308.241, 1303.661, 1303.846, 1302.874]
inst: std_all
operation: std_all(baseline_shift)
tarray: [23963.0, 24500.0, 25038.0, 25575.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.86925
varray: [1.632, 2.218, 1.748, 1.879]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [23963.0, 24500.0, 25038.0, 25575.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.253450027948
varray: [1.632, 2.218, 1.748, 1.879]
inst: std_all
operation: std_all(slope_shift)
tarray: [23963.0, 24500.0, 25038.0, 25575.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0379857429395
varray: [-6.24, -6.323, -6.247, -6.281]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23963.0, 24500.0, 25038.0, 25575.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.132164039486
varray: [-50.322, -50.493, -50.183, -50.258]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23963.0, 24500.0, 25038.0, 25575.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 117.61775
varray: [119.468, 117.021, 118.081, 115.901]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [23963.0, 24500.0, 25038.0, 25575.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.52111436673
varray: [119.468, 117.021, 118.081, 115.901]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [23963.0, 24500.0, 25038.0, 25575.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.0675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3375
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1304.6555
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.86925
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0291
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0291
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.939740360395
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.3735
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-6.24, -6.323, -6.247, -6.281]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [23963.0, 24500.0, 25038.0, 25575.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-50.322, -50.493, -50.183, -50.258]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [23963.0, 24500.0, 25038.0, 25575.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.27275
varray: [-6.24, -6.323, -6.247, -6.281]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23963.0, 24500.0, 25038.0, 25575.0]
source_col: d(18_16)Mean
dD_mean | -50.314
varray: [-50.322, -50.493, -50.183, -50.258]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23963.0, 24500.0, 25038.0, 25575.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0379857429395
varray: [-6.24, -6.323, -6.247, -6.281]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23963.0, 24500.0, 25038.0, 25575.0]
source_col: d(18_16)Mean
dD_stdev | 0.132164039486
varray: [-50.322, -50.493, -50.183, -50.258]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23963.0, 24500.0, 25038.0, 25575.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.27275
varray: [-6.24, -6.323, -6.247, -6.281]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23963.0, 24500.0, 25038.0, 25575.0]
source_col: d(18_16)Mean
measured_dD_mean | -50.314
varray: [-50.322, -50.493, -50.183, -50.258]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23963.0, 24500.0, 25038.0, 25575.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0379857429395
varray: [-6.24, -6.323, -6.247, -6.281]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23963.0, 24500.0, 25038.0, 25575.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.132164039486
varray: [-50.322, -50.493, -50.183, -50.258]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23963.0, 24500.0, 25038.0, 25575.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.0530649325
actual_array: [ 1.93 -15.84]
varray: [-6.24, -6.323, -6.247, -6.281]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.27275
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [23963.0, 24500.0, 25038.0, 25575.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -65.1799448328
actual_array: [ 18.1 -119. ]
varray: [-50.322, -50.493, -50.183, -50.258]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -50.314
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [23963.0, 24500.0, 25038.0, 25575.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.34675
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1.24125
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 7 - Sample 7 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 7
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.925
varray: [2.926, 2.928, 2.969, 2.877]
inst: mean_all
operation: mean_all(residuals)
tarray: [26165.0, 26702.0, 27239.0, 27776.0]
source_col: residuals
SAMP_residuals_STD | 0.0376386326355
varray: [2.926, 2.928, 2.969, 2.877]
inst: std_all
operation: std_all(residuals)
tarray: [26165.0, 26702.0, 27239.0, 27776.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.39825
varray: [0.397, 0.415, 0.456, 0.325]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [26165.0, 26702.0, 27239.0, 27776.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0547197404965
varray: [0.397, 0.415, 0.456, 0.325]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [26165.0, 26702.0, 27239.0, 27776.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1297.68375
varray: [1298.391, 1296.662, 1297.715, 1297.967]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [26165.0, 26702.0, 27239.0, 27776.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.736066742898
varray: [1298.391, 1296.662, 1297.715, 1297.967]
inst: std_all
operation: std_all(baseline_shift)
tarray: [26165.0, 26702.0, 27239.0, 27776.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.60525
varray: [1.539, 1.545, 1.883, 1.454]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [26165.0, 26702.0, 27239.0, 27776.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.189772451461
varray: [1.539, 1.545, 1.883, 1.454]
inst: std_all
operation: std_all(slope_shift)
tarray: [26165.0, 26702.0, 27239.0, 27776.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0316227766017
varray: [-6.102, -6.162, -6.132, -6.092]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26165.0, 26702.0, 27239.0, 27776.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.203419107592
varray: [-51.369, -51.867, -51.62, -51.632]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26165.0, 26702.0, 27239.0, 27776.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.16275
varray: [113.843, 112.275, 112.913, 113.62]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [26165.0, 26702.0, 27239.0, 27776.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.712341853794
varray: [113.843, 112.275, 112.913, 113.62]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [26165.0, 26702.0, 27239.0, 27776.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.925
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.39825
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1297.68375
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.60525
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1716
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1716
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.54156171285
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.8285
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-6.102, -6.162, -6.132, -6.092]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [26165.0, 26702.0, 27239.0, 27776.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-51.369, -51.867, -51.62, -51.632]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [26165.0, 26702.0, 27239.0, 27776.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.122
varray: [-6.102, -6.162, -6.132, -6.092]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26165.0, 26702.0, 27239.0, 27776.0]
source_col: d(18_16)Mean
dD_mean | -51.622
varray: [-51.369, -51.867, -51.62, -51.632]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26165.0, 26702.0, 27239.0, 27776.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0316227766017
varray: [-6.102, -6.162, -6.132, -6.092]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26165.0, 26702.0, 27239.0, 27776.0]
source_col: d(18_16)Mean
dD_stdev | 0.203419107592
varray: [-51.369, -51.867, -51.62, -51.632]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26165.0, 26702.0, 27239.0, 27776.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.122
varray: [-6.102, -6.162, -6.132, -6.092]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26165.0, 26702.0, 27239.0, 27776.0]
source_col: d(18_16)Mean
measured_dD_mean | -51.622
varray: [-51.369, -51.867, -51.62, -51.632]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26165.0, 26702.0, 27239.0, 27776.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0316227766017
varray: [-6.102, -6.162, -6.132, -6.092]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26165.0, 26702.0, 27239.0, 27776.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.203419107592
varray: [-51.369, -51.867, -51.62, -51.632]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26165.0, 26702.0, 27239.0, 27776.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.90207276724
actual_array: [ 1.93 -15.84]
varray: [-6.102, -6.162, -6.132, -6.092]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.122
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [26165.0, 26702.0, 27239.0, 27776.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -66.4076631584
actual_array: [ 18.1 -119. ]
varray: [-51.369, -51.867, -51.62, -51.632]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -51.622
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [26165.0, 26702.0, 27239.0, 27776.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.48925
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -5.7305
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 8 - Sample 8 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 8
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.027
varray: [3.061, 2.996, 3.011, 3.04]
inst: mean_all
operation: mean_all(residuals)
tarray: [28363.0, 28900.0, 29437.0, 29974.0]
source_col: residuals
SAMP_residuals_STD | 0.0291089905928
varray: [3.061, 2.996, 3.011, 3.04]
inst: std_all
operation: std_all(residuals)
tarray: [28363.0, 28900.0, 29437.0, 29974.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3785
varray: [0.249, 0.143, 0.299, 0.823]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [28363.0, 28900.0, 29437.0, 29974.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.303386991591
varray: [0.249, 0.143, 0.299, 0.823]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [28363.0, 28900.0, 29437.0, 29974.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1302.4225
varray: [1305.632, 1302.875, 1300.36, 1300.823]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [28363.0, 28900.0, 29437.0, 29974.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.4026341794
varray: [1305.632, 1302.875, 1300.36, 1300.823]
inst: std_all
operation: std_all(baseline_shift)
tarray: [28363.0, 28900.0, 29437.0, 29974.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.9635
varray: [1.858, 1.899, 1.762, 2.335]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [28363.0, 28900.0, 29437.0, 29974.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.254234144048
varray: [1.858, 1.899, 1.762, 2.335]
inst: std_all
operation: std_all(slope_shift)
tarray: [28363.0, 28900.0, 29437.0, 29974.0]
source_col: slope_shift
SAMP_d18O_STD | 0.101098549281
varray: [-7.8, -7.98, -7.994, -8.023]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28363.0, 28900.0, 29437.0, 29974.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.905010635297
varray: [-58.579, -59.743, -60.325, -60.63]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28363.0, 28900.0, 29437.0, 29974.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 116.59025
varray: [118.812, 115.828, 115.626, 116.095]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [28363.0, 28900.0, 29437.0, 29974.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.49356941028
varray: [118.812, 115.828, 115.626, 116.095]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [28363.0, 28900.0, 29437.0, 29974.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.027
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3785
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1302.4225
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.9635
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0696
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0696
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.24762642899
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.401
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-7.8, -7.98, -7.994, -8.023]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [28363.0, 28900.0, 29437.0, 29974.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-58.579, -59.743, -60.325, -60.63]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [28363.0, 28900.0, 29437.0, 29974.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.94925
varray: [-7.8, -7.98, -7.994, -8.023]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28363.0, 28900.0, 29437.0, 29974.0]
source_col: d(18_16)Mean
dD_mean | -59.81925
varray: [-58.579, -59.743, -60.325, -60.63]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28363.0, 28900.0, 29437.0, 29974.0]
source_col: d(D_H)Mean
d18O_stdev | 0.101098549281
varray: [-7.8, -7.98, -7.994, -8.023]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28363.0, 28900.0, 29437.0, 29974.0]
source_col: d(18_16)Mean
dD_stdev | 0.905010635297
varray: [-58.579, -59.743, -60.325, -60.63]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28363.0, 28900.0, 29437.0, 29974.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.94925
varray: [-7.8, -7.98, -7.994, -8.023]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28363.0, 28900.0, 29437.0, 29974.0]
source_col: d(18_16)Mean
measured_dD_mean | -59.81925
varray: [-58.579, -59.743, -60.325, -60.63]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28363.0, 28900.0, 29437.0, 29974.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.101098549281
varray: [-7.8, -7.98, -7.994, -8.023]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28363.0, 28900.0, 29437.0, 29974.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.905010635297
varray: [-58.579, -59.743, -60.325, -60.63]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28363.0, 28900.0, 29437.0, 29974.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.7322580672
actual_array: [ 1.93 -15.84]
varray: [-7.8, -7.98, -7.994, -8.023]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.94925
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [28363.0, 28900.0, 29437.0, 29974.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -74.1017870456
actual_array: [ 18.1 -119. ]
varray: [-58.579, -59.743, -60.325, -60.63]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -59.81925
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [28363.0, 28900.0, 29437.0, 29974.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.38725
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.99175
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 9 - Sample 9 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 9
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.73975
varray: [2.707, 2.731, 2.695, 2.826]
inst: mean_all
operation: mean_all(residuals)
tarray: [30562.0, 31099.0, 31637.0, 32174.0]
source_col: residuals
SAMP_residuals_STD | 0.059415906961
varray: [2.707, 2.731, 2.695, 2.826]
inst: std_all
operation: std_all(residuals)
tarray: [30562.0, 31099.0, 31637.0, 32174.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.44575
varray: [0.582, 0.521, 0.41, 0.27]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [30562.0, 31099.0, 31637.0, 32174.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.137104279048
varray: [0.582, 0.521, 0.41, 0.27]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [30562.0, 31099.0, 31637.0, 32174.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1288.3525
varray: [1284.684, 1287.236, 1288.285, 1293.205]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [30562.0, 31099.0, 31637.0, 32174.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.57098385136
varray: [1284.684, 1287.236, 1288.285, 1293.205]
inst: std_all
operation: std_all(baseline_shift)
tarray: [30562.0, 31099.0, 31637.0, 32174.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.719
varray: [1.655, 1.688, 1.748, 1.785]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [30562.0, 31099.0, 31637.0, 32174.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.058463663929
varray: [1.655, 1.688, 1.748, 1.785]
inst: std_all
operation: std_all(slope_shift)
tarray: [30562.0, 31099.0, 31637.0, 32174.0]
source_col: slope_shift
SAMP_d18O_STD | 0.106659504968
varray: [-6.684, -6.496, -6.532, -6.433]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30562.0, 31099.0, 31637.0, 32174.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.998683466703
varray: [-49.748, -48.403, -48.199, -47.334]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30562.0, 31099.0, 31637.0, 32174.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 105.683
varray: [103.516, 104.669, 105.035, 109.512]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [30562.0, 31099.0, 31637.0, 32174.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.63345337279
varray: [103.516, 104.669, 105.035, 109.512]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [30562.0, 31099.0, 31637.0, 32174.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.73975
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.44575
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1288.3525
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.719
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.35685
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.35685
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -11.523929471
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -12.30825
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-6.684, -6.496, -6.532, -6.433]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [30562.0, 31099.0, 31637.0, 32174.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-49.748, -48.403, -48.199, -47.334]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [30562.0, 31099.0, 31637.0, 32174.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.53625
varray: [-6.684, -6.496, -6.532, -6.433]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30562.0, 31099.0, 31637.0, 32174.0]
source_col: d(18_16)Mean
dD_mean | -48.421
varray: [-49.748, -48.403, -48.199, -47.334]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30562.0, 31099.0, 31637.0, 32174.0]
source_col: d(D_H)Mean
d18O_stdev | 0.106659504968
varray: [-6.684, -6.496, -6.532, -6.433]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30562.0, 31099.0, 31637.0, 32174.0]
source_col: d(18_16)Mean
dD_stdev | 0.998683466703
varray: [-49.748, -48.403, -48.199, -47.334]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30562.0, 31099.0, 31637.0, 32174.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.53625
varray: [-6.684, -6.496, -6.532, -6.433]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30562.0, 31099.0, 31637.0, 32174.0]
source_col: d(18_16)Mean
measured_dD_mean | -48.421
varray: [-49.748, -48.403, -48.199, -47.334]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30562.0, 31099.0, 31637.0, 32174.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.106659504968
varray: [-6.684, -6.496, -6.532, -6.433]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30562.0, 31099.0, 31637.0, 32174.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.998683466703
varray: [-49.748, -48.403, -48.199, -47.334]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30562.0, 31099.0, 31637.0, 32174.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.31698821971
actual_array: [ 1.93 -15.84]
varray: [-6.684, -6.496, -6.532, -6.433]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.53625
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [30562.0, 31099.0, 31637.0, 32174.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -63.403132302
actual_array: [ 18.1 -119. ]
varray: [-49.748, -48.403, -48.199, -47.334]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -48.421
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [30562.0, 31099.0, 31637.0, 32174.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.6745
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -15.06175
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 10 - Sample 10 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 10
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.96075
varray: [2.969, 3.028, 2.932, 2.914]
inst: mean_all
operation: mean_all(residuals)
tarray: [32763.0, 33299.0, 33836.0, 34374.0]
source_col: residuals
SAMP_residuals_STD | 0.0503413349048
varray: [2.969, 3.028, 2.932, 2.914]
inst: std_all
operation: std_all(residuals)
tarray: [32763.0, 33299.0, 33836.0, 34374.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.46525
varray: [0.411, 0.483, 0.725, 0.242]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [32763.0, 33299.0, 33836.0, 34374.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.200473398069
varray: [0.411, 0.483, 0.725, 0.242]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [32763.0, 33299.0, 33836.0, 34374.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1297.3395
varray: [1298.315, 1298.5, 1295.596, 1296.947]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [32763.0, 33299.0, 33836.0, 34374.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.35304705511
varray: [1298.315, 1298.5, 1295.596, 1296.947]
inst: std_all
operation: std_all(baseline_shift)
tarray: [32763.0, 33299.0, 33836.0, 34374.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.5425
varray: [1.237, 1.545, 1.582, 1.806]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [32763.0, 33299.0, 33836.0, 34374.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.234043443261
varray: [1.237, 1.545, 1.582, 1.806]
inst: std_all
operation: std_all(slope_shift)
tarray: [32763.0, 33299.0, 33836.0, 34374.0]
source_col: slope_shift
SAMP_d18O_STD | 0.142630933999
varray: [-2.761, -2.442, -2.506, -2.496]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32763.0, 33299.0, 33836.0, 34374.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.21771603148
varray: [-28.095, -25.908, -25.724, -25.439]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32763.0, 33299.0, 33836.0, 34374.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.66825
varray: [114.221, 115.103, 113.418, 111.931]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [32763.0, 33299.0, 33836.0, 34374.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.34718258475
varray: [114.221, 115.103, 113.418, 111.931]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [32763.0, 33299.0, 33836.0, 34374.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.96075
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.46525
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1297.3395
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.5425
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.13585
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.13585
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.38706968934
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.323
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-2.761, -2.442, -2.506, -2.496]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [32763.0, 33299.0, 33836.0, 34374.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-28.095, -25.908, -25.724, -25.439]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [32763.0, 33299.0, 33836.0, 34374.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.55125
varray: [-2.761, -2.442, -2.506, -2.496]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32763.0, 33299.0, 33836.0, 34374.0]
source_col: d(18_16)Mean
dD_mean | -26.2915
varray: [-28.095, -25.908, -25.724, -25.439]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32763.0, 33299.0, 33836.0, 34374.0]
source_col: d(D_H)Mean
d18O_stdev | 0.142630933999
varray: [-2.761, -2.442, -2.506, -2.496]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32763.0, 33299.0, 33836.0, 34374.0]
source_col: d(18_16)Mean
dD_stdev | 1.21771603148
varray: [-28.095, -25.908, -25.724, -25.439]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32763.0, 33299.0, 33836.0, 34374.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.55125
varray: [-2.761, -2.442, -2.506, -2.496]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32763.0, 33299.0, 33836.0, 34374.0]
source_col: d(18_16)Mean
measured_dD_mean | -26.2915
varray: [-28.095, -25.908, -25.724, -25.439]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32763.0, 33299.0, 33836.0, 34374.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.142630933999
varray: [-2.761, -2.442, -2.506, -2.496]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32763.0, 33299.0, 33836.0, 34374.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.21771603148
varray: [-28.095, -25.908, -25.724, -25.439]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32763.0, 33299.0, 33836.0, 34374.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.32558670349
actual_array: [ 1.93 -15.84]
varray: [-2.761, -2.442, -2.506, -2.496]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.55125
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [32763.0, 33299.0, 33836.0, 34374.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -42.6318840709
actual_array: [ 18.1 -119. ]
varray: [-28.095, -25.908, -25.724, -25.439]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -26.2915
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [32763.0, 33299.0, 33836.0, 34374.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.4535
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -6.07475
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 11 - Sample 11 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 11
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.84175
varray: [2.724, 2.936, 2.765, 2.942]
inst: mean_all
operation: mean_all(residuals)
tarray: [34963.0, 35500.0, 36037.0, 36574.0]
source_col: residuals
SAMP_residuals_STD | 0.113561657262
varray: [2.724, 2.936, 2.765, 2.942]
inst: std_all
operation: std_all(residuals)
tarray: [34963.0, 35500.0, 36037.0, 36574.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.43075
varray: [0.569, 0.406, 0.244, 0.504]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [34963.0, 35500.0, 36037.0, 36574.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.141382165306
varray: [0.569, 0.406, 0.244, 0.504]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [34963.0, 35500.0, 36037.0, 36574.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1293.26975
varray: [1287.144, 1296.915, 1291.969, 1297.051]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [34963.0, 35500.0, 36037.0, 36574.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 4.71884388207
varray: [1287.144, 1296.915, 1291.969, 1297.051]
inst: std_all
operation: std_all(baseline_shift)
tarray: [34963.0, 35500.0, 36037.0, 36574.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.96975
varray: [1.738, 1.857, 1.639, 2.645]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [34963.0, 35500.0, 36037.0, 36574.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.458904038625
varray: [1.738, 1.857, 1.639, 2.645]
inst: std_all
operation: std_all(slope_shift)
tarray: [34963.0, 35500.0, 36037.0, 36574.0]
source_col: slope_shift
SAMP_d18O_STD | 0.107846727659
varray: [-6.612, -6.527, -6.738, -6.754]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34963.0, 35500.0, 36037.0, 36574.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.04700760105
varray: [-50.686, -51.862, -53.072, -52.641]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34963.0, 35500.0, 36037.0, 36574.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.95875
varray: [105.957, 113.056, 108.293, 112.529]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [34963.0, 35500.0, 36037.0, 36574.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.41505825963
varray: [105.957, 113.056, 108.293, 112.529]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [34963.0, 35500.0, 36037.0, 36574.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.84175
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.43075
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1293.26975
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.96975
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.25485
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.25485
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -8.22999418717
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -8.0325
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-6.612, -6.527, -6.738, -6.754]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [34963.0, 35500.0, 36037.0, 36574.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-50.686, -51.862, -53.072, -52.641]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [34963.0, 35500.0, 36037.0, 36574.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.65775
varray: [-6.612, -6.527, -6.738, -6.754]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [34963.0, 35500.0, 36037.0, 36574.0]
source_col: d(18_16)Mean
dD_mean | -52.06525
varray: [-50.686, -51.862, -53.072, -52.641]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [34963.0, 35500.0, 36037.0, 36574.0]
source_col: d(D_H)Mean
d18O_stdev | 0.107846727659
varray: [-6.612, -6.527, -6.738, -6.754]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34963.0, 35500.0, 36037.0, 36574.0]
source_col: d(18_16)Mean
dD_stdev | 1.04700760105
varray: [-50.686, -51.862, -53.072, -52.641]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34963.0, 35500.0, 36037.0, 36574.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.65775
varray: [-6.612, -6.527, -6.738, -6.754]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [34963.0, 35500.0, 36037.0, 36574.0]
source_col: d(18_16)Mean
measured_dD_mean | -52.06525
varray: [-50.686, -51.862, -53.072, -52.641]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [34963.0, 35500.0, 36037.0, 36574.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.107846727659
varray: [-6.612, -6.527, -6.738, -6.754]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34963.0, 35500.0, 36037.0, 36574.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.04700760105
varray: [-50.686, -51.862, -53.072, -52.641]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34963.0, 35500.0, 36037.0, 36574.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.43868339768
actual_array: [ 1.93 -15.84]
varray: [-6.612, -6.527, -6.738, -6.754]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.65775
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [34963.0, 35500.0, 36037.0, 36574.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -66.8237076139
actual_array: [ 18.1 -119. ]
varray: [-50.686, -51.862, -53.072, -52.641]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -52.06525
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [34963.0, 35500.0, 36037.0, 36574.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.5725
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.1445
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 12 - Sample 12 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 12
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.62925
varray: [2.217, 2.747, 2.715, 2.838]
inst: mean_all
operation: mean_all(residuals)
tarray: [37164.0, 37701.0, 38239.0, 38775.0]
source_col: residuals
SAMP_residuals_STD | 0.279728886603
varray: [2.217, 2.747, 2.715, 2.838]
inst: std_all
operation: std_all(residuals)
tarray: [37164.0, 37701.0, 38239.0, 38775.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.298
varray: [0.138, 0.383, 0.453, 0.218]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [37164.0, 37701.0, 38239.0, 38775.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.145201010097
varray: [0.138, 0.383, 0.453, 0.218]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [37164.0, 37701.0, 38239.0, 38775.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1285.32675
varray: [1270.065, 1288.323, 1288.43, 1294.489]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [37164.0, 37701.0, 38239.0, 38775.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 10.5747421206
varray: [1270.065, 1288.323, 1288.43, 1294.489]
inst: std_all
operation: std_all(baseline_shift)
tarray: [37164.0, 37701.0, 38239.0, 38775.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.6135
varray: [1.812, 1.22, 1.899, 1.523]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [37164.0, 37701.0, 38239.0, 38775.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.307649692562
varray: [1.812, 1.22, 1.899, 1.523]
inst: std_all
operation: std_all(slope_shift)
tarray: [37164.0, 37701.0, 38239.0, 38775.0]
source_col: slope_shift
SAMP_d18O_STD | 0.295501833046
varray: [-4.542, -4.002, -3.986, -3.89]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37164.0, 37701.0, 38239.0, 38775.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.0190761914
varray: [-43.252, -39.884, -39.147, -38.881]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37164.0, 37701.0, 38239.0, 38775.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 102.72975
varray: [88.797, 106.598, 105.431, 110.093]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [37164.0, 37701.0, 38239.0, 38775.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 9.49735195989
varray: [88.797, 106.598, 105.431, 110.093]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [37164.0, 37701.0, 38239.0, 38775.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.62925
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.298
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1285.32675
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.6135
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.46735
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.46735
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -15.0923593619
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -15.2615
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-4.542, -4.002, -3.986, -3.89]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [37164.0, 37701.0, 38239.0, 38775.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-43.252, -39.884, -39.147, -38.881]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [37164.0, 37701.0, 38239.0, 38775.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.105
varray: [-4.542, -4.002, -3.986, -3.89]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37164.0, 37701.0, 38239.0, 38775.0]
source_col: d(18_16)Mean
dD_mean | -40.291
varray: [-43.252, -39.884, -39.147, -38.881]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37164.0, 37701.0, 38239.0, 38775.0]
source_col: d(D_H)Mean
d18O_stdev | 0.295501833046
varray: [-4.542, -4.002, -3.986, -3.89]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37164.0, 37701.0, 38239.0, 38775.0]
source_col: d(18_16)Mean
dD_stdev | 2.0190761914
varray: [-43.252, -39.884, -39.147, -38.881]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37164.0, 37701.0, 38239.0, 38775.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.105
varray: [-4.542, -4.002, -3.986, -3.89]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37164.0, 37701.0, 38239.0, 38775.0]
source_col: d(18_16)Mean
measured_dD_mean | -40.291
varray: [-43.252, -39.884, -39.147, -38.881]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37164.0, 37701.0, 38239.0, 38775.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.295501833046
varray: [-4.542, -4.002, -3.986, -3.89]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37164.0, 37701.0, 38239.0, 38775.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.0190761914
varray: [-43.252, -39.884, -39.147, -38.881]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37164.0, 37701.0, 38239.0, 38775.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.88183265226
actual_array: [ 1.93 -15.84]
varray: [-4.542, -4.002, -3.986, -3.89]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -4.105
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [37164.0, 37701.0, 38239.0, 38775.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -55.7721307828
actual_array: [ 18.1 -119. ]
varray: [-43.252, -39.884, -39.147, -38.881]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -40.291
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [37164.0, 37701.0, 38239.0, 38775.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.785
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -18.0875
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 13 - Sample 13 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 13
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.0135
varray: [2.974, 3.051, 2.994, 3.035]
inst: mean_all
operation: mean_all(residuals)
tarray: [39366.0, 39902.0, 40439.0, 40975.0]
source_col: residuals
SAMP_residuals_STD | 0.0356323822761
varray: [2.974, 3.051, 2.994, 3.035]
inst: std_all
operation: std_all(residuals)
tarray: [39366.0, 39902.0, 40439.0, 40975.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3525
varray: [-0.075, 0.507, 0.412, 0.566]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [39366.0, 39902.0, 40439.0, 40975.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.291975455589
varray: [-0.075, 0.507, 0.412, 0.566]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [39366.0, 39902.0, 40439.0, 40975.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1300.12625
varray: [1301.214, 1300.128, 1299.488, 1299.675]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [39366.0, 39902.0, 40439.0, 40975.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.773346138543
varray: [1301.214, 1300.128, 1299.488, 1299.675]
inst: std_all
operation: std_all(baseline_shift)
tarray: [39366.0, 39902.0, 40439.0, 40975.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.67375
varray: [2.161, 1.251, 2.0, 1.283]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [39366.0, 39902.0, 40439.0, 40975.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.474431150607
varray: [2.161, 1.251, 2.0, 1.283]
inst: std_all
operation: std_all(slope_shift)
tarray: [39366.0, 39902.0, 40439.0, 40975.0]
source_col: slope_shift
SAMP_d18O_STD | 0.134092940406
varray: [-7.201, -7.38, -7.457, -7.507]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39366.0, 39902.0, 40439.0, 40975.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.06973996061
varray: [-56.688, -58.454, -58.619, -59.154]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39366.0, 39902.0, 40439.0, 40975.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 115.565
varray: [114.237, 117.109, 115.026, 115.888]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [39366.0, 39902.0, 40439.0, 40975.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.23049718949
varray: [114.237, 117.109, 115.026, 115.888]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [39366.0, 39902.0, 40439.0, 40975.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.0135
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3525
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1300.12625
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.67375
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0831
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0831
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.68358845185
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.42625
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-7.201, -7.38, -7.457, -7.507]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [39366.0, 39902.0, 40439.0, 40975.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-56.688, -58.454, -58.619, -59.154]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [39366.0, 39902.0, 40439.0, 40975.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.38625
varray: [-7.201, -7.38, -7.457, -7.507]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39366.0, 39902.0, 40439.0, 40975.0]
source_col: d(18_16)Mean
dD_mean | -58.22875
varray: [-56.688, -58.454, -58.619, -59.154]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39366.0, 39902.0, 40439.0, 40975.0]
source_col: d(D_H)Mean
d18O_stdev | 0.134092940406
varray: [-7.201, -7.38, -7.457, -7.507]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39366.0, 39902.0, 40439.0, 40975.0]
source_col: d(18_16)Mean
dD_stdev | 1.06973996061
varray: [-56.688, -58.454, -58.619, -59.154]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39366.0, 39902.0, 40439.0, 40975.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.38625
varray: [-7.201, -7.38, -7.457, -7.507]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39366.0, 39902.0, 40439.0, 40975.0]
source_col: d(18_16)Mean
measured_dD_mean | -58.22875
varray: [-56.688, -58.454, -58.619, -59.154]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39366.0, 39902.0, 40439.0, 40975.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.134092940406
varray: [-7.201, -7.38, -7.457, -7.507]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39366.0, 39902.0, 40439.0, 40975.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.06973996061
varray: [-56.688, -58.454, -58.619, -59.154]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39366.0, 39902.0, 40439.0, 40975.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.1683536623
actual_array: [ 1.93 -15.84]
varray: [-7.201, -7.38, -7.457, -7.507]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.38625
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [39366.0, 39902.0, 40439.0, 40975.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -72.6089078431
actual_array: [ 18.1 -119. ]
varray: [-56.688, -58.454, -58.619, -59.154]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -58.22875
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [39366.0, 39902.0, 40439.0, 40975.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.40075
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -3.288
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 14 - Sample 14 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 14
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.11025
varray: [0.0, 2.708, 3.025, 2.708]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [41565.0, 42103.0, 42640.0, 43177.0]
source_col: residuals
SAMP_residuals_STD | 1.41474765123
varray: [0.0, 2.708, 3.025, 2.708]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [41565.0, 42103.0, 42640.0, 43177.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.41275
varray: [0.0, 0.36, 0.931, 0.36]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [41565.0, 42103.0, 42640.0, 43177.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.384928889537
varray: [0.0, 0.36, 0.931, 0.36]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [41565.0, 42103.0, 42640.0, 43177.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 966.1105
varray: [0.0, 1284.884, 1294.674, 1284.884]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [41565.0, 42103.0, 42640.0, 43177.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 644.090200812
varray: [0.0, 1284.884, 1294.674, 1284.884]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [41565.0, 42103.0, 42640.0, 43177.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.794
varray: [0.0, 1.01, 1.156, 1.01]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [41565.0, 42103.0, 42640.0, 43177.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.533788972035
varray: [0.0, 1.01, 1.156, 1.01]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [41565.0, 42103.0, 42640.0, 43177.0]
source_col: slope_shift
SAMP_d18O_STD | 7.47371408672
varray: [0.0, 14.907, 15.027, 14.907]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41565.0, 42103.0, 42640.0, 43177.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 31.9028862056
varray: [0.0, 63.694, 64.027, 63.694]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41565.0, 42103.0, 42640.0, 43177.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 80.27825
varray: [0.0, 103.178, 114.757, 103.178]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [41565.0, 42103.0, 42640.0, 43177.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 53.7964643347
varray: [0.0, 103.178, 114.757, 103.178]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [41565.0, 42103.0, 42640.0, 43177.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.11025
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.41275
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 966.1105
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.794
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.98635
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.98635
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -31.8526771298
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -37.713
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [0.0, 14.907, 15.027, 14.907]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [41565.0, 42103.0, 42640.0, 43177.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [0.0, 63.694, 64.027, 63.694]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [41565.0, 42103.0, 42640.0, 43177.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 11.21025
varray: [0.0, 14.907, 15.027, 14.907]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [41565.0, 42103.0, 42640.0, 43177.0]
source_col: d(18_16)Mean
dD_mean | 47.85375
varray: [0.0, 63.694, 64.027, 63.694]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [41565.0, 42103.0, 42640.0, 43177.0]
source_col: d(D_H)Mean
d18O_stdev | 7.47371408672
varray: [0.0, 14.907, 15.027, 14.907]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41565.0, 42103.0, 42640.0, 43177.0]
source_col: d(18_16)Mean
dD_stdev | 31.9028862056
varray: [0.0, 63.694, 64.027, 63.694]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41565.0, 42103.0, 42640.0, 43177.0]
source_col: d(D_H)Mean
measured_d18O_mean | 11.21025
varray: [0.0, 14.907, 15.027, 14.907]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [41565.0, 42103.0, 42640.0, 43177.0]
source_col: d(18_16)Mean
measured_dD_mean | 47.85375
varray: [0.0, 63.694, 64.027, 63.694]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [41565.0, 42103.0, 42640.0, 43177.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 7.47371408672
varray: [0.0, 14.907, 15.027, 14.907]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41565.0, 42103.0, 42640.0, 43177.0]
source_col: d(18_16)Mean
measured_dD_stdev | 31.9028862056
varray: [0.0, 63.694, 64.027, 63.694]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41565.0, 42103.0, 42640.0, 43177.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 8.4580198123
actual_array: [ 1.93 -15.84]
varray: [0.0, 14.907, 15.027, 14.907]
fit_rsq: 1.0
fit_intercept: 2.76579538548
except: True
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 11.21025
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [41565.0, 42103.0, 42640.0, 43177.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 26.9625212662
actual_array: [ 18.1 -119. ]
varray: [0.0, 63.694, 64.027, 63.694]
fit_rsq: 1.0
fit_intercept: 19.1281237053
except: True
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 47.85375
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [41565.0, 42103.0, 42640.0, 43177.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1301.304
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -337.30375
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 15 - Sample 15 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 15
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.7375
varray: [2.678, 2.994, 2.684, 2.594]
inst: mean_all
operation: mean_all(residuals)
tarray: [43764.0, 44300.0, 44837.0, 45374.0]
source_col: residuals
SAMP_residuals_STD | 0.175866426586
varray: [2.678, 2.994, 2.684, 2.594]
inst: std_all
operation: std_all(residuals)
tarray: [43764.0, 44300.0, 44837.0, 45374.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3405
varray: [0.196, 0.46, 0.572, 0.134]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [43764.0, 44300.0, 44837.0, 45374.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.209280513506
varray: [0.196, 0.46, 0.572, 0.134]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [43764.0, 44300.0, 44837.0, 45374.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1290.13725
varray: [1286.694, 1299.596, 1287.764, 1286.495]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [43764.0, 44300.0, 44837.0, 45374.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 6.33040869534
varray: [1286.694, 1299.596, 1287.764, 1286.495]
inst: std_all
operation: std_all(baseline_shift)
tarray: [43764.0, 44300.0, 44837.0, 45374.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.77775
varray: [1.362, 1.379, 2.329, 2.041]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [43764.0, 44300.0, 44837.0, 45374.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.484777182494
varray: [1.362, 1.379, 2.329, 2.041]
inst: std_all
operation: std_all(slope_shift)
tarray: [43764.0, 44300.0, 44837.0, 45374.0]
source_col: slope_shift
SAMP_d18O_STD | 0.60057222713
varray: [-5.504, -6.254, -6.731, -6.813]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43764.0, 44300.0, 44837.0, 45374.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 6.27003032954
varray: [-42.282, -51.614, -55.091, -56.007]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43764.0, 44300.0, 44837.0, 45374.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 107.245
varray: [105.002, 116.397, 104.671, 102.91]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [43764.0, 44300.0, 44837.0, 45374.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 6.17003117226
varray: [105.002, 116.397, 104.671, 102.91]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [43764.0, 44300.0, 44837.0, 45374.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.7375
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3405
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1290.13725
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.77775
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.3591
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.3591
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -11.5965898082
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -10.74625
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-5.504, -6.254, -6.731, -6.813]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [43764.0, 44300.0, 44837.0, 45374.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-42.282, -51.614, -55.091, -56.007]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [43764.0, 44300.0, 44837.0, 45374.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.3255
varray: [-5.504, -6.254, -6.731, -6.813]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [43764.0, 44300.0, 44837.0, 45374.0]
source_col: d(18_16)Mean
dD_mean | -51.2485
varray: [-42.282, -51.614, -55.091, -56.007]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [43764.0, 44300.0, 44837.0, 45374.0]
source_col: d(D_H)Mean
d18O_stdev | 0.60057222713
varray: [-5.504, -6.254, -6.731, -6.813]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43764.0, 44300.0, 44837.0, 45374.0]
source_col: d(18_16)Mean
dD_stdev | 6.27003032954
varray: [-42.282, -51.614, -55.091, -56.007]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43764.0, 44300.0, 44837.0, 45374.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.3255
varray: [-5.504, -6.254, -6.731, -6.813]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [43764.0, 44300.0, 44837.0, 45374.0]
source_col: d(18_16)Mean
measured_dD_mean | -51.2485
varray: [-42.282, -51.614, -55.091, -56.007]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [43764.0, 44300.0, 44837.0, 45374.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.60057222713
varray: [-5.504, -6.254, -6.731, -6.813]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43764.0, 44300.0, 44837.0, 45374.0]
source_col: d(18_16)Mean
measured_dD_stdev | 6.27003032954
varray: [-42.282, -51.614, -55.091, -56.007]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43764.0, 44300.0, 44837.0, 45374.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.10589967026
actual_array: [ 1.93 -15.84]
varray: [-5.504, -6.254, -6.731, -6.813]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.3255
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [43764.0, 44300.0, 44837.0, 45374.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -66.0570876274
actual_array: [ 18.1 -119. ]
varray: [-42.282, -51.614, -55.091, -56.007]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -51.2485
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [43764.0, 44300.0, 44837.0, 45374.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.67675
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -13.277
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 16 - Sample 16 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 16
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.753
varray: [2.724, 2.925, 2.701, 2.662]
inst: mean_all
operation: mean_all(residuals)
tarray: [45962.0, 46499.0, 47036.0, 47573.0]
source_col: residuals
SAMP_residuals_STD | 0.117487587997
varray: [2.724, 2.925, 2.701, 2.662]
inst: std_all
operation: std_all(residuals)
tarray: [45962.0, 46499.0, 47036.0, 47573.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.50075
varray: [0.633, 0.717, 0.332, 0.321]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [45962.0, 46499.0, 47036.0, 47573.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.204157414756
varray: [0.633, 0.717, 0.332, 0.321]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [45962.0, 46499.0, 47036.0, 47573.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1288.83125
varray: [1286.438, 1294.389, 1288.167, 1286.331]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [45962.0, 46499.0, 47036.0, 47573.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.79950473746
varray: [1286.438, 1294.389, 1288.167, 1286.331]
inst: std_all
operation: std_all(baseline_shift)
tarray: [45962.0, 46499.0, 47036.0, 47573.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.78475
varray: [1.857, 1.23, 2.041, 2.011]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [45962.0, 46499.0, 47036.0, 47573.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.378514971082
varray: [1.857, 1.23, 2.041, 2.011]
inst: std_all
operation: std_all(slope_shift)
tarray: [45962.0, 46499.0, 47036.0, 47573.0]
source_col: slope_shift
SAMP_d18O_STD | 0.168705858424
varray: [-4.358, -3.969, -4.181, -4.282]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45962.0, 46499.0, 47036.0, 47573.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.637743613584
varray: [-50.174, -48.663, -49.207, -49.603]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45962.0, 46499.0, 47036.0, 47573.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 107.11975
varray: [105.861, 112.821, 105.994, 103.803]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [45962.0, 46499.0, 47036.0, 47573.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.93093910535
varray: [105.861, 112.821, 105.994, 103.803]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [45962.0, 46499.0, 47036.0, 47573.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.753
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.50075
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1288.83125
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.78475
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.3436
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.3436
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -11.096040819
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -10.8715
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-4.358, -3.969, -4.181, -4.282]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [45962.0, 46499.0, 47036.0, 47573.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-50.174, -48.663, -49.207, -49.603]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [45962.0, 46499.0, 47036.0, 47573.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.1975
varray: [-4.358, -3.969, -4.181, -4.282]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45962.0, 46499.0, 47036.0, 47573.0]
source_col: d(18_16)Mean
dD_mean | -49.41175
varray: [-50.174, -48.663, -49.207, -49.603]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45962.0, 46499.0, 47036.0, 47573.0]
source_col: d(D_H)Mean
d18O_stdev | 0.168705858424
varray: [-4.358, -3.969, -4.181, -4.282]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45962.0, 46499.0, 47036.0, 47573.0]
source_col: d(18_16)Mean
dD_stdev | 0.637743613584
varray: [-50.174, -48.663, -49.207, -49.603]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45962.0, 46499.0, 47036.0, 47573.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.1975
varray: [-4.358, -3.969, -4.181, -4.282]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45962.0, 46499.0, 47036.0, 47573.0]
source_col: d(18_16)Mean
measured_dD_mean | -49.41175
varray: [-50.174, -48.663, -49.207, -49.603]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45962.0, 46499.0, 47036.0, 47573.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.168705858424
varray: [-4.358, -3.969, -4.181, -4.282]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45962.0, 46499.0, 47036.0, 47573.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.637743613584
varray: [-50.174, -48.663, -49.207, -49.603]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45962.0, 46499.0, 47036.0, 47573.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.97448124454
actual_array: [ 1.93 -15.84]
varray: [-4.358, -3.969, -4.181, -4.282]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -4.1975
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [45962.0, 46499.0, 47036.0, 47573.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -64.3330726163
actual_array: [ 18.1 -119. ]
varray: [-50.174, -48.663, -49.207, -49.603]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -49.41175
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [45962.0, 46499.0, 47036.0, 47573.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.66125
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -14.583
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 17 - Sample 17 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 17
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8
varray: [2.799, 2.8, 2.803, 2.798]
inst: mean_all
operation: mean_all(residuals)
tarray: [48161.0, 48697.0, 49234.0, 49770.0]
source_col: residuals
SAMP_residuals_STD | 0.00216024689947
varray: [2.799, 2.8, 2.803, 2.798]
inst: std_all
operation: std_all(residuals)
tarray: [48161.0, 48697.0, 49234.0, 49770.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.31275
varray: [0.459, 0.277, 0.397, 0.118]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [48161.0, 48697.0, 49234.0, 49770.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.150214013993
varray: [0.459, 0.277, 0.397, 0.118]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [48161.0, 48697.0, 49234.0, 49770.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1290.96725
varray: [1289.447, 1291.149, 1291.149, 1292.124]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [48161.0, 48697.0, 49234.0, 49770.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.11284870939
varray: [1289.447, 1291.149, 1291.149, 1292.124]
inst: std_all
operation: std_all(baseline_shift)
tarray: [48161.0, 48697.0, 49234.0, 49770.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.679
varray: [1.75, 1.479, 1.897, 1.59]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [48161.0, 48697.0, 49234.0, 49770.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.183017303371
varray: [1.75, 1.479, 1.897, 1.59]
inst: std_all
operation: std_all(slope_shift)
tarray: [48161.0, 48697.0, 49234.0, 49770.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0415
varray: [-3.312, -3.24, -3.222, -3.229]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48161.0, 48697.0, 49234.0, 49770.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.99676092085
varray: [-33.383, -31.682, -31.168, -31.453]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48161.0, 48697.0, 49234.0, 49770.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 107.823
varray: [107.17, 107.594, 108.016, 108.512]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [48161.0, 48697.0, 49234.0, 49770.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.574694121541
varray: [107.17, 107.594, 108.016, 108.512]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [48161.0, 48697.0, 49234.0, 49770.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.31275
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1290.96725
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.679
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.2966
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.2966
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -9.57824710973
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -10.16825
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-3.312, -3.24, -3.222, -3.229]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [48161.0, 48697.0, 49234.0, 49770.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-33.383, -31.682, -31.168, -31.453]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [48161.0, 48697.0, 49234.0, 49770.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.25075
varray: [-3.312, -3.24, -3.222, -3.229]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48161.0, 48697.0, 49234.0, 49770.0]
source_col: d(18_16)Mean
dD_mean | -31.9215
varray: [-33.383, -31.682, -31.168, -31.453]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48161.0, 48697.0, 49234.0, 49770.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0415
varray: [-3.312, -3.24, -3.222, -3.229]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48161.0, 48697.0, 49234.0, 49770.0]
source_col: d(18_16)Mean
dD_stdev | 0.99676092085
varray: [-33.383, -31.682, -31.168, -31.453]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48161.0, 48697.0, 49234.0, 49770.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.25075
varray: [-3.312, -3.24, -3.222, -3.229]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48161.0, 48697.0, 49234.0, 49770.0]
source_col: d(18_16)Mean
measured_dD_mean | -31.9215
varray: [-33.383, -31.682, -31.168, -31.453]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48161.0, 48697.0, 49234.0, 49770.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0415
varray: [-3.312, -3.24, -3.222, -3.229]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48161.0, 48697.0, 49234.0, 49770.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.99676092085
varray: [-33.383, -31.682, -31.168, -31.453]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48161.0, 48697.0, 49234.0, 49770.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.02621038244
actual_array: [ 1.93 -15.84]
varray: [-3.312, -3.24, -3.222, -3.229]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -3.25075
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [48161.0, 48697.0, 49234.0, 49770.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -47.9163291573
actual_array: [ 18.1 -119. ]
varray: [-33.383, -31.682, -31.168, -31.453]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -31.9215
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [48161.0, 48697.0, 49234.0, 49770.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.61425
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -12.447
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 18 - Sample 18 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 18
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.96075
varray: [2.871, 3.054, 2.953, 2.965]
inst: mean_all
operation: mean_all(residuals)
tarray: [50359.0, 50896.0, 51433.0, 51970.0]
source_col: residuals
SAMP_residuals_STD | 0.0748971517037
varray: [2.871, 3.054, 2.953, 2.965]
inst: std_all
operation: std_all(residuals)
tarray: [50359.0, 50896.0, 51433.0, 51970.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.36225
varray: [0.113, 0.506, 0.437, 0.393]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [50359.0, 50896.0, 51433.0, 51970.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.172552166025
varray: [0.113, 0.506, 0.437, 0.393]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [50359.0, 50896.0, 51433.0, 51970.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1298.607
varray: [1296.393, 1301.035, 1298.004, 1298.996]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [50359.0, 50896.0, 51433.0, 51970.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.94181272698
varray: [1296.393, 1301.035, 1298.004, 1298.996]
inst: std_all
operation: std_all(baseline_shift)
tarray: [50359.0, 50896.0, 51433.0, 51970.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.56925
varray: [2.203, 0.998, 1.7, 1.376]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [50359.0, 50896.0, 51433.0, 51970.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.510688016307
varray: [2.203, 0.998, 1.7, 1.376]
inst: std_all
operation: std_all(slope_shift)
tarray: [50359.0, 50896.0, 51433.0, 51970.0]
source_col: slope_shift
SAMP_d18O_STD | 0.107632089391
varray: [-2.454, -2.207, -2.399, -2.392]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50359.0, 50896.0, 51433.0, 51970.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.240233913787
varray: [-31.79, -31.349, -31.87, -31.525]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50359.0, 50896.0, 51433.0, 51970.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 114.544
varray: [110.441, 117.612, 114.672, 115.451]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [50359.0, 50896.0, 51433.0, 51970.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.00478873356
varray: [110.441, 117.612, 114.672, 115.451]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [50359.0, 50896.0, 51433.0, 51970.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.96075
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.36225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1298.607
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.56925
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.13585
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.13585
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.38706968934
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.44725
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-2.454, -2.207, -2.399, -2.392]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [50359.0, 50896.0, 51433.0, 51970.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-31.79, -31.349, -31.87, -31.525]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [50359.0, 50896.0, 51433.0, 51970.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.363
varray: [-2.454, -2.207, -2.399, -2.392]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50359.0, 50896.0, 51433.0, 51970.0]
source_col: d(18_16)Mean
dD_mean | -31.6335
varray: [-31.79, -31.349, -31.87, -31.525]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50359.0, 50896.0, 51433.0, 51970.0]
source_col: d(D_H)Mean
d18O_stdev | 0.107632089391
varray: [-2.454, -2.207, -2.399, -2.392]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50359.0, 50896.0, 51433.0, 51970.0]
source_col: d(18_16)Mean
dD_stdev | 0.240233913787
varray: [-31.79, -31.349, -31.87, -31.525]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50359.0, 50896.0, 51433.0, 51970.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.363
varray: [-2.454, -2.207, -2.399, -2.392]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50359.0, 50896.0, 51433.0, 51970.0]
source_col: d(18_16)Mean
measured_dD_mean | -31.6335
varray: [-31.79, -31.349, -31.87, -31.525]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50359.0, 50896.0, 51433.0, 51970.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.107632089391
varray: [-2.454, -2.207, -2.399, -2.392]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50359.0, 50896.0, 51433.0, 51970.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.240233913787
varray: [-31.79, -31.349, -31.87, -31.525]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50359.0, 50896.0, 51433.0, 51970.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.13703429811
actual_array: [ 1.93 -15.84]
varray: [-2.454, -2.207, -2.399, -2.392]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.363
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [50359.0, 50896.0, 51433.0, 51970.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -47.6460058563
actual_array: [ 18.1 -119. ]
varray: [-31.79, -31.349, -31.87, -31.525]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -31.6335
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [50359.0, 50896.0, 51433.0, 51970.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.4535
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -4.80725
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 19 - Sample 19 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 19
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.64325
varray: [2.566, 2.763, 2.582, 2.662]
inst: mean_all
operation: mean_all(residuals)
tarray: [52560.0, 53097.0, 53634.0, 54171.0]
source_col: residuals
SAMP_residuals_STD | 0.0902048594404
varray: [2.566, 2.763, 2.582, 2.662]
inst: std_all
operation: std_all(residuals)
tarray: [52560.0, 53097.0, 53634.0, 54171.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3615
varray: [0.173, 0.509, 0.174, 0.59]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [52560.0, 53097.0, 53634.0, 54171.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.219588251052
varray: [0.173, 0.509, 0.174, 0.59]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [52560.0, 53097.0, 53634.0, 54171.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1284.1515
varray: [1282.159, 1288.465, 1282.574, 1283.408]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [52560.0, 53097.0, 53634.0, 54171.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.92219306914
varray: [1282.159, 1288.465, 1282.574, 1283.408]
inst: std_all
operation: std_all(baseline_shift)
tarray: [52560.0, 53097.0, 53634.0, 54171.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.67425
varray: [1.762, 1.742, 1.637, 1.556]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [52560.0, 53097.0, 53634.0, 54171.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0960221328653
varray: [1.762, 1.742, 1.637, 1.556]
inst: std_all
operation: std_all(slope_shift)
tarray: [52560.0, 53097.0, 53634.0, 54171.0]
source_col: slope_shift
SAMP_d18O_STD | 0.133517789077
varray: [-6.831, -6.867, -7.064, -7.092]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52560.0, 53097.0, 53634.0, 54171.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.30877534996
varray: [-55.566, -57.028, -58.21, -58.409]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52560.0, 53097.0, 53634.0, 54171.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 102.78675
varray: [99.713, 107.522, 100.47, 103.442]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [52560.0, 53097.0, 53634.0, 54171.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.54340771716
varray: [99.713, 107.522, 100.47, 103.442]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [52560.0, 53097.0, 53634.0, 54171.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.64325
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3615
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1284.1515
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.67425
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.45335
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.45335
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -14.6402505974
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -15.2045
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-6.831, -6.867, -7.064, -7.092]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [52560.0, 53097.0, 53634.0, 54171.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-55.566, -57.028, -58.21, -58.409]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [52560.0, 53097.0, 53634.0, 54171.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.9635
varray: [-6.831, -6.867, -7.064, -7.092]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [52560.0, 53097.0, 53634.0, 54171.0]
source_col: d(18_16)Mean
dD_mean | -57.30325
varray: [-55.566, -57.028, -58.21, -58.409]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [52560.0, 53097.0, 53634.0, 54171.0]
source_col: d(D_H)Mean
d18O_stdev | 0.133517789077
varray: [-6.831, -6.867, -7.064, -7.092]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52560.0, 53097.0, 53634.0, 54171.0]
source_col: d(18_16)Mean
dD_stdev | 1.30877534996
varray: [-55.566, -57.028, -58.21, -58.409]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52560.0, 53097.0, 53634.0, 54171.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.9635
varray: [-6.831, -6.867, -7.064, -7.092]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [52560.0, 53097.0, 53634.0, 54171.0]
source_col: d(18_16)Mean
measured_dD_mean | -57.30325
varray: [-55.566, -57.028, -58.21, -58.409]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [52560.0, 53097.0, 53634.0, 54171.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.133517789077
varray: [-6.831, -6.867, -7.064, -7.092]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52560.0, 53097.0, 53634.0, 54171.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.30877534996
varray: [-55.566, -57.028, -58.21, -58.409]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52560.0, 53097.0, 53634.0, 54171.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.74492455542
actual_array: [ 1.93 -15.84]
varray: [-6.831, -6.867, -7.064, -7.092]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.9635
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [52560.0, 53097.0, 53634.0, 54171.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -71.7402126518
actual_array: [ 18.1 -119. ]
varray: [-55.566, -57.028, -58.21, -58.409]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -57.30325
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [52560.0, 53097.0, 53634.0, 54171.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.771
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -19.26275
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 20 - Sample 20 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 20
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.64
varray: [2.636, 2.713, 2.594, 2.617]
inst: mean_all
operation: mean_all(residuals)
tarray: [54760.0, 55297.0, 55835.0, 56372.0]
source_col: residuals
SAMP_residuals_STD | 0.051607492996
varray: [2.636, 2.713, 2.594, 2.617]
inst: std_all
operation: std_all(residuals)
tarray: [54760.0, 55297.0, 55835.0, 56372.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.44025
varray: [0.47, 0.28, 0.575, 0.436]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [54760.0, 55297.0, 55835.0, 56372.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.122121182984
varray: [0.47, 0.28, 0.575, 0.436]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [54760.0, 55297.0, 55835.0, 56372.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1283.9065
varray: [1283.051, 1287.724, 1282.124, 1282.727]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [54760.0, 55297.0, 55835.0, 56372.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.57382419757
varray: [1283.051, 1287.724, 1282.124, 1282.727]
inst: std_all
operation: std_all(baseline_shift)
tarray: [54760.0, 55297.0, 55835.0, 56372.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.76875
varray: [1.783, 1.801, 1.318, 2.173]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [54760.0, 55297.0, 55835.0, 56372.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.350160320425
varray: [1.783, 1.801, 1.318, 2.173]
inst: std_all
operation: std_all(slope_shift)
tarray: [54760.0, 55297.0, 55835.0, 56372.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0978412830387
varray: [-6.721, -6.503, -6.669, -6.692]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54760.0, 55297.0, 55835.0, 56372.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.807107799491
varray: [-49.391, -47.574, -48.141, -47.82]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54760.0, 55297.0, 55835.0, 56372.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 102.29475
varray: [102.513, 104.958, 101.146, 100.562]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [54760.0, 55297.0, 55835.0, 56372.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.9547022578
varray: [102.513, 104.958, 101.146, 100.562]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [54760.0, 55297.0, 55835.0, 56372.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.64
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.44025
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1283.9065
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.76875
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.4566
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.4566
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -14.7452044177
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -15.6965
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-6.721, -6.503, -6.669, -6.692]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [54760.0, 55297.0, 55835.0, 56372.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-49.391, -47.574, -48.141, -47.82]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [54760.0, 55297.0, 55835.0, 56372.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.64625
varray: [-6.721, -6.503, -6.669, -6.692]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [54760.0, 55297.0, 55835.0, 56372.0]
source_col: d(18_16)Mean
dD_mean | -48.2315
varray: [-49.391, -47.574, -48.141, -47.82]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [54760.0, 55297.0, 55835.0, 56372.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0978412830387
varray: [-6.721, -6.503, -6.669, -6.692]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54760.0, 55297.0, 55835.0, 56372.0]
source_col: d(18_16)Mean
dD_stdev | 0.807107799491
varray: [-49.391, -47.574, -48.141, -47.82]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54760.0, 55297.0, 55835.0, 56372.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.64625
varray: [-6.721, -6.503, -6.669, -6.692]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [54760.0, 55297.0, 55835.0, 56372.0]
source_col: d(18_16)Mean
measured_dD_mean | -48.2315
varray: [-49.391, -47.574, -48.141, -47.82]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [54760.0, 55297.0, 55835.0, 56372.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0978412830387
varray: [-6.721, -6.503, -6.669, -6.692]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54760.0, 55297.0, 55835.0, 56372.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.807107799491
varray: [-49.391, -47.574, -48.141, -47.82]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54760.0, 55297.0, 55835.0, 56372.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.42716492405
actual_array: [ 1.93 -15.84]
varray: [-6.721, -6.503, -6.669, -6.692]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.64625
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [54760.0, 55297.0, 55835.0, 56372.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -63.2252633244
actual_array: [ 18.1 -119. ]
varray: [-49.391, -47.574, -48.141, -47.82]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -48.2315
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [54760.0, 55297.0, 55835.0, 56372.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.77425
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -19.50775
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 21 - Sample 21 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 21
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.794
varray: [2.856, 2.892, 2.687, 2.741]
inst: mean_all
operation: mean_all(residuals)
tarray: [56961.0, 57498.0, 58035.0, 58572.0]
source_col: residuals
SAMP_residuals_STD | 0.0961006416905
varray: [2.856, 2.892, 2.687, 2.741]
inst: std_all
operation: std_all(residuals)
tarray: [56961.0, 57498.0, 58035.0, 58572.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.5005
varray: [0.416, 0.532, 0.448, 0.606]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [56961.0, 57498.0, 58035.0, 58572.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.085671854577
varray: [0.416, 0.532, 0.448, 0.606]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [56961.0, 57498.0, 58035.0, 58572.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1289.1095
varray: [1292.937, 1293.019, 1285.178, 1285.304]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [56961.0, 57498.0, 58035.0, 58572.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 4.46738062702
varray: [1292.937, 1293.019, 1285.178, 1285.304]
inst: std_all
operation: std_all(baseline_shift)
tarray: [56961.0, 57498.0, 58035.0, 58572.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.49625
varray: [1.678, 1.607, 1.215, 1.485]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [56961.0, 57498.0, 58035.0, 58572.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.203737371797
varray: [1.678, 1.607, 1.215, 1.485]
inst: std_all
operation: std_all(slope_shift)
tarray: [56961.0, 57498.0, 58035.0, 58572.0]
source_col: slope_shift
SAMP_d18O_STD | 0.296386431313
varray: [2.412, 3.042, 2.92, 3.023]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56961.0, 57498.0, 58035.0, 58572.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.9952335494
varray: [-11.651, -7.642, -8.016, -7.416]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56961.0, 57498.0, 58035.0, 58572.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 106.9085
varray: [109.563, 110.204, 103.531, 104.336]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [56961.0, 57498.0, 58035.0, 58572.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.46082615378
varray: [109.563, 110.204, 103.531, 104.336]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [56961.0, 57498.0, 58035.0, 58572.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.794
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.5005
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1289.1095
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.49625
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.3026
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.3026
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -9.77200800878
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -11.08275
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [2.412, 3.042, 2.92, 3.023]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [56961.0, 57498.0, 58035.0, 58572.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-11.651, -7.642, -8.016, -7.416]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [56961.0, 57498.0, 58035.0, 58572.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.84925
varray: [2.412, 3.042, 2.92, 3.023]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56961.0, 57498.0, 58035.0, 58572.0]
source_col: d(18_16)Mean
dD_mean | -8.68125
varray: [-11.651, -7.642, -8.016, -7.416]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56961.0, 57498.0, 58035.0, 58572.0]
source_col: d(D_H)Mean
d18O_stdev | 0.296386431313
varray: [2.412, 3.042, 2.92, 3.023]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56961.0, 57498.0, 58035.0, 58572.0]
source_col: d(18_16)Mean
dD_stdev | 1.9952335494
varray: [-11.651, -7.642, -8.016, -7.416]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56961.0, 57498.0, 58035.0, 58572.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.84925
varray: [2.412, 3.042, 2.92, 3.023]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56961.0, 57498.0, 58035.0, 58572.0]
source_col: d(18_16)Mean
measured_dD_mean | -8.68125
varray: [-11.651, -7.642, -8.016, -7.416]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56961.0, 57498.0, 58035.0, 58572.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.296386431313
varray: [2.412, 3.042, 2.92, 3.023]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56961.0, 57498.0, 58035.0, 58572.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.9952335494
varray: [-11.651, -7.642, -8.016, -7.416]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56961.0, 57498.0, 58035.0, 58572.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 0.0835886762675
actual_array: [ 1.93 -15.84]
varray: [2.412, 3.042, 2.92, 3.023]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 2.84925
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [56961.0, 57498.0, 58035.0, 58572.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -26.1025059032
actual_array: [ 18.1 -119. ]
varray: [-11.651, -7.642, -8.016, -7.416]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -8.68125
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [56961.0, 57498.0, 58035.0, 58572.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.62025
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -14.30475
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 22 - Sample 22 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 22
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.986
varray: [2.946, 2.993, 3.025, 2.98]
inst: mean_all
operation: mean_all(residuals)
tarray: [59158.0, 59695.0, 60232.0, 60769.0]
source_col: residuals
SAMP_residuals_STD | 0.0326904675199
varray: [2.946, 2.993, 3.025, 2.98]
inst: std_all
operation: std_all(residuals)
tarray: [59158.0, 59695.0, 60232.0, 60769.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.388
varray: [0.319, 0.335, 0.399, 0.499]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [59158.0, 59695.0, 60232.0, 60769.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0816741493171
varray: [0.319, 0.335, 0.399, 0.499]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [59158.0, 59695.0, 60232.0, 60769.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1299.48325
varray: [1299.422, 1299.64, 1299.894, 1298.977]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [59158.0, 59695.0, 60232.0, 60769.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.388727303732
varray: [1299.422, 1299.64, 1299.894, 1298.977]
inst: std_all
operation: std_all(baseline_shift)
tarray: [59158.0, 59695.0, 60232.0, 60769.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.89275
varray: [1.806, 1.909, 1.575, 2.281]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [59158.0, 59695.0, 60232.0, 60769.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.294104714912
varray: [1.806, 1.909, 1.575, 2.281]
inst: std_all
operation: std_all(slope_shift)
tarray: [59158.0, 59695.0, 60232.0, 60769.0]
source_col: slope_shift
SAMP_d18O_STD | 0.273163412631
varray: [-6.835, -7.273, -7.39, -7.431]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59158.0, 59695.0, 60232.0, 60769.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.51545163274
varray: [-53.081, -57.245, -58.231, -58.506]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59158.0, 59695.0, 60232.0, 60769.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 114.64175
varray: [114.036, 113.954, 115.65, 114.927]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [59158.0, 59695.0, 60232.0, 60769.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.803714034965
varray: [114.036, 113.954, 115.65, 114.927]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [59158.0, 59695.0, 60232.0, 60769.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.986
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.388
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1299.48325
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.89275
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1106
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1106
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.57165923917
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.3495
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-6.835, -7.273, -7.39, -7.431]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [59158.0, 59695.0, 60232.0, 60769.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-53.081, -57.245, -58.231, -58.506]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [59158.0, 59695.0, 60232.0, 60769.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.23225
varray: [-6.835, -7.273, -7.39, -7.431]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [59158.0, 59695.0, 60232.0, 60769.0]
source_col: d(18_16)Mean
dD_mean | -56.76575
varray: [-53.081, -57.245, -58.231, -58.506]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [59158.0, 59695.0, 60232.0, 60769.0]
source_col: d(D_H)Mean
d18O_stdev | 0.273163412631
varray: [-6.835, -7.273, -7.39, -7.431]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59158.0, 59695.0, 60232.0, 60769.0]
source_col: d(18_16)Mean
dD_stdev | 2.51545163274
varray: [-53.081, -57.245, -58.231, -58.506]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59158.0, 59695.0, 60232.0, 60769.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.23225
varray: [-6.835, -7.273, -7.39, -7.431]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [59158.0, 59695.0, 60232.0, 60769.0]
source_col: d(18_16)Mean
measured_dD_mean | -56.76575
varray: [-53.081, -57.245, -58.231, -58.506]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [59158.0, 59695.0, 60232.0, 60769.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.273163412631
varray: [-6.835, -7.273, -7.39, -7.431]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59158.0, 59695.0, 60232.0, 60769.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.51545163274
varray: [-53.081, -57.245, -58.231, -58.506]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59158.0, 59695.0, 60232.0, 60769.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.0141062762
actual_array: [ 1.93 -15.84]
varray: [-6.835, -7.273, -7.39, -7.431]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.23225
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [59158.0, 59695.0, 60232.0, 60769.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -71.2357030187
actual_array: [ 18.1 -119. ]
varray: [-53.081, -57.245, -58.231, -58.506]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -56.76575
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [59158.0, 59695.0, 60232.0, 60769.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.42825
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -3.931
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 23 - Sample 23 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 23
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.84075
varray: [2.887, 2.82, 2.822, 2.834]
inst: mean_all
operation: mean_all(residuals)
tarray: [61356.0, 61893.0, 62430.0, 62967.0]
source_col: residuals
SAMP_residuals_STD | 0.0314470454362
varray: [2.887, 2.82, 2.822, 2.834]
inst: std_all
operation: std_all(residuals)
tarray: [61356.0, 61893.0, 62430.0, 62967.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.355
varray: [0.507, 0.32, 0.243, 0.35]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [61356.0, 61893.0, 62430.0, 62967.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.110902359458
varray: [0.507, 0.32, 0.243, 0.35]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [61356.0, 61893.0, 62430.0, 62967.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1291.33125
varray: [1292.844, 1291.68, 1289.822, 1290.979]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [61356.0, 61893.0, 62430.0, 62967.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.26648526113
varray: [1292.844, 1291.68, 1289.822, 1290.979]
inst: std_all
operation: std_all(baseline_shift)
tarray: [61356.0, 61893.0, 62430.0, 62967.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.98175
varray: [1.218, 0.964, 0.879, 0.866]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [61356.0, 61893.0, 62430.0, 62967.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.163385790896
varray: [1.218, 0.964, 0.879, 0.866]
inst: std_all
operation: std_all(slope_shift)
tarray: [61356.0, 61893.0, 62430.0, 62967.0]
source_col: slope_shift
SAMP_d18O_STD | 0.56786999686
varray: [11.277, 12.322, 12.42, 12.475]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61356.0, 61893.0, 62430.0, 62967.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 5.76265367113
varray: [38.442, 48.366, 49.9, 51.016]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61356.0, 61893.0, 62430.0, 62967.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.2895
varray: [111.217, 109.469, 107.534, 108.938]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [61356.0, 61893.0, 62430.0, 62967.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.52236756403
varray: [111.217, 109.469, 107.534, 108.938]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [61356.0, 61893.0, 62430.0, 62967.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.84075
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.355
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1291.33125
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.98175
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.25585
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.25585
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -8.26228767035
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -8.70175
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [11.277, 12.322, 12.42, 12.475]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [61356.0, 61893.0, 62430.0, 62967.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [38.442, 48.366, 49.9, 51.016]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [61356.0, 61893.0, 62430.0, 62967.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 12.1235
varray: [11.277, 12.322, 12.42, 12.475]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [61356.0, 61893.0, 62430.0, 62967.0]
source_col: d(18_16)Mean
dD_mean | 46.931
varray: [38.442, 48.366, 49.9, 51.016]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [61356.0, 61893.0, 62430.0, 62967.0]
source_col: d(D_H)Mean
d18O_stdev | 0.56786999686
varray: [11.277, 12.322, 12.42, 12.475]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61356.0, 61893.0, 62430.0, 62967.0]
source_col: d(18_16)Mean
dD_stdev | 5.76265367113
varray: [38.442, 48.366, 49.9, 51.016]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61356.0, 61893.0, 62430.0, 62967.0]
source_col: d(D_H)Mean
measured_d18O_mean | 12.1235
varray: [11.277, 12.322, 12.42, 12.475]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [61356.0, 61893.0, 62430.0, 62967.0]
source_col: d(18_16)Mean
measured_dD_mean | 46.931
varray: [38.442, 48.366, 49.9, 51.016]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [61356.0, 61893.0, 62430.0, 62967.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.56786999686
varray: [11.277, 12.322, 12.42, 12.475]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61356.0, 61893.0, 62430.0, 62967.0]
source_col: d(18_16)Mean
measured_dD_stdev | 5.76265367113
varray: [38.442, 48.366, 49.9, 51.016]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61356.0, 61893.0, 62430.0, 62967.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 9.3727368599
actual_array: [ 1.93 -15.84]
varray: [11.277, 12.322, 12.42, 12.475]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 12.1235
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [61356.0, 61893.0, 62430.0, 62967.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 26.0964072869
actual_array: [ 18.1 -119. ]
varray: [38.442, 48.366, 49.9, 51.016]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 46.931
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [61356.0, 61893.0, 62430.0, 62967.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.5735
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -12.083
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 24 - Sample 24 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 24
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.17475
varray: [0.0, 3.025, 2.822, 2.852]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [63554.0, 64091.0, 64628.0, 65164.0]
source_col: residuals
SAMP_residuals_STD | 1.45259110443
varray: [0.0, 3.025, 2.822, 2.852]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [63554.0, 64091.0, 64628.0, 65164.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.33075
varray: [0.0, 0.235, 0.509, 0.579]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [63554.0, 64091.0, 64628.0, 65164.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.265809925824
varray: [0.0, 0.235, 0.509, 0.579]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [63554.0, 64091.0, 64628.0, 65164.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 971.18775
varray: [0.0, 1303.154, 1290.541, 1291.056]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [63554.0, 64091.0, 64628.0, 65164.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 647.484731489
varray: [0.0, 1303.154, 1290.541, 1291.056]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [63554.0, 64091.0, 64628.0, 65164.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.92
varray: [0.0, 1.012, 1.387, 1.281]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [63554.0, 64091.0, 64628.0, 65164.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.633317719106
varray: [0.0, 1.012, 1.387, 1.281]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [63554.0, 64091.0, 64628.0, 65164.0]
source_col: slope_shift
SAMP_d18O_STD | 1.16414503678
varray: [0.0, -2.027, -2.414, -2.449]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63554.0, 64091.0, 64628.0, 65164.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 10.8079525967
varray: [0.0, -19.615, -21.897, -22.828]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63554.0, 64091.0, 64628.0, 65164.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 84.26825
varray: [0.0, 117.747, 109.444, 109.882]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [63554.0, 64091.0, 64628.0, 65164.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 56.3082209414
varray: [0.0, 117.747, 109.444, 109.882]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [63554.0, 64091.0, 64628.0, 65164.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.17475
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.33075
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 971.18775
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.92
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.92185
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.92185
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -29.769747465
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -33.723
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [0.0, -2.027, -2.414, -2.449]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [63554.0, 64091.0, 64628.0, 65164.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [0.0, -19.615, -21.897, -22.828]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [63554.0, 64091.0, 64628.0, 65164.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.7225
varray: [0.0, -2.027, -2.414, -2.449]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63554.0, 64091.0, 64628.0, 65164.0]
source_col: d(18_16)Mean
dD_mean | -16.085
varray: [0.0, -19.615, -21.897, -22.828]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63554.0, 64091.0, 64628.0, 65164.0]
source_col: d(D_H)Mean
d18O_stdev | 1.16414503678
varray: [0.0, -2.027, -2.414, -2.449]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63554.0, 64091.0, 64628.0, 65164.0]
source_col: d(18_16)Mean
dD_stdev | 10.8079525967
varray: [0.0, -19.615, -21.897, -22.828]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63554.0, 64091.0, 64628.0, 65164.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.7225
varray: [0.0, -2.027, -2.414, -2.449]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63554.0, 64091.0, 64628.0, 65164.0]
source_col: d(18_16)Mean
measured_dD_mean | -16.085
varray: [0.0, -19.615, -21.897, -22.828]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63554.0, 64091.0, 64628.0, 65164.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 1.16414503678
varray: [0.0, -2.027, -2.414, -2.449]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63554.0, 64091.0, 64628.0, 65164.0]
source_col: d(18_16)Mean
measured_dD_stdev | 10.8079525967
varray: [0.0, -19.615, -21.897, -22.828]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63554.0, 64091.0, 64628.0, 65164.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.49550539695
actual_array: [ 1.93 -15.84]
varray: [0.0, -2.027, -2.414, -2.449]
fit_rsq: 1.0
fit_intercept: 2.76579538548
except: True
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -1.7225
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [63554.0, 64091.0, 64628.0, 65164.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -33.0518327787
actual_array: [ 18.1 -119. ]
varray: [0.0, -19.615, -21.897, -22.828]
fit_rsq: 1.0
fit_intercept: 19.1281237053
except: True
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -16.085
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [63554.0, 64091.0, 64628.0, 65164.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1301.2395
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -332.2265
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 25 - Sample 25 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 25
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.9755
varray: [2.685, 0.0, 2.522, 2.695]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [65752.0, 66289.0, 66825.0, 67362.0]
source_col: residuals
SAMP_residuals_STD | 1.31938533669
varray: [2.685, 0.0, 2.522, 2.695]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [65752.0, 66289.0, 66825.0, 67362.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.28575
varray: [0.376, 0.0, 0.174, 0.593]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [65752.0, 66289.0, 66825.0, 67362.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.256052566218
varray: [0.376, 0.0, 0.174, 0.593]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [65752.0, 66289.0, 66825.0, 67362.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 963.18375
varray: [1285.187, 0.0, 1282.227, 1285.321]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [65752.0, 66289.0, 66825.0, 67362.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 642.124087822
varray: [1285.187, 0.0, 1282.227, 1285.321]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [65752.0, 66289.0, 66825.0, 67362.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.04625
varray: [1.09, 0.0, 1.804, 1.291]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [65752.0, 66289.0, 66825.0, 67362.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.759526332657
varray: [1.09, 0.0, 1.804, 1.291]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [65752.0, 66289.0, 66825.0, 67362.0]
source_col: slope_shift
SAMP_d18O_STD | 0.518400745884
varray: [-1.063, 0.0, -0.988, -1.053]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65752.0, 66289.0, 66825.0, 67362.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 14.7929179227
varray: [-29.011, 0.0, -29.654, -30.055]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65752.0, 66289.0, 66825.0, 67362.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 77.2535
varray: [103.813, 0.0, 99.676, 105.525]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [65752.0, 66289.0, 66825.0, 67362.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 51.5608266064
varray: [103.813, 0.0, 99.676, 105.525]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [65752.0, 66289.0, 66825.0, 67362.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.9755
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.28575
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 963.18375
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.04625
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.1211
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.1211
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -36.2042239876
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -40.73775
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-1.063, 0.0, -0.988, -1.053]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [65752.0, 66289.0, 66825.0, 67362.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-29.011, 0.0, -29.654, -30.055]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [65752.0, 66289.0, 66825.0, 67362.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.776
varray: [-1.063, 0.0, -0.988, -1.053]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65752.0, 66289.0, 66825.0, 67362.0]
source_col: d(18_16)Mean
dD_mean | -22.18
varray: [-29.011, 0.0, -29.654, -30.055]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65752.0, 66289.0, 66825.0, 67362.0]
source_col: d(D_H)Mean
d18O_stdev | 0.518400745884
varray: [-1.063, 0.0, -0.988, -1.053]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65752.0, 66289.0, 66825.0, 67362.0]
source_col: d(18_16)Mean
dD_stdev | 14.7929179227
varray: [-29.011, 0.0, -29.654, -30.055]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65752.0, 66289.0, 66825.0, 67362.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.776
varray: [-1.063, 0.0, -0.988, -1.053]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65752.0, 66289.0, 66825.0, 67362.0]
source_col: d(18_16)Mean
measured_dD_mean | -22.18
varray: [-29.011, 0.0, -29.654, -30.055]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65752.0, 66289.0, 66825.0, 67362.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.518400745884
varray: [-1.063, 0.0, -0.988, -1.053]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65752.0, 66289.0, 66825.0, 67362.0]
source_col: d(18_16)Mean
measured_dD_stdev | 14.7929179227
varray: [-29.011, 0.0, -29.654, -30.055]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65752.0, 66289.0, 66825.0, 67362.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.54748493645
actual_array: [ 1.93 -15.84]
varray: [-1.063, 0.0, -0.988, -1.053]
fit_rsq: 1.0
fit_intercept: 2.76579538548
except: True
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -0.776
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [65752.0, 66289.0, 66825.0, 67362.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -38.7727373616
actual_array: [ 18.1 -119. ]
varray: [-29.011, 0.0, -29.654, -30.055]
fit_rsq: 1.0
fit_intercept: 19.1281237053
except: True
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -22.18
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [65752.0, 66289.0, 66825.0, 67362.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1301.43875
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -340.2305
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 26 - Sample 26 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 26
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.93025
varray: [2.755, 3.188, 2.932, 2.846]
inst: mean_all
operation: mean_all(residuals)
tarray: [67950.0, 68486.0, 69023.0, 69560.0]
source_col: residuals
SAMP_residuals_STD | 0.186412401233
varray: [2.755, 3.188, 2.932, 2.846]
inst: std_all
operation: std_all(residuals)
tarray: [67950.0, 68486.0, 69023.0, 69560.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.5695
varray: [0.478, 0.857, 0.442, 0.501]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [67950.0, 68486.0, 69023.0, 69560.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.193198516212
varray: [0.478, 0.857, 0.442, 0.501]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [67950.0, 68486.0, 69023.0, 69560.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1294.65325
varray: [1288.976, 1303.183, 1293.583, 1292.871]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [67950.0, 68486.0, 69023.0, 69560.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 6.03626961928
varray: [1288.976, 1303.183, 1293.583, 1292.871]
inst: std_all
operation: std_all(baseline_shift)
tarray: [67950.0, 68486.0, 69023.0, 69560.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.25775
varray: [1.428, 1.216, 1.003, 1.384]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [67950.0, 68486.0, 69023.0, 69560.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.192842552358
varray: [1.428, 1.216, 1.003, 1.384]
inst: std_all
operation: std_all(slope_shift)
tarray: [67950.0, 68486.0, 69023.0, 69560.0]
source_col: slope_shift
SAMP_d18O_STD | 0.350410592496
varray: [5.553, 6.411, 6.002, 5.993]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67950.0, 68486.0, 69023.0, 69560.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.81977237557
varray: [7.71, 14.018, 12.809, 12.896]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67950.0, 68486.0, 69023.0, 69560.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 112.68225
varray: [107.069, 121.276, 111.653, 110.731]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [67950.0, 68486.0, 69023.0, 69560.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 6.06156873947
varray: [107.069, 121.276, 111.653, 110.731]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [67950.0, 68486.0, 69023.0, 69560.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.93025
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.5695
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1294.65325
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.25775
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.16635
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.16635
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.37202092618
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.309
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [5.553, 6.411, 6.002, 5.993]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [67950.0, 68486.0, 69023.0, 69560.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [7.71, 14.018, 12.809, 12.896]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [67950.0, 68486.0, 69023.0, 69560.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 5.98975
varray: [5.553, 6.411, 6.002, 5.993]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [67950.0, 68486.0, 69023.0, 69560.0]
source_col: d(18_16)Mean
dD_mean | 11.85825
varray: [7.71, 14.018, 12.809, 12.896]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [67950.0, 68486.0, 69023.0, 69560.0]
source_col: d(D_H)Mean
d18O_stdev | 0.350410592496
varray: [5.553, 6.411, 6.002, 5.993]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67950.0, 68486.0, 69023.0, 69560.0]
source_col: d(18_16)Mean
dD_stdev | 2.81977237557
varray: [7.71, 14.018, 12.809, 12.896]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67950.0, 68486.0, 69023.0, 69560.0]
source_col: d(D_H)Mean
measured_d18O_mean | 5.98975
varray: [5.553, 6.411, 6.002, 5.993]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [67950.0, 68486.0, 69023.0, 69560.0]
source_col: d(18_16)Mean
measured_dD_mean | 11.85825
varray: [7.71, 14.018, 12.809, 12.896]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [67950.0, 68486.0, 69023.0, 69560.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.350410592496
varray: [5.553, 6.411, 6.002, 5.993]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67950.0, 68486.0, 69023.0, 69560.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.81977237557
varray: [7.71, 14.018, 12.809, 12.896]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67950.0, 68486.0, 69023.0, 69560.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 3.2291335851
actual_array: [ 1.93 -15.84]
varray: [5.553, 6.411, 6.002, 5.993]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 5.98975
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [67950.0, 68486.0, 69023.0, 69560.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -6.8236675633
actual_array: [ 18.1 -119. ]
varray: [7.71, 14.018, 12.809, 12.896]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 11.85825
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [67950.0, 68486.0, 69023.0, 69560.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.484
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -8.761
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 27 - Sample 27 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 27
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.7165
varray: [2.743, 2.943, 2.603, 2.577]
inst: mean_all
operation: mean_all(residuals)
tarray: [70149.0, 70686.0, 71222.0, 71760.0]
source_col: residuals
SAMP_residuals_STD | 0.16767726938
varray: [2.743, 2.943, 2.603, 2.577]
inst: std_all
operation: std_all(residuals)
tarray: [70149.0, 70686.0, 71222.0, 71760.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3565
varray: [0.262, 0.594, 0.28, 0.29]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [70149.0, 70686.0, 71222.0, 71760.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.158756627158
varray: [0.262, 0.594, 0.28, 0.29]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [70149.0, 70686.0, 71222.0, 71760.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1287.60425
varray: [1289.823, 1295.25, 1282.7, 1282.644]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [70149.0, 70686.0, 71222.0, 71760.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 6.11108535777
varray: [1289.823, 1295.25, 1282.7, 1282.644]
inst: std_all
operation: std_all(baseline_shift)
tarray: [70149.0, 70686.0, 71222.0, 71760.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.721
varray: [1.275, 1.724, 1.927, 1.958]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [70149.0, 70686.0, 71222.0, 71760.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.314923271078
varray: [1.275, 1.724, 1.927, 1.958]
inst: std_all
operation: std_all(slope_shift)
tarray: [70149.0, 70686.0, 71222.0, 71760.0]
source_col: slope_shift
SAMP_d18O_STD | 0.381675691305
varray: [-5.955, -6.365, -6.708, -6.794]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70149.0, 70686.0, 71222.0, 71760.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.94500674904
varray: [-47.445, -53.137, -55.39, -56.179]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70149.0, 70686.0, 71222.0, 71760.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 105.58675
varray: [107.239, 113.194, 101.008, 100.906]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [70149.0, 70686.0, 71222.0, 71760.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 5.87294800335
varray: [107.239, 113.194, 101.008, 100.906]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [70149.0, 70686.0, 71222.0, 71760.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.7165
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3565
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1287.60425
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.721
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.3801
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.3801
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -12.2747529549
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -12.4045
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-5.955, -6.365, -6.708, -6.794]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [70149.0, 70686.0, 71222.0, 71760.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-47.445, -53.137, -55.39, -56.179]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [70149.0, 70686.0, 71222.0, 71760.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.4555
varray: [-5.955, -6.365, -6.708, -6.794]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [70149.0, 70686.0, 71222.0, 71760.0]
source_col: d(18_16)Mean
dD_mean | -53.03775
varray: [-47.445, -53.137, -55.39, -56.179]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [70149.0, 70686.0, 71222.0, 71760.0]
source_col: d(D_H)Mean
d18O_stdev | 0.381675691305
varray: [-5.955, -6.365, -6.708, -6.794]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70149.0, 70686.0, 71222.0, 71760.0]
source_col: d(18_16)Mean
dD_stdev | 3.94500674904
varray: [-47.445, -53.137, -55.39, -56.179]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70149.0, 70686.0, 71222.0, 71760.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.4555
varray: [-5.955, -6.365, -6.708, -6.794]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [70149.0, 70686.0, 71222.0, 71760.0]
source_col: d(18_16)Mean
measured_dD_mean | -53.03775
varray: [-47.445, -53.137, -55.39, -56.179]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [70149.0, 70686.0, 71222.0, 71760.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.381675691305
varray: [-5.955, -6.365, -6.708, -6.794]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70149.0, 70686.0, 71222.0, 71760.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.94500674904
varray: [-47.445, -53.137, -55.39, -56.179]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70149.0, 70686.0, 71222.0, 71760.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.23610850266
actual_array: [ 1.93 -15.84]
varray: [-5.955, -6.365, -6.708, -6.794]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.4555
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [70149.0, 70686.0, 71222.0, 71760.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -67.7365180663
actual_array: [ 18.1 -119. ]
varray: [-47.445, -53.137, -55.39, -56.179]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -53.03775
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [70149.0, 70686.0, 71222.0, 71760.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.69775
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -15.81
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 28 - Sample 28 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 28
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.5875
varray: [2.767, 2.625, 2.472, 2.486]
inst: mean_all
operation: mean_all(residuals)
tarray: [72349.0, 72886.0, 73424.0, 73961.0]
source_col: residuals
SAMP_residuals_STD | 0.138165359865
varray: [2.767, 2.625, 2.472, 2.486]
inst: std_all
operation: std_all(residuals)
tarray: [72349.0, 72886.0, 73424.0, 73961.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.33025
varray: [0.513, 0.289, 0.276, 0.243]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [72349.0, 72886.0, 73424.0, 73961.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.12336227138
varray: [0.513, 0.289, 0.276, 0.243]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [72349.0, 72886.0, 73424.0, 73961.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1283.53525
varray: [1289.458, 1285.774, 1279.26, 1279.649]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [72349.0, 72886.0, 73424.0, 73961.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 4.94879348091
varray: [1289.458, 1285.774, 1279.26, 1279.649]
inst: std_all
operation: std_all(baseline_shift)
tarray: [72349.0, 72886.0, 73424.0, 73961.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.9765
varray: [1.87, 1.831, 2.309, 1.896]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [72349.0, 72886.0, 73424.0, 73961.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.22327038914
varray: [1.87, 1.831, 2.309, 1.896]
inst: std_all
operation: std_all(slope_shift)
tarray: [72349.0, 72886.0, 73424.0, 73961.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0501530989405
varray: [-6.022, -5.941, -6.051, -6.042]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72349.0, 72886.0, 73424.0, 73961.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.349413317815
varray: [-48.814, -48.251, -48.76, -49.089]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72349.0, 72886.0, 73424.0, 73961.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 101.234
varray: [107.557, 103.088, 96.779, 97.512]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [72349.0, 72886.0, 73424.0, 73961.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 5.07010696797
varray: [107.557, 103.088, 96.779, 97.512]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [72349.0, 72886.0, 73424.0, 73961.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.5875
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.33025
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1283.53525
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.9765
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.5091
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.5091
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -16.4406122844
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -16.75725
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-6.022, -5.941, -6.051, -6.042]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [72349.0, 72886.0, 73424.0, 73961.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-48.814, -48.251, -48.76, -49.089]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [72349.0, 72886.0, 73424.0, 73961.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.014
varray: [-6.022, -5.941, -6.051, -6.042]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72349.0, 72886.0, 73424.0, 73961.0]
source_col: d(18_16)Mean
dD_mean | -48.7285
varray: [-48.814, -48.251, -48.76, -49.089]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72349.0, 72886.0, 73424.0, 73961.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0501530989405
varray: [-6.022, -5.941, -6.051, -6.042]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72349.0, 72886.0, 73424.0, 73961.0]
source_col: d(18_16)Mean
dD_stdev | 0.349413317815
varray: [-48.814, -48.251, -48.76, -49.089]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72349.0, 72886.0, 73424.0, 73961.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.014
varray: [-6.022, -5.941, -6.051, -6.042]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72349.0, 72886.0, 73424.0, 73961.0]
source_col: d(18_16)Mean
measured_dD_mean | -48.7285
varray: [-48.814, -48.251, -48.76, -49.089]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72349.0, 72886.0, 73424.0, 73961.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0501530989405
varray: [-6.022, -5.941, -6.051, -6.042]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72349.0, 72886.0, 73424.0, 73961.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.349413317815
varray: [-48.814, -48.251, -48.76, -49.089]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72349.0, 72886.0, 73424.0, 73961.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.79389927571
actual_array: [ 1.93 -15.84]
varray: [-6.022, -5.941, -6.051, -6.042]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.014
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [72349.0, 72886.0, 73424.0, 73961.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -63.6917587432
actual_array: [ 18.1 -119. ]
varray: [-48.814, -48.251, -48.76, -49.089]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -48.7285
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [72349.0, 72886.0, 73424.0, 73961.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.82675
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -19.879
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 29 - Sample 29 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 29
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.572
varray: [2.489, 2.726, 2.505, 2.568]
inst: mean_all
operation: mean_all(residuals)
tarray: [74545.0, 75082.0, 75620.0, 76156.0]
source_col: residuals
SAMP_residuals_STD | 0.108181945505
varray: [2.489, 2.726, 2.505, 2.568]
inst: std_all
operation: std_all(residuals)
tarray: [74545.0, 75082.0, 75620.0, 76156.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.314
varray: [0.125, 0.341, 0.201, 0.589]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [74545.0, 75082.0, 75620.0, 76156.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.203996732
varray: [0.125, 0.341, 0.201, 0.589]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [74545.0, 75082.0, 75620.0, 76156.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1281.44425
varray: [1278.46, 1286.367, 1280.403, 1280.547]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [74545.0, 75082.0, 75620.0, 76156.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.41703939837
varray: [1278.46, 1286.367, 1280.403, 1280.547]
inst: std_all
operation: std_all(baseline_shift)
tarray: [74545.0, 75082.0, 75620.0, 76156.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.71775
varray: [2.295, 1.541, 1.344, 1.691]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [74545.0, 75082.0, 75620.0, 76156.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.41022869638
varray: [2.295, 1.541, 1.344, 1.691]
inst: std_all
operation: std_all(slope_shift)
tarray: [74545.0, 75082.0, 75620.0, 76156.0]
source_col: slope_shift
SAMP_d18O_STD | 0.115542200083
varray: [-6.286, -6.067, -6.328, -6.259]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74545.0, 75082.0, 75620.0, 76156.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.468032050185
varray: [-50.408, -49.357, -50.273, -50.078]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74545.0, 75082.0, 75620.0, 76156.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 100.1675
varray: [96.299, 104.583, 98.884, 100.904]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [74545.0, 75082.0, 75620.0, 76156.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.4953176775
varray: [96.299, 104.583, 98.884, 100.904]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [74545.0, 75082.0, 75620.0, 76156.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.572
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.314
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1281.44425
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.71775
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.5246
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.5246
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -16.9411612737
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -17.82375
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-6.286, -6.067, -6.328, -6.259]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [74545.0, 75082.0, 75620.0, 76156.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-50.408, -49.357, -50.273, -50.078]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [74545.0, 75082.0, 75620.0, 76156.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.235
varray: [-6.286, -6.067, -6.328, -6.259]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [74545.0, 75082.0, 75620.0, 76156.0]
source_col: d(18_16)Mean
dD_mean | -50.029
varray: [-50.408, -49.357, -50.273, -50.078]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [74545.0, 75082.0, 75620.0, 76156.0]
source_col: d(D_H)Mean
d18O_stdev | 0.115542200083
varray: [-6.286, -6.067, -6.328, -6.259]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74545.0, 75082.0, 75620.0, 76156.0]
source_col: d(18_16)Mean
dD_stdev | 0.468032050185
varray: [-50.408, -49.357, -50.273, -50.078]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74545.0, 75082.0, 75620.0, 76156.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.235
varray: [-6.286, -6.067, -6.328, -6.259]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [74545.0, 75082.0, 75620.0, 76156.0]
source_col: d(18_16)Mean
measured_dD_mean | -50.029
varray: [-50.408, -49.357, -50.273, -50.078]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [74545.0, 75082.0, 75620.0, 76156.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.115542200083
varray: [-6.286, -6.067, -6.328, -6.259]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74545.0, 75082.0, 75620.0, 76156.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.468032050185
varray: [-50.408, -49.357, -50.273, -50.078]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74545.0, 75082.0, 75620.0, 76156.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.01525429079
actual_array: [ 1.93 -15.84]
varray: [-6.286, -6.067, -6.328, -6.259]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.235
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [74545.0, 75082.0, 75620.0, 76156.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -64.9124373995
actual_array: [ 18.1 -119. ]
varray: [-50.408, -49.357, -50.273, -50.078]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -50.029
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [74545.0, 75082.0, 75620.0, 76156.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.84225
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -21.97
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 30 - Sample 30 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 30
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.87575
varray: [2.879, 2.895, 2.842, 2.887]
inst: mean_all
operation: mean_all(residuals)
tarray: [76744.0, 77280.0, 77816.0, 78353.0]
source_col: residuals
SAMP_residuals_STD | 0.0234289706702
varray: [2.879, 2.895, 2.842, 2.887]
inst: std_all
operation: std_all(residuals)
tarray: [76744.0, 77280.0, 77816.0, 78353.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.29975
varray: [0.336, 0.294, 0.245, 0.324]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [76744.0, 77280.0, 77816.0, 78353.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0405493526459
varray: [0.336, 0.294, 0.245, 0.324]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [76744.0, 77280.0, 77816.0, 78353.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1294.96275
varray: [1296.059, 1295.693, 1294.09, 1294.009]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [76744.0, 77280.0, 77816.0, 78353.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.06557664358
varray: [1296.059, 1295.693, 1294.09, 1294.009]
inst: std_all
operation: std_all(baseline_shift)
tarray: [76744.0, 77280.0, 77816.0, 78353.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.571
varray: [1.182, 1.877, 1.317, 1.908]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [76744.0, 77280.0, 77816.0, 78353.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.375518308475
varray: [1.182, 1.877, 1.317, 1.908]
inst: std_all
operation: std_all(slope_shift)
tarray: [76744.0, 77280.0, 77816.0, 78353.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0898758402835
varray: [-2.166, -1.999, -2.003, -1.966]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76744.0, 77280.0, 77816.0, 78353.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.08057469278
varray: [-27.532, -25.691, -25.33, -25.209]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76744.0, 77280.0, 77816.0, 78353.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.564
varray: [113.528, 111.788, 110.855, 110.085]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [76744.0, 77280.0, 77816.0, 78353.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.48296931863
varray: [113.528, 111.788, 110.855, 110.085]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [76744.0, 77280.0, 77816.0, 78353.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.87575
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.29975
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1294.96275
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.571
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.22085
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.22085
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -7.13201575922
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -6.42725
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-2.166, -1.999, -2.003, -1.966]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [76744.0, 77280.0, 77816.0, 78353.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-27.532, -25.691, -25.33, -25.209]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [76744.0, 77280.0, 77816.0, 78353.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.0335
varray: [-2.166, -1.999, -2.003, -1.966]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [76744.0, 77280.0, 77816.0, 78353.0]
source_col: d(18_16)Mean
dD_mean | -25.9405
varray: [-27.532, -25.691, -25.33, -25.209]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [76744.0, 77280.0, 77816.0, 78353.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0898758402835
varray: [-2.166, -1.999, -2.003, -1.966]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76744.0, 77280.0, 77816.0, 78353.0]
source_col: d(18_16)Mean
dD_stdev | 1.08057469278
varray: [-27.532, -25.691, -25.33, -25.209]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76744.0, 77280.0, 77816.0, 78353.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.0335
varray: [-2.166, -1.999, -2.003, -1.966]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [76744.0, 77280.0, 77816.0, 78353.0]
source_col: d(18_16)Mean
measured_dD_mean | -25.9405
varray: [-27.532, -25.691, -25.33, -25.209]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [76744.0, 77280.0, 77816.0, 78353.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0898758402835
varray: [-2.166, -1.999, -2.003, -1.966]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76744.0, 77280.0, 77816.0, 78353.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.08057469278
varray: [-27.532, -25.691, -25.33, -25.209]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76744.0, 77280.0, 77816.0, 78353.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.8070049883
actual_array: [ 1.93 -15.84]
varray: [-2.166, -1.999, -2.003, -1.966]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.0335
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [76744.0, 77280.0, 77816.0, 78353.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -42.3024275477
actual_array: [ 18.1 -119. ]
varray: [-27.532, -25.691, -25.33, -25.209]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -25.9405
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [76744.0, 77280.0, 77816.0, 78353.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.5385
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -8.4515
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 31 - Sample 31 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 31
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.91075
varray: [2.941, 2.856, 2.878, 2.968]
inst: mean_all
operation: mean_all(residuals)
tarray: [78939.0, 79476.0, 80012.0, 80549.0]
source_col: residuals
SAMP_residuals_STD | 0.0524809489243
varray: [2.941, 2.856, 2.878, 2.968]
inst: std_all
operation: std_all(residuals)
tarray: [78939.0, 79476.0, 80012.0, 80549.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.403
varray: [0.633, 0.226, 0.231, 0.522]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [78939.0, 79476.0, 80012.0, 80549.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.206538132073
varray: [0.633, 0.226, 0.231, 0.522]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [78939.0, 79476.0, 80012.0, 80549.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1295.5625
varray: [1294.735, 1295.856, 1295.337, 1296.322]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [78939.0, 79476.0, 80012.0, 80549.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.682785715531
varray: [1294.735, 1295.856, 1295.337, 1296.322]
inst: std_all
operation: std_all(baseline_shift)
tarray: [78939.0, 79476.0, 80012.0, 80549.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.97725
varray: [1.962, 2.204, 1.797, 1.946]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [78939.0, 79476.0, 80012.0, 80549.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.168438663416
varray: [1.962, 2.204, 1.797, 1.946]
inst: std_all
operation: std_all(slope_shift)
tarray: [78939.0, 79476.0, 80012.0, 80549.0]
source_col: slope_shift
SAMP_d18O_STD | 0.160485721068
varray: [-7.217, -7.491, -7.545, -7.561]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78939.0, 79476.0, 80012.0, 80549.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.8518733533
varray: [-55.14, -58.174, -58.785, -59.252]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78939.0, 79476.0, 80012.0, 80549.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 112.153
varray: [112.416, 110.845, 110.894, 114.457]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [78939.0, 79476.0, 80012.0, 80549.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.70034604321
varray: [112.416, 110.845, 110.894, 114.457]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [78939.0, 79476.0, 80012.0, 80549.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.91075
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.403
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1295.5625
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.97725
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.18585
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.18585
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.00174384809
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.83825
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-7.217, -7.491, -7.545, -7.561]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [78939.0, 79476.0, 80012.0, 80549.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-55.14, -58.174, -58.785, -59.252]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [78939.0, 79476.0, 80012.0, 80549.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.4535
varray: [-7.217, -7.491, -7.545, -7.561]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [78939.0, 79476.0, 80012.0, 80549.0]
source_col: d(18_16)Mean
dD_mean | -57.83775
varray: [-55.14, -58.174, -58.785, -59.252]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [78939.0, 79476.0, 80012.0, 80549.0]
source_col: d(D_H)Mean
d18O_stdev | 0.160485721068
varray: [-7.217, -7.491, -7.545, -7.561]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78939.0, 79476.0, 80012.0, 80549.0]
source_col: d(18_16)Mean
dD_stdev | 1.8518733533
varray: [-55.14, -58.174, -58.785, -59.252]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78939.0, 79476.0, 80012.0, 80549.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.4535
varray: [-7.217, -7.491, -7.545, -7.561]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [78939.0, 79476.0, 80012.0, 80549.0]
source_col: d(18_16)Mean
measured_dD_mean | -57.83775
varray: [-55.14, -58.174, -58.785, -59.252]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [78939.0, 79476.0, 80012.0, 80549.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.160485721068
varray: [-7.217, -7.491, -7.545, -7.561]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78939.0, 79476.0, 80012.0, 80549.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.8518733533
varray: [-55.14, -58.174, -58.785, -59.252]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78939.0, 79476.0, 80012.0, 80549.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.2357116929
actual_array: [ 1.93 -15.84]
varray: [-7.217, -7.491, -7.545, -7.561]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.4535
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [78939.0, 79476.0, 80012.0, 80549.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -72.2419064171
actual_array: [ 18.1 -119. ]
varray: [-55.14, -58.174, -58.785, -59.252]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -57.83775
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [78939.0, 79476.0, 80012.0, 80549.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.5035
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -7.85175
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 32 - Sample 32 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 32
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.45375
varray: [2.507, 2.579, 2.32, 2.409]
inst: mean_all
operation: mean_all(residuals)
tarray: [81135.0, 81673.0, 82209.0, 82747.0]
source_col: residuals
SAMP_residuals_STD | 0.113158811706
varray: [2.507, 2.579, 2.32, 2.409]
inst: std_all
operation: std_all(residuals)
tarray: [81135.0, 81673.0, 82209.0, 82747.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.23475
varray: [0.271, 0.378, 0.124, 0.166]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [81135.0, 81673.0, 82209.0, 82747.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.113764010126
varray: [0.271, 0.378, 0.124, 0.166]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [81135.0, 81673.0, 82209.0, 82747.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1277.23275
varray: [1279.71, 1281.518, 1271.765, 1275.938]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [81135.0, 81673.0, 82209.0, 82747.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 4.32330131381
varray: [1279.71, 1281.518, 1271.765, 1275.938]
inst: std_all
operation: std_all(baseline_shift)
tarray: [81135.0, 81673.0, 82209.0, 82747.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.6665
varray: [1.763, 1.466, 1.579, 1.858]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [81135.0, 81673.0, 82209.0, 82747.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.17686247011
varray: [1.763, 1.466, 1.579, 1.858]
inst: std_all
operation: std_all(slope_shift)
tarray: [81135.0, 81673.0, 82209.0, 82747.0]
source_col: slope_shift
SAMP_d18O_STD | 0.218037458555
varray: [-2.758, -2.251, -2.493, -2.364]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [81135.0, 81673.0, 82209.0, 82747.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.82215303053
varray: [-29.838, -26.156, -26.487, -26.005]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [81135.0, 81673.0, 82209.0, 82747.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 95.997
varray: [98.651, 100.913, 90.386, 94.038]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [81135.0, 81673.0, 82209.0, 82747.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 4.70927361702
varray: [98.651, 100.913, 90.386, 94.038]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [81135.0, 81673.0, 82209.0, 82747.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.45375
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.23475
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1277.23275
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.6665
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.64285
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.64285
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -20.7598656591
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -21.99425
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-2.758, -2.251, -2.493, -2.364]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [81135.0, 81673.0, 82209.0, 82747.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-29.838, -26.156, -26.487, -26.005]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [81135.0, 81673.0, 82209.0, 82747.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.4665
varray: [-2.758, -2.251, -2.493, -2.364]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [81135.0, 81673.0, 82209.0, 82747.0]
source_col: d(18_16)Mean
dD_mean | -27.1215
varray: [-29.838, -26.156, -26.487, -26.005]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [81135.0, 81673.0, 82209.0, 82747.0]
source_col: d(D_H)Mean
d18O_stdev | 0.218037458555
varray: [-2.758, -2.251, -2.493, -2.364]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [81135.0, 81673.0, 82209.0, 82747.0]
source_col: d(18_16)Mean
dD_stdev | 1.82215303053
varray: [-29.838, -26.156, -26.487, -26.005]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [81135.0, 81673.0, 82209.0, 82747.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.4665
varray: [-2.758, -2.251, -2.493, -2.364]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [81135.0, 81673.0, 82209.0, 82747.0]
source_col: d(18_16)Mean
measured_dD_mean | -27.1215
varray: [-29.838, -26.156, -26.487, -26.005]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [81135.0, 81673.0, 82209.0, 82747.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.218037458555
varray: [-2.758, -2.251, -2.493, -2.364]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [81135.0, 81673.0, 82209.0, 82747.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.82215303053
varray: [-29.838, -26.156, -26.487, -26.005]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [81135.0, 81673.0, 82209.0, 82747.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.24070056083
actual_array: [ 1.93 -15.84]
varray: [-2.758, -2.251, -2.493, -2.364]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.4665
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [81135.0, 81673.0, 82209.0, 82747.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -43.4109408065
actual_array: [ 18.1 -119. ]
varray: [-29.838, -26.156, -26.487, -26.005]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -27.1215
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [81135.0, 81673.0, 82209.0, 82747.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.9605
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -26.1815
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 33 - Sample 33 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 33
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.801
varray: [2.724, 2.761, 2.848, 2.871]
inst: mean_all
operation: mean_all(residuals)
tarray: [83334.0, 83870.0, 84407.0, 84944.0]
source_col: residuals
SAMP_residuals_STD | 0.0698522249705
varray: [2.724, 2.761, 2.848, 2.871]
inst: std_all
operation: std_all(residuals)
tarray: [83334.0, 83870.0, 84407.0, 84944.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.30125
varray: [0.049, 0.226, 0.35, 0.58]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [83334.0, 83870.0, 84407.0, 84944.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.223137289578
varray: [0.049, 0.226, 0.35, 0.58]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [83334.0, 83870.0, 84407.0, 84944.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1290.61375
varray: [1289.293, 1290.925, 1291.273, 1290.964]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [83334.0, 83870.0, 84407.0, 84944.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.894155607263
varray: [1289.293, 1290.925, 1291.273, 1290.964]
inst: std_all
operation: std_all(baseline_shift)
tarray: [83334.0, 83870.0, 84407.0, 84944.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.60075
varray: [1.662, 1.744, 1.413, 1.584]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [83334.0, 83870.0, 84407.0, 84944.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.141188703514
varray: [1.662, 1.744, 1.413, 1.584]
inst: std_all
operation: std_all(slope_shift)
tarray: [83334.0, 83870.0, 84407.0, 84944.0]
source_col: slope_shift
SAMP_d18O_STD | 0.20242261896
varray: [5.384, 5.77, 5.777, 5.814]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83334.0, 83870.0, 84407.0, 84944.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.01556168764
varray: [8.649, 12.038, 12.751, 13.001]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83334.0, 83870.0, 84407.0, 84944.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 107.5845
varray: [105.399, 107.524, 107.773, 109.642]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [83334.0, 83870.0, 84407.0, 84944.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.73675070174
varray: [105.399, 107.524, 107.773, 109.642]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [83334.0, 83870.0, 84407.0, 84944.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.801
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.30125
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1290.61375
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.60075
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.2956
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.2956
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -9.54595362656
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -10.40675
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [5.384, 5.77, 5.777, 5.814]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [83334.0, 83870.0, 84407.0, 84944.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [8.649, 12.038, 12.751, 13.001]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [83334.0, 83870.0, 84407.0, 84944.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 5.68625
varray: [5.384, 5.77, 5.777, 5.814]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [83334.0, 83870.0, 84407.0, 84944.0]
source_col: d(18_16)Mean
dD_mean | 11.60975
varray: [8.649, 12.038, 12.751, 13.001]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [83334.0, 83870.0, 84407.0, 84944.0]
source_col: d(D_H)Mean
d18O_stdev | 0.20242261896
varray: [5.384, 5.77, 5.777, 5.814]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83334.0, 83870.0, 84407.0, 84944.0]
source_col: d(18_16)Mean
dD_stdev | 2.01556168764
varray: [8.649, 12.038, 12.751, 13.001]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83334.0, 83870.0, 84407.0, 84944.0]
source_col: d(D_H)Mean
measured_d18O_mean | 5.68625
varray: [5.384, 5.77, 5.777, 5.814]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [83334.0, 83870.0, 84407.0, 84944.0]
source_col: d(18_16)Mean
measured_dD_mean | 11.60975
varray: [8.649, 12.038, 12.751, 13.001]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [83334.0, 83870.0, 84407.0, 84944.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.20242261896
varray: [5.384, 5.77, 5.777, 5.814]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83334.0, 83870.0, 84407.0, 84944.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.01556168764
varray: [8.649, 12.038, 12.751, 13.001]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83334.0, 83870.0, 84407.0, 84944.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 2.92514604177
actual_array: [ 1.93 -15.84]
varray: [5.384, 5.77, 5.777, 5.814]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 5.68625
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [83334.0, 83870.0, 84407.0, 84944.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -7.05691527271
actual_array: [ 18.1 -119. ]
varray: [8.649, 12.038, 12.751, 13.001]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 11.60975
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [83334.0, 83870.0, 84407.0, 84944.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.61325
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -12.8005
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 34 - Sample 34 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 34
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.71375
varray: [2.692, 2.754, 2.683, 2.726]
inst: mean_all
operation: mean_all(residuals)
tarray: [85531.0, 86069.0, 86605.0, 87141.0]
source_col: residuals
SAMP_residuals_STD | 0.0326024027744
varray: [2.692, 2.754, 2.683, 2.726]
inst: std_all
operation: std_all(residuals)
tarray: [85531.0, 86069.0, 86605.0, 87141.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.35425
varray: [0.35, 0.441, 0.327, 0.299]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [85531.0, 86069.0, 86605.0, 87141.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0614783159605
varray: [0.35, 0.441, 0.327, 0.299]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [85531.0, 86069.0, 86605.0, 87141.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1287.5355
varray: [1287.558, 1287.695, 1286.867, 1288.022]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [85531.0, 86069.0, 86605.0, 87141.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.486319853594
varray: [1287.558, 1287.695, 1286.867, 1288.022]
inst: std_all
operation: std_all(baseline_shift)
tarray: [85531.0, 86069.0, 86605.0, 87141.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.72975
varray: [1.681, 1.764, 1.836, 1.638]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [85531.0, 86069.0, 86605.0, 87141.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0880468625222
varray: [1.681, 1.764, 1.836, 1.638]
inst: std_all
operation: std_all(slope_shift)
tarray: [85531.0, 86069.0, 86605.0, 87141.0]
source_col: slope_shift
SAMP_d18O_STD | 0.244051224131
varray: [-1.735, -2.073, -2.227, -2.275]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85531.0, 86069.0, 86605.0, 87141.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.03934341967
varray: [-19.798, -22.93, -24.09, -24.141]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85531.0, 86069.0, 86605.0, 87141.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 105.24475
varray: [104.907, 105.72, 104.823, 105.529]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [85531.0, 86069.0, 86605.0, 87141.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.446694806328
varray: [104.907, 105.72, 104.823, 105.529]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [85531.0, 86069.0, 86605.0, 87141.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.71375
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.35425
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1287.5355
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.72975
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.38285
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.38285
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -12.3635600336
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -12.7465
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-1.735, -2.073, -2.227, -2.275]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [85531.0, 86069.0, 86605.0, 87141.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-19.798, -22.93, -24.09, -24.141]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [85531.0, 86069.0, 86605.0, 87141.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.0775
varray: [-1.735, -2.073, -2.227, -2.275]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [85531.0, 86069.0, 86605.0, 87141.0]
source_col: d(18_16)Mean
dD_mean | -22.73975
varray: [-19.798, -22.93, -24.09, -24.141]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [85531.0, 86069.0, 86605.0, 87141.0]
source_col: d(D_H)Mean
d18O_stdev | 0.244051224131
varray: [-1.735, -2.073, -2.227, -2.275]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85531.0, 86069.0, 86605.0, 87141.0]
source_col: d(18_16)Mean
dD_stdev | 2.03934341967
varray: [-19.798, -22.93, -24.09, -24.141]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85531.0, 86069.0, 86605.0, 87141.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.0775
varray: [-1.735, -2.073, -2.227, -2.275]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [85531.0, 86069.0, 86605.0, 87141.0]
source_col: d(18_16)Mean
measured_dD_mean | -22.73975
varray: [-19.798, -22.93, -24.09, -24.141]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [85531.0, 86069.0, 86605.0, 87141.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.244051224131
varray: [-1.735, -2.073, -2.227, -2.275]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85531.0, 86069.0, 86605.0, 87141.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.03934341967
varray: [-19.798, -22.93, -24.09, -24.141]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85531.0, 86069.0, 86605.0, 87141.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.85107567004
actual_array: [ 1.93 -15.84]
varray: [-1.735, -2.073, -2.227, -2.275]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.0775
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [85531.0, 86069.0, 86605.0, 87141.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -39.2981313469
actual_array: [ 18.1 -119. ]
varray: [-19.798, -22.93, -24.09, -24.141]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -22.73975
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [85531.0, 86069.0, 86605.0, 87141.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.7005
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -15.87875
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 35 - Sample 35 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 35
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.88675
varray: [2.899, 2.946, 2.882, 2.82]
inst: mean_all
operation: mean_all(residuals)
tarray: [87729.0, 88267.0, 88803.0, 89341.0]
source_col: residuals
SAMP_residuals_STD | 0.0520856666144
varray: [2.899, 2.946, 2.882, 2.82]
inst: std_all
operation: std_all(residuals)
tarray: [87729.0, 88267.0, 88803.0, 89341.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4215
varray: [0.365, 0.696, 0.438, 0.187]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [87729.0, 88267.0, 88803.0, 89341.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.21119106673
varray: [0.365, 0.696, 0.438, 0.187]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [87729.0, 88267.0, 88803.0, 89341.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1293.7305
varray: [1294.337, 1294.385, 1293.698, 1292.502]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [87729.0, 88267.0, 88803.0, 89341.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.876827805216
varray: [1294.337, 1294.385, 1293.698, 1292.502]
inst: std_all
operation: std_all(baseline_shift)
tarray: [87729.0, 88267.0, 88803.0, 89341.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.80875
varray: [1.497, 1.93, 2.151, 1.657]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [87729.0, 88267.0, 88803.0, 89341.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.289857867468
varray: [1.497, 1.93, 2.151, 1.657]
inst: std_all
operation: std_all(slope_shift)
tarray: [87729.0, 88267.0, 88803.0, 89341.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0691682730737
varray: [-3.0, -3.159, -3.126, -3.114]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87729.0, 88267.0, 88803.0, 89341.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.308932975471
varray: [-28.089, -28.666, -28.552, -28.8]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87729.0, 88267.0, 88803.0, 89341.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.041
varray: [111.345, 112.803, 110.718, 109.298]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [87729.0, 88267.0, 88803.0, 89341.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.45366639914
varray: [111.345, 112.803, 110.718, 109.298]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [87729.0, 88267.0, 88803.0, 89341.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.88675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4215
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1293.7305
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.80875
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.20985
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.20985
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.77678744429
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -6.95025
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-3.0, -3.159, -3.126, -3.114]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [87729.0, 88267.0, 88803.0, 89341.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-28.089, -28.666, -28.552, -28.8]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [87729.0, 88267.0, 88803.0, 89341.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.09975
varray: [-3.0, -3.159, -3.126, -3.114]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [87729.0, 88267.0, 88803.0, 89341.0]
source_col: d(18_16)Mean
dD_mean | -28.52675
varray: [-28.089, -28.666, -28.552, -28.8]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [87729.0, 88267.0, 88803.0, 89341.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0691682730737
varray: [-3.0, -3.159, -3.126, -3.114]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87729.0, 88267.0, 88803.0, 89341.0]
source_col: d(18_16)Mean
dD_stdev | 0.308932975471
varray: [-28.089, -28.666, -28.552, -28.8]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87729.0, 88267.0, 88803.0, 89341.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.09975
varray: [-3.0, -3.159, -3.126, -3.114]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [87729.0, 88267.0, 88803.0, 89341.0]
source_col: d(18_16)Mean
measured_dD_mean | -28.52675
varray: [-28.089, -28.666, -28.552, -28.8]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [87729.0, 88267.0, 88803.0, 89341.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0691682730737
varray: [-3.0, -3.159, -3.126, -3.114]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87729.0, 88267.0, 88803.0, 89341.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.308932975471
varray: [-28.089, -28.666, -28.552, -28.8]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87729.0, 88267.0, 88803.0, 89341.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.87496781557
actual_array: [ 1.93 -15.84]
varray: [-3.0, -3.159, -3.126, -3.114]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -3.09975
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [87729.0, 88267.0, 88803.0, 89341.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -44.7299401774
actual_array: [ 18.1 -119. ]
varray: [-28.089, -28.666, -28.552, -28.8]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -28.52675
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [87729.0, 88267.0, 88803.0, 89341.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.5275
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -9.68375
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 36 - Sample 36 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 36
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.548
varray: [2.524, 2.68, 2.505, 2.483]
inst: mean_all
operation: mean_all(residuals)
tarray: [89925.0, 90461.0, 90998.0, 91536.0]
source_col: residuals
SAMP_residuals_STD | 0.0895805038313
varray: [2.524, 2.68, 2.505, 2.483]
inst: std_all
operation: std_all(residuals)
tarray: [89925.0, 90461.0, 90998.0, 91536.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.324
varray: [0.183, 0.239, 0.45, 0.424]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [89925.0, 90461.0, 90998.0, 91536.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.132893441022
varray: [0.183, 0.239, 0.45, 0.424]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [89925.0, 90461.0, 90998.0, 91536.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1281.28275
varray: [1281.724, 1286.479, 1278.922, 1278.006]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [89925.0, 90461.0, 90998.0, 91536.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.80814682621
varray: [1281.724, 1286.479, 1278.922, 1278.006]
inst: std_all
operation: std_all(baseline_shift)
tarray: [89925.0, 90461.0, 90998.0, 91536.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.01725
varray: [2.106, 2.463, 1.721, 1.779]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [89925.0, 90461.0, 90998.0, 91536.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.34209976615
varray: [2.106, 2.463, 1.721, 1.779]
inst: std_all
operation: std_all(slope_shift)
tarray: [89925.0, 90461.0, 90998.0, 91536.0]
source_col: slope_shift
SAMP_d18O_STD | 0.139605396266
varray: [-6.892, -6.902, -7.125, -7.151]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89925.0, 90461.0, 90998.0, 91536.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.65810300746
varray: [-55.672, -57.699, -59.082, -59.26]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89925.0, 90461.0, 90998.0, 91536.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 99.3
varray: [98.633, 103.08, 98.029, 97.458]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [89925.0, 90461.0, 90998.0, 91536.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.56526113031
varray: [98.633, 103.08, 98.029, 97.458]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [89925.0, 90461.0, 90998.0, 91536.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.548
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.324
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1281.28275
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.01725
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.5486
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.5486
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -17.7162048699
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -18.69125
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-6.892, -6.902, -7.125, -7.151]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [89925.0, 90461.0, 90998.0, 91536.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-55.672, -57.699, -59.082, -59.26]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [89925.0, 90461.0, 90998.0, 91536.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.0175
varray: [-6.892, -6.902, -7.125, -7.151]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [89925.0, 90461.0, 90998.0, 91536.0]
source_col: d(18_16)Mean
dD_mean | -57.92825
varray: [-55.672, -57.699, -59.082, -59.26]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [89925.0, 90461.0, 90998.0, 91536.0]
source_col: d(D_H)Mean
d18O_stdev | 0.139605396266
varray: [-6.892, -6.902, -7.125, -7.151]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89925.0, 90461.0, 90998.0, 91536.0]
source_col: d(18_16)Mean
dD_stdev | 1.65810300746
varray: [-55.672, -57.699, -59.082, -59.26]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89925.0, 90461.0, 90998.0, 91536.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.0175
varray: [-6.892, -6.902, -7.125, -7.151]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [89925.0, 90461.0, 90998.0, 91536.0]
source_col: d(18_16)Mean
measured_dD_mean | -57.92825
varray: [-55.672, -57.699, -59.082, -59.26]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [89925.0, 90461.0, 90998.0, 91536.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.139605396266
varray: [-6.892, -6.902, -7.125, -7.151]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89925.0, 90461.0, 90998.0, 91536.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.65810300746
varray: [-55.672, -57.699, -59.082, -59.26]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89925.0, 90461.0, 90998.0, 91536.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.79901130119
actual_array: [ 1.93 -15.84]
varray: [-6.892, -6.902, -7.125, -7.151]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.0175
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [89925.0, 90461.0, 90998.0, 91536.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -72.3268517599
actual_array: [ 18.1 -119. ]
varray: [-55.672, -57.699, -59.082, -59.26]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -57.92825
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [89925.0, 90461.0, 90998.0, 91536.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.86625
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -22.1315
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 37 - Sample 37 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 37
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.221
varray: [0.0, 0.0, 2.442, 2.442]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [92120.0, 92656.0, 93193.0, 93730.0]
source_col: residuals
SAMP_residuals_STD | 1.40988935736
varray: [0.0, 0.0, 2.442, 2.442]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [92120.0, 92656.0, 93193.0, 93730.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.2245
varray: [0.0, 0.0, 0.449, 0.449]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [92120.0, 92656.0, 93193.0, 93730.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.259230270866
varray: [0.0, 0.0, 0.449, 0.449]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [92120.0, 92656.0, 93193.0, 93730.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 637.3865
varray: [0.0, 0.0, 1274.773, 1274.773]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [92120.0, 92656.0, 93193.0, 93730.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 735.990534706
varray: [0.0, 0.0, 1274.773, 1274.773]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [92120.0, 92656.0, 93193.0, 93730.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.394
varray: [0.0, 0.0, 0.788, 0.788]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [92120.0, 92656.0, 93193.0, 93730.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.454952012121
varray: [0.0, 0.0, 0.788, 0.788]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [92120.0, 92656.0, 93193.0, 93730.0]
source_col: slope_shift
SAMP_d18O_STD | 5.22328788536
varray: [0.0, 0.0, 9.047, 9.047]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [92120.0, 92656.0, 93193.0, 93730.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 24.0974455354
varray: [0.0, 0.0, 41.738, 41.738]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [92120.0, 92656.0, 93193.0, 93730.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 48.049
varray: [0.0, 0.0, 96.098, 96.098]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [92120.0, 92656.0, 93193.0, 93730.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 55.4822061686
varray: [0.0, 0.0, 96.098, 96.098]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [92120.0, 92656.0, 93193.0, 93730.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.221
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.2245
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 637.3865
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.394
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.8756
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.8756
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -60.5696570432
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -69.94225
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [0.0, 0.0, 9.047, 9.047]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [92120.0, 92656.0, 93193.0, 93730.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [0.0, 0.0, 41.738, 41.738]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [92120.0, 92656.0, 93193.0, 93730.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 4.5235
varray: [0.0, 0.0, 9.047, 9.047]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [92120.0, 92656.0, 93193.0, 93730.0]
source_col: d(18_16)Mean
dD_mean | 20.869
varray: [0.0, 0.0, 41.738, 41.738]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [92120.0, 92656.0, 93193.0, 93730.0]
source_col: d(D_H)Mean
d18O_stdev | 5.22328788536
varray: [0.0, 0.0, 9.047, 9.047]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [92120.0, 92656.0, 93193.0, 93730.0]
source_col: d(18_16)Mean
dD_stdev | 24.0974455354
varray: [0.0, 0.0, 41.738, 41.738]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [92120.0, 92656.0, 93193.0, 93730.0]
source_col: d(D_H)Mean
measured_d18O_mean | 4.5235
varray: [0.0, 0.0, 9.047, 9.047]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [92120.0, 92656.0, 93193.0, 93730.0]
source_col: d(18_16)Mean
measured_dD_mean | 20.869
varray: [0.0, 0.0, 41.738, 41.738]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [92120.0, 92656.0, 93193.0, 93730.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 5.22328788536
varray: [0.0, 0.0, 9.047, 9.047]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [92120.0, 92656.0, 93193.0, 93730.0]
source_col: d(18_16)Mean
measured_dD_stdev | 24.0974455354
varray: [0.0, 0.0, 41.738, 41.738]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [92120.0, 92656.0, 93193.0, 93730.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.7605281966
actual_array: [ 1.93 -15.84]
varray: [0.0, 0.0, 9.047, 9.047]
fit_rsq: 1.0
fit_intercept: 2.76579538548
except: True
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 4.5235
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [92120.0, 92656.0, 93193.0, 93730.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 1.63402578713
actual_array: [ 18.1 -119. ]
varray: [0.0, 0.0, 41.738, 41.738]
fit_rsq: 1.0
fit_intercept: 19.1281237053
except: True
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 20.869
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [92120.0, 92656.0, 93193.0, 93730.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1302.19325
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -666.02775
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 38 - Sample 38 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 38
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.374
varray: [2.308, 2.527, 2.325, 2.336]
inst: mean_all
operation: mean_all(residuals)
tarray: [94315.0, 94852.0, 95387.0, 95924.0]
source_col: residuals
SAMP_residuals_STD | 0.102648266749
varray: [2.308, 2.527, 2.325, 2.336]
inst: std_all
operation: std_all(residuals)
tarray: [94315.0, 94852.0, 95387.0, 95924.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.135
varray: [0.108, 0.302, 0.056, 0.074]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [94315.0, 94852.0, 95387.0, 95924.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.113401940019
varray: [0.108, 0.302, 0.056, 0.074]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [94315.0, 94852.0, 95387.0, 95924.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1274.98875
varray: [1271.705, 1279.858, 1273.778, 1274.614]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [94315.0, 94852.0, 95387.0, 95924.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.46885969496
varray: [1271.705, 1279.858, 1273.778, 1274.614]
inst: std_all
operation: std_all(baseline_shift)
tarray: [94315.0, 94852.0, 95387.0, 95924.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.56475
varray: [1.244, 1.467, 1.707, 1.841]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [94315.0, 94852.0, 95387.0, 95924.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.26393480887
varray: [1.244, 1.467, 1.707, 1.841]
inst: std_all
operation: std_all(slope_shift)
tarray: [94315.0, 94852.0, 95387.0, 95924.0]
source_col: slope_shift
SAMP_d18O_STD | 0.324966152084
varray: [-0.596, -1.043, -1.228, -1.331]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94315.0, 94852.0, 95387.0, 95924.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.43987814707
varray: [-11.692, -16.765, -18.683, -19.254]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94315.0, 94852.0, 95387.0, 95924.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 93.6815
varray: [91.111, 99.027, 91.947, 92.641]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [94315.0, 94852.0, 95387.0, 95924.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.61814736203
varray: [91.111, 99.027, 91.947, 92.641]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [94315.0, 94852.0, 95387.0, 95924.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.374
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.135
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1274.98875
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.56475
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.7226
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.7226
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -23.3352709423
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -24.30975
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-0.596, -1.043, -1.228, -1.331]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [94315.0, 94852.0, 95387.0, 95924.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-11.692, -16.765, -18.683, -19.254]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [94315.0, 94852.0, 95387.0, 95924.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.0495
varray: [-0.596, -1.043, -1.228, -1.331]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [94315.0, 94852.0, 95387.0, 95924.0]
source_col: d(18_16)Mean
dD_mean | -16.5985
varray: [-11.692, -16.765, -18.683, -19.254]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [94315.0, 94852.0, 95387.0, 95924.0]
source_col: d(D_H)Mean
d18O_stdev | 0.324966152084
varray: [-0.596, -1.043, -1.228, -1.331]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94315.0, 94852.0, 95387.0, 95924.0]
source_col: d(18_16)Mean
dD_stdev | 3.43987814707
varray: [-11.692, -16.765, -18.683, -19.254]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94315.0, 94852.0, 95387.0, 95924.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.0495
varray: [-0.596, -1.043, -1.228, -1.331]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [94315.0, 94852.0, 95387.0, 95924.0]
source_col: d(18_16)Mean
measured_dD_mean | -16.5985
varray: [-11.692, -16.765, -18.683, -19.254]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [94315.0, 94852.0, 95387.0, 95924.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.324966152084
varray: [-0.596, -1.043, -1.228, -1.331]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94315.0, 94852.0, 95387.0, 95924.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.43987814707
varray: [-11.692, -16.765, -18.683, -19.254]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94315.0, 94852.0, 95387.0, 95924.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.82142428769
actual_array: [ 1.93 -15.84]
varray: [-0.596, -1.043, -1.228, -1.331]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -1.0495
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [94315.0, 94852.0, 95387.0, 95924.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -33.53381547
actual_array: [ 18.1 -119. ]
varray: [-11.692, -16.765, -18.683, -19.254]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -16.5985
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [94315.0, 94852.0, 95387.0, 95924.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1301.04025
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -28.4255
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 39 - Sample 39 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 39
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.955
varray: [3.024, 2.96, 2.918, 2.918]
inst: mean_all
operation: mean_all(residuals)
tarray: [96510.0, 97048.0, 97585.0, 98122.0]
source_col: residuals
SAMP_residuals_STD | 0.0500799361022
varray: [3.024, 2.96, 2.918, 2.918]
inst: std_all
operation: std_all(residuals)
tarray: [96510.0, 97048.0, 97585.0, 98122.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.41775
varray: [0.609, 0.429, 0.186, 0.447]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [96510.0, 97048.0, 97585.0, 98122.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.174419752322
varray: [0.609, 0.429, 0.186, 0.447]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [96510.0, 97048.0, 97585.0, 98122.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1296.8755
varray: [1298.455, 1298.053, 1296.022, 1294.972]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [96510.0, 97048.0, 97585.0, 98122.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.65661311114
varray: [1298.455, 1298.053, 1296.022, 1294.972]
inst: std_all
operation: std_all(baseline_shift)
tarray: [96510.0, 97048.0, 97585.0, 98122.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.76225
varray: [1.976, 1.799, 1.525, 1.749]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [96510.0, 97048.0, 97585.0, 98122.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.185744223777
varray: [1.976, 1.799, 1.525, 1.749]
inst: std_all
operation: std_all(slope_shift)
tarray: [96510.0, 97048.0, 97585.0, 98122.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0921737489744
varray: [-2.598, -2.73, -2.778, -2.806]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96510.0, 97048.0, 97585.0, 98122.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.25401418785
varray: [-32.291, -33.945, -34.89, -35.003]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96510.0, 97048.0, 97585.0, 98122.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.508
varray: [114.777, 114.546, 112.24, 112.469]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [96510.0, 97048.0, 97585.0, 98122.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.33855021074
varray: [114.777, 114.546, 112.24, 112.469]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [96510.0, 97048.0, 97585.0, 98122.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.955
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.41775
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1296.8755
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.76225
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1416
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1416
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.57275721759
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.48325
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-2.598, -2.73, -2.778, -2.806]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [96510.0, 97048.0, 97585.0, 98122.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-32.291, -33.945, -34.89, -35.003]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [96510.0, 97048.0, 97585.0, 98122.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.728
varray: [-2.598, -2.73, -2.778, -2.806]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [96510.0, 97048.0, 97585.0, 98122.0]
source_col: d(18_16)Mean
dD_mean | -34.03225
varray: [-32.291, -33.945, -34.89, -35.003]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [96510.0, 97048.0, 97585.0, 98122.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0921737489744
varray: [-2.598, -2.73, -2.778, -2.806]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96510.0, 97048.0, 97585.0, 98122.0]
source_col: d(18_16)Mean
dD_stdev | 1.25401418785
varray: [-32.291, -33.945, -34.89, -35.003]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96510.0, 97048.0, 97585.0, 98122.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.728
varray: [-2.598, -2.73, -2.778, -2.806]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [96510.0, 97048.0, 97585.0, 98122.0]
source_col: d(18_16)Mean
measured_dD_mean | -34.03225
varray: [-32.291, -33.945, -34.89, -35.003]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [96510.0, 97048.0, 97585.0, 98122.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0921737489744
varray: [-2.598, -2.73, -2.778, -2.806]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96510.0, 97048.0, 97585.0, 98122.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.25401418785
varray: [-32.291, -33.945, -34.89, -35.003]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96510.0, 97048.0, 97585.0, 98122.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.50262063523
actual_array: [ 1.93 -15.84]
varray: [-2.598, -2.73, -2.778, -2.806]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.728
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [96510.0, 97048.0, 97585.0, 98122.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -49.8975267535
actual_array: [ 18.1 -119. ]
varray: [-32.291, -33.945, -34.89, -35.003]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -34.03225
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [96510.0, 97048.0, 97585.0, 98122.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.45925
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -6.53875
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 40 - Sample 40 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 40
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.84675
varray: [2.878, 2.836, 2.822, 2.851]
inst: mean_all
operation: mean_all(residuals)
tarray: [98707.0, 99243.0, 99780.0, 100317.0]
source_col: residuals
SAMP_residuals_STD | 0.0239635139326
varray: [2.878, 2.836, 2.822, 2.851]
inst: std_all
operation: std_all(residuals)
tarray: [98707.0, 99243.0, 99780.0, 100317.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.28725
varray: [0.257, 0.342, 0.241, 0.309]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [98707.0, 99243.0, 99780.0, 100317.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0466360018298
varray: [0.257, 0.342, 0.241, 0.309]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [98707.0, 99243.0, 99780.0, 100317.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1293.8615
varray: [1294.462, 1293.412, 1293.753, 1293.819]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [98707.0, 99243.0, 99780.0, 100317.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.43826514045
varray: [1294.462, 1293.412, 1293.753, 1293.819]
inst: std_all
operation: std_all(baseline_shift)
tarray: [98707.0, 99243.0, 99780.0, 100317.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.061
varray: [1.957, 2.07, 2.087, 2.13]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [98707.0, 99243.0, 99780.0, 100317.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0737879845684
varray: [1.957, 2.07, 2.087, 2.13]
inst: std_all
operation: std_all(slope_shift)
tarray: [98707.0, 99243.0, 99780.0, 100317.0]
source_col: slope_shift
SAMP_d18O_STD | 0.172233126895
varray: [-7.182, -7.485, -7.496, -7.572]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [98707.0, 99243.0, 99780.0, 100317.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.43217116645
varray: [-56.013, -58.389, -58.674, -59.277]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [98707.0, 99243.0, 99780.0, 100317.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.38575
varray: [111.012, 110.304, 110.09, 110.137]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [98707.0, 99243.0, 99780.0, 100317.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.427479726615
varray: [111.012, 110.304, 110.09, 110.137]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [98707.0, 99243.0, 99780.0, 100317.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.84675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.28725
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1293.8615
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.061
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.24985
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.24985
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -8.0685267713
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -7.6055
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-7.182, -7.485, -7.496, -7.572]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [98707.0, 99243.0, 99780.0, 100317.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-56.013, -58.389, -58.674, -59.277]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [98707.0, 99243.0, 99780.0, 100317.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.43375
varray: [-7.182, -7.485, -7.496, -7.572]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [98707.0, 99243.0, 99780.0, 100317.0]
source_col: d(18_16)Mean
dD_mean | -58.08825
varray: [-56.013, -58.389, -58.674, -59.277]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [98707.0, 99243.0, 99780.0, 100317.0]
source_col: d(D_H)Mean
d18O_stdev | 0.172233126895
varray: [-7.182, -7.485, -7.496, -7.572]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [98707.0, 99243.0, 99780.0, 100317.0]
source_col: d(18_16)Mean
dD_stdev | 1.43217116645
varray: [-56.013, -58.389, -58.674, -59.277]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [98707.0, 99243.0, 99780.0, 100317.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.43375
varray: [-7.182, -7.485, -7.496, -7.572]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [98707.0, 99243.0, 99780.0, 100317.0]
source_col: d(18_16)Mean
measured_dD_mean | -58.08825
varray: [-56.013, -58.389, -58.674, -59.277]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [98707.0, 99243.0, 99780.0, 100317.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.172233126895
varray: [-7.182, -7.485, -7.496, -7.572]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [98707.0, 99243.0, 99780.0, 100317.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.43217116645
varray: [-56.013, -58.389, -58.674, -59.277]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [98707.0, 99243.0, 99780.0, 100317.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.2159299665
actual_array: [ 1.93 -15.84]
varray: [-7.182, -7.485, -7.496, -7.572]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.43375
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [98707.0, 99243.0, 99780.0, 100317.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -72.4770313716
actual_array: [ 18.1 -119. ]
varray: [-56.013, -58.389, -58.674, -59.277]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -58.08825
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [98707.0, 99243.0, 99780.0, 100317.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.5675
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -9.55275
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 41 - Sample 41 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 41
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.982
varray: [2.696, 0.0, 2.696, 2.536]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [100903.0, 101440.0, 101976.0, 102514.0]
source_col: residuals
SAMP_residuals_STD | 1.3234842903
varray: [2.696, 0.0, 2.696, 2.536]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [100903.0, 101440.0, 101976.0, 102514.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.31525
varray: [0.321, 0.0, 0.321, 0.619]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [100903.0, 101440.0, 101976.0, 102514.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.252792899426
varray: [0.321, 0.0, 0.321, 0.619]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [100903.0, 101440.0, 101976.0, 102514.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 961.97875
varray: [1285.309, 0.0, 1285.309, 1277.297]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [100903.0, 101440.0, 101976.0, 102514.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 641.330288116
varray: [1285.309, 0.0, 1285.309, 1277.297]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [100903.0, 101440.0, 101976.0, 102514.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.913
varray: [1.218, 0.0, 1.218, 1.216]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [100903.0, 101440.0, 101976.0, 102514.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.60866739686
varray: [1.218, 0.0, 1.218, 1.216]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [100903.0, 101440.0, 101976.0, 102514.0]
source_col: slope_shift
SAMP_d18O_STD | 1.40608878335
varray: [-2.759, 0.0, -2.759, -2.908]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100903.0, 101440.0, 101976.0, 102514.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 15.6408363907
varray: [-30.987, 0.0, -30.987, -31.84]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100903.0, 101440.0, 101976.0, 102514.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 76.5595
varray: [103.932, 0.0, 103.932, 98.374]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [100903.0, 101440.0, 101976.0, 102514.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 51.1068715634
varray: [103.932, 0.0, 103.932, 98.374]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [100903.0, 101440.0, 101976.0, 102514.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.982
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.31525
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 961.97875
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.913
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.1146
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.1146
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -35.994316347
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -41.43175
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-2.759, 0.0, -2.759, -2.908]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [100903.0, 101440.0, 101976.0, 102514.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-30.987, 0.0, -30.987, -31.84]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [100903.0, 101440.0, 101976.0, 102514.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.1065
varray: [-2.759, 0.0, -2.759, -2.908]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [100903.0, 101440.0, 101976.0, 102514.0]
source_col: d(18_16)Mean
dD_mean | -23.4535
varray: [-30.987, 0.0, -30.987, -31.84]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [100903.0, 101440.0, 101976.0, 102514.0]
source_col: d(D_H)Mean
d18O_stdev | 1.40608878335
varray: [-2.759, 0.0, -2.759, -2.908]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100903.0, 101440.0, 101976.0, 102514.0]
source_col: d(18_16)Mean
dD_stdev | 15.6408363907
varray: [-30.987, 0.0, -30.987, -31.84]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100903.0, 101440.0, 101976.0, 102514.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.1065
varray: [-2.759, 0.0, -2.759, -2.908]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [100903.0, 101440.0, 101976.0, 102514.0]
source_col: d(18_16)Mean
measured_dD_mean | -23.4535
varray: [-30.987, 0.0, -30.987, -31.84]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [100903.0, 101440.0, 101976.0, 102514.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 1.40608878335
varray: [-2.759, 0.0, -2.759, -2.908]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100903.0, 101440.0, 101976.0, 102514.0]
source_col: d(18_16)Mean
measured_dD_stdev | 15.6408363907
varray: [-30.987, 0.0, -30.987, -31.84]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100903.0, 101440.0, 101976.0, 102514.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.88012225573
actual_array: [ 1.93 -15.84]
varray: [-2.759, 0.0, -2.759, -2.908]
fit_rsq: 1.0
fit_intercept: 2.76579538548
except: True
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.1065
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [100903.0, 101440.0, 101976.0, 102514.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -39.9680732085
actual_array: [ 18.1 -119. ]
varray: [-30.987, 0.0, -30.987, -31.84]
fit_rsq: 1.0
fit_intercept: 19.1281237053
except: True
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -23.4535
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [100903.0, 101440.0, 101976.0, 102514.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1301.43225
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -341.4355
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 42 - Sample 42 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 42
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.728
varray: [2.747, 2.746, 2.706, 2.713]
inst: mean_all
operation: mean_all(residuals)
tarray: [103099.0, 103636.0, 104174.0, 104710.0]
source_col: residuals
SAMP_residuals_STD | 0.0215561282856
varray: [2.747, 2.746, 2.706, 2.713]
inst: std_all
operation: std_all(residuals)
tarray: [103099.0, 103636.0, 104174.0, 104710.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4225
varray: [0.443, 0.41, 0.426, 0.411]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [103099.0, 103636.0, 104174.0, 104710.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.015502687939
varray: [0.443, 0.41, 0.426, 0.411]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [103099.0, 103636.0, 104174.0, 104710.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1287.463
varray: [1287.175, 1287.448, 1287.978, 1287.251]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [103099.0, 103636.0, 104174.0, 104710.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.362094831041
varray: [1287.175, 1287.448, 1287.978, 1287.251]
inst: std_all
operation: std_all(baseline_shift)
tarray: [103099.0, 103636.0, 104174.0, 104710.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.6285
varray: [1.542, 1.742, 1.957, 1.273]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [103099.0, 103636.0, 104174.0, 104710.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.291351448712
varray: [1.542, 1.742, 1.957, 1.273]
inst: std_all
operation: std_all(slope_shift)
tarray: [103099.0, 103636.0, 104174.0, 104710.0]
source_col: slope_shift
SAMP_d18O_STD | 0.129529919324
varray: [-6.615, -6.758, -6.858, -6.909]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [103099.0, 103636.0, 104174.0, 104710.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.34233341611
varray: [-53.961, -55.784, -56.705, -56.9]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [103099.0, 103636.0, 104174.0, 104710.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 105.68225
varray: [105.567, 105.144, 105.904, 106.114]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [103099.0, 103636.0, 104174.0, 104710.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.423704594421
varray: [105.567, 105.144, 105.904, 106.114]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [103099.0, 103636.0, 104174.0, 104710.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.728
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1287.463
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.6285
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.3686
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.3686
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -11.9033778983
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -12.309
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-6.615, -6.758, -6.858, -6.909]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [103099.0, 103636.0, 104174.0, 104710.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-53.961, -55.784, -56.705, -56.9]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [103099.0, 103636.0, 104174.0, 104710.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.785
varray: [-6.615, -6.758, -6.858, -6.909]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [103099.0, 103636.0, 104174.0, 104710.0]
source_col: d(18_16)Mean
dD_mean | -55.8375
varray: [-53.961, -55.784, -56.705, -56.9]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [103099.0, 103636.0, 104174.0, 104710.0]
source_col: d(D_H)Mean
d18O_stdev | 0.129529919324
varray: [-6.615, -6.758, -6.858, -6.909]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [103099.0, 103636.0, 104174.0, 104710.0]
source_col: d(18_16)Mean
dD_stdev | 1.34233341611
varray: [-53.961, -55.784, -56.705, -56.9]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [103099.0, 103636.0, 104174.0, 104710.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.785
varray: [-6.615, -6.758, -6.858, -6.909]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [103099.0, 103636.0, 104174.0, 104710.0]
source_col: d(18_16)Mean
measured_dD_mean | -55.8375
varray: [-53.961, -55.784, -56.705, -56.9]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [103099.0, 103636.0, 104174.0, 104710.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.129529919324
varray: [-6.615, -6.758, -6.858, -6.909]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [103099.0, 103636.0, 104174.0, 104710.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.34233341611
varray: [-53.961, -55.784, -56.705, -56.9]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [103099.0, 103636.0, 104174.0, 104710.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.56613781247
actual_array: [ 1.93 -15.84]
varray: [-6.615, -6.758, -6.858, -6.909]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.785
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [103099.0, 103636.0, 104174.0, 104710.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -70.3644266153
actual_array: [ 18.1 -119. ]
varray: [-53.961, -55.784, -56.705, -56.9]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -55.8375
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [103099.0, 103636.0, 104174.0, 104710.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.68625
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -15.95125
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 43 - Sample 43 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 43
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.79175
varray: [2.664, 2.883, 2.832, 2.788]
inst: mean_all
operation: mean_all(residuals)
tarray: [105295.0, 105832.0, 106369.0, 106905.0]
source_col: residuals
SAMP_residuals_STD | 0.0935962071881
varray: [2.664, 2.883, 2.832, 2.788]
inst: std_all
operation: std_all(residuals)
tarray: [105295.0, 105832.0, 106369.0, 106905.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.363
varray: [0.528, 0.324, 0.39, 0.21]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [105295.0, 105832.0, 106369.0, 106905.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.132770478646
varray: [0.528, 0.324, 0.39, 0.21]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [105295.0, 105832.0, 106369.0, 106905.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1290.249
varray: [1286.097, 1292.65, 1291.556, 1290.693]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [105295.0, 105832.0, 106369.0, 106905.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.8815094424
varray: [1286.097, 1292.65, 1291.556, 1290.693]
inst: std_all
operation: std_all(baseline_shift)
tarray: [105295.0, 105832.0, 106369.0, 106905.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.457
varray: [1.176, 1.525, 1.551, 1.576]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [105295.0, 105832.0, 106369.0, 106905.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.188486957993
varray: [1.176, 1.525, 1.551, 1.576]
inst: std_all
operation: std_all(slope_shift)
tarray: [105295.0, 105832.0, 106369.0, 106905.0]
source_col: slope_shift
SAMP_d18O_STD | 0.198095095009
varray: [-2.593, -2.142, -2.219, -2.276]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [105295.0, 105832.0, 106369.0, 106905.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.31321198049
varray: [-23.386, -18.81, -18.539, -18.971]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [105295.0, 105832.0, 106369.0, 106905.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 107.92775
varray: [104.761, 110.324, 108.877, 107.749]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [105295.0, 105832.0, 106369.0, 106905.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.35961485205
varray: [104.761, 110.324, 108.877, 107.749]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [105295.0, 105832.0, 106369.0, 106905.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.79175
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.363
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1290.249
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.457
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.30485
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.30485
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -9.84466834593
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -10.0635
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-2.593, -2.142, -2.219, -2.276]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [105295.0, 105832.0, 106369.0, 106905.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-23.386, -18.81, -18.539, -18.971]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [105295.0, 105832.0, 106369.0, 106905.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.3075
varray: [-2.593, -2.142, -2.219, -2.276]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [105295.0, 105832.0, 106369.0, 106905.0]
source_col: d(18_16)Mean
dD_mean | -19.9265
varray: [-23.386, -18.81, -18.539, -18.971]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [105295.0, 105832.0, 106369.0, 106905.0]
source_col: d(D_H)Mean
d18O_stdev | 0.198095095009
varray: [-2.593, -2.142, -2.219, -2.276]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [105295.0, 105832.0, 106369.0, 106905.0]
source_col: d(18_16)Mean
dD_stdev | 2.31321198049
varray: [-23.386, -18.81, -18.539, -18.971]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [105295.0, 105832.0, 106369.0, 106905.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.3075
varray: [-2.593, -2.142, -2.219, -2.276]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [105295.0, 105832.0, 106369.0, 106905.0]
source_col: d(18_16)Mean
measured_dD_mean | -19.9265
varray: [-23.386, -18.81, -18.539, -18.971]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [105295.0, 105832.0, 106369.0, 106905.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.198095095009
varray: [-2.593, -2.142, -2.219, -2.276]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [105295.0, 105832.0, 106369.0, 106905.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.31321198049
varray: [-23.386, -18.81, -18.539, -18.971]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [105295.0, 105832.0, 106369.0, 106905.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.08144514274
actual_array: [ 1.93 -15.84]
varray: [-2.593, -2.142, -2.219, -2.276]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.3075
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [105295.0, 105832.0, 106369.0, 106905.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -36.6575513932
actual_array: [ 18.1 -119. ]
varray: [-23.386, -18.81, -18.539, -18.971]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -19.9265
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [105295.0, 105832.0, 106369.0, 106905.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.6225
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -13.16525
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 44 - Sample 44 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 44
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.71025
varray: [2.668, 2.941, 2.586, 2.646]
inst: mean_all
operation: mean_all(residuals)
tarray: [107491.0, 108029.0, 108566.0, 109102.0]
source_col: residuals
SAMP_residuals_STD | 0.15768824729
varray: [2.668, 2.941, 2.586, 2.646]
inst: std_all
operation: std_all(residuals)
tarray: [107491.0, 108029.0, 108566.0, 109102.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.29575
varray: [0.267, 0.6, 0.079, 0.237]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [107491.0, 108029.0, 108566.0, 109102.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.218957187596
varray: [0.267, 0.6, 0.079, 0.237]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [107491.0, 108029.0, 108566.0, 109102.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1288.116
varray: [1287.4, 1294.356, 1284.472, 1286.236]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [107491.0, 108029.0, 108566.0, 109102.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 4.33064244657
varray: [1287.4, 1294.356, 1284.472, 1286.236]
inst: std_all
operation: std_all(baseline_shift)
tarray: [107491.0, 108029.0, 108566.0, 109102.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.8415
varray: [1.857, 1.439, 2.24, 1.83]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [107491.0, 108029.0, 108566.0, 109102.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.32720075387
varray: [1.857, 1.439, 2.24, 1.83]
inst: std_all
operation: std_all(slope_shift)
tarray: [107491.0, 108029.0, 108566.0, 109102.0]
source_col: slope_shift
SAMP_d18O_STD | 0.209490453243
varray: [-7.469, -7.448, -7.817, -7.825]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [107491.0, 108029.0, 108566.0, 109102.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.65895610782
varray: [-61.536, -64.846, -66.939, -67.354]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [107491.0, 108029.0, 108566.0, 109102.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 105.28475
varray: [104.163, 112.386, 101.09, 103.5]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [107491.0, 108029.0, 108566.0, 109102.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 4.91485282079
varray: [104.163, 112.386, 101.09, 103.5]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [107491.0, 108029.0, 108566.0, 109102.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.71025
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.29575
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1288.116
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.8415
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.38635
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.38635
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -12.4765872247
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -12.7065
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-7.469, -7.448, -7.817, -7.825]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [107491.0, 108029.0, 108566.0, 109102.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-61.536, -64.846, -66.939, -67.354]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [107491.0, 108029.0, 108566.0, 109102.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.63975
varray: [-7.469, -7.448, -7.817, -7.825]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [107491.0, 108029.0, 108566.0, 109102.0]
source_col: d(18_16)Mean
dD_mean | -65.16875
varray: [-61.536, -64.846, -66.939, -67.354]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [107491.0, 108029.0, 108566.0, 109102.0]
source_col: d(D_H)Mean
d18O_stdev | 0.209490453243
varray: [-7.469, -7.448, -7.817, -7.825]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [107491.0, 108029.0, 108566.0, 109102.0]
source_col: d(18_16)Mean
dD_stdev | 2.65895610782
varray: [-61.536, -64.846, -66.939, -67.354]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [107491.0, 108029.0, 108566.0, 109102.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.63975
varray: [-7.469, -7.448, -7.817, -7.825]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [107491.0, 108029.0, 108566.0, 109102.0]
source_col: d(18_16)Mean
measured_dD_mean | -65.16875
varray: [-61.536, -64.846, -66.939, -67.354]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [107491.0, 108029.0, 108566.0, 109102.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.209490453243
varray: [-7.469, -7.448, -7.817, -7.825]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [107491.0, 108029.0, 108566.0, 109102.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.65895610782
varray: [-61.536, -64.846, -66.939, -67.354]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [107491.0, 108029.0, 108566.0, 109102.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.4222608855
actual_array: [ 1.93 -15.84]
varray: [-7.469, -7.448, -7.817, -7.825]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.63975
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [107491.0, 108029.0, 108566.0, 109102.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -79.1229485004
actual_array: [ 18.1 -119. ]
varray: [-61.536, -64.846, -66.939, -67.354]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -65.16875
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [107491.0, 108029.0, 108566.0, 109102.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.704
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -15.29825
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 45 - Sample 45 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 45
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.908
varray: [2.926, 2.882, 2.863, 2.961]
inst: mean_all
operation: mean_all(residuals)
tarray: [173508.0, 174044.0, 174581.0, 175117.0]
source_col: residuals
SAMP_residuals_STD | 0.0440983748756
varray: [2.926, 2.882, 2.863, 2.961]
inst: std_all
operation: std_all(residuals)
tarray: [173508.0, 174044.0, 174581.0, 175117.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.41625
varray: [0.616, 0.212, 0.3, 0.537]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [173508.0, 174044.0, 174581.0, 175117.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.191235239082
varray: [0.616, 0.212, 0.3, 0.537]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [173508.0, 174044.0, 174581.0, 175117.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1297.5215
varray: [1297.622, 1299.157, 1297.112, 1296.195]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [173508.0, 174044.0, 174581.0, 175117.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.23992647631
varray: [1297.622, 1299.157, 1297.112, 1296.195]
inst: std_all
operation: std_all(baseline_shift)
tarray: [173508.0, 174044.0, 174581.0, 175117.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.7715
varray: [1.776, 2.02, 1.463, 1.827]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [173508.0, 174044.0, 174581.0, 175117.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.230958149167
varray: [1.776, 2.02, 1.463, 1.827]
inst: std_all
operation: std_all(slope_shift)
tarray: [173508.0, 174044.0, 174581.0, 175117.0]
source_col: slope_shift
SAMP_d18O_STD | 0.127934879268
varray: [-3.853, -3.577, -3.654, -3.588]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [173508.0, 174044.0, 174581.0, 175117.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.68092791834
varray: [-40.975, -37.955, -37.659, -37.339]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [173508.0, 174044.0, 174581.0, 175117.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.611
varray: [113.014, 110.593, 110.733, 112.104]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [173508.0, 174044.0, 174581.0, 175117.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.1573916076
varray: [113.014, 110.593, 110.733, 112.104]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [173508.0, 174044.0, 174581.0, 175117.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.908
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.41625
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1297.5215
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.7715
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1886
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1886
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.09055092682
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -6.38025
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-3.853, -3.577, -3.654, -3.588]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [173508.0, 174044.0, 174581.0, 175117.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-40.975, -37.955, -37.659, -37.339]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [173508.0, 174044.0, 174581.0, 175117.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.668
varray: [-3.853, -3.577, -3.654, -3.588]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [173508.0, 174044.0, 174581.0, 175117.0]
source_col: d(18_16)Mean
dD_mean | -38.482
varray: [-40.975, -37.955, -37.659, -37.339]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [173508.0, 174044.0, 174581.0, 175117.0]
source_col: d(D_H)Mean
d18O_stdev | 0.127934879268
varray: [-3.853, -3.577, -3.654, -3.588]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [173508.0, 174044.0, 174581.0, 175117.0]
source_col: d(18_16)Mean
dD_stdev | 1.68092791834
varray: [-40.975, -37.955, -37.659, -37.339]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [173508.0, 174044.0, 174581.0, 175117.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.668
varray: [-3.853, -3.577, -3.654, -3.588]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [173508.0, 174044.0, 174581.0, 175117.0]
source_col: d(18_16)Mean
measured_dD_mean | -38.482
varray: [-40.975, -37.955, -37.659, -37.339]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [173508.0, 174044.0, 174581.0, 175117.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.127934879268
varray: [-3.853, -3.577, -3.654, -3.588]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [173508.0, 174044.0, 174581.0, 175117.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.68092791834
varray: [-40.975, -37.955, -37.659, -37.339]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [173508.0, 174044.0, 174581.0, 175117.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.44413065412
actual_array: [ 1.93 -15.84]
varray: [-3.853, -3.577, -3.654, -3.588]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -3.668
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [173508.0, 174044.0, 174581.0, 175117.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -54.0741625481
actual_array: [ 18.1 -119. ]
varray: [-40.975, -37.955, -37.659, -37.339]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -38.482
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [173508.0, 174044.0, 174581.0, 175117.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.50625
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -5.89275
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 46 - Sample 46 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 46
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.91225
varray: [2.825, 2.987, 2.935, 2.902]
inst: mean_all
operation: mean_all(residuals)
tarray: [175703.0, 176239.0, 176776.0, 177312.0]
source_col: residuals
SAMP_residuals_STD | 0.0678791818847
varray: [2.825, 2.987, 2.935, 2.902]
inst: std_all
operation: std_all(residuals)
tarray: [175703.0, 176239.0, 176776.0, 177312.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.501
varray: [0.37, 0.526, 0.683, 0.425]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [175703.0, 176239.0, 176776.0, 177312.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.137460297298
varray: [0.37, 0.526, 0.683, 0.425]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [175703.0, 176239.0, 176776.0, 177312.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1296.01375
varray: [1293.523, 1298.084, 1296.615, 1295.833]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [175703.0, 176239.0, 176776.0, 177312.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.90472594267
varray: [1293.523, 1298.084, 1296.615, 1295.833]
inst: std_all
operation: std_all(baseline_shift)
tarray: [175703.0, 176239.0, 176776.0, 177312.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.9925
varray: [2.224, 2.284, 1.878, 1.584]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [175703.0, 176239.0, 176776.0, 177312.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.3258563078
varray: [2.224, 2.284, 1.878, 1.584]
inst: std_all
operation: std_all(slope_shift)
tarray: [175703.0, 176239.0, 176776.0, 177312.0]
source_col: slope_shift
SAMP_d18O_STD | 0.12544454286
varray: [-7.111, -7.154, -7.271, -7.39]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [175703.0, 176239.0, 176776.0, 177312.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.11663019841
varray: [-58.013, -59.354, -59.947, -60.644]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [175703.0, 176239.0, 176776.0, 177312.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 112.113
varray: [109.118, 113.902, 113.242, 112.19]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [175703.0, 176239.0, 176776.0, 177312.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.11747617066
varray: [109.118, 113.902, 113.242, 112.19]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [175703.0, 176239.0, 176776.0, 177312.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.91225
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.501
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1296.01375
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.9925
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.18435
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.18435
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.95330362333
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.87825
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-7.111, -7.154, -7.271, -7.39]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [175703.0, 176239.0, 176776.0, 177312.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-58.013, -59.354, -59.947, -60.644]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [175703.0, 176239.0, 176776.0, 177312.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.2315
varray: [-7.111, -7.154, -7.271, -7.39]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [175703.0, 176239.0, 176776.0, 177312.0]
source_col: d(18_16)Mean
dD_mean | -59.4895
varray: [-58.013, -59.354, -59.947, -60.644]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [175703.0, 176239.0, 176776.0, 177312.0]
source_col: d(D_H)Mean
d18O_stdev | 0.12544454286
varray: [-7.111, -7.154, -7.271, -7.39]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [175703.0, 176239.0, 176776.0, 177312.0]
source_col: d(18_16)Mean
dD_stdev | 1.11663019841
varray: [-58.013, -59.354, -59.947, -60.644]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [175703.0, 176239.0, 176776.0, 177312.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.2315
varray: [-7.111, -7.154, -7.271, -7.39]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [175703.0, 176239.0, 176776.0, 177312.0]
source_col: d(18_16)Mean
measured_dD_mean | -59.4895
varray: [-58.013, -59.354, -59.947, -60.644]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [175703.0, 176239.0, 176776.0, 177312.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.12544454286
varray: [-7.111, -7.154, -7.271, -7.39]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [175703.0, 176239.0, 176776.0, 177312.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.11663019841
varray: [-58.013, -59.354, -59.947, -60.644]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [175703.0, 176239.0, 176776.0, 177312.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.0133550714
actual_array: [ 1.93 -15.84]
varray: [-7.111, -7.154, -7.271, -7.39]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.2315
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [175703.0, 176239.0, 176776.0, 177312.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -73.7922762522
actual_array: [ 18.1 -119. ]
varray: [-58.013, -59.354, -59.947, -60.644]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -59.4895
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [175703.0, 176239.0, 176776.0, 177312.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.502
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -7.4005
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 47 - Sample 47 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 47
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8525
varray: [2.772, 2.889, 2.857, 2.892]
inst: mean_all
operation: mean_all(residuals)
tarray: [177898.0, 178433.0, 178969.0, 179505.0]
source_col: residuals
SAMP_residuals_STD | 0.0559553393342
varray: [2.772, 2.889, 2.857, 2.892]
inst: std_all
operation: std_all(residuals)
tarray: [177898.0, 178433.0, 178969.0, 179505.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.40125
varray: [0.343, 0.78, 0.289, 0.193]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [177898.0, 178433.0, 178969.0, 179505.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.260008172948
varray: [0.343, 0.78, 0.289, 0.193]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [177898.0, 178433.0, 178969.0, 179505.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1292.64625
varray: [1289.661, 1292.488, 1294.346, 1294.09]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [177898.0, 178433.0, 178969.0, 179505.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.15331718905
varray: [1289.661, 1292.488, 1294.346, 1294.09]
inst: std_all
operation: std_all(baseline_shift)
tarray: [177898.0, 178433.0, 178969.0, 179505.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.599
varray: [1.628, 1.709, 1.68, 1.379]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [177898.0, 178433.0, 178969.0, 179505.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.150446003603
varray: [1.628, 1.709, 1.68, 1.379]
inst: std_all
operation: std_all(slope_shift)
tarray: [177898.0, 178433.0, 178969.0, 179505.0]
source_col: slope_shift
SAMP_d18O_STD | 0.395522860696
varray: [4.495, 5.263, 5.261, 5.327]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [177898.0, 178433.0, 178969.0, 179505.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 4.66601689524
varray: [7.861, 15.828, 17.406, 17.855]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [177898.0, 178433.0, 178969.0, 179505.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 109.577
varray: [107.286, 110.702, 110.899, 109.421]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [177898.0, 178433.0, 178969.0, 179505.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.6619592855
varray: [107.286, 110.702, 110.899, 109.421]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [177898.0, 178433.0, 178969.0, 179505.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8525
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.40125
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1292.64625
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.599
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.2441
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.2441
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -7.88283924304
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -8.41425
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [4.495, 5.263, 5.261, 5.327]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [177898.0, 178433.0, 178969.0, 179505.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [7.861, 15.828, 17.406, 17.855]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [177898.0, 178433.0, 178969.0, 179505.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 5.0865
varray: [4.495, 5.263, 5.261, 5.327]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [177898.0, 178433.0, 178969.0, 179505.0]
source_col: d(18_16)Mean
dD_mean | 14.7375
varray: [7.861, 15.828, 17.406, 17.855]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [177898.0, 178433.0, 178969.0, 179505.0]
source_col: d(D_H)Mean
d18O_stdev | 0.395522860696
varray: [4.495, 5.263, 5.261, 5.327]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [177898.0, 178433.0, 178969.0, 179505.0]
source_col: d(18_16)Mean
dD_stdev | 4.66601689524
varray: [7.861, 15.828, 17.406, 17.855]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [177898.0, 178433.0, 178969.0, 179505.0]
source_col: d(D_H)Mean
measured_d18O_mean | 5.0865
varray: [4.495, 5.263, 5.261, 5.327]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [177898.0, 178433.0, 178969.0, 179505.0]
source_col: d(18_16)Mean
measured_dD_mean | 14.7375
varray: [7.861, 15.828, 17.406, 17.855]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [177898.0, 178433.0, 178969.0, 179505.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.395522860696
varray: [4.495, 5.263, 5.261, 5.327]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [177898.0, 178433.0, 178969.0, 179505.0]
source_col: d(18_16)Mean
measured_dD_stdev | 4.66601689524
varray: [7.861, 15.828, 17.406, 17.855]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [177898.0, 178433.0, 178969.0, 179505.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 2.32443260153
actual_array: [ 1.93 -15.84]
varray: [4.495, 5.263, 5.261, 5.327]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 5.0865
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [177898.0, 178433.0, 178969.0, 179505.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -4.12113851974
actual_array: [ 18.1 -119. ]
varray: [7.861, 15.828, 17.406, 17.855]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 14.7375
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [177898.0, 178433.0, 178969.0, 179505.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.56175
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.768
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 48 - Sample 48 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 48
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.76375
varray: [2.579, 2.827, 2.791, 2.858]
inst: mean_all
operation: mean_all(residuals)
tarray: [180092.0, 180629.0, 181165.0, 181702.0]
source_col: residuals
SAMP_residuals_STD | 0.126172831201
varray: [2.579, 2.827, 2.791, 2.858]
inst: std_all
operation: std_all(residuals)
tarray: [180092.0, 180629.0, 181165.0, 181702.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.29375
varray: [0.338, 0.634, 0.079, 0.124]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [180092.0, 180629.0, 181165.0, 181702.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.253417146223
varray: [0.338, 0.634, 0.079, 0.124]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [180092.0, 180629.0, 181165.0, 181702.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1290.961
varray: [1283.553, 1290.524, 1293.863, 1295.904]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [180092.0, 180629.0, 181165.0, 181702.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 5.41369516443
varray: [1283.553, 1290.524, 1293.863, 1295.904]
inst: std_all
operation: std_all(baseline_shift)
tarray: [180092.0, 180629.0, 181165.0, 181702.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.8885
varray: [1.614, 1.745, 2.243, 1.952]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [180092.0, 180629.0, 181165.0, 181702.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.274253532338
varray: [1.614, 1.745, 2.243, 1.952]
inst: std_all
operation: std_all(slope_shift)
tarray: [180092.0, 180629.0, 181165.0, 181702.0]
source_col: slope_shift
SAMP_d18O_STD | 0.28304578546
varray: [-5.594, -5.964, -6.187, -6.2]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [180092.0, 180629.0, 181165.0, 181702.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.33850650241
varray: [-41.996, -47.317, -48.866, -49.215]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [180092.0, 180629.0, 181165.0, 181702.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 107.1455
varray: [101.39, 108.586, 108.273, 110.333]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [180092.0, 180629.0, 181165.0, 181702.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.94259783223
varray: [101.39, 108.586, 108.273, 110.333]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [180092.0, 180629.0, 181165.0, 181702.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.76375
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.29375
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1290.961
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.8885
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.33285
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.33285
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -10.7488858748
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -10.84575
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-5.594, -5.964, -6.187, -6.2]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [180092.0, 180629.0, 181165.0, 181702.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-41.996, -47.317, -48.866, -49.215]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [180092.0, 180629.0, 181165.0, 181702.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.98625
varray: [-5.594, -5.964, -6.187, -6.2]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [180092.0, 180629.0, 181165.0, 181702.0]
source_col: d(18_16)Mean
dD_mean | -46.8485
varray: [-41.996, -47.317, -48.866, -49.215]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [180092.0, 180629.0, 181165.0, 181702.0]
source_col: d(D_H)Mean
d18O_stdev | 0.28304578546
varray: [-5.594, -5.964, -6.187, -6.2]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [180092.0, 180629.0, 181165.0, 181702.0]
source_col: d(18_16)Mean
dD_stdev | 3.33850650241
varray: [-41.996, -47.317, -48.866, -49.215]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [180092.0, 180629.0, 181165.0, 181702.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.98625
varray: [-5.594, -5.964, -6.187, -6.2]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [180092.0, 180629.0, 181165.0, 181702.0]
source_col: d(18_16)Mean
measured_dD_mean | -46.8485
varray: [-41.996, -47.317, -48.866, -49.215]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [180092.0, 180629.0, 181165.0, 181702.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.28304578546
varray: [-5.594, -5.964, -6.187, -6.2]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [180092.0, 180629.0, 181165.0, 181702.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.33850650241
varray: [-41.996, -47.317, -48.866, -49.215]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [180092.0, 180629.0, 181165.0, 181702.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.76610469802
actual_array: [ 1.93 -15.84]
varray: [-5.594, -5.964, -6.187, -6.2]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -5.98625
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [180092.0, 180629.0, 181165.0, 181702.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -61.9271483058
actual_array: [ 18.1 -119. ]
varray: [-41.996, -47.317, -48.866, -49.215]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -46.8485
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [180092.0, 180629.0, 181165.0, 181702.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.6505
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -12.45325
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 49 - Sample 49 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 49
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.94825
varray: [2.917, 3.0, 2.972, 2.904]
inst: mean_all
operation: mean_all(residuals)
tarray: [182288.0, 182824.0, 183361.0, 183897.0]
source_col: residuals
SAMP_residuals_STD | 0.0453752869596
varray: [2.917, 3.0, 2.972, 2.904]
inst: std_all
operation: std_all(residuals)
tarray: [182288.0, 182824.0, 183361.0, 183897.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3505
varray: [0.279, 0.472, 0.545, 0.106]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [182288.0, 182824.0, 183361.0, 183897.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.197893068769
varray: [0.279, 0.472, 0.545, 0.106]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [182288.0, 182824.0, 183361.0, 183897.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1298.00075
varray: [1297.031, 1299.435, 1297.353, 1298.184]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [182288.0, 182824.0, 183361.0, 183897.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.07248135182
varray: [1297.031, 1299.435, 1297.353, 1298.184]
inst: std_all
operation: std_all(baseline_shift)
tarray: [182288.0, 182824.0, 183361.0, 183897.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.91825
varray: [1.863, 1.859, 1.768, 2.183]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [182288.0, 182824.0, 183361.0, 183897.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.181870604185
varray: [1.863, 1.859, 1.768, 2.183]
inst: std_all
operation: std_all(slope_shift)
tarray: [182288.0, 182824.0, 183361.0, 183897.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0492950301755
varray: [-7.403, -7.454, -7.5, -7.511]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [182288.0, 182824.0, 183361.0, 183897.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.609100155968
varray: [-57.745, -58.502, -58.661, -59.222]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [182288.0, 182824.0, 183361.0, 183897.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.798
varray: [112.536, 115.316, 114.208, 113.132]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [182288.0, 182824.0, 183361.0, 183897.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.22591788741
varray: [112.536, 115.316, 114.208, 113.132]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [182288.0, 182824.0, 183361.0, 183897.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.94825
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3505
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1298.00075
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.91825
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.14835
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.14835
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.79073822903
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.19325
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-7.403, -7.454, -7.5, -7.511]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [182288.0, 182824.0, 183361.0, 183897.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-57.745, -58.502, -58.661, -59.222]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [182288.0, 182824.0, 183361.0, 183897.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.467
varray: [-7.403, -7.454, -7.5, -7.511]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [182288.0, 182824.0, 183361.0, 183897.0]
source_col: d(18_16)Mean
dD_mean | -58.5325
varray: [-57.745, -58.502, -58.661, -59.222]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [182288.0, 182824.0, 183361.0, 183897.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0492950301755
varray: [-7.403, -7.454, -7.5, -7.511]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [182288.0, 182824.0, 183361.0, 183897.0]
source_col: d(18_16)Mean
dD_stdev | 0.609100155968
varray: [-57.745, -58.502, -58.661, -59.222]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [182288.0, 182824.0, 183361.0, 183897.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.467
varray: [-7.403, -7.454, -7.5, -7.511]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [182288.0, 182824.0, 183361.0, 183897.0]
source_col: d(18_16)Mean
measured_dD_mean | -58.5325
varray: [-57.745, -58.502, -58.661, -59.222]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [182288.0, 182824.0, 183361.0, 183897.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0492950301755
varray: [-7.403, -7.454, -7.5, -7.511]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [182288.0, 182824.0, 183361.0, 183897.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.609100155968
varray: [-57.745, -58.502, -58.661, -59.222]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [182288.0, 182824.0, 183361.0, 183897.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.2492333794
actual_array: [ 1.93 -15.84]
varray: [-7.403, -7.454, -7.5, -7.511]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.467
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [182288.0, 182824.0, 183361.0, 183897.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -72.8940144497
actual_array: [ 18.1 -119. ]
varray: [-57.745, -58.502, -58.661, -59.222]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -58.5325
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [182288.0, 182824.0, 183361.0, 183897.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.466
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -5.4135
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 50 - Sample 50 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 50
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1865
varray: [0.0, 3.037, 2.832, 2.877]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [184481.0, 185017.0, 185552.0, 186088.0]
source_col: residuals
SAMP_residuals_STD | 1.46031880081
varray: [0.0, 3.037, 2.832, 2.877]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [184481.0, 185017.0, 185552.0, 186088.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.2805
varray: [0.0, 0.364, 0.383, 0.375]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [184481.0, 185017.0, 185552.0, 186088.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.187162140046
varray: [0.0, 0.364, 0.383, 0.375]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [184481.0, 185017.0, 185552.0, 186088.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 971.94875
varray: [0.0, 1301.242, 1293.108, 1293.445]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [184481.0, 185017.0, 185552.0, 186088.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 647.976727912
varray: [0.0, 1301.242, 1293.108, 1293.445]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [184481.0, 185017.0, 185552.0, 186088.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.048
varray: [0.0, 1.495, 1.662, 1.035]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [184481.0, 185017.0, 185552.0, 186088.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.747279064339
varray: [0.0, 1.495, 1.662, 1.035]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [184481.0, 185017.0, 185552.0, 186088.0]
source_col: slope_shift
SAMP_d18O_STD | 3.39446474671
varray: [0.0, 6.919, 6.724, 6.716]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [184481.0, 185017.0, 185552.0, 186088.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 8.54048886189
varray: [0.0, 16.727, 17.034, 17.451]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [184481.0, 185017.0, 185552.0, 186088.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 84.50975
varray: [0.0, 117.017, 110.354, 110.668]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [184481.0, 185017.0, 185552.0, 186088.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 56.4233948191
varray: [0.0, 117.017, 110.354, 110.668]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [184481.0, 185017.0, 185552.0, 186088.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1865
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.2805
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 971.94875
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.048
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.9101
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.9101
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -29.3902990377
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -33.4815
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [0.0, 6.919, 6.724, 6.716]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [184481.0, 185017.0, 185552.0, 186088.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [0.0, 16.727, 17.034, 17.451]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [184481.0, 185017.0, 185552.0, 186088.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 5.08975
varray: [0.0, 6.919, 6.724, 6.716]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [184481.0, 185017.0, 185552.0, 186088.0]
source_col: d(18_16)Mean
dD_mean | 12.803
varray: [0.0, 16.727, 17.034, 17.451]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [184481.0, 185017.0, 185552.0, 186088.0]
source_col: d(D_H)Mean
d18O_stdev | 3.39446474671
varray: [0.0, 6.919, 6.724, 6.716]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [184481.0, 185017.0, 185552.0, 186088.0]
source_col: d(18_16)Mean
dD_stdev | 8.54048886189
varray: [0.0, 16.727, 17.034, 17.451]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [184481.0, 185017.0, 185552.0, 186088.0]
source_col: d(D_H)Mean
measured_d18O_mean | 5.08975
varray: [0.0, 6.919, 6.724, 6.716]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [184481.0, 185017.0, 185552.0, 186088.0]
source_col: d(18_16)Mean
measured_dD_mean | 12.803
varray: [0.0, 16.727, 17.034, 17.451]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [184481.0, 185017.0, 185552.0, 186088.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 3.39446474671
varray: [0.0, 6.919, 6.724, 6.716]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [184481.0, 185017.0, 185552.0, 186088.0]
source_col: d(18_16)Mean
measured_dD_stdev | 8.54048886189
varray: [0.0, 16.727, 17.034, 17.451]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [184481.0, 185017.0, 185552.0, 186088.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 2.32768782234
actual_array: [ 1.93 -15.84]
varray: [0.0, 6.919, 6.724, 6.716]
fit_rsq: 1.0
fit_intercept: 2.76579538548
except: True
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 5.08975
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [184481.0, 185017.0, 185552.0, 186088.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -5.93690388738
actual_array: [ 18.1 -119. ]
varray: [0.0, 16.727, 17.034, 17.451]
fit_rsq: 1.0
fit_intercept: 19.1281237053
except: True
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 12.803
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [184481.0, 185017.0, 185552.0, 186088.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1301.22775
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -331.4655
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 51 - Sample 51 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 51
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.86675
varray: [0.0, 2.489, 2.489, 2.489]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [186674.0, 187210.0, 187745.0, 188282.0]
source_col: residuals
SAMP_residuals_STD | 1.2445
varray: [0.0, 2.489, 2.489, 2.489]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [186674.0, 187210.0, 187745.0, 188282.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.15375
varray: [0.0, 0.205, 0.205, 0.205]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [186674.0, 187210.0, 187745.0, 188282.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.1025
varray: [0.0, 0.205, 0.205, 0.205]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [186674.0, 187210.0, 187745.0, 188282.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 960.02775
varray: [0.0, 1280.037, 1280.037, 1280.037]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [186674.0, 187210.0, 187745.0, 188282.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 640.0185
varray: [0.0, 1280.037, 1280.037, 1280.037]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [186674.0, 187210.0, 187745.0, 188282.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.1655
varray: [0.0, 1.554, 1.554, 1.554]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [186674.0, 187210.0, 187745.0, 188282.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.777
varray: [0.0, 1.554, 1.554, 1.554]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [186674.0, 187210.0, 187745.0, 188282.0]
source_col: slope_shift
SAMP_d18O_STD | 2.1915
varray: [0.0, -4.383, -4.383, -4.383]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [186674.0, 187210.0, 187745.0, 188282.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 20.4315
varray: [0.0, -40.863, -40.863, -40.863]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [186674.0, 187210.0, 187745.0, 188282.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 73.3965
varray: [0.0, 97.862, 97.862, 97.862]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [186674.0, 187210.0, 187745.0, 188282.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 48.931
varray: [0.0, 97.862, 97.862, 97.862]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [186674.0, 187210.0, 187745.0, 188282.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.86675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.15375
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 960.02775
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.1655
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.22985
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.22985
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -39.7161402829
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -44.59475
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [0.0, -4.383, -4.383, -4.383]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [186674.0, 187210.0, 187745.0, 188282.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [0.0, -40.863, -40.863, -40.863]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [186674.0, 187210.0, 187745.0, 188282.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.28725
varray: [0.0, -4.383, -4.383, -4.383]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [186674.0, 187210.0, 187745.0, 188282.0]
source_col: d(18_16)Mean
dD_mean | -30.64725
varray: [0.0, -40.863, -40.863, -40.863]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [186674.0, 187210.0, 187745.0, 188282.0]
source_col: d(D_H)Mean
d18O_stdev | 2.1915
varray: [0.0, -4.383, -4.383, -4.383]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [186674.0, 187210.0, 187745.0, 188282.0]
source_col: d(18_16)Mean
dD_stdev | 20.4315
varray: [0.0, -40.863, -40.863, -40.863]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [186674.0, 187210.0, 187745.0, 188282.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.28725
varray: [0.0, -4.383, -4.383, -4.383]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [186674.0, 187210.0, 187745.0, 188282.0]
source_col: d(18_16)Mean
measured_dD_mean | -30.64725
varray: [0.0, -40.863, -40.863, -40.863]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [186674.0, 187210.0, 187745.0, 188282.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 2.1915
varray: [0.0, -4.383, -4.383, -4.383]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [186674.0, 187210.0, 187745.0, 188282.0]
source_col: d(18_16)Mean
measured_dD_stdev | 20.4315
varray: [0.0, -40.863, -40.863, -40.863]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [186674.0, 187210.0, 187745.0, 188282.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.06276901615
actual_array: [ 1.93 -15.84]
varray: [0.0, -4.383, -4.383, -4.383]
fit_rsq: 1.0
fit_intercept: 2.76579538548
except: True
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -3.28725
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [186674.0, 187210.0, 187745.0, 188282.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -46.7202893436
actual_array: [ 18.1 -119. ]
varray: [0.0, -40.863, -40.863, -40.863]
fit_rsq: 1.0
fit_intercept: 19.1281237053
except: True
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -30.64725
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [186674.0, 187210.0, 187745.0, 188282.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1301.5475
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -343.3865
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 52 - Sample 52 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 52
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.87875
varray: [2.857, 2.905, 2.838, 2.915]
inst: mean_all
operation: mean_all(residuals)
tarray: [188868.0, 189403.0, 189939.0, 190475.0]
source_col: residuals
SAMP_residuals_STD | 0.0371337672027
varray: [2.857, 2.905, 2.838, 2.915]
inst: std_all
operation: std_all(residuals)
tarray: [188868.0, 189403.0, 189939.0, 190475.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.466
varray: [0.557, 0.367, 0.269, 0.671]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [188868.0, 189403.0, 189939.0, 190475.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.181581937428
varray: [0.557, 0.367, 0.269, 0.671]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [188868.0, 189403.0, 189939.0, 190475.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1292.2575
varray: [1291.079, 1293.164, 1292.156, 1292.631]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [188868.0, 189403.0, 189939.0, 190475.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.88701916552
varray: [1291.079, 1293.164, 1292.156, 1292.631]
inst: std_all
operation: std_all(baseline_shift)
tarray: [188868.0, 189403.0, 189939.0, 190475.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.833
varray: [1.001, 0.704, 0.819, 0.808]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [188868.0, 189403.0, 189939.0, 190475.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.123404483981
varray: [1.001, 0.704, 0.819, 0.808]
inst: std_all
operation: std_all(slope_shift)
tarray: [188868.0, 189403.0, 189939.0, 190475.0]
source_col: slope_shift
SAMP_d18O_STD | 0.552393805782
varray: [12.429, 13.461, 13.456, 13.643]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [188868.0, 189403.0, 189939.0, 190475.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 5.36294894003
varray: [42.888, 51.812, 53.544, 54.692]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [188868.0, 189403.0, 189939.0, 190475.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.20175
varray: [109.262, 110.927, 108.618, 112.0]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [188868.0, 189403.0, 189939.0, 190475.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.543921927
varray: [109.262, 110.927, 108.618, 112.0]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [188868.0, 189403.0, 189939.0, 190475.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.87875
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.466
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1292.2575
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.833
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.21785
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.21785
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -7.03513530969
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -7.7895
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [12.429, 13.461, 13.456, 13.643]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [188868.0, 189403.0, 189939.0, 190475.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [42.888, 51.812, 53.544, 54.692]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [188868.0, 189403.0, 189939.0, 190475.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 13.24725
varray: [12.429, 13.461, 13.456, 13.643]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [188868.0, 189403.0, 189939.0, 190475.0]
source_col: d(18_16)Mean
dD_mean | 50.734
varray: [42.888, 51.812, 53.544, 54.692]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [188868.0, 189403.0, 189939.0, 190475.0]
source_col: d(D_H)Mean
d18O_stdev | 0.552393805782
varray: [12.429, 13.461, 13.456, 13.643]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [188868.0, 189403.0, 189939.0, 190475.0]
source_col: d(18_16)Mean
dD_stdev | 5.36294894003
varray: [42.888, 51.812, 53.544, 54.692]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [188868.0, 189403.0, 189939.0, 190475.0]
source_col: d(D_H)Mean
measured_d18O_mean | 13.24725
varray: [12.429, 13.461, 13.456, 13.643]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [188868.0, 189403.0, 189939.0, 190475.0]
source_col: d(18_16)Mean
measured_dD_mean | 50.734
varray: [42.888, 51.812, 53.544, 54.692]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [188868.0, 189403.0, 189939.0, 190475.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.552393805782
varray: [12.429, 13.461, 13.456, 13.643]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [188868.0, 189403.0, 189939.0, 190475.0]
source_col: d(18_16)Mean
measured_dD_stdev | 5.36294894003
varray: [42.888, 51.812, 53.544, 54.692]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [188868.0, 189403.0, 189939.0, 190475.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 10.4982920554
actual_array: [ 1.93 -15.84]
varray: [12.429, 13.461, 13.456, 13.643]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 13.24725
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [188868.0, 189403.0, 189939.0, 190475.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 29.6659889323
actual_array: [ 18.1 -119. ]
varray: [42.888, 51.812, 53.544, 54.692]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 50.734
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [188868.0, 189403.0, 189939.0, 190475.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.5355
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -11.15675
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 53 - Sample 53 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 53
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.776
varray: [2.728, 2.812, 2.8, 2.764]
inst: mean_all
operation: mean_all(residuals)
tarray: [191061.0, 191598.0, 192133.0, 192669.0]
source_col: residuals
SAMP_residuals_STD | 0.037947331922
varray: [2.728, 2.812, 2.8, 2.764]
inst: std_all
operation: std_all(residuals)
tarray: [191061.0, 191598.0, 192133.0, 192669.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.34925
varray: [0.172, 0.338, 0.476, 0.411]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [191061.0, 191598.0, 192133.0, 192669.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.130923323616
varray: [0.172, 0.338, 0.476, 0.411]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [191061.0, 191598.0, 192133.0, 192669.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1291.4775
varray: [1289.896, 1293.585, 1291.709, 1290.72]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [191061.0, 191598.0, 192133.0, 192669.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.58851051827
varray: [1289.896, 1293.585, 1291.709, 1290.72]
inst: std_all
operation: std_all(baseline_shift)
tarray: [191061.0, 191598.0, 192133.0, 192669.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.838
varray: [1.182, 2.568, 1.707, 1.895]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [191061.0, 191598.0, 192133.0, 192669.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.572609814795
varray: [1.182, 2.568, 1.707, 1.895]
inst: std_all
operation: std_all(slope_shift)
tarray: [191061.0, 191598.0, 192133.0, 192669.0]
source_col: slope_shift
SAMP_d18O_STD | 0.600752028711
varray: [-5.75, -6.629, -6.944, -7.091]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [191061.0, 191598.0, 192133.0, 192669.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 5.86842309881
varray: [-43.865, -53.188, -55.764, -56.705]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [191061.0, 191598.0, 192133.0, 192669.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.05475
varray: [106.829, 108.836, 109.48, 107.074]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [191061.0, 191598.0, 192133.0, 192669.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.30461012695
varray: [106.829, 108.836, 109.48, 107.074]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [191061.0, 191598.0, 192133.0, 192669.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.776
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.34925
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1291.4775
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.838
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.3206
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.3206
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -10.3532907059
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -9.9365
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-5.75, -6.629, -6.944, -7.091]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [191061.0, 191598.0, 192133.0, 192669.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-43.865, -53.188, -55.764, -56.705]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [191061.0, 191598.0, 192133.0, 192669.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.6035
varray: [-5.75, -6.629, -6.944, -7.091]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [191061.0, 191598.0, 192133.0, 192669.0]
source_col: d(18_16)Mean
dD_mean | -52.3805
varray: [-43.865, -53.188, -55.764, -56.705]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [191061.0, 191598.0, 192133.0, 192669.0]
source_col: d(D_H)Mean
d18O_stdev | 0.600752028711
varray: [-5.75, -6.629, -6.944, -7.091]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [191061.0, 191598.0, 192133.0, 192669.0]
source_col: d(18_16)Mean
dD_stdev | 5.86842309881
varray: [-43.865, -53.188, -55.764, -56.705]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [191061.0, 191598.0, 192133.0, 192669.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.6035
varray: [-5.75, -6.629, -6.944, -7.091]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [191061.0, 191598.0, 192133.0, 192669.0]
source_col: d(18_16)Mean
measured_dD_mean | -52.3805
varray: [-43.865, -53.188, -55.764, -56.705]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [191061.0, 191598.0, 192133.0, 192669.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.600752028711
varray: [-5.75, -6.629, -6.944, -7.091]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [191061.0, 191598.0, 192133.0, 192669.0]
source_col: d(18_16)Mean
measured_dD_stdev | 5.86842309881
varray: [-43.865, -53.188, -55.764, -56.705]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [191061.0, 191598.0, 192133.0, 192669.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.38434625032
actual_array: [ 1.93 -15.84]
varray: [-5.75, -6.629, -6.944, -7.091]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.6035
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [191061.0, 191598.0, 192133.0, 192669.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -67.1196083801
actual_array: [ 18.1 -119. ]
varray: [-43.865, -53.188, -55.764, -56.705]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -52.3805
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [191061.0, 191598.0, 192133.0, 192669.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.63825
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -11.93675
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 54 - Sample 54 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 54
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8745
varray: [2.839, 2.864, 2.887, 2.908]
inst: mean_all
operation: mean_all(residuals)
tarray: [193256.0, 193792.0, 194327.0, 194863.0]
source_col: residuals
SAMP_residuals_STD | 0.0297153159162
varray: [2.839, 2.864, 2.887, 2.908]
inst: std_all
operation: std_all(residuals)
tarray: [193256.0, 193792.0, 194327.0, 194863.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.50775
varray: [0.472, 0.481, 0.514, 0.564]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [193256.0, 193792.0, 194327.0, 194863.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0416203075433
varray: [0.472, 0.481, 0.514, 0.564]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [193256.0, 193792.0, 194327.0, 194863.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1293.215
varray: [1293.239, 1292.311, 1293.637, 1293.673]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [193256.0, 193792.0, 194327.0, 194863.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.633940060258
varray: [1293.239, 1292.311, 1293.637, 1293.673]
inst: std_all
operation: std_all(baseline_shift)
tarray: [193256.0, 193792.0, 194327.0, 194863.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.47775
varray: [1.626, 2.022, 1.135, 1.128]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [193256.0, 193792.0, 194327.0, 194863.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.431272825483
varray: [1.626, 2.022, 1.135, 1.128]
inst: std_all
operation: std_all(slope_shift)
tarray: [193256.0, 193792.0, 194327.0, 194863.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0544854720698
varray: [-4.437, -4.321, -4.324, -4.346]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [193256.0, 193792.0, 194327.0, 194863.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.656172423986
varray: [-42.758, -41.596, -41.318, -41.483]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [193256.0, 193792.0, 194327.0, 194863.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.07325
varray: [111.055, 109.611, 111.558, 112.069]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [193256.0, 193792.0, 194327.0, 194863.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.05908903466
varray: [111.055, 109.611, 111.558, 112.069]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [193256.0, 193792.0, 194327.0, 194863.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8745
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.50775
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1293.215
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.47775
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.2221
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.2221
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -7.17238261319
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -6.918
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-4.437, -4.321, -4.324, -4.346]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [193256.0, 193792.0, 194327.0, 194863.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-42.758, -41.596, -41.318, -41.483]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [193256.0, 193792.0, 194327.0, 194863.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.357
varray: [-4.437, -4.321, -4.324, -4.346]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [193256.0, 193792.0, 194327.0, 194863.0]
source_col: d(18_16)Mean
dD_mean | -41.78875
varray: [-42.758, -41.596, -41.318, -41.483]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [193256.0, 193792.0, 194327.0, 194863.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0544854720698
varray: [-4.437, -4.321, -4.324, -4.346]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [193256.0, 193792.0, 194327.0, 194863.0]
source_col: d(18_16)Mean
dD_stdev | 0.656172423986
varray: [-42.758, -41.596, -41.318, -41.483]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [193256.0, 193792.0, 194327.0, 194863.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.357
varray: [-4.437, -4.321, -4.324, -4.346]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [193256.0, 193792.0, 194327.0, 194863.0]
source_col: d(18_16)Mean
measured_dD_mean | -41.78875
varray: [-42.758, -41.596, -41.318, -41.483]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [193256.0, 193792.0, 194327.0, 194863.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0544854720698
varray: [-4.437, -4.321, -4.324, -4.346]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [193256.0, 193792.0, 194327.0, 194863.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.656172423986
varray: [-42.758, -41.596, -41.318, -41.483]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [193256.0, 193792.0, 194327.0, 194863.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.13423746583
actual_array: [ 1.93 -15.84]
varray: [-4.437, -4.321, -4.324, -4.346]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -4.357
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [193256.0, 193792.0, 194327.0, 194863.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -57.1779527416
actual_array: [ 18.1 -119. ]
varray: [-42.758, -41.596, -41.318, -41.483]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -41.78875
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [193256.0, 193792.0, 194327.0, 194863.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.53975
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.19925
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 55 - Sample 55 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 55
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.80025
varray: [2.811, 2.78, 2.784, 2.826]
inst: mean_all
operation: mean_all(residuals)
tarray: [195451.0, 195987.0, 196523.0, 197059.0]
source_col: residuals
SAMP_residuals_STD | 0.0220056810847
varray: [2.811, 2.78, 2.784, 2.826]
inst: std_all
operation: std_all(residuals)
tarray: [195451.0, 195987.0, 196523.0, 197059.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.37825
varray: [0.348, 0.375, 0.329, 0.461]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [195451.0, 195987.0, 196523.0, 197059.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0583059459518
varray: [0.348, 0.375, 0.329, 0.461]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [195451.0, 195987.0, 196523.0, 197059.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1291.76375
varray: [1291.326, 1291.274, 1291.821, 1292.634]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [195451.0, 195987.0, 196523.0, 197059.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.6303683447
varray: [1291.326, 1291.274, 1291.821, 1292.634]
inst: std_all
operation: std_all(baseline_shift)
tarray: [195451.0, 195987.0, 196523.0, 197059.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.91775
varray: [1.809, 2.208, 1.606, 2.048]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [195451.0, 195987.0, 196523.0, 197059.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.264716672438
varray: [1.809, 2.208, 1.606, 2.048]
inst: std_all
operation: std_all(slope_shift)
tarray: [195451.0, 195987.0, 196523.0, 197059.0]
source_col: slope_shift
SAMP_d18O_STD | 0.112896338884
varray: [-6.695, -6.926, -6.918, -6.918]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [195451.0, 195987.0, 196523.0, 197059.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.2407516539
varray: [-56.318, -58.759, -58.908, -58.715]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [195451.0, 195987.0, 196523.0, 197059.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.34775
varray: [108.011, 107.721, 108.254, 109.405]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [195451.0, 195987.0, 196523.0, 197059.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.737740412792
varray: [108.011, 107.721, 108.254, 109.405]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [195451.0, 195987.0, 196523.0, 197059.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.80025
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.37825
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1291.76375
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.91775
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.29635
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.29635
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -9.57017373894
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -9.6435
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-6.695, -6.926, -6.918, -6.918]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [195451.0, 195987.0, 196523.0, 197059.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-56.318, -58.759, -58.908, -58.715]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [195451.0, 195987.0, 196523.0, 197059.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.86425
varray: [-6.695, -6.926, -6.918, -6.918]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [195451.0, 195987.0, 196523.0, 197059.0]
source_col: d(18_16)Mean
dD_mean | -58.175
varray: [-56.318, -58.759, -58.908, -58.715]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [195451.0, 195987.0, 196523.0, 197059.0]
source_col: d(D_H)Mean
d18O_stdev | 0.112896338884
varray: [-6.695, -6.926, -6.918, -6.918]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [195451.0, 195987.0, 196523.0, 197059.0]
source_col: d(18_16)Mean
dD_stdev | 1.2407516539
varray: [-56.318, -58.759, -58.908, -58.715]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [195451.0, 195987.0, 196523.0, 197059.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.86425
varray: [-6.695, -6.926, -6.918, -6.918]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [195451.0, 195987.0, 196523.0, 197059.0]
source_col: d(18_16)Mean
measured_dD_mean | -58.175
varray: [-56.318, -58.759, -58.908, -58.715]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [195451.0, 195987.0, 196523.0, 197059.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.112896338884
varray: [-6.695, -6.926, -6.918, -6.918]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [195451.0, 195987.0, 196523.0, 197059.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.2407516539
varray: [-56.318, -58.759, -58.908, -58.715]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [195451.0, 195987.0, 196523.0, 197059.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.64551511992
actual_array: [ 1.93 -15.84]
varray: [-6.695, -6.926, -6.918, -6.918]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.86425
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [195451.0, 195987.0, 196523.0, 197059.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -72.5584568798
actual_array: [ 18.1 -119. ]
varray: [-56.318, -58.759, -58.908, -58.715]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -58.175
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [195451.0, 195987.0, 196523.0, 197059.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.614
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -11.6505
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 56 - Sample 56 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 56
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.5565
varray: [2.505, 2.534, 2.65, 2.537]
inst: mean_all
operation: mean_all(residuals)
tarray: [197646.0, 198182.0, 198718.0, 199254.0]
source_col: residuals
SAMP_residuals_STD | 0.0639817682365
varray: [2.505, 2.534, 2.65, 2.537]
inst: std_all
operation: std_all(residuals)
tarray: [197646.0, 198182.0, 198718.0, 199254.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.93575
varray: [-0.922, -1.205, -0.591, -1.025]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [197646.0, 198182.0, 198718.0, 199254.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.257877716499
varray: [-0.922, -1.205, -0.591, -1.025]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [197646.0, 198182.0, 198718.0, 199254.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1940.066
varray: [1924.243, 1948.377, 1949.033, 1938.611]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [197646.0, 198182.0, 198718.0, 199254.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 11.5753209315
varray: [1924.243, 1948.377, 1949.033, 1938.611]
inst: std_all
operation: std_all(baseline_shift)
tarray: [197646.0, 198182.0, 198718.0, 199254.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -23.4175
varray: [-22.285, -24.02, -23.586, -23.779]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [197646.0, 198182.0, 198718.0, 199254.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.775593751049
varray: [-22.285, -24.02, -23.586, -23.779]
inst: std_all
operation: std_all(slope_shift)
tarray: [197646.0, 198182.0, 198718.0, 199254.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0894632140417
varray: [-3.761, -3.593, -3.585, -3.571]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [197646.0, 198182.0, 198718.0, 199254.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.17721663257
varray: [-38.704, -36.757, -36.563, -35.998]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [197646.0, 198182.0, 198718.0, 199254.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 96.02025
varray: [93.834, 96.612, 97.645, 95.99]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [197646.0, 198182.0, 198718.0, 199254.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.60940825047
varray: [93.834, 96.612, 97.645, 95.99]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [197646.0, 198182.0, 198718.0, 199254.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.5565
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.93575
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1940.066
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 23.4175
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.5401
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.5401
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -17.4417102629
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -21.971
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-3.761, -3.593, -3.585, -3.571]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [197646.0, 198182.0, 198718.0, 199254.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-38.704, -36.757, -36.563, -35.998]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [197646.0, 198182.0, 198718.0, 199254.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.6275
varray: [-3.761, -3.593, -3.585, -3.571]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [197646.0, 198182.0, 198718.0, 199254.0]
source_col: d(18_16)Mean
dD_mean | -37.0055
varray: [-38.704, -36.757, -36.563, -35.998]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [197646.0, 198182.0, 198718.0, 199254.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0894632140417
varray: [-3.761, -3.593, -3.585, -3.571]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [197646.0, 198182.0, 198718.0, 199254.0]
source_col: d(18_16)Mean
dD_stdev | 1.17721663257
varray: [-38.704, -36.757, -36.563, -35.998]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [197646.0, 198182.0, 198718.0, 199254.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.6275
varray: [-3.761, -3.593, -3.585, -3.571]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [197646.0, 198182.0, 198718.0, 199254.0]
source_col: d(18_16)Mean
measured_dD_mean | -37.0055
varray: [-38.704, -36.757, -36.563, -35.998]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [197646.0, 198182.0, 198718.0, 199254.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0894632140417
varray: [-3.761, -3.593, -3.585, -3.571]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [197646.0, 198182.0, 198718.0, 199254.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.17721663257
varray: [-38.704, -36.757, -36.563, -35.998]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [197646.0, 198182.0, 198718.0, 199254.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.40356559479
actual_array: [ 1.93 -15.84]
varray: [-3.761, -3.593, -3.585, -3.571]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -3.6275
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [197646.0, 198182.0, 198718.0, 199254.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -52.6882863189
actual_array: [ 18.1 -119. ]
varray: [-38.704, -36.757, -36.563, -35.998]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -37.0055
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [197646.0, 198182.0, 198718.0, 199254.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.85775
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 636.65175
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 57 - Sample 57 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 57
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.9475
varray: [0.0, 2.54, 2.462, 2.788]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [199840.0, 200376.0, 200912.0, 201447.0]
source_col: residuals
SAMP_residuals_STD | 1.30575176303
varray: [0.0, 2.54, 2.462, 2.788]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [199840.0, 200376.0, 200912.0, 201447.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.56475
varray: [0.0, -0.807, -0.935, -0.517]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [199840.0, 200376.0, 200912.0, 201447.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.415127590828
varray: [0.0, -0.807, -0.935, -0.517]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [199840.0, 200376.0, 200912.0, 201447.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1295.24425
varray: [0.0, 1728.92, 1697.275, 1754.782]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [199840.0, 200376.0, 200912.0, 201447.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 863.816336752
varray: [0.0, 1728.92, 1697.275, 1754.782]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [199840.0, 200376.0, 200912.0, 201447.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -12.5335
varray: [0.0, -16.626, -15.561, -17.947]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [199840.0, 200376.0, 200912.0, 201447.0]
source_col: slope_shift
SAMP_slope_shift_STD | 8.412469257
varray: [0.0, -16.626, -15.561, -17.947]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [199840.0, 200376.0, 200912.0, 201447.0]
source_col: slope_shift
SAMP_d18O_STD | 1.41714839849
varray: [0.0, -2.846, -2.889, -2.762]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [199840.0, 200376.0, 200912.0, 201447.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 17.9312035546
varray: [0.0, -35.663, -36.458, -35.434]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [199840.0, 200376.0, 200912.0, 201447.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 75.612
varray: [0.0, 100.037, 94.401, 108.01]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [199840.0, 200376.0, 200912.0, 201447.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 50.7162435978
varray: [0.0, 100.037, 94.401, 108.01]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [199840.0, 200376.0, 200912.0, 201447.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.9475
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.56475
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1295.24425
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 12.5335
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.1491
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.1491
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -37.1084415165
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -42.37925
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [0.0, -2.846, -2.889, -2.762]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [199840.0, 200376.0, 200912.0, 201447.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [0.0, -35.663, -36.458, -35.434]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [199840.0, 200376.0, 200912.0, 201447.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.12425
varray: [0.0, -2.846, -2.889, -2.762]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [199840.0, 200376.0, 200912.0, 201447.0]
source_col: d(18_16)Mean
dD_mean | -26.88875
varray: [0.0, -35.663, -36.458, -35.434]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [199840.0, 200376.0, 200912.0, 201447.0]
source_col: d(D_H)Mean
d18O_stdev | 1.41714839849
varray: [0.0, -2.846, -2.889, -2.762]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [199840.0, 200376.0, 200912.0, 201447.0]
source_col: d(18_16)Mean
dD_stdev | 17.9312035546
varray: [0.0, -35.663, -36.458, -35.434]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [199840.0, 200376.0, 200912.0, 201447.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.12425
varray: [0.0, -2.846, -2.889, -2.762]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [199840.0, 200376.0, 200912.0, 201447.0]
source_col: d(18_16)Mean
measured_dD_mean | -26.88875
varray: [0.0, -35.663, -36.458, -35.434]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [199840.0, 200376.0, 200912.0, 201447.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 1.41714839849
varray: [0.0, -2.846, -2.889, -2.762]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [199840.0, 200376.0, 200912.0, 201447.0]
source_col: d(18_16)Mean
measured_dD_stdev | 17.9312035546
varray: [0.0, -35.663, -36.458, -35.434]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [199840.0, 200376.0, 200912.0, 201447.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.89790076938
actual_array: [ 1.93 -15.84]
varray: [0.0, -2.846, -2.889, -2.762]
fit_rsq: 1.0
fit_intercept: 2.76579538548
except: True
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.12425
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [199840.0, 200376.0, 200912.0, 201447.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -43.1924764026
actual_array: [ 18.1 -119. ]
varray: [0.0, -35.663, -36.458, -35.434]
fit_rsq: 1.0
fit_intercept: 19.1281237053
except: True
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -26.88875
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [199840.0, 200376.0, 200912.0, 201447.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1301.46675
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -8.17
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 58 - Sample 58 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 58
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.878
varray: [2.869, 2.865, 2.826, 2.952]
inst: mean_all
operation: mean_all(residuals)
tarray: [202033.0, 202569.0, 203104.0, 203640.0]
source_col: residuals
SAMP_residuals_STD | 0.0530094331228
varray: [2.869, 2.865, 2.826, 2.952]
inst: std_all
operation: std_all(residuals)
tarray: [202033.0, 202569.0, 203104.0, 203640.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.439
varray: [0.51, 0.479, 0.105, 0.662]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [202033.0, 202569.0, 203104.0, 203640.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.236591067738
varray: [0.51, 0.479, 0.105, 0.662]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [202033.0, 202569.0, 203104.0, 203640.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1299.01225
varray: [1303.56, 1298.865, 1297.078, 1296.546]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [202033.0, 202569.0, 203104.0, 203640.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.18995427083
varray: [1303.56, 1298.865, 1297.078, 1296.546]
inst: std_all
operation: std_all(baseline_shift)
tarray: [202033.0, 202569.0, 203104.0, 203640.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.689
varray: [1.587, 1.801, 2.015, 1.353]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [202033.0, 202569.0, 203104.0, 203640.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.284089187874
varray: [1.587, 1.801, 2.015, 1.353]
inst: std_all
operation: std_all(slope_shift)
tarray: [202033.0, 202569.0, 203104.0, 203640.0]
source_col: slope_shift
SAMP_d18O_STD | 0.106315176088
varray: [-7.264, -7.422, -7.511, -7.458]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [202033.0, 202569.0, 203104.0, 203640.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.29608821716
varray: [-56.223, -58.18, -58.898, -59.035]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [202033.0, 202569.0, 203104.0, 203640.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 110.776
varray: [110.657, 110.911, 108.629, 112.907]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [202033.0, 202569.0, 203104.0, 203640.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.7495862368
varray: [110.657, 110.911, 108.629, 112.907]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [202033.0, 202569.0, 203104.0, 203640.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.878
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.439
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1299.01225
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.689
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.2186
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.2186
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -7.05935542208
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -7.21525
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-7.264, -7.422, -7.511, -7.458]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [202033.0, 202569.0, 203104.0, 203640.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-56.223, -58.18, -58.898, -59.035]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [202033.0, 202569.0, 203104.0, 203640.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.41375
varray: [-7.264, -7.422, -7.511, -7.458]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [202033.0, 202569.0, 203104.0, 203640.0]
source_col: d(18_16)Mean
dD_mean | -58.084
varray: [-56.223, -58.18, -58.898, -59.035]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [202033.0, 202569.0, 203104.0, 203640.0]
source_col: d(D_H)Mean
d18O_stdev | 0.106315176088
varray: [-7.264, -7.422, -7.511, -7.458]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [202033.0, 202569.0, 203104.0, 203640.0]
source_col: d(18_16)Mean
dD_stdev | 1.29608821716
varray: [-56.223, -58.18, -58.898, -59.035]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [202033.0, 202569.0, 203104.0, 203640.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.41375
varray: [-7.264, -7.422, -7.511, -7.458]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [202033.0, 202569.0, 203104.0, 203640.0]
source_col: d(18_16)Mean
measured_dD_mean | -58.084
varray: [-56.223, -58.18, -58.898, -59.035]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [202033.0, 202569.0, 203104.0, 203640.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.106315176088
varray: [-7.264, -7.422, -7.511, -7.458]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [202033.0, 202569.0, 203104.0, 203640.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.29608821716
varray: [-56.223, -58.18, -58.898, -59.035]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [202033.0, 202569.0, 203104.0, 203640.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.1958978384
actual_array: [ 1.93 -15.84]
varray: [-7.264, -7.422, -7.511, -7.458]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.41375
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [202033.0, 202569.0, 203104.0, 203640.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -72.4730422257
actual_array: [ 18.1 -119. ]
varray: [-56.223, -58.18, -58.898, -59.035]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -58.084
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [202033.0, 202569.0, 203104.0, 203640.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.53625
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -4.402
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 59 - Sample 59 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 59
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.88875
varray: [2.72, 3.144, 2.832, 2.859]
inst: mean_all
operation: mean_all(residuals)
tarray: [204227.0, 204763.0, 205298.0, 205834.0]
source_col: residuals
SAMP_residuals_STD | 0.180494459749
varray: [2.72, 3.144, 2.832, 2.859]
inst: std_all
operation: std_all(residuals)
tarray: [204227.0, 204763.0, 205298.0, 205834.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.428
varray: [0.405, 0.527, 0.213, 0.567]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [204227.0, 204763.0, 205298.0, 205834.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.159034587433
varray: [0.405, 0.527, 0.213, 0.567]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [204227.0, 204763.0, 205298.0, 205834.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1297.04075
varray: [1289.631, 1308.455, 1294.765, 1295.312]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [204227.0, 204763.0, 205298.0, 205834.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 8.02822339728
varray: [1289.631, 1308.455, 1294.765, 1295.312]
inst: std_all
operation: std_all(baseline_shift)
tarray: [204227.0, 204763.0, 205298.0, 205834.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.57825
varray: [2.787, 2.386, 2.628, 2.512]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [204227.0, 204763.0, 205298.0, 205834.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.170685627202
varray: [2.787, 2.386, 2.628, 2.512]
inst: std_all
operation: std_all(slope_shift)
tarray: [204227.0, 204763.0, 205298.0, 205834.0]
source_col: slope_shift
SAMP_d18O_STD | 0.432089786194
varray: [-19.131, -19.071, -19.792, -19.897]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [204227.0, 204763.0, 205298.0, 205834.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 5.47798350977
varray: [-149.208, -155.98, -160.197, -161.299]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [204227.0, 204763.0, 205298.0, 205834.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.55425
varray: [104.946, 121.682, 109.023, 110.566]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [204227.0, 204763.0, 205298.0, 205834.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 7.15598008079
varray: [104.946, 121.682, 109.023, 110.566]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [204227.0, 204763.0, 205298.0, 205834.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.88875
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.428
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1297.04075
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.57825
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.20785
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.20785
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.71220047794
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -6.437
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-19.131, -19.071, -19.792, -19.897]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [204227.0, 204763.0, 205298.0, 205834.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-149.208, -155.98, -160.197, -161.299]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [204227.0, 204763.0, 205298.0, 205834.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -19.47275
varray: [-19.131, -19.071, -19.792, -19.897]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [204227.0, 204763.0, 205298.0, 205834.0]
source_col: d(18_16)Mean
dD_mean | -156.671
varray: [-149.208, -155.98, -160.197, -161.299]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [204227.0, 204763.0, 205298.0, 205834.0]
source_col: d(D_H)Mean
d18O_stdev | 0.432089786194
varray: [-19.131, -19.071, -19.792, -19.897]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [204227.0, 204763.0, 205298.0, 205834.0]
source_col: d(18_16)Mean
dD_stdev | 5.47798350977
varray: [-149.208, -155.98, -160.197, -161.299]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [204227.0, 204763.0, 205298.0, 205834.0]
source_col: d(D_H)Mean
measured_d18O_mean | -19.47275
varray: [-19.131, -19.071, -19.792, -19.897]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [204227.0, 204763.0, 205298.0, 205834.0]
source_col: d(18_16)Mean
measured_dD_mean | -156.671
varray: [-149.208, -155.98, -160.197, -161.299]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [204227.0, 204763.0, 205298.0, 205834.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.432089786194
varray: [-19.131, -19.071, -19.792, -19.897]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [204227.0, 204763.0, 205298.0, 205834.0]
source_col: d(18_16)Mean
measured_dD_stdev | 5.47798350977
varray: [-149.208, -155.98, -160.197, -161.299]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [204227.0, 204763.0, 205298.0, 205834.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -22.274269453
actual_array: [ 1.93 -15.84]
varray: [-19.131, -19.071, -19.792, -19.897]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -19.47275
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [204227.0, 204763.0, 205298.0, 205834.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -165.009025838
actual_array: [ 18.1 -119. ]
varray: [-149.208, -155.98, -160.197, -161.299]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -156.671
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [204227.0, 204763.0, 205298.0, 205834.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.5255
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -6.3735
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 60 - Sample 60 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Sample Variables
Sample 60
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.893
varray: [2.871, 2.872, 2.92, 2.909]
inst: mean_all
operation: mean_all(residuals)
tarray: [206421.0, 206957.0, 207493.0, 208029.0]
source_col: residuals
SAMP_residuals_STD | 0.0252322544904
varray: [2.871, 2.872, 2.92, 2.909]
inst: std_all
operation: std_all(residuals)
tarray: [206421.0, 206957.0, 207493.0, 208029.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.431
varray: [0.457, 0.423, 0.54, 0.304]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [206421.0, 206957.0, 207493.0, 208029.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0978945010373
varray: [0.457, 0.423, 0.54, 0.304]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [206421.0, 206957.0, 207493.0, 208029.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1296.2955
varray: [1296.07, 1295.658, 1296.861, 1296.593]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [206421.0, 206957.0, 207493.0, 208029.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.537138405503
varray: [1296.07, 1295.658, 1296.861, 1296.593]
inst: std_all
operation: std_all(baseline_shift)
tarray: [206421.0, 206957.0, 207493.0, 208029.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.96525
varray: [1.909, 2.081, 2.031, 1.84]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [206421.0, 206957.0, 207493.0, 208029.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.110412484198
varray: [1.909, 2.081, 2.031, 1.84]
inst: std_all
operation: std_all(slope_shift)
tarray: [206421.0, 206957.0, 207493.0, 208029.0]
source_col: slope_shift
SAMP_d18O_STD | 0.325539936106
varray: [-8.348, -7.794, -7.67, -7.661]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [206421.0, 206957.0, 207493.0, 208029.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 5.38088334601
varray: [-72.198, -63.258, -61.378, -60.445]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [206421.0, 206957.0, 207493.0, 208029.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.95275
varray: [111.583, 110.816, 113.528, 111.884]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [206421.0, 206957.0, 207493.0, 208029.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.14237395366
varray: [111.583, 110.816, 113.528, 111.884]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [206421.0, 206957.0, 207493.0, 208029.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.893
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.431
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1296.2955
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.96525
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.2036
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.2036
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.57495317445
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -6.0385
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0488
varray: [-8.348, -7.794, -7.67, -7.661]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [206421.0, 206957.0, 207493.0, 208029.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.6534
varray: [-72.198, -63.258, -61.378, -60.445]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [206421.0, 206957.0, 207493.0, 208029.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.86825
varray: [-8.348, -7.794, -7.67, -7.661]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [206421.0, 206957.0, 207493.0, 208029.0]
source_col: d(18_16)Mean
dD_mean | -64.31975
varray: [-72.198, -63.258, -61.378, -60.445]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [206421.0, 206957.0, 207493.0, 208029.0]
source_col: d(D_H)Mean
d18O_stdev | 0.325539936106
varray: [-8.348, -7.794, -7.67, -7.661]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [206421.0, 206957.0, 207493.0, 208029.0]
source_col: d(18_16)Mean
dD_stdev | 5.38088334601
varray: [-72.198, -63.258, -61.378, -60.445]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [206421.0, 206957.0, 207493.0, 208029.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.86825
varray: [-8.348, -7.794, -7.67, -7.661]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [206421.0, 206957.0, 207493.0, 208029.0]
source_col: d(18_16)Mean
measured_dD_mean | -64.31975
varray: [-72.198, -63.258, -61.378, -60.445]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [206421.0, 206957.0, 207493.0, 208029.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.325539936106
varray: [-8.348, -7.794, -7.67, -7.661]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [206421.0, 206957.0, 207493.0, 208029.0]
source_col: d(18_16)Mean
measured_dD_stdev | 5.38088334601
varray: [-72.198, -63.258, -61.378, -60.445]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [206421.0, 206957.0, 207493.0, 208029.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.6511279486
actual_array: [ 1.93 -15.84]
varray: [-8.348, -7.794, -7.67, -7.661]
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.86825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: [206421.0, 206957.0, 207493.0, 208029.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -78.3260579358
actual_array: [ 18.1 -119. ]
varray: [-72.198, -63.258, -61.378, -60.445]
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -64.31975
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: [206421.0, 206957.0, 207493.0, 208029.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1300.52125
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -7.11875
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 1 - Sample 1 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 1 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.77023836767
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.9540886906
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 1 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.54
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.804
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.3239570499
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -8.54
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -87.2281890746
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -73.804
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 1 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.996
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.476
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.7790831666
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.996
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.349962859
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -66.476
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 1 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.81
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.642
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5927843756
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.81
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.6899064869
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -63.642
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 1 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.663
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.793
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4455482344
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.663
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.9543933493
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -61.793
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 1 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.596
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.773
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3784406054
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.596
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.9969983247
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -60.773
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 1 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.629
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.621
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4114936167
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.629
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.8543276936
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -60.621
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 1 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.592
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.485
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3744341797
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.592
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.7266750237
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -60.485
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 1 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.572
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.328
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3544020517
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.572
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.5793112797
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -60.328
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 1 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.542
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.173
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3243538596
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.542
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.4338247809
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -60.173
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 2 - Sample 2 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 2 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.921
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.951
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6927
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 38.4117
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.002
operation: col(d(18_16)Mean)
inst: col
dD_actual | 27.931
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.2381904574
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 4.002
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 8.26257839826
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 27.931
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 2 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.921
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.951
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6927
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 38.4117
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.502
operation: col(d(18_16)Mean)
inst: col
dD_actual | 35.95
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.73899365894
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 4.502
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 15.7893928118
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 35.95
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 2 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.921
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.951
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6927
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 38.4117
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.634
operation: col(d(18_16)Mean)
inst: col
dD_actual | 38.323
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.87120570414
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 4.634
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 18.0167441778
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 38.323
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 2 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.921
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.951
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6927
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 38.4117
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.702
operation: col(d(18_16)Mean)
inst: col
dD_actual | 39.038
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.93931493955
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 4.702
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 18.6878593175
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 39.038
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 2 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.921
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.951
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6927
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 38.4117
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.734
operation: col(d(18_16)Mean)
inst: col
dD_actual | 39.387
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.97136634445
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 4.734
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.0154385955
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 39.387
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 2 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.921
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.951
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6927
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 38.4117
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.845
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.433
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.08254465519
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 4.845
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.997237807
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 40.433
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 2 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.921
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.951
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6927
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 38.4117
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.825
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.454
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.06251252713
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 4.825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.016948881
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 40.454
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 2 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.921
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.951
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6927
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 38.4117
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.881
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.7
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.1186024857
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 4.881
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.247850034
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 40.7
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 2 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.921
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.951
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6927
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 38.4117
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.881
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.95
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.1186024857
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 4.881
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.4825056773
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 40.95
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 2 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.921
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 40.951
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6927
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 38.4117
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.921
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.951
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.15866674182
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 4.921
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.4834442998
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 40.951
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 3 - Sample 3 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 3 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.045
operation: col(d(18_16)Mean)
inst: col
dD_actual | -92.364
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.8345874926
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -12.045
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -104.649024031
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -92.364
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 3 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.85
operation: col(d(18_16)Mean)
inst: col
dD_actual | -104.802
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.6408806471
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -12.85
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -116.323611595
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -104.802
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 3 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.025
operation: col(d(18_16)Mean)
inst: col
dD_actual | -107.803
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.8161617676
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -13.025
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -119.140417937
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -107.803
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 3 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.068
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.829
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.8592308429
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -13.068
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.103444697
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -108.829
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 3 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.188
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.722
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9794236113
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -13.188
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.941634655
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -109.722
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 3 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.205
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.03
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9964509202
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -13.205
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.230730407
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -110.03
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 3 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.267
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.336
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0585505171
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -13.267
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.517948915
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -110.336
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 3 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.27
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.728
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0615553364
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -13.27
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.885888963
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -110.728
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 3 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.217
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.81
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.008470197
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -13.217
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.962856014
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -110.81
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 3 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.353
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.11
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.1446886678
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -13.353
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.244442786
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -111.11
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 4 - Sample 4 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 4 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.83
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.169
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.6128165037
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.83
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.123183156
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -65.169
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 4 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.649
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.279
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4315257447
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.649
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.4719413467
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -61.279
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 4 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.59
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.391
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3724309669
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.59
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.6384445018
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -60.391
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 4 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.575
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.146
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3574068709
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.575
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.4084819714
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -60.146
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 4 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.582
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.969
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3644181157
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.582
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.242345776
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -59.969
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 4 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.549
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.972
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3313651044
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.549
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.2451616437
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -59.972
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 4 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.567
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.801
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3493940197
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.567
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.0846571837
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -59.801
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 4 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.588
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.029
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3704277541
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.588
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.2986631303
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -60.029
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 4 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.54
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.642
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3223506468
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.54
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.9354161946
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -59.642
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 4 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.535
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.547
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3173426148
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.535
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.8462470501
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -59.547
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 5 - Sample 5 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 5 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.132
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.458
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.91369523434
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.132
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.0082193488
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -55.458
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 5 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.205
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.637
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.98681250176
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.205
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.1762327894
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -55.637
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 5 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.164
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.339
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.94574663924
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.164
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.8965232626
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -55.339
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 5 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.226
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.742
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0078462362
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.226
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.2747881595
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -55.742
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 6 - Sample 6 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 6 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.24
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.322
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.0202623228
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.24
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -65.1874538134
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -50.322
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 6 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.323
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.493
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.10339565426
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.323
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -65.3479582734
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -50.493
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 6 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.247
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.183
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.02727356762
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.247
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -65.0569852757
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -50.183
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 6 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.281
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.258
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.06132818533
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.281
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -65.1273819687
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -50.258
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 7 - Sample 7 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 7 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.102
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.369
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.88204063918
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.102
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.1701916474
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -51.369
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 7 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.162
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.867
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.94213702336
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.162
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.6376256888
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -51.867
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 7 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.132
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.62
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.91208883127
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.132
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.4057859133
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -51.62
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 7 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.092
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.632
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.87202457515
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.092
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.4170493841
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -51.632
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 8 - Sample 8 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 8 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.8
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.579
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5827683116
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.8
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.9376603994
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -58.579
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 8 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.98
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.743
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.7630574641
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.98
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.0302170744
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -59.743
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 8 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.994
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.325
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.7770799538
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.994
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.576495412
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -60.325
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 8 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.023
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.63
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.8061265395
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -8.023
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.8627752968
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -60.63
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 9 - Sample 9 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 9 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.684
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.748
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.46497556576
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.684
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.6486844565
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -49.748
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 9 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.496
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.403
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.27667356199
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.496
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -63.3862370957
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -48.403
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 9 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.532
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.199
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.3127313925
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.532
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -63.1947580907
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -48.199
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 9 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.433
operation: col(d(18_16)Mean)
inst: col
dD_actual | -47.334
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.21357235859
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.433
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.382849565
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -47.334
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 10 - Sample 10 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 10 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.761
operation: col(d(18_16)Mean)
inst: col
dD_actual | -28.095
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.53567364653
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.761
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -44.3246898814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -28.095
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 10 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.442
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.908
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.21616120396
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.442
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -42.2719223141
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -25.908
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 10 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.506
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.724
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.28026401375
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.506
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -42.0992157606
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -25.724
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 10 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.496
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.439
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.27024794972
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.496
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.8317083273
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -25.439
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 11 - Sample 11 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 11 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.612
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.686
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.39285990474
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.612
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -65.52911243
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -50.686
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 11 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.527
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.862
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.30772336048
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.527
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.632932576
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -51.862
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 11 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.738
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.072
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.51906231153
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.738
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.7686658894
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -53.072
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 11 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.754
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.641
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.53508801398
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.754
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.3641195604
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -52.641
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 12 - Sample 12 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 12 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.542
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.252
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.3195346504
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -4.542
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.5513922217
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -43.252
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 12 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.002
operation: col(d(18_16)Mean)
inst: col
dD_actual | -39.884
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.77866719274
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -4.002
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.3901113955
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -39.884
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 12 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.986
operation: col(d(18_16)Mean)
inst: col
dD_actual | -39.147
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.76264149029
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -3.986
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.6983465592
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -39.147
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 12 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.89
operation: col(d(18_16)Mean)
inst: col
dD_actual | -38.881
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.6664872756
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -3.89
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.4486729547
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -38.881
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 13 - Sample 13 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 13 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.201
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.688
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.98280607615
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.201
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.1627251137
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -56.688
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 13 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.38
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.454
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1620936223
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.38
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.8203325777
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -58.454
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 13 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.457
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.619
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2392173153
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.457
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.9752053023
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -58.619
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 13 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.507
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.154
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2892976355
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.507
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.4773683789
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -59.154
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 14 - Sample 14 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 14 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.77023836767
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.9540886906
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 14 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 14.907
operation: col(d(18_16)Mean)
inst: col
dD_actual | 63.694
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 12.1607082828
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 14.907
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 41.8305374795
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 63.694
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 14 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 15.027
operation: col(d(18_16)Mean)
inst: col
dD_actual | 64.027
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 12.2809010512
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 15.027
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 42.1430987964
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 64.027
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 14 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 14.907
operation: col(d(18_16)Mean)
inst: col
dD_actual | 63.694
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 12.1607082828
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 14.907
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 41.8305374795
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 63.694
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 15 - Sample 15 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 15 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.504
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.282
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.28308001015
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -5.504
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.6409283258
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -42.282
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 15 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.254
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.614
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.03428481245
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.254
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.4001541778
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -51.614
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 15 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.731
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.091
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.51205106671
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.731
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.6637448644
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -55.091
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 15 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.813
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.007
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.59418279176
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.813
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.5235231414
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -56.007
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 16 - Sample 16 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 16 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.358
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.174
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.13523907223
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -4.358
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -65.0485376726
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -50.174
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 16 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.969
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.663
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.74561418144
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -3.969
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -63.6302789647
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -48.663
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 16 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.181
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.207
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.95795473889
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -4.181
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.1408896444
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -49.207
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 16 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.282
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.603
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.0591169856
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -4.282
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.5125841834
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -49.603
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 17 - Sample 17 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 17 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.312
operation: col(d(18_16)Mean)
inst: col
dD_actual | -33.383
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.08755877462
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -3.312
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -49.2881260479
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -33.383
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 17 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.24
operation: col(d(18_16)Mean)
inst: col
dD_actual | -31.682
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.0154431136
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -3.24
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -47.6915290511
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -31.682
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 17 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.222
operation: col(d(18_16)Mean)
inst: col
dD_actual | -31.168
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.99741419835
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -3.222
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -47.2090770485
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -31.168
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 17 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.229
operation: col(d(18_16)Mean)
inst: col
dD_actual | -31.453
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.00442544317
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -3.229
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -47.4765844819
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -31.453
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 18 - Sample 18 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 18 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.454
operation: col(d(18_16)Mean)
inst: col
dD_actual | -31.79
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.22818048079
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.454
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -47.792900289
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -31.79
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 18 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.207
operation: col(d(18_16)Mean)
inst: col
dD_actual | -31.349
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.98078369924
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.207
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -47.3789677343
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -31.349
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 18 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.399
operation: col(d(18_16)Mean)
inst: col
dD_actual | -31.87
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.17309212862
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.399
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -47.8679900948
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -31.87
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 18 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.392
operation: col(d(18_16)Mean)
inst: col
dD_actual | -31.525
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.1660808838
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.392
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -47.5441653071
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -31.525
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 19 - Sample 19 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 19 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.831
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.566
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.61221170701
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.831
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.1095905867
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -55.566
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 19 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.867
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.028
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.64826953753
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.867
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.4818567885
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -57.028
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 19 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.064
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.21
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.84558599893
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.064
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.5913086699
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -58.21
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 19 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.092
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.409
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.87363097821
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.092
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.7780945619
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -58.409
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 20 - Sample 20 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 20 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.721
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.391
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.50203500268
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.721
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.3135961979
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -49.391
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 20 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.503
operation: col(d(18_16)Mean)
inst: col
dD_actual | -47.574
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.28368480681
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.503
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.6081189826
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -47.574
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 20 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.669
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.141
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.44995146972
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.669
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -63.1403179815
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -48.141
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 20 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.692
operation: col(d(18_16)Mean)
inst: col
dD_actual | -47.82
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.47298841699
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.692
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.8390201355
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -47.82
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 21 - Sample 21 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 21 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.412
operation: col(d(18_16)Mean)
inst: col
dD_actual | -11.651
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.354363723473
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 2.412
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.8899802896
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -11.651
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 21 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.042
operation: col(d(18_16)Mean)
inst: col
dD_actual | -7.642
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.276648310459
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 3.042
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.1270423941
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -7.642
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 21 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.92
operation: col(d(18_16)Mean)
inst: col
dD_actual | -8.016
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.154452329284
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 2.92
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.4780872364
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -8.016
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 21 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.023
operation: col(d(18_16)Mean)
inst: col
dD_actual | -7.416
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.2576177888
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 3.023
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -24.9149136926
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -7.416
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 22 - Sample 22 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 22 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.835
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.081
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.61621813263
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.835
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.7771134925
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -53.081
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 22 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.273
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.245
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0549217372
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.273
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.6855378869
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -57.245
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 22 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.39
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.231
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1721096863
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.39
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.6110197439
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -58.231
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 22 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.431
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.506
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2131755489
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.431
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.8691409515
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -58.506
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 23 - Sample 23 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 23 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 11.277
operation: col(d(18_16)Mean)
inst: col
dD_actual | 38.442
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 8.52487703971
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 11.277
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 18.128440264
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 38.442
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 23 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 12.322
operation: col(d(18_16)Mean)
inst: col
dD_actual | 48.366
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 9.57155573091
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 12.322
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 27.4433306793
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 48.366
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 23 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 12.42
operation: col(d(18_16)Mean)
inst: col
dD_actual | 49.9
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 9.66971315841
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 12.42
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 28.8831777064
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 49.9
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 23 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 12.475
operation: col(d(18_16)Mean)
inst: col
dD_actual | 51.016
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 9.72480151058
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 12.475
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 29.9306804979
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 51.016
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 24 - Sample 24 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 24 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.77023836767
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.9540886906
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 24 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.027
operation: col(d(18_16)Mean)
inst: col
dD_actual | -19.615
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.80049454668
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.027
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -36.3651704617
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -19.615
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 24 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.414
operation: col(d(18_16)Mean)
inst: col
dD_actual | -21.897
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.18811622467
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.414
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -38.5071071734
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -21.897
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 24 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.449
operation: col(d(18_16)Mean)
inst: col
dD_actual | -22.828
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.22317244878
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.449
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -39.380964789
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -22.828
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 25 - Sample 25 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 25 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.063
operation: col(d(18_16)Mean)
inst: col
dD_actual | -29.011
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.83494597413
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -1.063
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -45.1844681584
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -29.011
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 25 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.77023836767
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.9540886906
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 25 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.988
operation: col(d(18_16)Mean)
inst: col
dD_actual | -29.654
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.7598254939
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -0.988
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -45.7880024729
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -29.654
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 25 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.053
operation: col(d(18_16)Mean)
inst: col
dD_actual | -30.055
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.8249299101
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -1.053
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -46.1643901247
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -30.055
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 26 - Sample 26 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 26 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 5.553
operation: col(d(18_16)Mean)
inst: col
dD_actual | 7.71
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.79168198856
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 5.553
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -10.7173086521
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 7.71
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 26 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 6.411
operation: col(d(18_16)Mean)
inst: col
dD_actual | 14.018
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.65106028239
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 6.411
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -4.79647746108
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 14.018
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 26 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 6.002
operation: col(d(18_16)Mean)
inst: col
dD_actual | 12.809
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.24140326353
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 6.002
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -5.93127215194
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 12.809
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 26 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 5.993
operation: col(d(18_16)Mean)
inst: col
dD_actual | 12.896
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.23238880591
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 5.993
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -5.84961198808
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 12.896
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 27 - Sample 27 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 27 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.955
operation: col(d(18_16)Mean)
inst: col
dD_actual | -47.445
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.73480449793
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -5.955
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.4870366706
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -47.445
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 27 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.365
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.137
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.14546312319
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.365
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.8296763566
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -53.137
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 27 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.708
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.39
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.48901411944
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.708
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.9443930138
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -55.39
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 27 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.794
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.179
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.5751522701
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.794
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.684966224
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -56.179
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 28 - Sample 28 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 28 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.022
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.814
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.80191212693
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.022
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -63.7720109732
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -48.814
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 28 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.941
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.251
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.72078200829
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -5.941
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -63.2435664645
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -48.251
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 28 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.051
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.76
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.83095871262
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.051
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -63.7213253542
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -48.76
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 28 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.042
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.089
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.821944255
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.042
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.0301321808
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -49.089
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 29 - Sample 29 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 29 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.286
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.408
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.06633621734
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.286
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -65.2681753547
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -50.408
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 29 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.067
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.357
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.84698441507
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.067
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.2816830304
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -49.357
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 29 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.328
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.273
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.10840368627
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.328
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -65.1414613073
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -50.273
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 29 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.259
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.078
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.03929284446
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.259
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.9584299056
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -50.078
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 30 - Sample 30 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 30 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.166
operation: col(d(18_16)Mean)
inst: col
dD_actual | -27.532
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.93971783671
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.166
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -43.7962453728
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -27.532
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 30 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.999
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.691
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.7724495674
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -1.999
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -42.0682412157
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -25.691
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 30 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.003
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.33
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.77645599301
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.003
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.7293984668
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -25.33
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 30 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.966
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.209
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.7393965561
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -1.966
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -41.6158251355
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -25.209
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 31 - Sample 31 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 31 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.217
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.14
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.9988317786
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.217
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.7097373705
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -55.14
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 31 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.491
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.174
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.273271933
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.491
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.5575182573
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -58.174
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 31 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.545
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.785
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3273586788
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.545
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.1310166494
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -58.785
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 31 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.561
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.252
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3433843813
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.561
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.5693533911
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -59.252
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 32 - Sample 32 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 32 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.758
operation: col(d(18_16)Mean)
inst: col
dD_actual | -29.838
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.53266882733
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.758
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -45.9607090263
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -29.838
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 32 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.251
operation: col(d(18_16)Mean)
inst: col
dD_actual | -26.156
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.02485438097
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.251
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -42.5047007122
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -26.156
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 32 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.493
operation: col(d(18_16)Mean)
inst: col
dD_actual | -26.487
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.26724313051
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.493
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -42.8153847839
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -26.487
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 32 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.364
operation: col(d(18_16)Mean)
inst: col
dD_actual | -26.005
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.13803590452
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.364
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -42.3629687037
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -26.005
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 33 - Sample 33 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 33 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 5.384
operation: col(d(18_16)Mean)
inst: col
dD_actual | 8.649
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.62241050644
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 5.384
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -9.83594205597
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 8.649
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 33 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 5.77
operation: col(d(18_16)Mean)
inst: col
dD_actual | 12.038
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.00903057802
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 5.77
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -6.65495015579
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 12.038
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 33 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 5.777
operation: col(d(18_16)Mean)
inst: col
dD_actual | 12.751
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.01604182284
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 5.777
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -5.98571226118
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 12.751
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 33 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 5.814
operation: col(d(18_16)Mean)
inst: col
dD_actual | 13.001
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.05310125976
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 5.814
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -5.75105661791
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 13.001
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 34 - Sample 34 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 34 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.735
operation: col(d(18_16)Mean)
inst: col
dD_actual | -19.798
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.50802547699
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -1.735
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -36.5369383925
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -19.798
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 34 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.073
operation: col(d(18_16)Mean)
inst: col
dD_actual | -22.93
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.84656844123
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.073
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -39.4767042914
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -22.93
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 34 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.227
operation: col(d(18_16)Mean)
inst: col
dD_actual | -24.09
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.0008158273
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.227
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.5655064762
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -24.09
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 34 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.275
operation: col(d(18_16)Mean)
inst: col
dD_actual | -24.141
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.04889293464
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.275
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.6133762274
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -24.141
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 35 - Sample 35 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 35 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.0
operation: col(d(18_16)Mean)
inst: col
dD_actual | -28.089
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.77505757687
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -3.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -44.319058146
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -28.089
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 35 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.159
operation: col(d(18_16)Mean)
inst: col
dD_actual | -28.666
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.93431299496
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -3.159
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -44.8606433707
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -28.666
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 35 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.126
operation: col(d(18_16)Mean)
inst: col
dD_actual | -28.552
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.90125998365
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -3.126
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -44.7536403973
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -28.552
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 35 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.114
operation: col(d(18_16)Mean)
inst: col
dD_actual | -28.8
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.88924070682
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -3.114
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -44.9864187955
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -28.8
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 36 - Sample 36 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 36 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.892
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.672
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.6733096976
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.892
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.2090845794
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -55.672
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 36 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.902
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.699
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.68332576163
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.902
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.1116725351
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -57.699
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 36 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.125
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.082
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.90668398952
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.125
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.4097875536
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -59.082
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 36 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.151
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.26
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.932725756
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.151
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.5768623716
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -59.26
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 37 - Sample 37 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 37 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.77023836767
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.9540886906
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 37 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.77023836767
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.9540886906
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 37 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 9.047
operation: col(d(18_16)Mean)
inst: col
dD_actual | 41.738
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 6.29129476087
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 9.047
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 21.2221402649
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 41.738
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 37 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 9.047
operation: col(d(18_16)Mean)
inst: col
dD_actual | 41.738
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 6.29129476087
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 9.047
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 21.2221402649
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 41.738
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 38 - Sample 38 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 38 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.596
operation: col(d(18_16)Mean)
inst: col
dD_actual | -11.692
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.3671957839
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -0.596
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.9284638151
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -11.692
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 38 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.043
operation: col(d(18_16)Mean)
inst: col
dD_actual | -16.765
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.81491384607
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -1.043
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -33.6900961284
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -16.765
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 38 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.228
operation: col(d(18_16)Mean)
inst: col
dD_actual | -18.683
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.00021103063
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -1.228
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -35.4903742235
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -18.683
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 38 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.331
operation: col(d(18_16)Mean)
inst: col
dD_actual | -19.254
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.10337649015
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -1.331
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -36.0263277128
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -19.254
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 39 - Sample 39 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 39 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.598
operation: col(d(18_16)Mean)
inst: col
dD_actual | -32.291
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.37241180284
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.598
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -48.2631501981
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -32.291
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 39 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.73
operation: col(d(18_16)Mean)
inst: col
dD_actual | -33.945
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.50462384804
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.73
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -49.815631934
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -33.945
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 39 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.778
operation: col(d(18_16)Mean)
inst: col
dD_actual | -34.89
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.55270095539
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.778
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -50.7026302655
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -34.89
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 39 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.806
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.003
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.58074593467
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.806
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -50.8086946163
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -35.003
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 40 - Sample 40 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 40 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.182
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.013
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.96377555449
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.182
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.5291548768
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -56.013
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 40 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.485
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.389
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2672622946
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.485
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.7593221105
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -58.389
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 40 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.496
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.674
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2782799651
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.496
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.0268295438
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -58.674
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 40 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.572
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.277
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3544020517
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.572
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.5928189554
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -59.277
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 41 - Sample 41 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 41 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.759
operation: col(d(18_16)Mean)
inst: col
dD_actual | -30.987
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.53367043373
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.759
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -47.0391863628
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -30.987
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 41 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.77023836767
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.9540886906
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 41 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.759
operation: col(d(18_16)Mean)
inst: col
dD_actual | -30.987
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.53367043373
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.759
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -47.0391863628
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -30.987
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 41 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.908
operation: col(d(18_16)Mean)
inst: col
dD_actual | -31.84
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.68290978779
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.908
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -47.8398314176
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -31.84
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 42 - Sample 42 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 42 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.615
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.961
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.39586472395
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.615
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -68.6031013569
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -53.961
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 42 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.758
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.784
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.53909443959
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.758
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.3142103076
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -55.784
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 42 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.858
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.705
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.6392550799
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.858
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.1786816974
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -56.705
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 42 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.909
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.9
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.69033700645
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.909
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.3617130992
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -56.9
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 43 - Sample 43 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 43 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.593
operation: col(d(18_16)Mean)
inst: col
dD_actual | -23.386
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.36740377082
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.593
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -39.9047161848
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -23.386
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 43 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.142
operation: col(d(18_16)Mean)
inst: col
dD_actual | -18.81
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.91567928304
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.142
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -35.6095792903
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -18.81
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 43 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.219
operation: col(d(18_16)Mean)
inst: col
dD_actual | -18.539
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.99280297607
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.219
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -35.355212573
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -18.539
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 43 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.276
operation: col(d(18_16)Mean)
inst: col
dD_actual | -18.971
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.04989454105
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.276
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -35.7606975246
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -18.971
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 44 - Sample 44 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 44 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.469
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.536
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2512365922
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.469
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.713167348
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -61.536
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 44 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.448
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.846
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2302028577
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.448
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.8200080649
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -64.846
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 44 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.817
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.939
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5997956204
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.817
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.7845451104
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -66.939
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 44 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.825
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.354
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.6078084717
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.1740734782
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -67.354
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 45 - Sample 45 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 45 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.853
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.975
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.62942783868
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -3.853
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.4141486228
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -40.975
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 45 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.577
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.955
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.35298447144
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -3.577
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.5795084521
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -37.955
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 45 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.654
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.659
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.43010816447
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -3.654
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.3016761704
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -37.659
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 45 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.588
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.339
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.36400214187
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -3.588
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.001316947
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -37.339
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 46 - Sample 46 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 46 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.111
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.013
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.89266149987
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.111
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.406400023
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -58.013
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 46 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.154
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.354
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.93573057521
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.154
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.6650928935
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -59.354
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 46 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.271
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.947
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0529185244
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.271
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.2216960794
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -59.947
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 46 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.39
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.644
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1721096863
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.39
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.8759160128
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -60.644
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 47 - Sample 47 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 47 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.495
operation: col(d(18_16)Mean)
inst: col
dD_actual | 7.861
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.73198241411
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 4.495
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -10.5755766436
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 7.861
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 47 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 5.263
operation: col(d(18_16)Mean)
inst: col
dD_actual | 15.828
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.50121613167
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 5.263
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -3.09757060379
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 15.828
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 47 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 5.261
operation: col(d(18_16)Mean)
inst: col
dD_actual | 17.406
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.49921291886
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 5.261
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -1.61642418346
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 17.406
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 47 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 5.327
operation: col(d(18_16)Mean)
inst: col
dD_actual | 17.855
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.56531894146
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 5.327
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -1.19498264815
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 17.855
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 48 - Sample 48 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 48 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.594
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.996
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.37322458642
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -5.594
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.3724822699
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -41.996
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 48 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.964
operation: col(d(18_16)Mean)
inst: col
dD_actual | -47.317
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.74381895556
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -5.964
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.3668929813
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -47.317
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 48 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.187
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.866
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.96717718344
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.187
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -63.820819347
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -48.866
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 48 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.2
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.215
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.98019806668
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.2
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.148398625
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -49.215
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 49 - Sample 49 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 49 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.403
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.745
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1851305696
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.403
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.1548491734
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -57.745
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 49 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.454
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.502
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2362124961
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.454
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.8653864612
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -58.502
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 49 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.5
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.661
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2822863907
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.5
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.0146274504
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -58.661
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 49 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.511
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.222
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2933040611
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.511
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.5411947139
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -59.222
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 50 - Sample 50 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 50 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.77023836767
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.9540886906
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 50 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 6.919
operation: col(d(18_16)Mean)
inst: col
dD_actual | 16.727
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 4.15987633515
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 6.919
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -2.25374891059
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 16.727
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 50 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 6.724
operation: col(d(18_16)Mean)
inst: col
dD_actual | 17.034
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.96456308655
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 6.724
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -1.96559178065
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 17.034
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 50 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 6.716
operation: col(d(18_16)Mean)
inst: col
dD_actual | 17.451
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.95655023532
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 6.716
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -1.57418616767
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 17.451
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 51 - Sample 51 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 51 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.77023836767
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.9540886906
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 51 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.383
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.863
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.16027923231
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -4.383
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.3090228946
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -40.863
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 51 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.383
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.863
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.16027923231
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -4.383
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.3090228946
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -40.863
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 51 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.383
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.863
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.16027923231
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -4.383
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.3090228946
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -40.863
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 52 - Sample 52 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 52 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 12.429
operation: col(d(18_16)Mean)
inst: col
dD_actual | 42.888
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 9.67872761604
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 12.429
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 22.3015562239
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 42.888
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 52 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 13.461
operation: col(d(18_16)Mean)
inst: col
dD_actual | 51.812
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 10.712385424
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 13.461
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 30.6778240661
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 51.812
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 52 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 13.456
operation: col(d(18_16)Mean)
inst: col
dD_actual | 53.544
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 10.707377392
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 13.456
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 32.3035183627
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 53.544
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 52 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 13.643
operation: col(d(18_16)Mean)
inst: col
dD_actual | 54.692
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 10.8946777894
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 13.643
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 33.3810570766
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 54.692
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 53 - Sample 53 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 53 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.75
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.865
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.5294751853
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -5.75
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -59.126767859
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -43.865
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 53 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.629
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.188
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.4098872136
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.629
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.8775461079
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -53.188
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 53 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.944
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.764
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.72539323056
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.944
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.2954378561
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -55.764
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 53 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.091
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.705
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.87262937181
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.091
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.1786816974
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -56.705
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 54 - Sample 54 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 54 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.437
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.758
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.21436597807
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -4.437
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.0877126706
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -42.758
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 54 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.321
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.596
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.09817963532
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -4.321
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.9970332407
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -41.596
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 54 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.324
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.318
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.10118445453
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -4.324
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.7360961653
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -41.318
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 54 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.346
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.483
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.12321979539
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -4.346
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.8909688899
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -41.483
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 55 - Sample 55 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 55 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.695
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.318
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.4759932362
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.695
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.8154347616
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -56.318
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 55 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.926
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.759
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.70736431531
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.926
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.1066124625
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -58.759
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 55 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.918
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.908
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.69935146408
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.918
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.2464672259
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -58.908
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 55 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.918
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.715
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.69935146408
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -6.918
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.0653130693
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -58.715
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 56 - Sample 56 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 56 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.761
operation: col(d(18_16)Mean)
inst: col
dD_actual | -38.704
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.5372800496
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -3.761
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.2825367593
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -38.704
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 56 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.593
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.757
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.36901017389
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -3.593
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.4550386095
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -36.757
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 56 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.585
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.563
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.36099732266
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -3.585
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.2729458303
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -36.563
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 56 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.571
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.998
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.34697483302
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -3.571
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.7426240765
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -35.998
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 57 - Sample 57 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 57 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.77023836767
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.9540886906
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 57 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.846
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.663
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.6208101908
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.846
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.4281855145
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -35.663
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 57 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.889
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.458
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.66387926613
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.889
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.1743904601
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -36.458
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 57 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.762
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.434
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.53667525294
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -2.762
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.2132409453
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -35.434
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 58 - Sample 58 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 58 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.264
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.223
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0459072795
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.264
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.7262656172
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -56.223
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 58 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.422
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.18
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2041610912
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.422
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.5631499927
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -58.18
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 58 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.511
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.898
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2933040611
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.511
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.2370810002
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -58.898
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 58 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.458
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.035
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2402189217
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.458
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.3656722927
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -59.035
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 59 - Sample 59 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 59 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -19.131
operation: col(d(18_16)Mean)
inst: col
dD_actual | -149.208
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -21.9319704647
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -19.131
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -158.004085576
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -149.208
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 59 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -19.071
operation: col(d(18_16)Mean)
inst: col
dD_actual | -155.98
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -21.8718740805
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -19.071
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -164.36043764
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -155.98
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 59 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -19.792
operation: col(d(18_16)Mean)
inst: col
dD_actual | -160.197
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -22.5940322972
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -19.792
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -168.318609031
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -160.197
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 59 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -19.897
operation: col(d(18_16)Mean)
inst: col
dD_actual | -161.299
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -22.6992009695
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -19.897
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -169.352971107
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -161.299
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean

## 60 - Sample 60 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:15:09
Detail Variables
Detail 60 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.348
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.198
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.1316486205
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -8.348
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -85.7207612222
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -72.198
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 60 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.794
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.258
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5767586732
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.794
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.3294754188
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -63.258
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 60 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.67
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.378
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4525594792
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.67
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.5648649814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -61.378
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean
Detail 60 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6927
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.4117
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.353
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.11
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0488
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.6534
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.661
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.445
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4435450216
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.76579538548
measure_array: [ 4.6927 -13.0488]
inst: fit_type1
actual_mean: -7.661
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.998396173326
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.6891301207
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.1281237053
measure_array: [ 38.4117 -107.6534]
inst: fit_type1
actual_mean: -60.445
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06539095551
tarray: []
source_col: d(D_H)Mean