## Summary
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
230224_23-011_HIDS2046
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -10.247 | -75.9998 |  | 20.0747 |  | 0.24081 | 2.6854 |  | -0.9844
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.2433 | 3.5606
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.33627 | 5.0362
4 |  | EV | -10.3587 | -75.2061 |  |  |  | 0.10743 | 1.6395
5 |  | TALL5-3040-3 | -4.7924 | -35.877 |  |  |  | 0.15284 | 1.953
6 |  | TALL5-010-2 | -3.0509 | -26.4311 |  | 22.8067 |  | 0.067072 | 0.67406 |  | -1.3493
7 |  | TALL5-2030-2 | -3.4796 | -27.2551 |  | 12.2234 |  | 0.033808 | 0.20914
8 |  | TALL5-2030-3 | -4.1663 | -31.451 |  | 21.0389 |  | 0.016021 | 0.12576 |  | -1.3038
9 |  | TALL5-3040-1 | -4.7942 | -29.5815 |  |  |  | 0.017597 | 0.29906
10 |  | TALL5-1020-1 | -3.8452 | -31.2311 |  | 7.9819 |  | 0.028135 | 0.25643
11 |  | TALL5-2030-1 | -3.6195 | -31.0901 |  | 14.9291 |  | 0.037408 | 0.18745
12 |  | TALL5-1020-2 | -4.188 | -29.2049 |  |  |  | 0.029269 | 0.22123
13 |  | EV | -10.0886 | -70.7869 |  |  |  | 0.14494 | 2.0587
14 |  | TALL5-3040-2 | -3.3223 | -26.4351 |  |  |  | 0.18074 | 2.2378
15 |  | TALL5-deep-3 | -5.129 | -36.1391 |  | 8.4346 |  | 0.058418 | 0.54833
16 |  | TALL5-B-4 | -0.67252 | -23.5566 | -3.0267 | 290.3662 | 18.4116 | 0.03573 | 0.25519 |  | -3.9523
17 |  | TALL5-B-5 | -1.8831 | -23.1007 | -1.9207 | 189.5094 | 7.7593 | 0.02645 | 0.26861 |  | -2.6268
18 |  | TALL5-C-2 | -3.0602 | -27.9279 | -1.6287 | 126.0571 |  | 0.041409 | 0.19863 |  | -3.182
19 |  | TALL5-B-3 | 2.0206 | -20.4431 | -4.0755 | 163.9631 | 33.0268 | 0.068862 | 0.338 |  | -0.33705
20 |  | TALL5-A-3 | -0.35187 | -28.2153 | -2.592 | 5.6004 | 15.6041 | 0.027837 | 0.33893 |  | 1.3142
21 |  | TALL5-B-1 | -2.6807 | -24.9968 | -1.5853 | 240.1626 |  | 0.022485 | 0.34632 |  | -5.2278
22 |  | EV | -10.0373 | -70.5515 |  |  |  | 0.16453 | 2.1361
23 |  | TALL5-A-4 | -2.0498 | -32.64 |  |  |  | 0.17965 | 1.7677
24 |  | TALL5-A-1 | 0.92994 | -24.2726 | -2.935 | 6.8136 | 19.9829 | 0.024792 | 0.21127 |  | 2.0767
25 |  | TALL5-010-1 | -3.5836 | -25.7153 |  | 2.8769 |  | 0.034189 | 0.24178
26 |  | TALL5-C-4 | -3.7711 | -29.0759 | -1.5123 | 270.7434 |  | 0.032588 | 0.35183 |  | -6.102
27 |  | TALL5-C-3 | -2.4826 | -28.2352 | -1.6965 | 161.7929 |  | 0.042035 | 0.11733 |  | -3.7415
28 |  | TALL5-A-5 | -1.9161 | -29.9748 |  | 4.3811 |  | 0.023514 | 0.31194
29 |  | TALL5-A-2 | -1.8094 | -31.2074 |  | 3.4001 |  | 0.028849 | 0.21099
30 |  | TALL5-D-1 | -2.3864 | -27.9682 | -1.6365 | 3.2132 |  | 0.058844 | 0.10296
31 |  | EV | -10.0257 | -70.4984 |  |  |  | 0.18833 | 2.1818
32 |  | TALL5-C-5 | -3.7928 | -32.9465 | -1.446 | 258.2014 |  | 0.16404 | 2.1687 |  | -5.3073
33 |  | TALL5-D-4 | -1.9585 | -27.3145 | -1.6102 | 94.5864 |  | 0.0631 | 0.55128 |  | -3.273
34 |  | TALL5-D-2 | -3.0419 | -30.8473 |  | 225.9761 |  | 0.0095 | 0.13701 |  | -17.395
35 |  | TALL5-1020-3 | -3.6391 | -27.6449 |  | 15.7241 |  | 0.019664 | 0.27152
36 |  | TALL5-010-3 | -3.6068 | -26.6417 |  | 12.4427 |  | 0.022706 | 0.15085
37 |  | TALL5-D-5 | -2.8185 | -28.5392 |  | 7.0151 |  | 0.031009 | 0.18278
38 |  | TALL5-B-2 | -2.0343 | -27.4761 | -2.0728 | 216.2034 | 7.5629 | 0.037322 | 0.17257 |  | -3.5573
39 |  | TALL5-C-1 | -3.095 | -28.4172 | -1.7532 | 106.3137 |  | 0.032877 | 0.1913 |  | -1.3558
40 |  | EV | -10.0711 | -70.6767 |  |  |  | 0.17113 | 2.1757
41 |  | TALL5-D-3 | -3.1293 | -30.4176 | -1.6007 | 282.9621 |  | 0.13502 | 1.7742 |  | -4.8568
42 |  | EV | -10.078 | -70.8503 |  |  |  | 0.17429 | 1.9809

## Detail
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
1 |  | EV | -9.604 | -75.854 | -10.864 | -83.3177 | -1.267 | 85.273 | 2023/02/24 10:33:59
1 |  | EV | -9.214 | -71.414 | -10.4616 | -78.4843 | -1.116 | 127.087 | 2023/02/24 10:42:56
1 |  | EV | -9.081 | -69.657 | -10.3244 | -76.5716 | -1.035 | 103.398 | 2023/02/24 10:51:52
1 |  | EV | -8.931 | -68.615 | -10.1696 | -75.4372 | -0.923 | 55.828 | 2023/02/24 11:00:49
1 |  | EV | -8.891 | -68.058 | -10.1283 | -74.8309 | -1.093 | 53.453 | 2023/02/24 11:09:46
1 |  | EV | -8.877 | -67.881 | -10.1139 | -74.6382 | -1.111 | 52.949 | 2023/02/24 11:18:42
1 |  | EV | -8.858 | -67.55 | -10.0943 | -74.2778 | -1.04 | 51.784 | 2023/02/24 11:27:39
1 |  | EV | -8.891 | -67.602 | -10.1283 | -74.3344 | -0.854 | 51.415 | 2023/02/24 11:36:36
1 |  | EV | -8.864 | -67.401 | -10.1005 | -74.1156 | -1.086 | 51.447 | 2023/02/24 11:45:33
1 |  | EV | -8.849 | -67.286 | -10.085 | -73.9904 | -1.019 | 51.414 | 2023/02/24 11:54:29
Sample 1 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -10.247 | -75.9998 |  | 20.0747 |  | 0.24081 | 2.6854 |  | -0.9844
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
2 |  | PZ | 2.157 | 8.037 | 1.2715 | 8.0077 | -1.084 | 47.956 | 2023/02/24 12:04:18
2 |  | PZ | 2.644 | 14.966 | 1.774 | 15.5508 | -1.014 | 49.524 | 2023/02/24 12:13:14
2 |  | PZ | 2.812 | 17.033 | 1.9473 | 17.801 | -1.06 | 49.904 | 2023/02/24 12:22:10
2 |  | PZ | 2.813 | 17.922 | 1.9484 | 18.7687 | -1.173 | 49.833 | 2023/02/24 12:31:06
2 |  | PZ | 2.909 | 18.799 | 2.0474 | 19.7235 | -1.091 | 49.505 | 2023/02/24 12:40:03
2 |  | PZ | 2.857 | 18.668 | 1.9938 | 19.5809 | -0.918 | 48.915 | 2023/02/24 12:49:00
2 |  | PZ | 2.916 | 18.973 | 2.0546 | 19.9129 | -1.16 | 48.893 | 2023/02/24 12:57:57
2 |  | PZ | 2.932 | 19.294 | 2.0712 | 20.2623 | -1.257 | 49.086 | 2023/02/24 13:06:54
2 |  | PZ | 2.946 | 19.536 | 2.0856 | 20.5258 | -1.104 | 48.77 | 2023/02/24 13:15:50
2 |  | PZ | 2.966 | 19.849 | 2.1062 | 20.8665 | -0.982 | 48.551 | 2023/02/24 13:24:46
Sample 2 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.2433 | 3.5606
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
3 |  | UT2 | -13.531 | -95.286 | -14.9161 | -104.4718 | -1.173 | 47.968 | 2023/02/24 13:34:34
3 |  | UT2 | -14.234 | -105.733 | -15.6415 | -115.8446 | -0.825 | 47.642 | 2023/02/24 13:43:30
3 |  | UT2 | -14.445 | -108.436 | -15.8592 | -118.7872 | -0.782 | 47.529 | 2023/02/24 13:52:27
3 |  | UT2 | -14.5 | -109.698 | -15.9159 | -120.161 | -0.95 | 48.072 | 2023/02/24 14:01:22
3 |  | UT2 | -14.536 | -110.396 | -15.9531 | -120.9209 | -0.947 | 47.744 | 2023/02/24 14:10:19
3 |  | UT2 | -14.597 | -110.983 | -16.016 | -121.5599 | -0.878 | 47.762 | 2023/02/24 14:19:15
3 |  | UT2 | -14.551 | -111.057 | -15.9686 | -121.6404 | -0.832 | 47.219 | 2023/02/24 14:28:11
3 |  | UT2 | -14.615 | -111.389 | -16.0346 | -122.0019 | -0.906 | 47.43 | 2023/02/24 14:37:07
3 |  | UT2 | -14.621 | -111.584 | -16.0408 | -122.2142 | -0.827 | 47.04 | 2023/02/24 14:46:04
3 |  | UT2 | -14.634 | -111.753 | -16.0542 | -122.3981 | -0.916 | 47.259 | 2023/02/24 14:55:01
Sample 3 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.33627 | 5.0362
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
4 |  | EV | -9.367 | -72.664 | -10.6195 | -79.845 | -0.895 | 47.074 | 2023/02/24 15:04:50
4 |  | EV | -9.2 | -69.402 | -10.4472 | -76.294 | -0.725 | 46.884 | 2023/02/24 15:13:47
4 |  | EV | -9.154 | -68.538 | -10.3997 | -75.3534 | -0.691 | 46.598 | 2023/02/24 15:22:42
4 |  | EV | -9.139 | -68.205 | -10.3842 | -74.9909 | -0.86 | 46.482 | 2023/02/24 15:31:39
4 |  | EV | -9.09 | -67.911 | -10.3337 | -74.6708 | -0.93 | 46.793 | 2023/02/24 15:40:35
4 |  | EV | -9.019 | -67.542 | -10.2604 | -74.2691 | -0.962 | 46.957 | 2023/02/24 15:49:31
4 |  | EV | -9.064 | -67.633 | -10.3068 | -74.3682 | -0.991 | 46.793 | 2023/02/24 15:58:28
4 |  | EV | -9.032 | -67.748 | -10.2738 | -74.4934 | -0.865 | 46.962 | 2023/02/24 16:07:24
4 |  | EV | -9.03 | -67.048 | -10.2718 | -73.7313 | -0.926 | 46.844 | 2023/02/24 16:16:21
4 |  | EV | -9.048 | -67.336 | -10.2903 | -74.0449 | -0.871 | 46.429 | 2023/02/24 16:25:16
Sample 4 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
4 |  | EV | -10.3587 | -75.2061 |  |  |  | 0.10743 | 1.6395
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
5 |  | TALL5-3040-3 | -3.936 | -35.132 | -5.0155 | -38.9869 | -1.029 | 49.994 | 2023/02/24 16:35:05
5 |  | TALL5-3040-3 | -3.704 | -31.895 | -4.7762 | -35.4631 | -1.099 | 49.843 | 2023/02/24 16:44:01
5 |  | TALL5-3040-3 | -3.658 | -31.224 | -4.7287 | -34.7326 | -0.991 | 49.864 | 2023/02/24 16:52:57
5 |  | TALL5-3040-3 | -3.581 | -30.85 | -4.6492 | -34.3255 | -1.177 | 50.566 | 2023/02/24 17:01:52
Sample 5 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
5 |  | TALL5-3040-3 | -4.7924 | -35.877 |  |  |  | 0.15284 | 1.953
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
6 |  | TALL5-010-2 | -2.132 | -24.608 | -3.1541 | -27.5303 | -1.262 | 71.545 | 2023/02/24 17:11:42
6 |  | TALL5-010-2 | -1.99 | -23.275 | -3.0076 | -26.0792 | -1.222 | 70.986 | 2023/02/24 17:20:39
6 |  | TALL5-010-2 | -2.008 | -23.296 | -3.0262 | -26.102 | -1.087 | 70.862 | 2023/02/24 17:29:35
6 |  | TALL5-010-2 | -1.998 | -23.214 | -3.0158 | -26.0128 | -1.131 | 71.154 | 2023/02/24 17:38:32
Sample 6 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
6 |  | TALL5-010-2 | -3.0509 | -26.4311 |  | 22.8067 |  | 0.067072 | 0.67406 |  | -1.3493
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
7 |  | TALL5-2030-2 | -2.404 | -24.281 | -3.4348 | -27.1743 | -1.055 | 60.881 | 2023/02/24 17:48:21
7 |  | TALL5-2030-2 | -2.485 | -24.484 | -3.5183 | -27.3953 | -1.09 | 60.541 | 2023/02/24 17:57:18
7 |  | TALL5-2030-2 | -2.458 | -24.56 | -3.4905 | -27.478 | -1.288 | 60.174 | 2023/02/24 18:06:14
7 |  | TALL5-2030-2 | -2.443 | -24.096 | -3.475 | -26.9729 | -1.09 | 60.618 | 2023/02/24 18:15:11
Sample 7 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
7 |  | TALL5-2030-2 | -3.4796 | -27.2551 |  | 12.2234 |  | 0.033808 | 0.20914
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
8 |  | TALL5-2030-3 | -3.089 | -28.14 | -4.1416 | -31.3753 | -1.204 | 70.025 | 2023/02/24 18:24:57
8 |  | TALL5-2030-3 | -3.12 | -28.398 | -4.1736 | -31.6562 | -1.318 | 69.281 | 2023/02/24 18:33:53
8 |  | TALL5-2030-3 | -3.122 | -28.15 | -4.1756 | -31.3862 | -1.116 | 69.194 | 2023/02/24 18:42:48
8 |  | TALL5-2030-3 | -3.121 | -28.15 | -4.1746 | -31.3862 | -1.114 | 68.976 | 2023/02/24 18:51:44
Sample 8 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
8 |  | TALL5-2030-3 | -4.1663 | -31.451 |  | 21.0389 |  | 0.016021 | 0.12576 |  | -1.3038
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
9 |  | TALL5-3040-1 | -3.719 | -26.917 | -4.7916 | -30.0439 | -0.949 | 47.076 | 2023/02/24 19:01:33
9 |  | TALL5-3040-1 | -3.712 | -26.457 | -4.7844 | -29.5432 | -0.883 | 46.25 | 2023/02/24 19:10:29
9 |  | TALL5-3040-1 | -3.747 | -26.371 | -4.8205 | -29.4495 | -0.922 | 46.446 | 2023/02/24 19:19:26
9 |  | TALL5-3040-1 | -3.708 | -26.224 | -4.7803 | -29.2895 | -1.096 | 46.521 | 2023/02/24 19:28:22
Sample 9 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
9 |  | TALL5-3040-1 | -4.7942 | -29.5815 |  |  |  | 0.017597 | 0.29906
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
10 |  | TALL5-1020-1 | -2.803 | -27.658 | -3.8465 | -30.8506 | -1.09 | 56.374 | 2023/02/24 19:38:09
10 |  | TALL5-1020-1 | -2.779 | -28.082 | -3.8217 | -31.3122 | -1.12 | 56.45 | 2023/02/24 19:47:06
10 |  | TALL5-1020-1 | -2.784 | -28.019 | -3.8269 | -31.2436 | -1.229 | 56.05 | 2023/02/24 19:56:02
10 |  | TALL5-1020-1 | -2.841 | -28.271 | -3.8857 | -31.5179 | -1.248 | 56.374 | 2023/02/24 20:04:58
Sample 10 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
10 |  | TALL5-1020-1 | -3.8452 | -31.2311 |  | 7.9819 |  | 0.028135 | 0.25643
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
11 |  | TALL5-2030-1 | -2.56 | -27.684 | -3.5957 | -30.8789 | -1.018 | 63.916 | 2023/02/24 20:14:46
11 |  | TALL5-2030-1 | -2.585 | -27.875 | -3.6215 | -31.0868 | -1.186 | 63.03 | 2023/02/24 20:23:42
11 |  | TALL5-2030-1 | -2.635 | -28.132 | -3.6731 | -31.3666 | -1.014 | 62.983 | 2023/02/24 20:32:37
11 |  | TALL5-2030-1 | -2.552 | -27.821 | -3.5875 | -31.028 | -1.24 | 63.108 | 2023/02/24 20:41:34
Sample 11 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
11 |  | TALL5-2030-1 | -3.6195 | -31.0901 |  | 14.9291 |  | 0.037408 | 0.18745
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
12 |  | TALL5-1020-2 | -3.102 | -26.273 | -4.155 | -29.3428 | -1.026 | 47.26 | 2023/02/24 20:51:21
12 |  | TALL5-1020-2 | -3.131 | -26.293 | -4.1849 | -29.3646 | -0.971 | 46.901 | 2023/02/24 21:00:17
12 |  | TALL5-1020-2 | -3.13 | -25.82 | -4.1839 | -28.8497 | -1.023 | 46.721 | 2023/02/24 21:09:13
12 |  | TALL5-1020-2 | -3.173 | -26.199 | -4.2282 | -29.2623 | -1.097 | 47.151 | 2023/02/24 21:18:09
Sample 12 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
12 |  | TALL5-1020-2 | -4.188 | -29.2049 |  |  |  | 0.029269 | 0.22123
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
13 |  | EV | -8.643 | -61.346 | -9.8724 | -67.524 | -1.112 | 47.199 | 2023/02/24 21:27:58
13 |  | EV | -8.905 | -64.711 | -10.1428 | -71.1872 | -1.029 | 46.435 | 2023/02/24 21:36:55
13 |  | EV | -8.886 | -65.395 | -10.1232 | -71.9319 | -0.913 | 46.26 | 2023/02/24 21:45:51
13 |  | EV | -8.976 | -65.921 | -10.216 | -72.5045 | -1.033 | 46.274 | 2023/02/24 21:54:47
Sample 13 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
13 |  | EV | -10.0886 | -70.7869 |  |  |  | 0.14494 | 2.0587
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
14 |  | TALL5-3040-2 | -2.565 | -26.886 | -3.6009 | -30.0102 | -0.951 | 48.018 | 2023/02/24 22:04:37
14 |  | TALL5-3040-2 | -2.228 | -23.126 | -3.2532 | -25.917 | -1.012 | 47.238 | 2023/02/24 22:13:33
14 |  | TALL5-3040-2 | -2.195 | -22.384 | -3.2191 | -25.1092 | -0.906 | 47.134 | 2023/02/24 22:22:30
14 |  | TALL5-3040-2 | -2.192 | -22.012 | -3.216 | -24.7042 | -1.111 | 47.639 | 2023/02/24 22:31:26
Sample 14 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
14 |  | TALL5-3040-2 | -3.3223 | -26.4351 |  |  |  | 0.18074 | 2.2378
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
15 |  | TALL5-deep-3 | -3.979 | -31.749 | -5.0599 | -35.3041 | -1.065 | 57.06 | 2023/02/24 22:41:12
15 |  | TALL5-deep-3 | -4.016 | -32.496 | -5.0981 | -36.1173 | -0.963 | 56.477 | 2023/02/24 22:50:09
15 |  | TALL5-deep-3 | -4.086 | -32.867 | -5.1703 | -36.5212 | -1.262 | 56.536 | 2023/02/24 22:59:05
15 |  | TALL5-deep-3 | -4.103 | -32.952 | -5.1879 | -36.6137 | -1.242 | 56.986 | 2023/02/24 23:08:01
Sample 15 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
15 |  | TALL5-deep-3 | -5.129 | -36.1391 |  | 8.4346 |  | 0.058418 | 0.54833
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
16 |  | TALL5-B-4 | 0.236 | -21.258 | -0.71069 | -23.8834 | -3.017 | 345.064 | 2023/02/24 23:17:48
16 |  | TALL5-B-4 | 0.3 | -20.949 | -0.64466 | -23.547 | -2.946 | 333.42 | 2023/02/24 23:26:43
16 |  | TALL5-B-4 | 0.249 | -20.989 | -0.69728 | -23.5906 | -3.112 | 338.415 | 2023/02/24 23:35:38
16 |  | TALL5-B-4 | 0.307 | -20.635 | -0.63743 | -23.2052 | -3.032 | 337.886 | 2023/02/24 23:44:34
Sample 16 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
16 |  | TALL5-B-4 | -0.67252 | -23.5566 | -3.0267 | 290.3662 | 18.4116 | 0.03573 | 0.25519 |  | -3.9523
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
17 |  | TALL5-B-5 | -0.932 | -20.935 | -1.9159 | -23.5318 | -2.009 | 243.492 | 2023/02/24 23:54:21
17 |  | TALL5-B-5 | -0.873 | -20.477 | -1.855 | -23.0332 | -1.941 | 235.504 | 2023/02/25 00:03:17
17 |  | TALL5-B-5 | -0.885 | -20.37 | -1.8674 | -22.9167 | -1.966 | 236.679 | 2023/02/25 00:12:13
17 |  | TALL5-B-5 | -0.911 | -20.374 | -1.8942 | -22.9211 | -1.767 | 235.683 | 2023/02/25 00:21:09
Sample 17 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
17 |  | TALL5-B-5 | -1.8831 | -23.1007 | -1.9207 | 189.5094 | 7.7593 | 0.02645 | 0.26861 |  | -2.6268
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
18 |  | TALL5-C-2 | -2.015 | -24.766 | -3.0334 | -27.7023 | -1.702 | 179.191 | 2023/02/25 00:30:56
18 |  | TALL5-C-2 | -2.097 | -25.228 | -3.118 | -28.2052 | -1.513 | 173.146 | 2023/02/25 00:39:52
18 |  | TALL5-C-2 | -2.005 | -25.018 | -3.0231 | -27.9766 | -1.834 | 172.933 | 2023/02/25 00:48:48
18 |  | TALL5-C-2 | -2.047 | -24.881 | -3.0664 | -27.8275 | -1.466 | 172.279 | 2023/02/25 00:57:45
Sample 18 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
18 |  | TALL5-C-2 | -3.0602 | -27.9279 | -1.6287 | 126.0571 |  | 0.041409 | 0.19863 |  | -3.182
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
19 |  | TALL5-B-3 | 2.784 | -18.509 | 1.9184 | -20.8908 | -4.12 | 216.487 | 2023/02/25 01:07:33
19 |  | TALL5-B-3 | 2.914 | -17.941 | 2.0526 | -20.2725 | -4.089 | 210.478 | 2023/02/25 01:16:29
19 |  | TALL5-B-3 | 2.893 | -18.212 | 2.0309 | -20.5675 | -3.964 | 211.101 | 2023/02/25 01:25:25
19 |  | TALL5-B-3 | 2.941 | -17.729 | 2.0804 | -20.0417 | -4.129 | 211.107 | 2023/02/25 01:34:21
Sample 19 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
19 |  | TALL5-B-3 | 2.0206 | -20.4431 | -4.0755 | 163.9631 | 33.0268 | 0.068862 | 0.338 |  | -0.33705
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
20 |  | TALL5-A-3 | 0.609 | -24.748 | -0.32582 | -27.6827 | -2.74 | 55.8 | 2023/02/25 01:44:09
20 |  | TALL5-A-3 | 0.544 | -25.528 | -0.39289 | -28.5318 | -2.3 | 53.468 | 2023/02/25 01:53:06
20 |  | TALL5-A-3 | 0.592 | -25.359 | -0.34336 | -28.3478 | -2.69 | 52.823 | 2023/02/25 02:02:02
20 |  | TALL5-A-3 | 0.59 | -25.314 | -0.34542 | -28.2989 | -2.638 | 53.631 | 2023/02/25 02:10:57
Sample 20 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
20 |  | TALL5-A-3 | -0.35187 | -28.2153 | -2.592 | 5.6004 | 15.6041 | 0.027837 | 0.33893 |  | 1.3142
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
21 |  | TALL5-B-1 | -1.705 | -22.755 | -2.7135 | -25.5131 | -1.511 | 292.732 | 2023/02/25 02:20:46
21 |  | TALL5-B-1 | -1.655 | -22.317 | -2.6619 | -25.0363 | -1.614 | 287.143 | 2023/02/25 02:29:42
21 |  | TALL5-B-1 | -1.673 | -21.985 | -2.6805 | -24.6748 | -1.66 | 286.463 | 2023/02/25 02:38:38
21 |  | TALL5-B-1 | -1.66 | -22.066 | -2.6671 | -24.763 | -1.556 | 287.633 | 2023/02/25 02:47:34
Sample 21 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
21 |  | TALL5-B-1 | -2.6807 | -24.9968 | -1.5853 | 240.1626 |  | 0.022485 | 0.34632 |  | -5.2278
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
22 |  | EV | -8.562 | -61.056 | -9.7889 | -67.2083 | -1.086 | 50.546 | 2023/02/25 02:57:19
22 |  | EV | -8.842 | -64.329 | -10.0778 | -70.7714 | -1.147 | 48.064 | 2023/02/25 03:06:15
22 |  | EV | -8.877 | -65.337 | -10.1139 | -71.8687 | -1.14 | 47.628 | 2023/02/25 03:15:11
22 |  | EV | -8.93 | -65.786 | -10.1686 | -72.3575 | -1.005 | 47.022 | 2023/02/25 03:24:06
Sample 22 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
22 |  | EV | -10.0373 | -70.5515 |  |  |  | 0.16453 | 2.1361
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
23 |  | TALL5-A-4 | -1.316 | -31.836 | -2.3121 | -35.3988 | -1.318 | 48.989 | 2023/02/25 03:33:53
23 |  | TALL5-A-4 | -1.061 | -29.189 | -2.049 | -32.5173 | -1.309 | 48.414 | 2023/02/25 03:42:50
23 |  | TALL5-A-4 | -0.93 | -28.152 | -1.9138 | -31.3884 | -1.285 | 48.389 | 2023/02/25 03:51:46
23 |  | TALL5-A-4 | -0.94 | -28.03 | -1.9241 | -31.2556 | -1.331 | 48.2 | 2023/02/25 04:00:41
Sample 23 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
23 |  | TALL5-A-4 | -2.0498 | -32.64 |  |  |  | 0.17965 | 1.7677
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
24 |  | TALL5-A-1 | 1.821 | -21.838 | 0.92478 | -24.5148 | -3.173 | 55.756 | 2023/02/25 04:10:28
24 |  | TALL5-A-1 | 1.843 | -21.43 | 0.94748 | -24.0707 | -2.845 | 54.865 | 2023/02/25 04:19:23
24 |  | TALL5-A-1 | 1.847 | -21.44 | 0.95161 | -24.0815 | -2.798 | 54.838 | 2023/02/25 04:28:19
24 |  | TALL5-A-1 | 1.793 | -21.754 | 0.89589 | -24.4234 | -2.924 | 55.116 | 2023/02/25 04:37:14
Sample 24 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
24 |  | TALL5-A-1 | 0.92994 | -24.2726 | -2.935 | 6.8136 | 19.9829 | 0.024792 | 0.21127 |  | 2.0767
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
25 |  | TALL5-010-1 | -2.52 | -23.266 | -3.5545 | -26.0694 | -1.373 | 51.335 | 2023/02/25 04:47:01
25 |  | TALL5-010-1 | -2.519 | -22.682 | -3.5534 | -25.4336 | -1.235 | 51.228 | 2023/02/25 04:55:58
25 |  | TALL5-010-1 | -2.587 | -22.893 | -3.6236 | -25.6633 | -1.189 | 51.168 | 2023/02/25 05:04:53
25 |  | TALL5-010-1 | -2.567 | -22.922 | -3.603 | -25.6949 | -1.138 | 51.097 | 2023/02/25 05:13:49
Sample 25 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
25 |  | TALL5-010-1 | -3.5836 | -25.7153 |  | 2.8769 |  | 0.034189 | 0.24178
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
26 |  | TALL5-C-4 | -2.683 | -25.57 | -3.7226 | -28.5775 | -1.443 | 324.478 | 2023/02/25 05:23:36
26 |  | TALL5-C-4 | -2.758 | -25.987 | -3.8 | -29.0315 | -1.446 | 316.745 | 2023/02/25 05:32:33
26 |  | TALL5-C-4 | -2.742 | -26.411 | -3.7835 | -29.4931 | -1.554 | 316.088 | 2023/02/25 05:41:28
26 |  | TALL5-C-4 | -2.737 | -26.143 | -3.7784 | -29.2013 | -1.606 | 318.983 | 2023/02/25 05:50:23
Sample 26 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
26 |  | TALL5-C-4 | -3.7711 | -29.0759 | -1.5123 | 270.7434 |  | 0.032588 | 0.35183 |  | -6.102
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
27 |  | TALL5-C-3 | -1.536 | -25.278 | -2.5391 | -28.2597 | -1.795 | 215.408 | 2023/02/25 06:00:10
27 |  | TALL5-C-3 | -1.485 | -25.197 | -2.4865 | -28.1715 | -1.704 | 208.802 | 2023/02/25 06:09:06
27 |  | TALL5-C-3 | -1.435 | -25.138 | -2.4349 | -28.1073 | -1.655 | 208.13 | 2023/02/25 06:18:02
27 |  | TALL5-C-3 | -1.469 | -25.409 | -2.47 | -28.4023 | -1.632 | 208.152 | 2023/02/25 06:26:58
Sample 27 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
27 |  | TALL5-C-3 | -2.4826 | -28.2352 | -1.6965 | 161.7929 |  | 0.042035 | 0.11733 |  | -3.7415
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
28 |  | TALL5-A-5 | -0.945 | -26.595 | -1.9293 | -29.6934 | -1.356 | 54.372 | 2023/02/25 06:36:46
28 |  | TALL5-A-5 | -0.913 | -26.696 | -1.8963 | -29.8033 | -1.363 | 52.521 | 2023/02/25 06:45:41
28 |  | TALL5-A-5 | -0.912 | -26.823 | -1.8953 | -29.9416 | -1.445 | 52.15 | 2023/02/25 06:54:36
28 |  | TALL5-A-5 | -0.959 | -27.3 | -1.9437 | -30.4609 | -1.31 | 51.802 | 2023/02/25 07:03:31
Sample 28 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
28 |  | TALL5-A-5 | -1.9161 | -29.9748 |  | 4.3811 |  | 0.023514 | 0.31194
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
29 |  | TALL5-A-2 | -0.86 | -28.194 | -1.8416 | -31.4341 | -1.27 | 52.379 | 2023/02/25 07:13:15
29 |  | TALL5-A-2 | -0.797 | -27.731 | -1.7766 | -30.9301 | -1.411 | 51.63 | 2023/02/25 07:22:11
29 |  | TALL5-A-2 | -0.813 | -27.899 | -1.7931 | -31.1129 | -1.241 | 51.368 | 2023/02/25 07:31:07
29 |  | TALL5-A-2 | -0.845 | -28.119 | -1.8261 | -31.3524 | -1.377 | 51.544 | 2023/02/25 07:40:03
Sample 29 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
29 |  | TALL5-A-2 | -1.8094 | -31.2074 |  | 3.4001 |  | 0.028849 | 0.21099
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
30 |  | TALL5-D-1 | -1.309 | -24.89 | -2.3049 | -27.8373 | -1.666 | 51.932 | 2023/02/25 07:49:48
30 |  | TALL5-D-1 | -1.391 | -25.035 | -2.3895 | -27.9951 | -1.626 | 51.154 | 2023/02/25 07:58:43
30 |  | TALL5-D-1 | -1.401 | -24.98 | -2.3998 | -27.9353 | -1.522 | 51.627 | 2023/02/25 08:07:38
30 |  | TALL5-D-1 | -1.451 | -25.136 | -2.4514 | -28.1051 | -1.732 | 51.46 | 2023/02/25 08:16:35
Sample 30 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
30 |  | TALL5-D-1 | -2.3864 | -27.9682 | -1.6365 | 3.2132 |  | 0.058844 | 0.10296
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
31 |  | EV | -8.522 | -60.941 | -9.7476 | -67.0831 | -1.17 | 47.446 | 2023/02/25 08:26:20
31 |  | EV | -8.803 | -64.283 | -10.0375 | -70.7213 | -1.167 | 46.696 | 2023/02/25 08:35:16
31 |  | EV | -8.907 | -65.325 | -10.1448 | -71.8557 | -1.09 | 46.523 | 2023/02/25 08:44:11
31 |  | EV | -8.934 | -65.764 | -10.1727 | -72.3336 | -0.892 | 46.402 | 2023/02/25 08:53:08
Sample 31 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
31 |  | EV | -10.0257 | -70.4984 |  |  |  | 0.18833 | 2.1818
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
32 |  | TALL5-C-5 | -2.996 | -32.785 | -4.0456 | -36.4319 | -1.713 | 310.038 | 2023/02/25 09:02:53
32 |  | TALL5-C-5 | -2.69 | -29.046 | -3.7299 | -32.3616 | -1.335 | 305.673 | 2023/02/25 09:11:49
32 |  | TALL5-C-5 | -2.664 | -28.353 | -3.703 | -31.6072 | -1.43 | 305.965 | 2023/02/25 09:20:44
32 |  | TALL5-C-5 | -2.654 | -28.149 | -3.6927 | -31.3851 | -1.306 | 304.45 | 2023/02/25 09:29:39
Sample 32 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
32 |  | TALL5-C-5 | -3.7928 | -32.9465 | -1.446 | 258.2014 |  | 0.16404 | 2.1687 |  | -5.3073
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
33 |  | TALL5-D-4 | -1.063 | -25.141 | -2.0511 | -28.1105 | -1.632 | 145.549 | 2023/02/25 09:39:25
33 |  | TALL5-D-4 | -0.957 | -24.344 | -1.9417 | -27.2429 | -1.601 | 140.235 | 2023/02/25 09:48:21
33 |  | TALL5-D-4 | -0.958 | -24.352 | -1.9427 | -27.2516 | -1.497 | 143.022 | 2023/02/25 09:57:16
33 |  | TALL5-D-4 | -0.915 | -23.802 | -1.8983 | -26.6529 | -1.711 | 142.86 | 2023/02/25 10:06:12
Sample 33 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
33 |  | TALL5-D-4 | -1.9585 | -27.3145 | -1.6102 | 94.5864 |  | 0.0631 | 0.55128 |  | -3.273
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
34 |  | TALL5-D-2 | -2.037 | -27.704 | -3.0561 | -30.9007 | -1.11 | 238.825 | 2023/02/25 10:15:58
34 |  | TALL5-D-2 | -2.022 | -27.453 | -3.0406 | -30.6274 | -1.364 | 266.803 | 2023/02/25 10:24:54
34 |  | TALL5-D-2 | -2.016 | -27.705 | -3.0344 | -30.9017 | -1.338 | 291.136 | 2023/02/25 10:33:50
34 |  | TALL5-D-2 | -2.018 | -27.758 | -3.0365 | -30.9594 | -1.33 | 300.461 | 2023/02/25 10:42:46
Sample 34 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
34 |  | TALL5-D-2 | -3.0419 | -30.8473 |  | 225.9761 |  | 0.0095 | 0.13701 |  | -17.395
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
35 |  | TALL5-1020-3 | -2.594 | -24.961 | -3.6308 | -27.9146 | -1.166 | 64.987 | 2023/02/25 10:52:33
35 |  | TALL5-1020-3 | -2.608 | -24.751 | -3.6453 | -27.686 | -1.233 | 63.926 | 2023/02/25 11:01:29
35 |  | TALL5-1020-3 | -2.626 | -24.813 | -3.6638 | -27.7535 | -1.203 | 63.839 | 2023/02/25 11:10:25
35 |  | TALL5-1020-3 | -2.58 | -24.328 | -3.6164 | -27.2255 | -1.093 | 63.465 | 2023/02/25 11:19:21
Sample 35 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
35 |  | TALL5-1020-3 | -3.6391 | -27.6449 |  | 15.7241 |  | 0.019664 | 0.27152
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
36 |  | TALL5-010-3 | -2.574 | -23.717 | -3.6102 | -26.5603 | -1.144 | 60.483 | 2023/02/25 11:29:04
36 |  | TALL5-010-3 | -2.544 | -23.984 | -3.5792 | -26.851 | -1.152 | 60.712 | 2023/02/25 11:37:58
36 |  | TALL5-010-3 | -2.566 | -23.83 | -3.6019 | -26.6833 | -1.331 | 60.947 | 2023/02/25 11:46:55
36 |  | TALL5-010-3 | -2.599 | -23.636 | -3.636 | -26.4722 | -1.164 | 60.949 | 2023/02/25 11:55:50
Sample 36 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
36 |  | TALL5-010-3 | -3.6068 | -26.6417 |  | 12.4427 |  | 0.022706 | 0.15085
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
37 |  | TALL5-D-5 | -1.849 | -25.372 | -2.8621 | -28.362 | -0.997 | 55.868 | 2023/02/25 12:05:33
37 |  | TALL5-D-5 | -1.784 | -25.751 | -2.795 | -28.7746 | -1.075 | 55.333 | 2023/02/25 12:14:29
37 |  | TALL5-D-5 | -1.811 | -25.621 | -2.8229 | -28.6331 | -0.96 | 55.074 | 2023/02/25 12:23:25
37 |  | TALL5-D-5 | -1.783 | -25.395 | -2.794 | -28.387 | -0.953 | 55.106 | 2023/02/25 12:32:21
Sample 37 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
37 |  | TALL5-D-5 | -2.8185 | -28.5392 |  | 7.0151 |  | 0.031009 | 0.18278
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
38 |  | TALL5-B-2 | -1.024 | -24.57 | -2.0108 | -27.4889 | -1.964 | 265.782 | 2023/02/25 12:42:04
38 |  | TALL5-B-2 | -1.086 | -24.689 | -2.0748 | -27.6185 | -2.17 | 264.16 | 2023/02/25 12:51:00
38 |  | TALL5-B-2 | -1.007 | -24.311 | -1.9933 | -27.207 | -2.178 | 264.768 | 2023/02/25 12:59:55
38 |  | TALL5-B-2 | -1.07 | -24.663 | -2.0583 | -27.5902 | -1.979 | 263.424 | 2023/02/25 13:08:50
Sample 38 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
38 |  | TALL5-B-2 | -2.0343 | -27.4761 | -2.0728 | 216.2034 | 7.5629 | 0.037322 | 0.17257 |  | -3.5573
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
39 |  | TALL5-C-1 | -2.032 | -25.313 | -3.0509 | -28.2978 | -1.714 | 157.127 | 2023/02/25 13:18:35
39 |  | TALL5-C-1 | -2.091 | -25.401 | -3.1118 | -28.3936 | -1.674 | 154.189 | 2023/02/25 13:27:31
39 |  | TALL5-C-1 | -2.068 | -25.278 | -3.0881 | -28.2597 | -1.797 | 154.451 | 2023/02/25 13:36:27
39 |  | TALL5-C-1 | -2.108 | -25.699 | -3.1293 | -28.718 | -1.828 | 152.808 | 2023/02/25 13:45:22
Sample 39 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
39 |  | TALL5-C-1 | -3.095 | -28.4172 | -1.7532 | 106.3137 |  | 0.032877 | 0.1913 |  | -1.3558
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
40 |  | EV | -8.581 | -61.022 | -9.8085 | -67.1713 | -1.097 | 49.11 | 2023/02/25 13:55:06
40 |  | EV | -8.915 | -64.822 | -10.1531 | -71.3081 | -0.953 | 47.13 | 2023/02/25 14:04:01
40 |  | EV | -8.95 | -65.485 | -10.1892 | -72.0298 | -0.9 | 47.164 | 2023/02/25 14:12:57
40 |  | EV | -8.896 | -65.639 | -10.1335 | -72.1975 | -0.915 | 47.007 | 2023/02/25 14:21:52
Sample 40 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
40 |  | EV | -10.0711 | -70.6767 |  |  |  | 0.17113 | 2.1757
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
41 |  | TALL5-D-3 | -2.281 | -29.773 | -3.3078 | -33.153 | -1.503 | 333.933 | 2023/02/25 14:31:36
41 |  | TALL5-D-3 | -2.082 | -27.029 | -3.1025 | -30.1658 | -1.566 | 330.394 | 2023/02/25 14:40:31
41 |  | TALL5-D-3 | -1.953 | -25.633 | -2.9694 | -28.6461 | -1.752 | 328.157 | 2023/02/25 14:49:26
41 |  | TALL5-D-3 | -2.116 | -26.606 | -3.1376 | -29.7054 | -1.582 | 332.685 | 2023/02/25 14:58:21
Sample 41 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
41 |  | TALL5-D-3 | -3.1293 | -30.4176 | -1.6007 | 282.9621 |  | 0.13502 | 1.7742 |  | -4.8568
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
42 |  | EV | -8.607 | -61.561 | -9.8353 | -67.7581 | -0.993 | 50.977 | 2023/02/25 15:08:06
42 |  | EV | -8.825 | -64.595 | -10.0602 | -71.061 | -0.937 | 48.161 | 2023/02/25 15:17:02
42 |  | EV | -8.926 | -65.44 | -10.1644 | -71.9808 | -0.891 | 47.573 | 2023/02/25 15:25:57
42 |  | EV | -9.011 | -66.01 | -10.2522 | -72.6014 | -0.78 | 46.877 | 2023/02/25 15:34:53
Sample 42 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
42 |  | EV | -10.078 | -70.8503 |  |  |  | 0.17429 | 1.9809

## Source
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Line | Analysis | Time Code | Port | Inj Nr | d(18_16)Mean | d(D_H)Mean | H2O_Mean | Ignore | Good | Identifier 1 | Identifier 2 | Gas Configuration | Timestamp Mean | d(18_16)_SD | d(D_H)_SD | H2O_SD | d(18_16)_Sl | d(D_H)_Sl | H2O_Sl | baseline_shift | slope_shift | residuals | baseline_curvature | interval | ch4_ppm | h16od_adjust | h16od_shift | n2_flag | DAS Temp | Tray | Sample | Job | Method | Error Code
1 | P-23700 | 2023/02/24 10:33:59 | 1-01 | 1 | -9.604 | -75.854 | 14023 | -1 | 1 | EV | standard | H2O | 1677260344.39 | 0.205 | 0.665 | 73.008 | 0.001 | -0.001 | 1.784 | 85.273 | 31.819 | 2.368 | 67.539 | 0.857 | -1.267 | 0 | 0 | 1 | 42.312 | 1 | 1 | 1 | GJB_Rinse | 0
2 | P-23700 | 2023/02/24 10:42:56 | 1-01 | 2 | -9.214 | -71.414 | 12962 | -1 | 1 | EV | standard | H2O | 1677260881.25 | 0.198 | 0.851 | 70.983 | 0.001 | -0.005 | 1.743 | 127.087 | 27.747 | 2.208 | 62.203 | 0.856 | -1.116 | 0 | 0 | 1 | 41.562 | 1 | 1 | 1 | GJB_Rinse | 0
3 | P-23700 | 2023/02/24 10:51:52 | 1-01 | 3 | -9.081 | -69.657 | 12678 | -1 | 1 | EV | standard | H2O | 1677261417.86 | 0.218 | 0.743 | 69.433 | 0.001 | -0.001 | 1.718 | 103.398 | 27.691 | 2.165 | 61.666 | 0.856 | -1.035 | 0 | 0 | 1 | 41.25 | 1 | 1 | 1 | GJB_Rinse | 0
4 | P-23700 | 2023/02/24 11:00:49 | 1-01 | 4 | -8.931 | -68.615 | 12204 | 0 | 1 | EV | standard | H2O | 1677261954.45 | 0.179 | 0.82 | 66.735 | 0.001 | -0.004 | 1.651 | 55.828 | 27.761 | 2.088 | 60.598 | 0.855 | -0.923 | 0 | 0 | 1 | 41.188 | 1 | 1 | 1 | GJB_Rinse | 0
5 | P-23700 | 2023/02/24 11:09:46 | 1-01 | 5 | -8.891 | -68.058 | 12321 | 0 | 1 | EV | standard | H2O | 1677262491.53 | 0.223 | 0.818 | 64.852 | 0.001 | 0.001 | 1.593 | 53.453 | 28.146 | 2.103 | 60.773 | 0.855 | -1.093 | 0 | 0 | 1 | 41.188 | 1 | 1 | 1 | GJB_Rinse | 0
6 | P-23700 | 2023/02/24 11:18:42 | 1-01 | 6 | -8.877 | -67.881 | 12392 | 0 | 1 | EV | standard | H2O | 1677263028.5 | 0.191 | 0.757 | 63.067 | 0 | 0 | 1.558 | 52.949 | 28.427 | 2.112 | 60.933 | 0.856 | -1.111 | 0 | 0 | 1 | 41.25 | 1 | 1 | 1 | GJB_Rinse | 0
7 | P-23700 | 2023/02/24 11:27:39 | 1-01 | 7 | -8.858 | -67.55 | 12184 | 0 | 1 | EV | standard | H2O | 1677263564.35 | 0.201 | 0.812 | 60.61 | 0.001 | 0.001 | 1.497 | 51.784 | 27.656 | 2.089 | 60.738 | 0.855 | -1.04 | 0 | 0 | 1 | 41.312 | 1 | 1 | 1 | GJB_Rinse | 0
8 | P-23700 | 2023/02/24 11:36:36 | 1-01 | 8 | -8.891 | -67.602 | 12309 | 0 | 1 | EV | standard | H2O | 1677264101.29 | 0.21 | 0.867 | 61.364 | 0 | 0 | 1.518 | 51.415 | 28.131 | 2.11 | 62.076 | 0.856 | -0.854 | 0 | 0 | 1 | 41.312 | 1 | 1 | 1 | GJB_Rinse | 0
9 | P-23700 | 2023/02/24 11:45:33 | 1-01 | 9 | -8.864 | -67.401 | 12359 | 0 | 1 | EV | standard | H2O | 1677264638.66 | 0.205 | 0.797 | 64.1 | 0.001 | -0.001 | 1.586 | 51.447 | 28.326 | 2.1 | 60.956 | 0.856 | -1.086 | 0 | 0 | 1 | 41.312 | 1 | 1 | 1 | GJB_Rinse | 0
10 | P-23700 | 2023/02/24 11:54:29 | 1-01 | 10 | -8.849 | -67.286 | 12424 | 0 | 1 | EV | standard | H2O | 1677265175.43 | 0.188 | 0.786 | 60.63 | 0.001 | -0.001 | 1.498 | 51.414 | 28.801 | 2.123 | 61.242 | 0.856 | -1.019 | 0 | 0 | 1 | 41.375 | 1 | 1 | 1 | GJB_Rinse | 0
11 | P-23701 | 2023/02/24 12:04:18 | 1-02 | 1 | 2.157 | 8.037 | 12705 | -1 | 1 | PZ | standard | H2O | 1677265763.44 | 0.176 | 0.851 | 62.025 | 0 | -0.004 | 1.535 | 47.956 | 29.464 | 2.274 | 64.517 | 0.856 | -1.084 | 0 | 0 | 1 | 41.375 | 1 | 2 | 1 | GJB_Rinse | 0
12 | P-23701 | 2023/02/24 12:13:14 | 1-02 | 2 | 2.644 | 14.966 | 12355 | -1 | 1 | PZ | standard | H2O | 1677266299.81 | 0.196 | 0.884 | 62.698 | 0 | -0.007 | 1.546 | 49.524 | 28.407 | 2.244 | 64.052 | 0.855 | -1.014 | 0 | 0 | 1 | 41.375 | 1 | 2 | 1 | GJB_Rinse | 0
13 | P-23701 | 2023/02/24 12:22:10 | 1-02 | 3 | 2.812 | 17.033 | 12471 | -1 | 1 | PZ | standard | H2O | 1677266835.34 | 0.209 | 0.886 | 62.708 | 0 | -0.008 | 1.55 | 49.904 | 28.442 | 2.274 | 65.537 | 0.856 | -1.06 | 0 | 0 | 1 | 41.438 | 1 | 2 | 1 | GJB_Rinse | 0
14 | P-23701 | 2023/02/24 12:31:06 | 1-02 | 4 | 2.813 | 17.922 | 12435 | 0 | 1 | PZ | standard | H2O | 1677267372.17 | 0.208 | 0.952 | 61.123 | 0 | -0.006 | 1.511 | 49.833 | 28.707 | 2.261 | 63.601 | 0.856 | -1.173 | 0 | 0 | 1 | 41.438 | 1 | 2 | 1 | GJB_Rinse | 0
15 | P-23701 | 2023/02/24 12:40:03 | 1-02 | 5 | 2.909 | 18.799 | 12454 | 0 | 1 | PZ | standard | H2O | 1677267908.42 | 0.196 | 0.858 | 63.789 | 0 | -0.007 | 1.579 | 49.505 | 28.683 | 2.261 | 64.274 | 0.856 | -1.091 | 0 | 0 | 1 | 41.438 | 1 | 2 | 1 | GJB_Rinse | 0
16 | P-23701 | 2023/02/24 12:49:00 | 1-02 | 6 | 2.857 | 18.668 | 12303 | 0 | 1 | PZ | standard | H2O | 1677268445.41 | 0.204 | 0.974 | 62.871 | 0.001 | -0.007 | 1.552 | 48.915 | 27.986 | 2.231 | 64.034 | 0.856 | -0.918 | 0 | 0 | 1 | 41.438 | 1 | 2 | 1 | GJB_Rinse | 0
17 | P-23701 | 2023/02/24 12:57:57 | 1-02 | 7 | 2.916 | 18.973 | 12431 | 0 | 1 | PZ | standard | H2O | 1677268982.62 | 0.205 | 0.816 | 60.697 | 0.001 | -0.002 | 1.508 | 48.893 | 28.938 | 2.257 | 64.273 | 0.856 | -1.16 | 0 | 0 | 1 | 41.5 | 1 | 2 | 1 | GJB_Rinse | 0
18 | P-23701 | 2023/02/24 13:06:54 | 1-02 | 8 | 2.932 | 19.294 | 12516 | 0 | 1 | PZ | standard | H2O | 1677269519.34 | 0.2 | 0.898 | 63.805 | 0 | -0.004 | 1.573 | 49.086 | 28.879 | 2.258 | 64.044 | 0.856 | -1.257 | 0 | 0 | 1 | 41.5 | 1 | 2 | 1 | GJB_Rinse | 0
19 | P-23701 | 2023/02/24 13:15:50 | 1-02 | 9 | 2.946 | 19.536 | 12520 | 0 | 1 | PZ | standard | H2O | 1677270055.81 | 0.19 | 0.788 | 62.256 | 0.001 | -0.002 | 1.534 | 48.77 | 28.805 | 2.275 | 65.086 | 0.856 | -1.104 | 0 | 0 | 1 | 41.562 | 1 | 2 | 1 | GJB_Rinse | 0
20 | P-23701 | 2023/02/24 13:24:46 | 1-02 | 10 | 2.966 | 19.849 | 12504 | 0 | 1 | PZ | standard | H2O | 1677270591.73 | 0.205 | 0.892 | 59.92 | 0.001 | -0.002 | 1.47 | 48.551 | 28.656 | 2.281 | 64.921 | 0.856 | -0.982 | 0 | 0 | 1 | 41.688 | 1 | 2 | 1 | GJB_Rinse | 0
21 | P-23702 | 2023/02/24 13:34:34 | 1-03 | 1 | -13.531 | -95.286 | 12829 | -1 | 1 | UT2 | standard | H2O | 1677271179.56 | 0.194 | 0.864 | 60.972 | 0.001 | -0.001 | 1.5 | 47.968 | 29.719 | 2.134 | 62.049 | 0.855 | -1.173 | 0 | 0 | 1 | 41.625 | 1 | 3 | 1 | GJB_Rinse | 0
22 | P-23702 | 2023/02/24 13:43:30 | 1-03 | 2 | -14.234 | -105.733 | 12338 | -1 | 1 | UT2 | standard | H2O | 1677271716.29 | 0.213 | 0.853 | 60.136 | 0.002 | 0.005 | 1.48 | 47.642 | 28.336 | 2.043 | 60.708 | 0.856 | -0.825 | 0 | 0 | 1 | 41.688 | 1 | 3 | 1 | GJB_Rinse | 0
23 | P-23702 | 2023/02/24 13:52:27 | 1-03 | 3 | -14.445 | -108.436 | 12336 | -1 | 1 | UT2 | standard | H2O | 1677272252.15 | 0.194 | 0.739 | 66.811 | 0.002 | 0.004 | 1.654 | 47.529 | 28.402 | 2.062 | 61.468 | 0.855 | -0.782 | 0 | 0 | 1 | 41.688 | 1 | 3 | 1 | GJB_Rinse | 0
24 | P-23702 | 2023/02/24 14:01:22 | 1-03 | 4 | -14.5 | -109.698 | 12491 | 0 | 1 | UT2 | standard | H2O | 1677272788.18 | 0.204 | 0.772 | 64.622 | 0.001 | 0.003 | 1.6 | 48.072 | 28.66 | 2.09 | 61.372 | 0.856 | -0.95 | 0 | 0 | 1 | 41.812 | 1 | 3 | 1 | GJB_Rinse | 0
25 | P-23702 | 2023/02/24 14:10:19 | 1-03 | 5 | -14.536 | -110.396 | 12466 | 0 | 1 | UT2 | standard | H2O | 1677273324.55 | 0.211 | 0.779 | 62.524 | 0.002 | 0.004 | 1.546 | 47.744 | 28.535 | 2.064 | 61.322 | 0.856 | -0.947 | 0 | 0 | 1 | 41.812 | 1 | 3 | 1 | GJB_Rinse | 0
26 | P-23702 | 2023/02/24 14:19:15 | 1-03 | 6 | -14.597 | -110.983 | 12451 | 0 | 1 | UT2 | standard | H2O | 1677273860.45 | 0.204 | 0.844 | 61.621 | 0 | 0.006 | 1.523 | 47.762 | 28.241 | 2.055 | 61.571 | 0.856 | -0.878 | 0 | 0 | 1 | 41.875 | 1 | 3 | 1 | GJB_Rinse | 0
27 | P-23702 | 2023/02/24 14:28:11 | 1-03 | 7 | -14.551 | -111.057 | 12328 | 0 | 1 | UT2 | standard | H2O | 1677274396.35 | 0.218 | 0.762 | 59.823 | 0.001 | 0.002 | 1.478 | 47.219 | 28.474 | 2.033 | 60.287 | 0.856 | -0.832 | 0 | 0 | 1 | 41.875 | 1 | 3 | 1 | GJB_Rinse | 0
28 | P-23702 | 2023/02/24 14:37:07 | 1-03 | 8 | -14.615 | -111.389 | 12439 | 0 | 1 | UT2 | standard | H2O | 1677274932.54 | 0.187 | 0.838 | 63.194 | 0.002 | 0.006 | 1.557 | 47.43 | 28.744 | 2.066 | 61.086 | 0.856 | -0.906 | 0 | 0 | 1 | 41.875 | 1 | 3 | 1 | GJB_Rinse | 0
29 | P-23702 | 2023/02/24 14:46:04 | 1-03 | 9 | -14.621 | -111.584 | 12327 | 0 | 1 | UT2 | standard | H2O | 1677275469.68 | 0.197 | 0.796 | 62.249 | 0.001 | 0.002 | 1.549 | 47.04 | 28.064 | 2.057 | 61.354 | 0.856 | -0.827 | 0 | 0 | 1 | 41.938 | 1 | 3 | 1 | GJB_Rinse | 0
30 | P-23702 | 2023/02/24 14:55:01 | 1-03 | 10 | -14.634 | -111.753 | 12461 | 0 | 1 | UT2 | standard | H2O | 1677276007.19 | 0.203 | 0.711 | 60.715 | 0.001 | 0.004 | 1.497 | 47.259 | 28.729 | 2.057 | 61.58 | 0.856 | -0.916 | 0 | 0 | 1 | 42 | 1 | 3 | 1 | GJB_Rinse | 0
31 | P-23703 | 2023/02/24 15:04:50 | 1-04 | 1 | -9.367 | -72.664 | 12603 | -1 | 1 | EV | standard | H2O | 1677276596.09 | 0.197 | 0.815 | 65.015 | 0 | 0.004 | 1.595 | 47.074 | 29.174 | 2.144 | 62.728 | 0.856 | -0.895 | 0 | 0 | 1 | 42 | 1 | 4 | 1 | GJB_Rinse | 0
32 | P-23703 | 2023/02/24 15:13:47 | 1-04 | 2 | -9.2 | -69.402 | 12393 | -1 | 1 | EV | standard | H2O | 1677277132.47 | 0.211 | 0.809 | 61.774 | 0.001 | 0 | 1.527 | 46.884 | 28.13 | 2.125 | 63.382 | 0.856 | -0.725 | 0 | 0 | 1 | 41.938 | 1 | 4 | 1 | GJB_Rinse | 0
33 | P-23703 | 2023/02/24 15:22:42 | 1-04 | 3 | -9.154 | -68.538 | 12299 | -1 | 1 | EV | standard | H2O | 1677277668.11 | 0.182 | 0.776 | 60.204 | 0 | 0 | 1.496 | 46.598 | 28.218 | 2.112 | 63.052 | 0.856 | -0.691 | 0 | 0 | 1 | 41.812 | 1 | 4 | 1 | GJB_Rinse | 0
34 | P-23703 | 2023/02/24 15:31:39 | 1-04 | 4 | -9.139 | -68.205 | 12229 | 0 | 1 | EV | standard | H2O | 1677278204.28 | 0.183 | 0.848 | 56.988 | 0.001 | -0.001 | 1.414 | 46.482 | 28.188 | 2.09 | 61.231 | 0.856 | -0.86 | 0 | 0 | 1 | 41.875 | 1 | 4 | 1 | GJB_Rinse | 0
35 | P-23703 | 2023/02/24 15:40:35 | 1-04 | 5 | -9.09 | -67.911 | 12326 | 0 | 1 | EV | standard | H2O | 1677278740.86 | 0.21 | 0.736 | 60.224 | 0.001 | -0.001 | 1.495 | 46.793 | 28.437 | 2.108 | 61.181 | 0.855 | -0.93 | 0 | 0 | 1 | 41.812 | 1 | 4 | 1 | GJB_Rinse | 0
36 | P-23703 | 2023/02/24 15:49:31 | 1-04 | 6 | -9.019 | -67.542 | 12352 | 0 | 1 | EV | standard | H2O | 1677279277.29 | 0.202 | 0.789 | 60.55 | 0.001 | -0.001 | 1.498 | 46.957 | 28.318 | 2.097 | 62.028 | 0.856 | -0.962 | 0 | 0 | 1 | 41.75 | 1 | 4 | 1 | GJB_Rinse | 0
37 | P-23703 | 2023/02/24 15:58:28 | 1-04 | 7 | -9.064 | -67.633 | 12333 | 0 | 1 | EV | standard | H2O | 1677279814.17 | 0.186 | 0.932 | 59.98 | 0.001 | 0.005 | 1.488 | 46.793 | 28.687 | 2.099 | 60.758 | 0.856 | -0.991 | 0 | 0 | 1 | 41.75 | 1 | 4 | 1 | GJB_Rinse | 0
38 | P-23703 | 2023/02/24 16:07:24 | 1-04 | 8 | -9.032 | -67.748 | 12465 | 0 | 1 | EV | standard | H2O | 1677280349.64 | 0.2 | 0.793 | 62.973 | 0.001 | -0.001 | 1.557 | 46.962 | 28.545 | 2.125 | 63.178 | 0.856 | -0.865 | 0 | 0 | 1 | 41.75 | 1 | 4 | 1 | GJB_Rinse | 0
39 | P-23703 | 2023/02/24 16:16:21 | 1-04 | 9 | -9.03 | -67.048 | 12413 | 0 | 1 | EV | standard | H2O | 1677280886.58 | 0.197 | 0.756 | 57.81 | 0 | 0 | 1.428 | 46.844 | 28.454 | 2.123 | 62.074 | 0.856 | -0.926 | 0 | 0 | 1 | 41.688 | 1 | 4 | 1 | GJB_Rinse | 0
40 | P-23703 | 2023/02/24 16:25:16 | 1-04 | 10 | -9.048 | -67.336 | 12239 | 0 | 1 | EV | standard | H2O | 1677281422.28 | 0.178 | 0.845 | 60.59 | 0.001 | 0.001 | 1.504 | 46.429 | 28.255 | 2.096 | 61.87 | 0.856 | -0.871 | 0 | 0 | 1 | 41.562 | 1 | 4 | 1 | GJB_Rinse | 0
41 | P-23704 | 2023/02/24 16:35:05 | 1-05 | 1 | -3.936 | -35.132 | 12230 | -1 | 1 | TALL5-3040-3 | 23-011 | H2O | 1677282011.23 | 0.219 | 0.867 | 63.931 | 0.001 | 0.001 | 1.579 | 49.994 | 27.979 | 2.147 | 61.848 | 0.856 | -1.029 | 0 | 0 | 1 | 41.438 | 1 | 5 | 1 | GJB_Rinse | 0
42 | P-23704 | 2023/02/24 16:44:01 | 1-05 | 2 | -3.704 | -31.895 | 12059 | -1 | 1 | TALL5-3040-3 | 23-011 | H2O | 1677282547.03 | 0.207 | 0.816 | 67.561 | 0.001 | -0.003 | 1.671 | 49.843 | 27.174 | 2.127 | 61.779 | 0.856 | -1.099 | 0 | 0 | 1 | 41.438 | 1 | 5 | 1 | GJB_Rinse | 0
43 | P-23704 | 2023/02/24 16:52:57 | 1-05 | 3 | -3.658 | -31.224 | 12096 | -1 | 1 | TALL5-3040-3 | 23-011 | H2O | 1677283082.89 | 0.198 | 0.826 | 64.216 | 0 | -0.001 | 1.589 | 49.864 | 27.201 | 2.131 | 62.378 | 0.855 | -0.991 | 0 | 0 | 1 | 41.438 | 1 | 5 | 1 | GJB_Rinse | 0
44 | P-23704 | 2023/02/24 17:01:52 | 1-05 | 4 | -3.581 | -30.85 | 12342 | 0 | 1 | TALL5-3040-3 | 23-011 | H2O | 1677283618.19 | 0.205 | 0.827 | 58.718 | 0.001 | 0.001 | 1.46 | 50.566 | 28.159 | 2.176 | 62.029 | 0.856 | -1.177 | 0 | 0 | 1 | 41.375 | 1 | 5 | 1 | GJB_Rinse | 0
45 | P-23705 | 2023/02/24 17:11:42 | 1-06 | 1 | -2.132 | -24.608 | 12191 | -1 | 1 | TALL5-010-2 | 23-011 | H2O | 1677284208.09 | 0.209 | 0.77 | 57.366 | 0 | -0.004 | 1.416 | 71.545 | 26.951 | 2.185 | 61.704 | 0.856 | -1.262 | 0 | 0 | 1 | 41.438 | 1 | 6 | 1 | GJB_Rinse | 0
46 | P-23705 | 2023/02/24 17:20:39 | 1-06 | 2 | -1.99 | -23.275 | 11987 | -1 | 1 | TALL5-010-2 | 23-011 | H2O | 1677284744.34 | 0.226 | 0.833 | 62.955 | 0 | -0.002 | 1.566 | 70.986 | 26.36 | 2.15 | 61.321 | 0.855 | -1.222 | 0 | 0 | 1 | 41.438 | 1 | 6 | 1 | GJB_Rinse | 0
47 | P-23705 | 2023/02/24 17:29:35 | 1-06 | 3 | -2.008 | -23.296 | 11973 | -1 | 1 | TALL5-010-2 | 23-011 | H2O | 1677285281.13 | 0.218 | 0.869 | 57.923 | 0.001 | -0.001 | 1.442 | 70.862 | 26.5 | 2.155 | 60.884 | 0.856 | -1.087 | 0 | 0 | 1 | 41.5 | 1 | 6 | 1 | GJB_Rinse | 0
48 | P-23705 | 2023/02/24 17:38:32 | 1-06 | 4 | -1.998 | -23.214 | 12046 | 0 | 1 | TALL5-010-2 | 23-011 | H2O | 1677285817.39 | 0.199 | 0.694 | 62.871 | 0.001 | -0.001 | 1.556 | 71.154 | 26.529 | 2.163 | 62.121 | 0.855 | -1.131 | 0 | 0 | 1 | 41.5 | 1 | 6 | 1 | GJB_Rinse | 0
49 | P-23706 | 2023/02/24 17:48:21 | 1-07 | 1 | -2.404 | -24.281 | 12260 | -1 | 1 | TALL5-2030-2 | 23-011 | H2O | 1677286406.88 | 0.19 | 0.77 | 62.278 | 0.001 | -0.001 | 1.55 | 60.881 | 27.7 | 2.184 | 61.866 | 0.856 | -1.055 | 0 | 0 | 1 | 41.438 | 1 | 7 | 1 | GJB_Rinse | 0
50 | P-23706 | 2023/02/24 17:57:18 | 1-07 | 2 | -2.485 | -24.484 | 12136 | -1 | 1 | TALL5-2030-2 | 23-011 | H2O | 1677286943.55 | 0.214 | 0.801 | 57.879 | 0 | 0 | 1.44 | 60.541 | 27.063 | 2.169 | 61.857 | 0.855 | -1.09 | 0 | 0 | 1 | 41.438 | 1 | 7 | 1 | GJB_Rinse | 0
51 | P-23706 | 2023/02/24 18:06:14 | 1-07 | 3 | -2.458 | -24.56 | 12026 | -1 | 1 | TALL5-2030-2 | 23-011 | H2O | 1677287479.43 | 0.196 | 0.732 | 58.874 | 0.001 | 0 | 1.464 | 60.174 | 26.915 | 2.13 | 60.756 | 0.856 | -1.288 | 0 | 0 | 1 | 41.5 | 1 | 7 | 1 | GJB_Rinse | 0
52 | P-23706 | 2023/02/24 18:15:11 | 1-07 | 4 | -2.443 | -24.096 | 12203 | 0 | 1 | TALL5-2030-2 | 23-011 | H2O | 1677288016.49 | 0.198 | 0.785 | 57.939 | 0.001 | 0 | 1.435 | 60.618 | 27.588 | 2.174 | 62.01 | 0.856 | -1.09 | 0 | 0 | 1 | 41.5 | 1 | 7 | 1 | GJB_Rinse | 0
53 | P-23707 | 2023/02/24 18:24:57 | 1-08 | 1 | -3.089 | -28.14 | 12328 | -1 | 1 | TALL5-2030-3 | 23-011 | H2O | 1677288602.88 | 0.178 | 0.848 | 61 | 0.001 | 0 | 1.511 | 70.025 | 27.45 | 2.197 | 61.817 | 0.855 | -1.204 | 0 | 0 | 1 | 41.5 | 1 | 8 | 1 | GJB_Rinse | 0
54 | P-23707 | 2023/02/24 18:33:53 | 1-08 | 2 | -3.12 | -28.398 | 12120 | -1 | 1 | TALL5-2030-3 | 23-011 | H2O | 1677289138.56 | 0.214 | 0.789 | 62.413 | 0.001 | -0.002 | 1.554 | 69.281 | 27.032 | 2.156 | 61.05 | 0.856 | -1.318 | 0 | 0 | 1 | 41.5 | 1 | 8 | 1 | GJB_Rinse | 0
55 | P-23707 | 2023/02/24 18:42:48 | 1-08 | 3 | -3.122 | -28.15 | 12111 | -1 | 1 | TALL5-2030-3 | 23-011 | H2O | 1677289673.95 | 0.2 | 0.843 | 61.226 | 0 | -0.004 | 1.515 | 69.194 | 26.741 | 2.169 | 62.259 | 0.856 | -1.116 | 0 | 0 | 1 | 41.562 | 1 | 8 | 1 | GJB_Rinse | 0
56 | P-23707 | 2023/02/24 18:51:44 | 1-08 | 4 | -3.121 | -28.15 | 12023 | 0 | 1 | TALL5-2030-3 | 23-011 | H2O | 1677290210.06 | 0.201 | 0.917 | 60.36 | 0.001 | -0.001 | 1.498 | 68.976 | 26.476 | 2.145 | 61.086 | 0.856 | -1.114 | 0 | 0 | 1 | 41.5 | 1 | 8 | 1 | GJB_Rinse | 0
57 | P-23708 | 2023/02/24 19:01:33 | 1-09 | 1 | -3.719 | -26.917 | 12190 | -1 | 1 | TALL5-3040-1 | 23-011 | H2O | 1677290798.39 | 0.203 | 0.9 | 60.104 | 0 | 0.001 | 1.489 | 47.076 | 27.752 | 2.147 | 62.71 | 0.855 | -0.949 | 0 | 0 | 1 | 41.562 | 1 | 9 | 1 | GJB_Rinse | 0
58 | P-23708 | 2023/02/24 19:10:29 | 1-09 | 2 | -3.712 | -26.457 | 11938 | -1 | 1 | TALL5-3040-1 | 23-011 | H2O | 1677291334.4 | 0.189 | 0.777 | 63.739 | 0.001 | 0 | 1.584 | 46.25 | 27.317 | 2.114 | 61.768 | 0.856 | -0.883 | 0 | 0 | 1 | 41.5 | 1 | 9 | 1 | GJB_Rinse | 0
59 | P-23708 | 2023/02/24 19:19:26 | 1-09 | 3 | -3.747 | -26.371 | 11990 | -1 | 1 | TALL5-3040-1 | 23-011 | H2O | 1677291871.36 | 0.209 | 0.831 | 57.898 | 0.001 | 0 | 1.441 | 46.446 | 27.223 | 2.118 | 61.878 | 0.856 | -0.922 | 0 | 0 | 1 | 41.5 | 1 | 9 | 1 | GJB_Rinse | 0
60 | P-23708 | 2023/02/24 19:28:22 | 1-09 | 4 | -3.708 | -26.224 | 12020 | 0 | 1 | TALL5-3040-1 | 23-011 | H2O | 1677292407.79 | 0.209 | 0.925 | 57.121 | 0.001 | 0.001 | 1.423 | 46.521 | 27.717 | 2.118 | 60.483 | 0.855 | -1.096 | 0 | 0 | 1 | 41.5 | 1 | 9 | 1 | GJB_Rinse | 0
61 | P-23709 | 2023/02/24 19:38:09 | 1-10 | 1 | -2.803 | -27.658 | 11848 | -1 | 1 | TALL5-1020-1 | 23-011 | H2O | 1677292994.84 | 0.202 | 0.812 | 57.447 | 0.001 | 0.002 | 1.421 | 56.374 | 26.301 | 2.116 | 61.065 | 0.856 | -1.09 | 0 | 0 | 1 | 41.438 | 1 | 10 | 1 | GJB_Rinse | 0
62 | P-23709 | 2023/02/24 19:47:06 | 1-10 | 2 | -2.779 | -28.082 | 11833 | -1 | 1 | TALL5-1020-1 | 23-011 | H2O | 1677293531.78 | 0.176 | 0.909 | 59.5 | 0 | -0.001 | 1.482 | 56.45 | 26.398 | 2.108 | 60.354 | 0.856 | -1.12 | 0 | 0 | 1 | 41.438 | 1 | 10 | 1 | GJB_Rinse | 0
63 | P-23709 | 2023/02/24 19:56:02 | 1-10 | 3 | -2.784 | -28.019 | 11671 | -1 | 1 | TALL5-1020-1 | 23-011 | H2O | 1677294067.5 | 0.203 | 0.821 | 58.637 | 0.001 | -0.003 | 1.46 | 56.05 | 26.054 | 2.076 | 59.054 | 0.855 | -1.229 | 0 | 0 | 1 | 41.375 | 1 | 10 | 1 | GJB_Rinse | 0
64 | P-23709 | 2023/02/24 20:04:58 | 1-10 | 4 | -2.841 | -28.271 | 11754 | 0 | 1 | TALL5-1020-1 | 23-011 | H2O | 1677294603.77 | 0.178 | 0.877 | 60.343 | 0.001 | 0.003 | 1.511 | 56.374 | 26.078 | 2.092 | 60.2 | 0.856 | -1.248 | 0 | 0 | 1 | 41.375 | 1 | 10 | 1 | GJB_Rinse | 0
65 | P-23710 | 2023/02/24 20:14:46 | 1-11 | 1 | -2.56 | -27.684 | 12288 | -1 | 1 | TALL5-2030-1 | 23-011 | H2O | 1677295191.42 | 0.198 | 0.816 | 62.278 | 0.001 | -0.001 | 1.549 | 63.916 | 27.379 | 2.184 | 62.837 | 0.856 | -1.018 | 0 | 0 | 1 | 41.375 | 1 | 11 | 1 | GJB_Rinse | 0
66 | P-23710 | 2023/02/24 20:23:42 | 1-11 | 2 | -2.585 | -27.875 | 11981 | -1 | 1 | TALL5-2030-1 | 23-011 | H2O | 1677295727.48 | 0.194 | 0.87 | 60.996 | 0.001 | 0.001 | 1.519 | 63.03 | 26.557 | 2.129 | 60.912 | 0.856 | -1.186 | 0 | 0 | 1 | 41.312 | 1 | 11 | 1 | GJB_Rinse | 0
67 | P-23710 | 2023/02/24 20:32:37 | 1-11 | 3 | -2.635 | -28.132 | 11988 | -1 | 1 | TALL5-2030-1 | 23-011 | H2O | 1677296263.27 | 0.212 | 0.799 | 59.278 | 0.001 | -0.001 | 1.471 | 62.983 | 26.43 | 2.139 | 62.409 | 0.855 | -1.014 | 0 | 0 | 1 | 41.375 | 1 | 11 | 1 | GJB_Rinse | 0
68 | P-23710 | 2023/02/24 20:41:34 | 1-11 | 4 | -2.552 | -27.821 | 12020 | 0 | 1 | TALL5-2030-1 | 23-011 | H2O | 1677296799.46 | 0.216 | 0.751 | 61.349 | 0.002 | -0.003 | 1.527 | 63.108 | 26.725 | 2.147 | 60.787 | 0.856 | -1.24 | 0 | 0 | 1 | 41.312 | 1 | 11 | 1 | GJB_Rinse | 0
69 | P-23711 | 2023/02/24 20:51:21 | 1-12 | 1 | -3.102 | -26.273 | 12269 | -1 | 1 | TALL5-1020-2 | 23-011 | H2O | 1677297387.17 | 0.219 | 0.85 | 60.169 | 0.002 | -0.001 | 1.485 | 47.26 | 28.164 | 2.169 | 62.915 | 0.856 | -1.026 | 0 | 0 | 1 | 41.375 | 1 | 12 | 1 | GJB_Rinse | 0
70 | P-23711 | 2023/02/24 21:00:17 | 1-12 | 2 | -3.131 | -26.293 | 12130 | -1 | 1 | TALL5-1020-2 | 23-011 | H2O | 1677297922.18 | 0.186 | 0.799 | 58.922 | 0.001 | 0.001 | 1.453 | 46.901 | 27.625 | 2.149 | 62.779 | 0.856 | -0.971 | 0 | 0 | 1 | 41.375 | 1 | 12 | 1 | GJB_Rinse | 0
71 | P-23711 | 2023/02/24 21:09:13 | 1-12 | 3 | -3.13 | -25.82 | 12092 | -1 | 1 | TALL5-1020-2 | 23-011 | H2O | 1677298458.8 | 0.186 | 0.765 | 65.421 | 0.001 | 0.002 | 1.617 | 46.721 | 27.84 | 2.134 | 61.786 | 0.856 | -1.023 | 0 | 0 | 1 | 41.375 | 1 | 12 | 1 | GJB_Rinse | 0
72 | P-23711 | 2023/02/24 21:18:09 | 1-12 | 4 | -3.173 | -26.199 | 12205 | 0 | 1 | TALL5-1020-2 | 23-011 | H2O | 1677298994.78 | 0.205 | 1.004 | 59.974 | 0.001 | -0.003 | 1.488 | 47.151 | 27.742 | 2.166 | 61.878 | 0.855 | -1.097 | 0 | 0 | 1 | 41.375 | 1 | 12 | 1 | GJB_Rinse | 0
73 | P-23712 | 2023/02/24 21:27:58 | 1-13 | 1 | -8.643 | -61.346 | 12570 | -1 | 1 | EV | standard | H2O | 1677299583.5 | 0.204 | 0.739 | 61.672 | 0.002 | -0.001 | 1.517 | 47.199 | 28.927 | 2.15 | 62.119 | 0.855 | -1.112 | 0 | 0 | 1 | 41.312 | 1 | 13 | 1 | GJB_Rinse | 0
74 | P-23712 | 2023/02/24 21:36:55 | 1-13 | 2 | -8.905 | -64.711 | 12265 | -1 | 1 | EV | standard | H2O | 1677300121.23 | 0.201 | 0.827 | 63.362 | 0.002 | 0 | 1.569 | 46.435 | 28.246 | 2.098 | 61.414 | 0.855 | -1.029 | 0 | 0 | 1 | 41.312 | 1 | 13 | 1 | GJB_Rinse | 0
75 | P-23712 | 2023/02/24 21:45:51 | 1-13 | 3 | -8.886 | -65.395 | 12193 | -1 | 1 | EV | standard | H2O | 1677300656.66 | 0.209 | 0.814 | 60.886 | 0.002 | -0.001 | 1.513 | 46.26 | 28.001 | 2.079 | 61.365 | 0.855 | -0.913 | 0 | 0 | 1 | 41.375 | 1 | 13 | 1 | GJB_Rinse | 0
76 | P-23712 | 2023/02/24 21:54:47 | 1-13 | 4 | -8.976 | -65.921 | 12177 | 0 | 1 | EV | standard | H2O | 1677301192.99 | 0.204 | 0.881 | 61.724 | 0.001 | 0.004 | 1.522 | 46.274 | 28.075 | 2.085 | 60.393 | 0.855 | -1.033 | 0 | 0 | 1 | 41.312 | 1 | 13 | 1 | GJB_Rinse | 0
77 | P-23713 | 2023/02/24 22:04:37 | 1-14 | 1 | -2.565 | -26.886 | 12638 | -1 | 1 | TALL5-3040-2 | 23-011 | H2O | 1677301782.93 | 0.192 | 0.793 | 58.577 | 0 | 0 | 1.44 | 48.018 | 28.991 | 2.234 | 64.85 | 0.855 | -0.951 | 0 | 0 | 1 | 41.312 | 1 | 14 | 1 | GJB_Rinse | 0
78 | P-23713 | 2023/02/24 22:13:33 | 1-14 | 2 | -2.228 | -23.126 | 12284 | -1 | 1 | TALL5-3040-2 | 23-011 | H2O | 1677302318.61 | 0.195 | 0.838 | 59.863 | 0.001 | -0.004 | 1.486 | 47.238 | 28.324 | 2.18 | 62.863 | 0.856 | -1.012 | 0 | 0 | 1 | 41.25 | 1 | 14 | 1 | GJB_Rinse | 0
79 | P-23713 | 2023/02/24 22:22:30 | 1-14 | 3 | -2.195 | -22.384 | 12195 | -1 | 1 | TALL5-3040-2 | 23-011 | H2O | 1677302855.46 | 0.199 | 0.837 | 57.759 | 0.001 | -0.001 | 1.435 | 47.134 | 27.888 | 2.171 | 62.663 | 0.856 | -0.906 | 0 | 0 | 1 | 41.125 | 1 | 14 | 1 | GJB_Rinse | 0
80 | P-23713 | 2023/02/24 22:31:26 | 1-14 | 4 | -2.192 | -22.012 | 12438 | 0 | 1 | TALL5-3040-2 | 23-011 | H2O | 1677303390.9 | 0.212 | 0.881 | 62.098 | 0.002 | -0.004 | 1.538 | 47.639 | 28.627 | 2.206 | 63.339 | 0.855 | -1.111 | 0 | 0 | 1 | 41.25 | 1 | 14 | 1 | GJB_Rinse | 0
81 | P-23714 | 2023/02/24 22:41:12 | 1-15 | 1 | -3.979 | -31.749 | 12355 | -1 | 1 | TALL5-deep-3 | 23-011 | H2O | 1677303977.89 | 0.223 | 0.828 | 65.243 | 0.001 | -0.001 | 1.604 | 57.06 | 27.568 | 2.192 | 62.934 | 0.856 | -1.065 | 0 | 0 | 1 | 41.125 | 1 | 15 | 1 | GJB_Rinse | 0
82 | P-23714 | 2023/02/24 22:50:09 | 1-15 | 2 | -4.016 | -32.496 | 12161 | -1 | 1 | TALL5-deep-3 | 23-011 | H2O | 1677304514.37 | 0.191 | 0.713 | 61.134 | 0 | 0 | 1.521 | 56.477 | 27.32 | 2.154 | 62.869 | 0.856 | -0.963 | 0 | 0 | 1 | 41.125 | 1 | 15 | 1 | GJB_Rinse | 0
83 | P-23714 | 2023/02/24 22:59:05 | 1-15 | 3 | -4.086 | -32.867 | 12079 | -1 | 1 | TALL5-deep-3 | 23-011 | H2O | 1677305050.69 | 0.196 | 0.785 | 60.871 | 0 | 0 | 1.512 | 56.536 | 27.203 | 2.125 | 60.553 | 0.856 | -1.262 | 0 | 0 | 1 | 41.125 | 1 | 15 | 1 | GJB_Rinse | 0
84 | P-23714 | 2023/02/24 23:08:01 | 1-15 | 4 | -4.103 | -32.952 | 12237 | 0 | 1 | TALL5-deep-3 | 23-011 | H2O | 1677305587.05 | 0.204 | 0.84 | 63.789 | 0 | -0.003 | 1.571 | 56.986 | 27.706 | 2.146 | 60.742 | 0.856 | -1.242 | 0 | 0 | 1 | 41.125 | 1 | 15 | 1 | GJB_Rinse | 0
85 | P-23715 | 2023/02/24 23:17:48 | 1-16 | 1 | 0.236 | -21.258 | 12600 | -1 | 1 | TALL5-B-4 | 23-011 | H2O | 1677306173.58 | 0.212 | 1.043 | 63.365 | 0.001 | 0.004 | 1.555 | 345.064 | 16.2 | 2.63 | 60.437 | 0.856 | -3.017 | 0 | 0 | 1 | 41.125 | 1 | 16 | 1 | GJB_Rinse | 0
86 | P-23715 | 2023/02/24 23:26:43 | 1-16 | 2 | 0.3 | -20.949 | 12066 | -1 | 1 | TALL5-B-4 | 23-011 | H2O | 1677306708.75 | 0.228 | 0.953 | 62.309 | 0.001 | -0.002 | 1.55 | 333.42 | 15.455 | 2.525 | 57.641 | 0.857 | -2.946 | 0 | 0 | 1 | 41.062 | 1 | 16 | 1 | GJB_Rinse | 0
87 | P-23715 | 2023/02/24 23:35:38 | 1-16 | 3 | 0.249 | -20.989 | 12244 | -1 | 1 | TALL5-B-4 | 23-011 | H2O | 1677307243.92 | 0.255 | 0.967 | 61.614 | 0.001 | -0.003 | 1.528 | 338.415 | 15.816 | 2.551 | 58.748 | 0.856 | -3.112 | 0 | 0 | 1 | 41.062 | 1 | 16 | 1 | GJB_Rinse | 0
88 | P-23715 | 2023/02/24 23:44:34 | 1-16 | 4 | 0.307 | -20.635 | 12195 | 0 | 1 | TALL5-B-4 | 23-011 | H2O | 1677307779.92 | 0.231 | 0.913 | 66.304 | 0.001 | -0.001 | 1.65 | 337.886 | 15.408 | 2.543 | 58.792 | 0.856 | -3.032 | 0 | 0 | 1 | 41.062 | 1 | 16 | 1 | GJB_Rinse | 0
89 | P-23716 | 2023/02/24 23:54:21 | 1-17 | 1 | -0.932 | -20.935 | 12531 | -1 | 1 | TALL5-B-5 | 23-011 | H2O | 1677308367.23 | 0.201 | 0.995 | 61.997 | 0 | 0.004 | 1.531 | 243.492 | 21.107 | 2.396 | 61.106 | 0.857 | -2.009 | 0 | 0 | 1 | 41 | 1 | 17 | 1 | GJB_Rinse | 0
90 | P-23716 | 2023/02/25 00:03:17 | 1-17 | 2 | -0.873 | -20.477 | 12090 | -1 | 1 | TALL5-B-5 | 23-011 | H2O | 1677308902.92 | 0.223 | 0.802 | 61.072 | 0.001 | -0.004 | 1.509 | 235.504 | 19.862 | 2.294 | 59.863 | 0.856 | -1.941 | 0 | 0 | 1 | 41 | 1 | 17 | 1 | GJB_Rinse | 0
91 | P-23716 | 2023/02/25 00:12:13 | 1-17 | 3 | -0.885 | -20.37 | 12193 | -1 | 1 | TALL5-B-5 | 23-011 | H2O | 1677309438.45 | 0.217 | 0.894 | 59.509 | 0.001 | -0.003 | 1.478 | 236.679 | 20.194 | 2.329 | 60.778 | 0.857 | -1.966 | 0 | 0 | 1 | 41 | 1 | 17 | 1 | GJB_Rinse | 0
92 | P-23716 | 2023/02/25 00:21:09 | 1-17 | 4 | -0.911 | -20.374 | 12104 | 0 | 1 | TALL5-B-5 | 23-011 | H2O | 1677309974.76 | 0.214 | 0.96 | 65.08 | 0.001 | 0 | 1.618 | 235.683 | 20.217 | 2.308 | 59.173 | 0.856 | -1.767 | 0 | 0 | 1 | 41 | 1 | 17 | 1 | GJB_Rinse | 0
93 | P-23717 | 2023/02/25 00:30:56 | 1-18 | 1 | -2.015 | -24.766 | 12625 | -1 | 1 | TALL5-C-2 | 23-011 | H2O | 1677310561.49 | 0.203 | 0.88 | 62.25 | 0.001 | -0.006 | 1.537 | 179.191 | 24.696 | 2.302 | 60.652 | 0.856 | -1.702 | 0 | 0 | 1 | 41 | 1 | 18 | 1 | GJB_Rinse | 0
94 | P-23717 | 2023/02/25 00:39:52 | 1-18 | 2 | -2.097 | -25.228 | 12174 | -1 | 1 | TALL5-C-2 | 23-011 | H2O | 1677311097.84 | 0.197 | 0.894 | 66.426 | 0.002 | 0 | 1.65 | 173.146 | 23.557 | 2.232 | 59.439 | 0.856 | -1.513 | 0 | 0 | 1 | 41 | 1 | 18 | 1 | GJB_Rinse | 0
95 | P-23717 | 2023/02/25 00:48:48 | 1-18 | 3 | -2.005 | -25.018 | 12155 | -1 | 1 | TALL5-C-2 | 23-011 | H2O | 1677311634.46 | 0.199 | 0.851 | 59.947 | 0.001 | 0.001 | 1.483 | 172.933 | 23.627 | 2.211 | 58.662 | 0.856 | -1.834 | 0 | 0 | 1 | 40.938 | 1 | 18 | 1 | GJB_Rinse | 0
96 | P-23717 | 2023/02/25 00:57:45 | 1-18 | 4 | -2.047 | -24.881 | 12121 | 0 | 1 | TALL5-C-2 | 23-011 | H2O | 1677312171.04 | 0.229 | 0.901 | 63.048 | 0 | -0.001 | 1.556 | 172.279 | 23.168 | 2.211 | 59.946 | 0.856 | -1.466 | 0 | 0 | 1 | 40.875 | 1 | 18 | 1 | GJB_Rinse | 0
97 | P-23718 | 2023/02/25 01:07:33 | 1-19 | 1 | 2.784 | -18.509 | 12444 | -1 | 1 | TALL5-B-3 | 23-011 | H2O | 1677312758.59 | 0.197 | 0.974 | 61.97 | 0 | 0 | 1.542 | 216.487 | 17.424 | 2.938 | 63.876 | 0.856 | -4.12 | 0 | 0 | 1 | 40.812 | 1 | 19 | 1 | GJB_Rinse | 0
98 | P-23718 | 2023/02/25 01:16:29 | 1-19 | 2 | 2.914 | -17.941 | 12080 | -1 | 1 | TALL5-B-3 | 23-011 | H2O | 1677313294.82 | 0.22 | 0.93 | 65.192 | 0 | -0.002 | 1.622 | 210.478 | 16.807 | 2.858 | 62.28 | 0.856 | -4.089 | 0 | 0 | 1 | 40.75 | 1 | 19 | 1 | GJB_Rinse | 0
99 | P-23718 | 2023/02/25 01:25:25 | 1-19 | 3 | 2.893 | -18.212 | 12069 | -1 | 1 | TALL5-B-3 | 23-011 | H2O | 1677313830.25 | 0.198 | 0.902 | 63.009 | 0 | 0.002 | 1.567 | 211.101 | 16.403 | 2.866 | 62.894 | 0.855 | -3.964 | 0 | 0 | 1 | 40.625 | 1 | 19 | 1 | GJB_Rinse | 0
100 | P-23718 | 2023/02/25 01:34:21 | 1-19 | 4 | 2.941 | -17.729 | 12119 | 0 | 1 | TALL5-B-3 | 23-011 | H2O | 1677314366.57 | 0.184 | 0.912 | 62.211 | 0 | -0.005 | 1.538 | 211.107 | 17.028 | 2.852 | 61.029 | 0.856 | -4.129 | 0 | 0 | 1 | 40.625 | 1 | 19 | 1 | GJB_Rinse | 0
101 | P-23719 | 2023/02/25 01:44:09 | 1-20 | 1 | 0.609 | -24.748 | 12629 | -1 | 1 | TALL5-A-3 | 23-011 | H2O | 1677314954.56 | 0.219 | 0.818 | 65.236 | 0.001 | 0.004 | 1.612 | 55.8 | 25.832 | 2.557 | 65.547 | 0.856 | -2.74 | 0 | 0 | 1 | 40.562 | 1 | 20 | 1 | GJB_Rinse | 0
102 | P-23719 | 2023/02/25 01:53:06 | 1-20 | 2 | 0.544 | -25.528 | 12221 | -1 | 1 | TALL5-A-3 | 23-011 | H2O | 1677315492.07 | 0.206 | 0.897 | 62.05 | 0.001 | -0.001 | 1.534 | 53.468 | 24.627 | 2.48 | 64.594 | 0.856 | -2.3 | 0 | 0 | 1 | 40.5 | 1 | 20 | 1 | GJB_Rinse | 0
103 | P-23719 | 2023/02/25 02:02:02 | 1-20 | 3 | 0.592 | -25.359 | 12101 | -1 | 1 | TALL5-A-3 | 23-011 | H2O | 1677316028.06 | 0.191 | 0.806 | 61.161 | 0.001 | -0.001 | 1.508 | 52.823 | 24.845 | 2.439 | 61.804 | 0.855 | -2.69 | 0 | 0 | 1 | 40.5 | 1 | 20 | 1 | GJB_Rinse | 0
104 | P-23719 | 2023/02/25 02:10:57 | 1-20 | 4 | 0.59 | -25.314 | 12469 | 0 | 1 | TALL5-A-3 | 23-011 | H2O | 1677316562.86 | 0.196 | 0.885 | 63.465 | 0.001 | -0.003 | 1.57 | 53.631 | 25.256 | 2.53 | 64.739 | 0.855 | -2.638 | 0 | 0 | 1 | 40.562 | 1 | 20 | 1 | GJB_Rinse | 0
105 | P-23720 | 2023/02/25 02:20:46 | 1-21 | 1 | -1.705 | -22.755 | 12348 | -1 | 1 | TALL5-B-1 | 23-011 | H2O | 1677317151.79 | 0.209 | 1.034 | 62.802 | 0.001 | -0.001 | 1.553 | 292.732 | 19.826 | 2.274 | 59.433 | 0.857 | -1.511 | 0 | 0 | 1 | 40.5 | 1 | 21 | 1 | GJB_Rinse | 0
106 | P-23720 | 2023/02/25 02:29:42 | 1-21 | 2 | -1.655 | -22.317 | 12032 | -1 | 1 | TALL5-B-1 | 23-011 | H2O | 1677317687.6 | 0.209 | 0.835 | 64.468 | 0.001 | 0 | 1.594 | 287.143 | 19.404 | 2.219 | 57.586 | 0.856 | -1.614 | 0 | 0 | 1 | 40.5 | 1 | 21 | 1 | GJB_Rinse | 0
107 | P-23720 | 2023/02/25 02:38:38 | 1-21 | 3 | -1.673 | -21.985 | 11981 | -1 | 1 | TALL5-B-1 | 23-011 | H2O | 1677318223.4 | 0.226 | 0.928 | 64.744 | 0.001 | -0.005 | 1.609 | 286.463 | 19.624 | 2.202 | 56.391 | 0.856 | -1.66 | 0 | 0 | 1 | 40.5 | 1 | 21 | 1 | GJB_Rinse | 0
108 | P-23720 | 2023/02/25 02:47:34 | 1-21 | 4 | -1.66 | -22.066 | 12028 | 0 | 1 | TALL5-B-1 | 23-011 | H2O | 1677318760.14 | 0.237 | 0.876 | 62.774 | 0.001 | 0 | 1.559 | 287.633 | 19.314 | 2.217 | 57.106 | 0.857 | -1.556 | 0 | 0 | 1 | 40.562 | 1 | 21 | 1 | GJB_Rinse | 0
109 | P-23721 | 2023/02/25 02:57:19 | 1-22 | 1 | -8.562 | -61.056 | 12625 | -1 | 1 | EV | standard | H2O | 1677319344.24 | 0.198 | 0.788 | 60.893 | 0.001 | -0.002 | 1.501 | 50.546 | 29.021 | 2.165 | 62.703 | 0.856 | -1.086 | 0 | 0 | 1 | 40.562 | 1 | 22 | 1 | GJB_Rinse | 0
110 | P-23721 | 2023/02/25 03:06:15 | 1-22 | 2 | -8.842 | -64.329 | 12299 | -1 | 1 | EV | standard | H2O | 1677319880.52 | 0.205 | 0.789 | 63.252 | 0.002 | 0.002 | 1.562 | 48.064 | 28.375 | 2.104 | 61.038 | 0.855 | -1.147 | 0 | 0 | 1 | 40.625 | 1 | 22 | 1 | GJB_Rinse | 0
111 | P-23721 | 2023/02/25 03:15:11 | 1-22 | 3 | -8.877 | -65.337 | 12339 | -1 | 1 | EV | standard | H2O | 1677320416.38 | 0.204 | 0.817 | 62.947 | 0.002 | 0.003 | 1.559 | 47.628 | 28.398 | 2.098 | 60.781 | 0.855 | -1.14 | 0 | 0 | 1 | 40.625 | 1 | 22 | 1 | GJB_Rinse | 0
112 | P-23721 | 2023/02/25 03:24:06 | 1-22 | 4 | -8.93 | -65.786 | 12219 | 0 | 1 | EV | standard | H2O | 1677320952.21 | 0.217 | 0.94 | 64.099 | 0.001 | 0.005 | 1.586 | 47.022 | 27.866 | 2.085 | 61.35 | 0.856 | -1.005 | 0 | 0 | 1 | 40.688 | 1 | 22 | 1 | GJB_Rinse | 0
113 | P-23722 | 2023/02/25 03:33:53 | 1-23 | 1 | -1.316 | -31.836 | 12573 | -1 | 1 | TALL5-A-4 | 23-011 | H2O | 1677321538.8 | 0.209 | 0.747 | 61.438 | 0 | -0.001 | 1.508 | 48.989 | 28.392 | 2.24 | 64.206 | 0.856 | -1.318 | 0 | 0 | 1 | 40.625 | 1 | 23 | 1 | GJB_Rinse | 0
114 | P-23722 | 2023/02/25 03:42:50 | 1-23 | 2 | -1.061 | -29.189 | 12227 | -1 | 1 | TALL5-A-4 | 23-011 | H2O | 1677322075.57 | 0.195 | 0.926 | 63.781 | 0.001 | 0 | 1.579 | 48.414 | 27.332 | 2.174 | 62.232 | 0.855 | -1.309 | 0 | 0 | 1 | 40.688 | 1 | 23 | 1 | GJB_Rinse | 0
115 | P-23722 | 2023/02/25 03:51:46 | 1-23 | 3 | -0.93 | -28.152 | 12260 | -1 | 1 | TALL5-A-4 | 23-011 | H2O | 1677322611.3 | 0.193 | 0.834 | 61.354 | 0.001 | -0.003 | 1.516 | 48.389 | 27.473 | 2.188 | 63.122 | 0.855 | -1.285 | 0 | 0 | 1 | 40.75 | 1 | 23 | 1 | GJB_Rinse | 0
116 | P-23722 | 2023/02/25 04:00:41 | 1-23 | 4 | -0.94 | -28.03 | 12209 | 0 | 1 | TALL5-A-4 | 23-011 | H2O | 1677323146.64 | 0.201 | 0.79 | 58.646 | 0.001 | 0.001 | 1.457 | 48.2 | 27.567 | 2.185 | 62.298 | 0.855 | -1.331 | 0 | 0 | 1 | 40.75 | 1 | 23 | 1 | GJB_Rinse | 0
117 | P-23723 | 2023/02/25 04:10:28 | 1-24 | 1 | 1.821 | -21.838 | 12497 | -1 | 1 | TALL5-A-1 | 23-011 | H2O | 1677323733.34 | 0.19 | 0.841 | 58.818 | 0.001 | -0.003 | 1.447 | 55.756 | 24.699 | 2.633 | 65.61 | 0.855 | -3.173 | 0 | 0 | 1 | 40.75 | 1 | 24 | 1 | GJB_Rinse | 0
118 | P-23723 | 2023/02/25 04:19:23 | 1-24 | 2 | 1.843 | -21.43 | 12147 | -1 | 1 | TALL5-A-1 | 23-011 | H2O | 1677324268.8 | 0.22 | 0.897 | 61.262 | 0.001 | -0.004 | 1.515 | 54.865 | 23.642 | 2.572 | 64.654 | 0.856 | -2.845 | 0 | 0 | 1 | 40.75 | 1 | 24 | 1 | GJB_Rinse | 0
119 | P-23723 | 2023/02/25 04:28:19 | 1-24 | 3 | 1.847 | -21.44 | 12138 | -1 | 1 | TALL5-A-1 | 23-011 | H2O | 1677324804.25 | 0.183 | 0.944 | 59.127 | 0.001 | -0.003 | 1.462 | 54.838 | 23.598 | 2.593 | 64.892 | 0.856 | -2.798 | 0 | 0 | 1 | 40.75 | 1 | 24 | 1 | GJB_Rinse | 0
120 | P-23723 | 2023/02/25 04:37:14 | 1-24 | 4 | 1.793 | -21.754 | 12215 | 0 | 1 | TALL5-A-1 | 23-011 | H2O | 1677325339.73 | 0.208 | 0.909 | 61.936 | 0 | -0.001 | 1.534 | 55.116 | 23.916 | 2.587 | 64.578 | 0.855 | -2.924 | 0 | 0 | 1 | 40.75 | 1 | 24 | 1 | GJB_Rinse | 0
121 | P-23724 | 2023/02/25 04:47:01 | 1-25 | 1 | -2.52 | -23.266 | 12344 | -1 | 1 | TALL5-010-1 | 23-011 | H2O | 1677325926.77 | 0.228 | 0.797 | 60.953 | 0.001 | 0 | 1.506 | 51.335 | 28.22 | 2.183 | 61.946 | 0.855 | -1.373 | 0 | 0 | 1 | 40.75 | 1 | 25 | 1 | GJB_Rinse | 0
122 | P-23724 | 2023/02/25 04:55:58 | 1-25 | 2 | -2.519 | -22.682 | 12313 | -1 | 1 | TALL5-010-1 | 23-011 | H2O | 1677326463.36 | 0.181 | 0.721 | 65.423 | 0.001 | 0.003 | 1.619 | 51.228 | 28.014 | 2.182 | 62.817 | 0.855 | -1.235 | 0 | 0 | 1 | 40.688 | 1 | 25 | 1 | GJB_Rinse | 0
123 | P-23724 | 2023/02/25 05:04:53 | 1-25 | 3 | -2.587 | -22.893 | 12301 | -1 | 1 | TALL5-010-1 | 23-011 | H2O | 1677326999.11 | 0.202 | 0.864 | 67.642 | 0.001 | 0.001 | 1.674 | 51.168 | 27.991 | 2.178 | 62.638 | 0.855 | -1.189 | 0 | 0 | 1 | 40.688 | 1 | 25 | 1 | GJB_Rinse | 0
124 | P-23724 | 2023/02/25 05:13:49 | 1-25 | 4 | -2.567 | -22.922 | 12325 | 0 | 1 | TALL5-010-1 | 23-011 | H2O | 1677327534.47 | 0.202 | 0.808 | 64.171 | 0 | 0.002 | 1.586 | 51.097 | 28.375 | 2.192 | 62.471 | 0.855 | -1.138 | 0 | 0 | 1 | 40.75 | 1 | 25 | 1 | GJB_Rinse | 0
125 | P-23725 | 2023/02/25 05:23:36 | 1-26 | 1 | -2.683 | -25.57 | 12506 | -1 | 1 | TALL5-C-4 | 23-011 | H2O | 1677328121.97 | 0.231 | 0.941 | 63.257 | 0.001 | 0.001 | 1.564 | 324.478 | 19.611 | 2.265 | 58.502 | 0.856 | -1.443 | 0 | 0 | 1 | 40.75 | 1 | 26 | 1 | GJB_Rinse | 0
126 | P-23725 | 2023/02/25 05:32:33 | 1-26 | 2 | -2.758 | -25.987 | 12104 | -1 | 1 | TALL5-C-4 | 23-011 | H2O | 1677328658.36 | 0.216 | 0.875 | 60.277 | 0.001 | -0.004 | 1.496 | 316.745 | 19.007 | 2.192 | 56.556 | 0.856 | -1.446 | 0 | 0 | 1 | 40.75 | 1 | 26 | 1 | GJB_Rinse | 0
127 | P-23725 | 2023/02/25 05:41:28 | 1-26 | 3 | -2.742 | -26.411 | 12025 | -1 | 1 | TALL5-C-4 | 23-011 | H2O | 1677329193.65 | 0.238 | 0.928 | 62.236 | 0.001 | 0.004 | 1.543 | 316.088 | 18.551 | 2.187 | 56.103 | 0.856 | -1.554 | 0 | 0 | 1 | 40.75 | 1 | 26 | 1 | GJB_Rinse | 0
128 | P-23725 | 2023/02/25 05:50:23 | 1-26 | 4 | -2.737 | -26.143 | 12157 | 0 | 1 | TALL5-C-4 | 23-011 | H2O | 1677329728.76 | 0.22 | 0.935 | 59.288 | 0.001 | 0.003 | 1.463 | 318.983 | 19.32 | 2.208 | 55.858 | 0.856 | -1.606 | 0 | 0 | 1 | 40.812 | 1 | 26 | 1 | GJB_Rinse | 0
129 | P-23726 | 2023/02/25 06:00:10 | 1-27 | 1 | -1.536 | -25.278 | 12570 | -1 | 1 | TALL5-C-3 | 23-011 | H2O | 1677330315.32 | 0.232 | 0.947 | 62.26 | 0.001 | 0 | 1.546 | 215.408 | 23.001 | 2.298 | 60.259 | 0.856 | -1.795 | 0 | 0 | 1 | 40.812 | 1 | 27 | 1 | GJB_Rinse | 0
130 | P-23726 | 2023/02/25 06:09:06 | 1-27 | 2 | -1.485 | -25.197 | 12183 | -1 | 1 | TALL5-C-3 | 23-011 | H2O | 1677330851.78 | 0.221 | 0.837 | 63.673 | 0.001 | 0 | 1.574 | 208.802 | 22.213 | 2.22 | 58.315 | 0.856 | -1.704 | 0 | 0 | 1 | 40.75 | 1 | 27 | 1 | GJB_Rinse | 0
131 | P-23726 | 2023/02/25 06:18:02 | 1-27 | 3 | -1.435 | -25.138 | 12168 | -1 | 1 | TALL5-C-3 | 23-011 | H2O | 1677331387.21 | 0.215 | 0.981 | 61.359 | 0.001 | -0.003 | 1.528 | 208.13 | 22.227 | 2.211 | 58.778 | 0.855 | -1.655 | 0 | 0 | 1 | 40.812 | 1 | 27 | 1 | GJB_Rinse | 0
132 | P-23726 | 2023/02/25 06:26:58 | 1-27 | 4 | -1.469 | -25.409 | 12174 | 0 | 1 | TALL5-C-3 | 23-011 | H2O | 1677331923.46 | 0.216 | 0.851 | 57.246 | 0.001 | -0.001 | 1.42 | 208.152 | 22.305 | 2.238 | 59.109 | 0.856 | -1.632 | 0 | 0 | 1 | 40.875 | 1 | 27 | 1 | GJB_Rinse | 0
133 | P-23727 | 2023/02/25 06:36:46 | 1-28 | 1 | -0.945 | -26.595 | 12632 | -1 | 1 | TALL5-A-5 | 23-011 | H2O | 1677332511.37 | 0.202 | 0.842 | 63.556 | 0.001 | 0.004 | 1.573 | 54.372 | 28.653 | 2.268 | 64.122 | 0.856 | -1.356 | 0 | 0 | 1 | 40.875 | 1 | 28 | 1 | GJB_Rinse | 0
134 | P-23727 | 2023/02/25 06:45:41 | 1-28 | 2 | -0.913 | -26.696 | 12365 | -1 | 1 | TALL5-A-5 | 23-011 | H2O | 1677333046.81 | 0.181 | 0.878 | 61.472 | 0.001 | 0 | 1.531 | 52.521 | 28.012 | 2.223 | 62.425 | 0.856 | -1.363 | 0 | 0 | 1 | 40.938 | 1 | 28 | 1 | GJB_Rinse | 0
135 | P-23727 | 2023/02/25 06:54:36 | 1-28 | 3 | -0.912 | -26.823 | 12333 | -1 | 1 | TALL5-A-5 | 23-011 | H2O | 1677333581.53 | 0.188 | 0.755 | 63.592 | 0.001 | -0.001 | 1.572 | 52.15 | 28.195 | 2.215 | 61.788 | 0.856 | -1.445 | 0 | 0 | 1 | 40.875 | 1 | 28 | 1 | GJB_Rinse | 0
136 | P-23727 | 2023/02/25 07:03:31 | 1-28 | 4 | -0.959 | -27.3 | 12300 | 0 | 1 | TALL5-A-5 | 23-011 | H2O | 1677334116.4 | 0.205 | 0.874 | 64.878 | 0.001 | 0 | 1.607 | 51.802 | 27.859 | 2.218 | 62.339 | 0.855 | -1.31 | 0 | 0 | 1 | 40.875 | 1 | 28 | 1 | GJB_Rinse | 0
137 | P-23728 | 2023/02/25 07:13:15 | 1-29 | 1 | -0.86 | -28.194 | 12649 | -1 | 1 | TALL5-A-2 | 23-011 | H2O | 1677334700.44 | 0.198 | 0.845 | 63.663 | 0.001 | 0.001 | 1.577 | 52.379 | 28.99 | 2.261 | 63.836 | 0.855 | -1.27 | 0 | 0 | 1 | 40.938 | 1 | 29 | 1 | GJB_Rinse | 0
138 | P-23728 | 2023/02/25 07:22:11 | 1-29 | 2 | -0.797 | -27.731 | 12233 | -1 | 1 | TALL5-A-2 | 23-011 | H2O | 1677335237.07 | 0.2 | 0.913 | 60.312 | 0.001 | 0.003 | 1.486 | 51.63 | 27.509 | 2.191 | 61.053 | 0.856 | -1.411 | 0 | 0 | 1 | 40.875 | 1 | 29 | 1 | GJB_Rinse | 0
139 | P-23728 | 2023/02/25 07:31:07 | 1-29 | 3 | -0.813 | -27.899 | 12217 | -1 | 1 | TALL5-A-2 | 23-011 | H2O | 1677335772.84 | 0.204 | 0.835 | 59.055 | 0.001 | -0.002 | 1.457 | 51.368 | 27.611 | 2.192 | 62.269 | 0.855 | -1.241 | 0 | 0 | 1 | 40.875 | 1 | 29 | 1 | GJB_Rinse | 0
140 | P-23728 | 2023/02/25 07:40:03 | 1-29 | 4 | -0.845 | -28.119 | 12231 | 0 | 1 | TALL5-A-2 | 23-011 | H2O | 1677336309.16 | 0.195 | 0.755 | 58.704 | 0 | -0.003 | 1.451 | 51.544 | 27.688 | 2.179 | 61.793 | 0.856 | -1.377 | 0 | 0 | 1 | 40.875 | 1 | 29 | 1 | GJB_Rinse | 0
141 | P-23729 | 2023/02/25 07:49:48 | 1-30 | 1 | -1.309 | -24.89 | 12648 | -1 | 1 | TALL5-D-1 | 23-011 | H2O | 1677336893.4 | 0.188 | 0.75 | 59.092 | 0 | 0 | 1.458 | 51.932 | 28.451 | 2.306 | 63.341 | 0.856 | -1.666 | 0 | 0 | 1 | 40.938 | 1 | 30 | 1 | GJB_Rinse | 0
142 | P-23729 | 2023/02/25 07:58:43 | 1-30 | 2 | -1.391 | -25.035 | 12276 | -1 | 1 | TALL5-D-1 | 23-011 | H2O | 1677337428.24 | 0.194 | 0.867 | 64.479 | 0.001 | 0 | 1.592 | 51.154 | 27.1 | 2.248 | 62.122 | 0.856 | -1.626 | 0 | 0 | 1 | 40.938 | 1 | 30 | 1 | GJB_Rinse | 0
143 | P-23729 | 2023/02/25 08:07:38 | 1-30 | 3 | -1.401 | -24.98 | 12526 | -1 | 1 | TALL5-D-1 | 23-011 | H2O | 1677337964.08 | 0.189 | 0.881 | 63.992 | 0.001 | 0.002 | 1.584 | 51.627 | 27.724 | 2.291 | 64.239 | 0.856 | -1.522 | 0 | 0 | 1 | 40.938 | 1 | 30 | 1 | GJB_Rinse | 0
144 | P-23729 | 2023/02/25 08:16:35 | 1-30 | 4 | -1.451 | -25.136 | 12459 | 0 | 1 | TALL5-D-1 | 23-011 | H2O | 1677338500.31 | 0.199 | 0.727 | 62.867 | 0.001 | 0.001 | 1.555 | 51.46 | 27.869 | 2.278 | 63.24 | 0.855 | -1.732 | 0 | 0 | 1 | 40.938 | 1 | 30 | 1 | GJB_Rinse | 0
145 | P-23730 | 2023/02/25 08:26:20 | 1-31 | 1 | -8.522 | -60.941 | 12638 | -1 | 1 | EV | standard | H2O | 1677339085.54 | 0.221 | 0.809 | 63.045 | 0.001 | 0.002 | 1.547 | 47.446 | 29.091 | 2.168 | 62.21 | 0.855 | -1.17 | 0 | 0 | 1 | 40.938 | 1 | 31 | 1 | GJB_Rinse | 0
146 | P-23730 | 2023/02/25 08:35:16 | 1-31 | 2 | -8.803 | -64.283 | 12262 | -1 | 1 | EV | standard | H2O | 1677339621.33 | 0.209 | 0.863 | 63.292 | 0.002 | 0.002 | 1.567 | 46.696 | 28.2 | 2.1 | 60.771 | 0.855 | -1.167 | 0 | 0 | 1 | 40.938 | 1 | 31 | 1 | GJB_Rinse | 0
147 | P-23730 | 2023/02/25 08:44:11 | 1-31 | 3 | -8.907 | -65.325 | 12204 | -1 | 1 | EV | standard | H2O | 1677340157.08 | 0.194 | 0.808 | 63.937 | 0.001 | 0.004 | 1.58 | 46.523 | 27.982 | 2.081 | 60.795 | 0.855 | -1.09 | 0 | 0 | 1 | 41 | 1 | 31 | 1 | GJB_Rinse | 0
148 | P-23730 | 2023/02/25 08:53:08 | 1-31 | 4 | -8.934 | -65.764 | 12218 | 0 | 1 | EV | standard | H2O | 1677340693.35 | 0.211 | 0.843 | 61.611 | 0.001 | 0.002 | 1.521 | 46.402 | 27.961 | 2.094 | 61.477 | 0.855 | -0.892 | 0 | 0 | 1 | 40.938 | 1 | 31 | 1 | GJB_Rinse | 0
149 | P-23731 | 2023/02/25 09:02:53 | 1-32 | 1 | -2.996 | -32.785 | 12458 | -1 | 1 | TALL5-C-5 | 23-011 | H2O | 1677341278.43 | 0.221 | 0.895 | 60 | 0.001 | 0 | 1.483 | 310.038 | 20.438 | 2.241 | 56.968 | 0.856 | -1.713 | 0 | 0 | 1 | 40.938 | 1 | 32 | 1 | GJB_Rinse | 0
150 | P-23731 | 2023/02/25 09:11:49 | 1-32 | 2 | -2.69 | -29.046 | 12195 | -1 | 1 | TALL5-C-5 | 23-011 | H2O | 1677341814.72 | 0.226 | 0.892 | 62.482 | 0.002 | 0 | 1.548 | 305.673 | 19.193 | 2.215 | 58.288 | 0.857 | -1.335 | 0 | 0 | 1 | 40.938 | 1 | 32 | 1 | GJB_Rinse | 0
151 | P-23731 | 2023/02/25 09:20:44 | 1-32 | 3 | -2.664 | -28.353 | 12183 | -1 | 1 | TALL5-C-5 | 23-011 | H2O | 1677342349.9 | 0.237 | 0.915 | 56.379 | 0.001 | -0.002 | 1.396 | 305.965 | 19.587 | 2.201 | 56.735 | 0.856 | -1.43 | 0 | 0 | 1 | 40.875 | 1 | 32 | 1 | GJB_Rinse | 0
152 | P-23731 | 2023/02/25 09:29:39 | 1-32 | 4 | -2.654 | -28.149 | 12083 | 0 | 1 | TALL5-C-5 | 23-011 | H2O | 1677342885.1 | 0.227 | 0.855 | 58.764 | 0.001 | -0.001 | 1.457 | 304.45 | 18.853 | 2.204 | 58.207 | 0.856 | -1.306 | 0 | 0 | 1 | 40.938 | 1 | 32 | 1 | GJB_Rinse | 0
153 | P-23732 | 2023/02/25 09:39:25 | 1-33 | 1 | -1.063 | -25.141 | 11903 | -1 | 1 | TALL5-D-4 | 23-011 | H2O | 1677343471.22 | 0.204 | 0.863 | 62.692 | 0 | 0.004 | 1.557 | 145.549 | 22.89 | 2.226 | 60.082 | 0.855 | -1.632 | 0 | 0 | 1 | 40.875 | 1 | 33 | 1 | GJB_Rinse | 0
154 | P-23732 | 2023/02/25 09:48:21 | 1-33 | 2 | -0.957 | -24.344 | 11477 | -1 | 1 | TALL5-D-4 | 23-011 | H2O | 1677344007.3 | 0.224 | 0.814 | 56.876 | 0.002 | -0.001 | 1.419 | 140.235 | 21.524 | 2.16 | 59.164 | 0.856 | -1.601 | 0 | 0 | 1 | 40.812 | 1 | 33 | 1 | GJB_Rinse | 0
155 | P-23732 | 2023/02/25 09:57:16 | 1-33 | 3 | -0.958 | -24.352 | 11782 | -1 | 1 | TALL5-D-4 | 23-011 | H2O | 1677344542.62 | 0.209 | 0.8 | 64.21 | 0.001 | 0.001 | 1.596 | 143.022 | 22.352 | 2.213 | 59.766 | 0.856 | -1.497 | 0 | 0 | 1 | 40.938 | 1 | 33 | 1 | GJB_Rinse | 0
156 | P-23732 | 2023/02/25 10:06:12 | 1-33 | 4 | -0.915 | -23.802 | 11748 | 0 | 1 | TALL5-D-4 | 23-011 | H2O | 1677345077.7 | 0.215 | 0.764 | 59.305 | 0 | 0 | 1.477 | 142.86 | 22.164 | 2.197 | 59.323 | 0.856 | -1.711 | 0 | 0 | 1 | 40.938 | 1 | 33 | 1 | GJB_Rinse | 0
157 | P-23733 | 2023/02/25 10:15:58 | 1-34 | 1 | -2.037 | -27.704 | 7812 | -1 | 1 | TALL5-D-2 | 23-011 | H2O | 1677345670.37 | 0.242 | 1.141 | 45.918 | 0 | 0 | 1.308 | 238.825 | 8.536 | 1.48 | 41.157 | 0.855 | -1.11 | 0 | 0 | 1 | 40.875 | 1 | 34 | 1 | GJB_Rinse | 0
158 | P-23733 | 2023/02/25 10:24:54 | 1-34 | 2 | -2.022 | -27.453 | 8784 | -1 | 1 | TALL5-D-2 | 23-011 | H2O | 1677346203.4 | 0.248 | 0.993 | 51.071 | 0.002 | -0.001 | 1.358 | 266.803 | 10.517 | 1.636 | 44.081 | 0.855 | -1.364 | 0 | 0 | 1 | 40.812 | 1 | 34 | 1 | GJB_Rinse | 0
159 | P-23733 | 2023/02/25 10:33:50 | 1-34 | 3 | -2.016 | -27.705 | 9653 | -1 | 1 | TALL5-D-2 | 23-011 | H2O | 1677346737.98 | 0.219 | 1.021 | 52.912 | 0.001 | 0.002 | 1.369 | 291.136 | 12.173 | 1.8 | 47.531 | 0.856 | -1.338 | 0 | 0 | 1 | 40.812 | 1 | 34 | 1 | GJB_Rinse | 0
160 | P-23733 | 2023/02/25 10:42:46 | 1-34 | 4 | -2.018 | -27.758 | 9985 | 0 | 1 | TALL5-D-2 | 23-011 | H2O | 1677347273.21 | 0.227 | 1.071 | 52.333 | 0 | 0.002 | 1.345 | 300.461 | 12.874 | 1.86 | 49.078 | 0.855 | -1.33 | 0 | 0 | 1 | 40.875 | 1 | 34 | 1 | GJB_Rinse | 0
161 | P-23734 | 2023/02/25 10:52:33 | 1-35 | 1 | -2.594 | -24.961 | 11676 | -1 | 1 | TALL5-1020-3 | 23-011 | H2O | 1677347858.82 | 0.198 | 0.847 | 59.648 | 0.001 | 0 | 1.5 | 64.987 | 25.741 | 2.073 | 59.814 | 0.856 | -1.166 | 0 | 0 | 1 | 40.938 | 1 | 35 | 1 | GJB_Rinse | 0
162 | P-23734 | 2023/02/25 11:01:29 | 1-35 | 2 | -2.608 | -24.751 | 11969 | -1 | 1 | TALL5-1020-3 | 23-011 | H2O | 1677348394.95 | 0.192 | 0.838 | 61.198 | 0.001 | 0 | 1.517 | 63.926 | 26.487 | 2.126 | 61.077 | 0.855 | -1.233 | 0 | 0 | 1 | 40.875 | 1 | 35 | 1 | GJB_Rinse | 0
163 | P-23734 | 2023/02/25 11:10:25 | 1-35 | 3 | -2.626 | -24.813 | 12094 | -1 | 1 | TALL5-1020-3 | 23-011 | H2O | 1677348931.01 | 0.205 | 0.808 | 59.247 | 0.001 | -0.002 | 1.462 | 63.839 | 26.82 | 2.144 | 61.709 | 0.855 | -1.203 | 0 | 0 | 1 | 40.875 | 1 | 35 | 1 | GJB_Rinse | 0
164 | P-23734 | 2023/02/25 11:19:21 | 1-35 | 4 | -2.58 | -24.328 | 12130 | 0 | 1 | TALL5-1020-3 | 23-011 | H2O | 1677349466.69 | 0.208 | 0.825 | 63.008 | 0.001 | 0.001 | 1.557 | 63.465 | 27.257 | 2.162 | 61.702 | 0.856 | -1.093 | 0 | 0 | 1 | 40.938 | 1 | 35 | 1 | GJB_Rinse | 0
165 | P-23735 | 2023/02/25 11:29:04 | 1-36 | 1 | -2.574 | -23.717 | 11812 | -1 | 1 | TALL5-010-3 | 23-011 | H2O | 1677350049.6 | 0.195 | 0.79 | 57.573 | 0.001 | -0.002 | 1.438 | 60.483 | 26.254 | 2.105 | 59.638 | 0.856 | -1.144 | 0 | 0 | 1 | 41 | 1 | 36 | 1 | GJB_Rinse | 0
166 | P-23735 | 2023/02/25 11:37:58 | 1-36 | 2 | -2.544 | -23.984 | 11870 | -1 | 1 | TALL5-010-3 | 23-011 | H2O | 1677350584.29 | 0.203 | 0.838 | 61.527 | 0.001 | -0.001 | 1.532 | 60.712 | 26.46 | 2.115 | 60.336 | 0.855 | -1.152 | 0 | 0 | 1 | 41 | 1 | 36 | 1 | GJB_Rinse | 0
167 | P-23735 | 2023/02/25 11:46:55 | 1-36 | 3 | -2.566 | -23.83 | 11921 | -1 | 1 | TALL5-010-3 | 23-011 | H2O | 1677351120.38 | 0.205 | 0.761 | 60.167 | 0 | 0 | 1.505 | 60.947 | 26.737 | 2.123 | 59.739 | 0.855 | -1.331 | 0 | 0 | 1 | 41.062 | 1 | 36 | 1 | GJB_Rinse | 0
168 | P-23735 | 2023/02/25 11:55:50 | 1-36 | 4 | -2.599 | -23.636 | 11964 | 0 | 1 | TALL5-010-3 | 23-011 | H2O | 1677351656.1 | 0.205 | 0.887 | 60.963 | 0.001 | -0.002 | 1.514 | 60.949 | 26.473 | 2.133 | 61.249 | 0.856 | -1.164 | 0 | 0 | 1 | 41 | 1 | 36 | 1 | GJB_Rinse | 0
169 | P-23736 | 2023/02/25 12:05:33 | 1-37 | 1 | -1.849 | -25.372 | 11398 | -1 | 1 | TALL5-D-5 | 23-011 | H2O | 1677352239.32 | 0.211 | 0.708 | 53.435 | 0.001 | -0.002 | 1.337 | 55.868 | 25.02 | 2.043 | 59.908 | 0.855 | -0.997 | 0 | 0 | 1 | 41.125 | 1 | 37 | 1 | GJB_Rinse | 0
170 | P-23736 | 2023/02/25 12:14:29 | 1-37 | 2 | -1.784 | -25.751 | 11190 | -1 | 1 | TALL5-D-5 | 23-011 | H2O | 1677352775.1 | 0.193 | 0.94 | 62.111 | 0.002 | -0.001 | 1.556 | 55.333 | 24.501 | 1.989 | 58.429 | 0.856 | -1.075 | 0 | 0 | 1 | 41.188 | 1 | 37 | 1 | GJB_Rinse | 0
171 | P-23736 | 2023/02/25 12:23:25 | 1-37 | 3 | -1.811 | -25.621 | 11093 | -1 | 1 | TALL5-D-5 | 23-011 | H2O | 1677353310.28 | 0.202 | 0.831 | 60.172 | 0.001 | 0.002 | 1.516 | 55.074 | 24.153 | 1.984 | 58.085 | 0.855 | -0.96 | 0 | 0 | 1 | 41.188 | 1 | 37 | 1 | GJB_Rinse | 0
172 | P-23736 | 2023/02/25 12:32:21 | 1-37 | 4 | -1.783 | -25.395 | 11170 | 0 | 1 | TALL5-D-5 | 23-011 | H2O | 1677353846.31 | 0.217 | 0.848 | 57.133 | 0.001 | -0.003 | 1.436 | 55.106 | 24.639 | 1.992 | 58.286 | 0.855 | -0.953 | 0 | 0 | 1 | 41.188 | 1 | 37 | 1 | GJB_Rinse | 0
173 | P-23737 | 2023/02/25 12:42:04 | 1-38 | 1 | -1.024 | -24.57 | 12425 | -1 | 1 | TALL5-B-2 | 23-011 | H2O | 1677354429.44 | 0.223 | 1.079 | 62.405 | 0.001 | 0.001 | 1.537 | 265.782 | 20.289 | 2.349 | 59.031 | 0.856 | -1.964 | 0 | 0 | 1 | 41.25 | 1 | 38 | 1 | GJB_Rinse | 0
174 | P-23737 | 2023/02/25 12:51:00 | 1-38 | 2 | -1.086 | -24.689 | 12227 | -1 | 1 | TALL5-B-2 | 23-011 | H2O | 1677354965.17 | 0.214 | 0.779 | 62.702 | 0.001 | 0.001 | 1.55 | 264.16 | 19.449 | 2.325 | 59.367 | 0.856 | -2.17 | 0 | 0 | 1 | 41.25 | 1 | 38 | 1 | GJB_Rinse | 0
175 | P-23737 | 2023/02/25 12:59:55 | 1-38 | 3 | -1.007 | -24.311 | 12257 | -1 | 1 | TALL5-B-2 | 23-011 | H2O | 1677355500.29 | 0.236 | 0.845 | 58.314 | 0.001 | 0.001 | 1.437 | 264.768 | 19.723 | 2.325 | 59.55 | 0.856 | -2.178 | 0 | 0 | 1 | 41.312 | 1 | 38 | 1 | GJB_Rinse | 0
176 | P-23737 | 2023/02/25 13:08:50 | 1-38 | 4 | -1.07 | -24.663 | 12167 | 0 | 1 | TALL5-B-2 | 23-011 | H2O | 1677356036.41 | 0.238 | 0.902 | 57.224 | 0.001 | 0 | 1.421 | 263.424 | 19.381 | 2.311 | 59.25 | 0.856 | -1.979 | 0 | 0 | 1 | 41.438 | 1 | 38 | 1 | GJB_Rinse | 0
177 | P-23738 | 2023/02/25 13:18:35 | 1-39 | 1 | -2.032 | -25.313 | 12549 | -1 | 1 | TALL5-C-1 | 23-011 | H2O | 1677356620.46 | 0.209 | 0.842 | 64.608 | 0.002 | 0.006 | 1.591 | 157.127 | 24.352 | 2.326 | 62.226 | 0.857 | -1.714 | 0 | 0 | 1 | 41.438 | 1 | 39 | 1 | GJB_Rinse | 0
178 | P-23738 | 2023/02/25 13:27:31 | 1-39 | 2 | -2.091 | -25.401 | 12345 | -1 | 1 | TALL5-C-1 | 23-011 | H2O | 1677357156.47 | 0.199 | 0.936 | 63.761 | 0 | 0.005 | 1.576 | 154.189 | 23.481 | 2.315 | 62.265 | 0.856 | -1.674 | 0 | 0 | 1 | 41.438 | 1 | 39 | 1 | GJB_Rinse | 0
179 | P-23738 | 2023/02/25 13:36:27 | 1-39 | 3 | -2.068 | -25.278 | 12393 | -1 | 1 | TALL5-C-1 | 23-011 | H2O | 1677357691.95 | 0.211 | 0.774 | 64.051 | 0 | 0.003 | 1.593 | 154.451 | 23.909 | 2.302 | 61.298 | 0.856 | -1.797 | 0 | 0 | 1 | 41.5 | 1 | 39 | 1 | GJB_Rinse | 0
180 | P-23738 | 2023/02/25 13:45:22 | 1-39 | 4 | -2.108 | -25.699 | 12226 | 0 | 1 | TALL5-C-1 | 23-011 | H2O | 1677358227.49 | 0.202 | 0.938 | 65.405 | 0 | 0.003 | 1.628 | 152.808 | 23.485 | 2.269 | 60.215 | 0.856 | -1.828 | 0 | 0 | 1 | 41.5 | 1 | 39 | 1 | GJB_Rinse | 0
181 | P-23739 | 2023/02/25 13:55:06 | 1-40 | 1 | -8.581 | -61.022 | 12712 | -1 | 1 | EV | standard | H2O | 1677358811.16 | 0.207 | 0.822 | 60.942 | 0.001 | -0.002 | 1.502 | 49.11 | 29.386 | 2.175 | 62.922 | 0.856 | -1.097 | 0 | 0 | 1 | 41.5 | 1 | 40 | 1 | GJB_Rinse | 0
182 | P-23739 | 2023/02/25 14:04:01 | 1-40 | 2 | -8.915 | -64.822 | 12232 | -1 | 1 | EV | standard | H2O | 1677359347.08 | 0.2 | 0.76 | 61.985 | 0.001 | 0.003 | 1.541 | 47.13 | 28.192 | 2.087 | 61.26 | 0.856 | -0.953 | 0 | 0 | 1 | 41.562 | 1 | 40 | 1 | GJB_Rinse | 0
183 | P-23739 | 2023/02/25 14:12:57 | 1-40 | 3 | -8.95 | -65.485 | 12365 | -1 | 1 | EV | standard | H2O | 1677359882.41 | 0.208 | 0.651 | 63.134 | 0.001 | 0.004 | 1.559 | 47.164 | 28.327 | 2.119 | 62.802 | 0.856 | -0.9 | 0 | 0 | 1 | 41.625 | 1 | 40 | 1 | GJB_Rinse | 0
184 | P-23739 | 2023/02/25 14:21:52 | 1-40 | 4 | -8.896 | -65.639 | 12330 | 0 | 1 | EV | standard | H2O | 1677360417.45 | 0.193 | 0.777 | 60.939 | 0.001 | 0.004 | 1.504 | 47.007 | 28.25 | 2.11 | 61.987 | 0.856 | -0.915 | 0 | 0 | 1 | 41.562 | 1 | 40 | 1 | GJB_Rinse | 0
185 | P-23740 | 2023/02/25 14:31:36 | 1-41 | 1 | -2.281 | -29.773 | 12429 | -1 | 1 | TALL5-D-3 | 23-011 | H2O | 1677361000.97 | 0.224 | 0.971 | 60.517 | 0.001 | 0 | 1.497 | 333.933 | 18.269 | 2.295 | 59.642 | 0.856 | -1.503 | 0 | 0 | 1 | 41.625 | 1 | 41 | 1 | GJB_Rinse | 0
186 | P-23740 | 2023/02/25 14:40:31 | 1-41 | 2 | -2.082 | -27.029 | 12195 | -1 | 1 | TALL5-D-3 | 23-011 | H2O | 1677361537.27 | 0.229 | 0.956 | 60.498 | 0 | -0.002 | 1.491 | 330.394 | 18.325 | 2.263 | 57.838 | 0.857 | -1.566 | 0 | 0 | 1 | 41.625 | 1 | 41 | 1 | GJB_Rinse | 0
187 | P-23740 | 2023/02/25 14:49:26 | 1-41 | 3 | -1.953 | -25.633 | 12074 | -1 | 1 | TALL5-D-3 | 23-011 | H2O | 1677362071.3 | 0.244 | 0.861 | 59.92 | 0.001 | -0.003 | 1.481 | 328.157 | 18.057 | 2.222 | 56.9 | 0.857 | -1.752 | 0 | 0 | 1 | 41.688 | 1 | 41 | 1 | GJB_Rinse | 0
188 | P-23740 | 2023/02/25 14:58:21 | 1-41 | 4 | -2.116 | -26.606 | 12130 | 0 | 1 | TALL5-D-3 | 23-011 | H2O | 1677362606.61 | 0.245 | 0.911 | 58.568 | 0.001 | -0.001 | 1.451 | 332.685 | 17.962 | 2.246 | 57.62 | 0.857 | -1.582 | 0 | 0 | 1 | 41.75 | 1 | 41 | 1 | GJB_Rinse | 0
189 | P-23741 | 2023/02/25 15:08:06 | 1-42 | 1 | -8.607 | -61.561 | 12681 | -1 | 1 | EV | standard | H2O | 1677363191.71 | 0.197 | 0.8 | 61.725 | 0.001 | -0.003 | 1.525 | 50.977 | 29.113 | 2.181 | 63.596 | 0.856 | -0.993 | 0 | 0 | 1 | 41.75 | 1 | 42 | 1 | GJB_Rinse | 0
190 | P-23741 | 2023/02/25 15:17:02 | 1-42 | 2 | -8.825 | -64.595 | 12355 | -1 | 1 | EV | standard | H2O | 1677363727.47 | 0.205 | 0.84 | 61.27 | 0.001 | 0 | 1.525 | 48.161 | 28.214 | 2.11 | 61.801 | 0.855 | -0.937 | 0 | 0 | 1 | 41.75 | 1 | 42 | 1 | GJB_Rinse | 0
191 | P-23741 | 2023/02/25 15:25:57 | 1-42 | 3 | -8.926 | -65.44 | 12340 | -1 | 1 | EV | standard | H2O | 1677364262.31 | 0.204 | 0.831 | 60.79 | 0.002 | 0.001 | 1.497 | 47.573 | 27.978 | 2.112 | 62.973 | 0.856 | -0.891 | 0 | 0 | 1 | 41.812 | 1 | 42 | 1 | GJB_Rinse | 0
192 | P-23741 | 2023/02/25 15:34:53 | 1-42 | 4 | -9.011 | -66.01 | 12200 | 0 | 1 | EV | standard | H2O | 1677364798.15 | 0.205 | 0.812 | 59.668 | 0.001 | 0.001 | 1.473 | 46.877 | 27.661 | 2.093 | 62.375 | 0.856 | -0.78 | 0 | 0 | 1 | 41.812 | 1 | 42 | 1 | GJB_Rinse | 0

## Instructions
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Variable | Instruction | Comment
StandardsFile | standards file.csv
OutputBase | ChemCorrect
CorrectionType | M
timecode | Time Code
identifier | Identifier 1
sample | Sample
injection | Inj Nr | Inj Nr
line | Line
job | Job
Variable | Instruction | Comment
baseline_shift_MEAN | mean_standards(baseline_shift) | The mean of the organic baseline of the standards.
baseline_shift_STD | std_standards(baseline_shift) | The standard deviation of the organic baseline of the standards.
ch4_ppm_MEAN | mean_standards(ch4_ppm) | The mean methane concentration of the standards.
ch4_ppm_STD | std_standards(ch4_ppm) | The standard deviation of the methane concentrations of the standards.
residuals_MEAN | mean_standards(residuals) | The mean difference between the measured and modeled spectra of the standards.
residuals_STD | std_standards(residuals) | The standard deviation of the difference between measured and modeled spectra of the standards.
slope_shift_MEAN | mean_standards(slope_shift) | The mean organic slope of the standards.
slope_shift_STD | std_standards(slope_shift) | The standard deviation of the organic slope of the standards.
baseline_curvature_MEAN | mean_standards(baseline_curvature) | The mean of the baseline curvature of the standards.
baseline_curvature_STD | std_standards(baseline_curvature) | The standard deviation of the baseline curvature of the standards.
residuals_MEAN_ABS | abs(residuals_MEAN) | The absolute value of the mean difference between the measured and modeled spectra of the standards.
baseline_shift_MEAN_ABS | abs(baseline_shift_MEAN) | The absolute value of the mean of the organic baseline of the standards.
ch4_ppm_MEAN_ABS | abs(ch4_ppm_MEAN) | The absolute value of the mean methane concentration of the standards.
slope_shift_MEAN_ABS | abs(slope_shift_MEAN) | The absolute value of the mean organic slope of the standards.
residuals_15SIGMA | evaluation(residuals_STD*1.5) | One and half times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_15SIGMA | evaluation(residuals_MEAN+residuals_15SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_2SIGMA | evaluation(residuals_STD*2) | Twice the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_2SIGMA | evaluation(residuals_MEAN+residuals_2SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_3SIGMA | evaluation(residuals_STD*3) | Three times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_3SIGMA | evaluation(residuals_MEAN+residuals_3SIGMA) | The mean difference between measured and modeled spectra of the standards plus three standard deviation.
residuals_10PCT | evaluation(residuals_MEAN*0.1) | 30% of the difference between measured and modeled spectra of the standards.
ch4_ppm_3SIGMA | evaluation(ch4_ppm_STD*3) | Three times the standard deviation of the methanol amplitudes of the standards.
ch4_ppm_PLUS_3SIGMA | evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA) | The mean methanol amplitude of the standards plus three standard deviation.
HALF | float(0.5) | A fixed value of 0.5 or half.
baseline_shift_HALFSIGMA | evaluation(baseline_shift_STD*HALF) | Half the standard deviation of the organic baseline of the standards.
baseline_shift_PLUS_HALFSIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA) | The absolute value of the mean organic baseline of the standards plus one-half standard deviation.
baseline_shift_PLUS_1SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD) | The absolute value of the mean organic baseline of the standards plus one standard deviation.
baseline_shift_3SIGMA | evaluation(baseline_shift_STD*3) | Three times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_3SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA) | The absolute value of the mean organic baseline of the standards plus three standard deviation.
baseline_shift_18SIGMA | evaluation(baseline_shift_STD*18) | Eighteen times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_18SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA) | The absolute value of the mean organic baseline of the standards plus eighteen standard deviation.
slope_shift_5SIGMA | evaluation(slope_shift_STD*5) | Five times the standard deviation of the organic slope of the standards.
slope_shift_PLUS_5SIGMA | evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA) | The absolute value of the mean organic slope of the standards plus five standard deviation.
baseline_curvature_3SIGMA | evaluation(baseline_curvature_STD*3) | Three times the standard deviation of the baseline curvature of the standards.
baseline_curvature_PLUS_3SIGMA | evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
TRUE | evaluation(1>0)
FALSE | evaluation(1<0)
PROC_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the uncalibrated d18O value in order to calibrate it.
PROC_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the uncalibrated dD value in order to calibrate it.
Variable | Instruction | Comment
IS_STANDARD | is_standard_sample() | Standard recognition.
IS_NOT_STANDARD | bool_xor(IS_STANDARD,TRUE)
SAMP_residuals_MEAN | mean_all(residuals) | The mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_residuals_STD | std_all(residuals) | One standard deviation of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN | mean_all(ch4_ppm) | The mean methane concentration of a sample.
SAMP_ch4_ppm_STD | std_all(ch4_ppm) | The standard deviation of the methane concentrations of a sample.
SAMP_baseline_shift_MEAN | mean_all(baseline_shift) | The mean of the organic baseline of a sample.
SAMP_baseline_shift_STD | std_all(baseline_shift) | The standard deviation of the organic baseline of a sample.
SAMP_slope_shift_MEAN | mean_all(slope_shift) | The mean organic slope of a sample.
SAMP_slope_shift_STD | std_all(slope_shift) | The standard deviation of the organic slope of a sample.
SAMP_d18O_STD | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
SAMP_dD_STD | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_baseline_curvature_MEAN | mean_all(baseline_curvature) | The mean of the baseline curvature of the ssamples
SAMP_baseline_curvature_STD | std_all(baseline_curvature) | The standard deviation of the baseline curvature of the samples
SAMP_residuals_MEAN_ABS | abs(SAMP_residuals_MEAN) | Absolute value of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN_ABS | abs(SAMP_ch4_ppm_MEAN) | Absolute value of the mean methane concentration of a sample.
SAMP_baseline_shift_MEAN_ABS | abs(SAMP_baseline_shift_MEAN) | Absolute value of the mean organic baseline of a sample.
SAMP_slope_shift_MEAN_ABS | abs(SAMP_slope_shift_MEAN) | Absolute value of the mean organic slope of a sample.
SAMP_residuals_DIFFERENTIAL | evaluation(SAMP_residuals_MEAN-residuals_MEAN)
SAMP_residuals_DIFFERENTIAL_ABS | abs(SAMP_residuals_DIFFERENTIAL) | Absolute value of the difference between the mean of the sample residual and the mean of all standards residuals
SAMP_residuals_DIFFERENTIAL_PCT | evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
SAMP_baseline_curvature_DIFFERENTIAL | evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
SAMP_STD_DELTA_18_16O_CURR_MEAN | current_standards_mean(d(18_16)Mean) | The d18O mean of the last run standard.
SAMP_STD_DELTA_DH_CURR_MEAN | current_standards_mean(d(D_H)Mean) | The dD mean of the last run standard.
known_dD | known_standard(d(D_H)Mean) | The dD value of the known Standard.
known_d18O | known_standard(d(18_16)Mean) | The d18O value of the known Standard.
d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
measured_d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
measured_dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
measured_d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
measured_dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the mean uncalibrated d18O value in order to calibrate it.
SAMP_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the mean uncalibrated dD value in order to calibrate it.
SAMP_TEST_VALUE | evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
residuals_GT_3SIGMA | evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA) | Threshold of red trigger. If the mean difference between the measured spectra of the sample and the modeled spectra of the standard is greater than that of the mean difference of the standard plus three standard deviations, this will trigger.
residuals_GE_2SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GE_15SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GT_10PCT | evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
ch4_ppm_GE_3SIGMA | evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA) | True if the mean methane concentration of the sample is greater than or equal to that of the standard plus three standard deviations, otherwise false.
baseline_shift_GT_1SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus one standard deviation, otherwise false.
baseline_shift_GT_3SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus three standard deviation, otherwise false. Sets yellow trigger for "other", trigger threshold is approximately sample to sample dD precision specification
baseline_shift_GT_18SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus eighteen standard deviation, otherwise false. Sets red trigger for "other", trigger threshold is approximately 24 hour dD drift specification
slope_shift_GT_5SIGMA | evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA) | True if the mean organic slope of the sample is greater than that of the standard plus five standard deviations, otherwise false.
baseline_curvature_GT_3SIGMA | evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
RED | truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA)) | Red trigger. If the sample is not a standard and the threshold is triggered (measured-modeled > standard+3stdev's), then RED.
NOT_RED | bool_xor(RED,TRUE) | If the sample is determined to be "red" than it can not be "not red."
YELLOW | truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA))) | Truth evaluation used to determine whether or not a given sample is yellow.
NOT_YELLOW | bool_xor(YELLOW,TRUE) | If sample is determined to be "yellow" than it can not be "not yellow."
GREEN | truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW)) | This is the pre-requisite for a sample to be colored green. It must not be a standard, must not be yellow, and must not be red.
NOT_GREEN | bool_xor(GREEN,TRUE) | Pre-requisite for being "not-green."
CYAN | truth_evaluation((IS_STANDARD)) | This is the pre-requisite for a sample to be a standard and therefore cyan colored.
NOT_CYAN | bool_xor(CYAN,TRUE) | Pre-requisite for being not a standard.
METHANE | truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA)) | This is the threshold for amounts of methane being included in the main ChemCorrect spreadsheet. Methane will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured methane concentration must be greater than or equal to that of the standard plus three standard deviations.
OTHER1 | truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA)) | This is the threshold for amounts of "OTHER1" being included in the main ChemCorrect spreadsheet. This column (OTHER1) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic baseline must be greater than that of the standard plus one standard deviation.
OTHER2 | truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA)) | This is the threshold for amounts of "OTHER2" being included in the main ChemCorrect spreadsheet. This column (OTHER2) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic slope must be greater than that of the standard plus three standard deviations.
OTHER3 | truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | truth_evaluation(RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN)) | Organic baseline differential showing the difference between the measured baseline and the standard baseline.
Variable | Instruction | Comment
DTL_TEST | math_evaluation((1 + 2)) | Test evaluation of 1+2 being equal to 3.
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample | FALSE | FALSE | FALSE | FALSE | Testing Sample
B | identifier | TRUE | Name | RED | YELLOW | GREEN | CYAN | Testing Name
C | SAMP_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d<sup>18</sup>O Mean | RED | YELLOW | GREEN | CYAN | Testing18 16 Mean
D | SAMP_DELTA_DH_CALIBRATE | TRUE | Calibrated d<sup>2</sup>H Mean | RED | YELLOW | GREEN | CYAN | This is a test to see that I can use tool tips
E | SAMP_ch4_ppm_MEAN | METHANE | CH<sub>4</sub> | RED | YELLOW | FALSE | FALSE
F | baseline_shift_DIFF | OTHER1 | C<sub>2+</sub> alcohols | RED | YELLOW | FALSE | FALSE
G | SAMP_residuals_DIFFERENTIAL_PCT | OTHER3 | relative %deviation | RED | YELLOW | FALSE | FALSE
H | SAMP_d18O_STD | TRUE | Uncalibrated d<sup>18</sup>O precision | RED | YELLOW | GREEN | CYAN
I | SAMP_dD_STD | TRUE | Uncalibrated d<sup>2</sup>H precision | RED | YELLOW | GREEN | CYAN
J | SAMP_slope_shift_MEAN | OTHER2 | slope | RED | YELLOW | FALSE | FALSE
K | SAMP_baseline_curvature_DIFFERENTIAL | OTHER4 | curvature | RED | YELLOW | FALSE | FALSE
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample
B | identifier | TRUE | Name
C | col(d(18_16)Mean) | TRUE | Measured d(18_16) Mean
D | col(d(D_H)Mean) | TRUE | Measured d(D_H) Mean
E | DTL_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d(18_16) Mean
F | DTL_DELTA_DH_CALIBRATE | TRUE | Calibrated d(D_H) Mean
G | col(ch4_ppm) | METHANE | CH<sub>4</sub>
H | col(baseline_shift) | OTHER1 | Other
I | col(Time Code) | TRUE | Time Code
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | d18O_mean | TRUE | d(18_16) Mean | RED | YELLOW | GREEN | CYAN
B | dD_mean | TRUE | d(D_H) Mean | RED | YELLOW | GREEN | CYAN
C | XY_X(d18O_mean) XY_Y(dD_mean) | TRUE | d18O x dD | gmwl | best_fit

## Process Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Process Variables
Variable | Value | Operations
baseline_shift_MEAN | 48.3301
varray: [47.956, 49.524, 49.904, 49.833, 49.505, 48.915, 48.893, 49.086, 48.77, 48.551, 47.968, 47.642, 47.529, 48.072, 47.744, 47.762, 47.219, 47.43, 47.04, 47.259]
inst: mean_standards
operation: mean_standards(baseline_shift)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0, 10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: baseline_shift
baseline_shift_STD | 0.917371181262
varray: [47.956, 49.524, 49.904, 49.833, 49.505, 48.915, 48.893, 49.086, 48.77, 48.551, 47.968, 47.642, 47.529, 48.072, 47.744, 47.762, 47.219, 47.43, 47.04, 47.259]
inst: std_standards
operation: std_standards(baseline_shift)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0, 10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: baseline_shift
ch4_ppm_MEAN | -0.99395
varray: [-1.084, -1.014, -1.06, -1.173, -1.091, -0.918, -1.16, -1.257, -1.104, -0.982, -1.173, -0.825, -0.782, -0.95, -0.947, -0.878, -0.832, -0.906, -0.827, -0.916]
inst: mean_standards
operation: mean_standards(ch4_ppm)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0, 10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: ch4_ppm
ch4_ppm_STD | 0.137679824995
varray: [-1.084, -1.014, -1.06, -1.173, -1.091, -0.918, -1.16, -1.257, -1.104, -0.982, -1.173, -0.825, -0.782, -0.95, -0.947, -0.878, -0.832, -0.906, -0.827, -0.916]
inst: std_standards
operation: std_standards(ch4_ppm)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0, 10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: ch4_ppm
residuals_MEAN | 2.16385
varray: [2.274, 2.244, 2.274, 2.261, 2.261, 2.231, 2.257, 2.258, 2.275, 2.281, 2.134, 2.043, 2.062, 2.09, 2.064, 2.055, 2.033, 2.066, 2.057, 2.057]
inst: mean_standards
operation: mean_standards(residuals)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0, 10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: residuals
residuals_STD | 0.102691810261
varray: [2.274, 2.244, 2.274, 2.261, 2.261, 2.231, 2.257, 2.258, 2.275, 2.281, 2.134, 2.043, 2.062, 2.09, 2.064, 2.055, 2.033, 2.066, 2.057, 2.057]
inst: std_standards
operation: std_standards(residuals)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0, 10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: residuals
slope_shift_MEAN | 28.64355
varray: [29.464, 28.407, 28.442, 28.707, 28.683, 27.986, 28.938, 28.879, 28.805, 28.656, 29.719, 28.336, 28.402, 28.66, 28.535, 28.241, 28.474, 28.744, 28.064, 28.729]
inst: mean_standards
operation: mean_standards(slope_shift)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0, 10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: slope_shift
slope_shift_STD | 0.413054918612
varray: [29.464, 28.407, 28.442, 28.707, 28.683, 27.986, 28.938, 28.879, 28.805, 28.656, 29.719, 28.336, 28.402, 28.66, 28.535, 28.241, 28.474, 28.744, 28.064, 28.729]
inst: std_standards
operation: std_standards(slope_shift)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0, 10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: slope_shift
baseline_curvature_MEAN | 62.8568
varray: [64.517, 64.052, 65.537, 63.601, 64.274, 64.034, 64.273, 64.044, 65.086, 64.921, 62.049, 60.708, 61.468, 61.372, 61.322, 61.571, 60.287, 61.086, 61.354, 61.58]
inst: mean_standards
operation: mean_standards(baseline_curvature)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0, 10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: baseline_curvature
baseline_curvature_STD | 1.70142434448
varray: [64.517, 64.052, 65.537, 63.601, 64.274, 64.034, 64.273, 64.044, 65.086, 64.921, 62.049, 60.708, 61.468, 61.372, 61.322, 61.571, 60.287, 61.086, 61.354, 61.58]
inst: std_standards
operation: std_standards(baseline_curvature)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0, 10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: baseline_curvature
residuals_MEAN_ABS | 2.16385
operation: abs(residuals_MEAN)
inst: abs
source_col: residuals_MEAN
baseline_shift_MEAN_ABS | 48.3301
operation: abs(baseline_shift_MEAN)
inst: abs
source_col: baseline_shift_MEAN
ch4_ppm_MEAN_ABS | 0.99395
operation: abs(ch4_ppm_MEAN)
inst: abs
source_col: ch4_ppm_MEAN
slope_shift_MEAN_ABS | 28.64355
operation: abs(slope_shift_MEAN)
inst: abs
source_col: slope_shift_MEAN
residuals_15SIGMA | 0.154037715392
operation: evaluation(residuals_STD*1.5)
inst: evaluation
factors: residuals_STD*1.5
residuals_PLUS_15SIGMA | 2.31788771539
operation: evaluation(residuals_MEAN+residuals_15SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_15SIGMA
residuals_2SIGMA | 0.205383620523
operation: evaluation(residuals_STD*2)
inst: evaluation
factors: residuals_STD*2
residuals_PLUS_2SIGMA | 2.36923362052
operation: evaluation(residuals_MEAN+residuals_2SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_2SIGMA
residuals_3SIGMA | 0.308075430784
operation: evaluation(residuals_STD*3)
inst: evaluation
factors: residuals_STD*3
residuals_PLUS_3SIGMA | 2.47192543078
operation: evaluation(residuals_MEAN+residuals_3SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_3SIGMA
residuals_10PCT | 0.216385
operation: evaluation(residuals_MEAN*0.1)
inst: evaluation
factors: residuals_MEAN*0.1
ch4_ppm_3SIGMA | 0.413039474984
operation: evaluation(ch4_ppm_STD*3)
inst: evaluation
factors: ch4_ppm_STD*3
ch4_ppm_PLUS_3SIGMA | 1.40698947498
operation: evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA)
inst: evaluation
factors: ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA
HALF | 0.5
operation: float(0.5)
inst: float
baseline_shift_HALFSIGMA | 0.458685590631
operation: evaluation(baseline_shift_STD*HALF)
inst: evaluation
factors: baseline_shift_STD*HALF
baseline_shift_PLUS_HALFSIGMA | 48.7887855906
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA
baseline_shift_PLUS_1SIGMA | 49.2474711813
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_STD
baseline_shift_3SIGMA | 2.75211354379
operation: evaluation(baseline_shift_STD*3)
inst: evaluation
factors: baseline_shift_STD*3
baseline_shift_PLUS_3SIGMA | 51.0822135438
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_3SIGMA
baseline_shift_18SIGMA | 16.5126812627
operation: evaluation(baseline_shift_STD*18)
inst: evaluation
factors: baseline_shift_STD*18
baseline_shift_PLUS_18SIGMA | 64.8427812627
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_18SIGMA
slope_shift_5SIGMA | 2.06527459306
operation: evaluation(slope_shift_STD*5)
inst: evaluation
factors: slope_shift_STD*5
slope_shift_PLUS_5SIGMA | 30.7088245931
operation: evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA)
inst: evaluation
factors: slope_shift_MEAN_ABS+slope_shift_5SIGMA
baseline_curvature_3SIGMA | 5.10427303345
operation: evaluation(baseline_curvature_STD*3)
inst: evaluation
factors: baseline_curvature_STD*3
baseline_curvature_PLUS_3SIGMA | 67.9610730334
operation: evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
inst: evaluation
factors: baseline_curvature_3SIGMA+baseline_curvature_MEAN
TRUE | True
operation: evaluation(1>0)
inst: evaluation
factors: 1>0
FALSE | False
operation: evaluation(1<0)
inst: evaluation
factors: 1<0
PROC_DELTA_18_16O_CALIBRATE | -4.74563877449
actual_array: [ 1.93 -15.84]
varray: [-9.604, -9.214, -9.081, -8.931, -8.891, -8.877, -8.858, -8.891, -8.864, -8.849, 2.157, 2.644, 2.812, 2.813, 2.909, 2.857, 2.916, 2.932, 2.946, 2.966, -13.531, -14.234, -14.445, -14.5, -14.536, -14.597, -14.551, -14.615, -14.621, -14.634, -9.367, -9.2, -9.154, -9.139, -9.09, -9.019, -9.064, -9.032, -9.03, -9.048, -3.936, -3.704, -3.658, -3.581, -2.132, -1.99, -2.008, -1.998, -2.404, -2.485, -2.458, -2.443, -3.089, -3.12, -3.122, -3.121, -3.719, -3.712, -3.747, -3.708, -2.803, -2.779, -2.784, -2.841, -2.56, -2.585, -2.635, -2.552, -3.102, -3.131, -3.13, -3.173, -8.643, -8.905, -8.886, -8.976, -2.565, -2.228, -2.195, -2.192, -3.979, -4.016, -4.086, -4.103, 0.236, 0.3, 0.249, 0.307, -0.932, -0.873, -0.885, -0.911, -2.015, -2.097, -2.005, -2.047, 2.784, 2.914, 2.893, 2.941, 0.609, 0.544, 0.592, 0.59, -1.705, -1.655, -1.673, -1.66, -8.562, -8.842, -8.877, -8.93, -1.316, -1.061, -0.93, -0.94, 1.821, 1.843, 1.847, 1.793, -2.52, -2.519, -2.587, -2.567, -2.683, -2.758, -2.742, -2.737, -1.536, -1.485, -1.435, -1.469, -0.945, -0.913, -0.912, -0.959, -0.86, -0.797, -0.813, -0.845, -1.309, -1.391, -1.401, -1.451, -8.522, -8.803, -8.907, -8.934, -2.996, -2.69, -2.664, -2.654, -1.063, -0.957, -0.958, -0.915, -2.037, -2.022, -2.016, -2.018, -2.594, -2.608, -2.626, -2.58, -2.574, -2.544, -2.566, -2.599, -1.849, -1.784, -1.811, -1.783, -1.024, -1.086, -1.007, -1.07, -2.032, -2.091, -2.068, -2.108, -8.581, -8.915, -8.95, -8.896, -2.281, -2.082, -1.953, -2.116, -8.607, -8.825, -8.926, -9.011]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -3.674421875
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3758.0, 4295.0, 4831.0, 5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0, 10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0, 16252.0, 16789.0, 17324.0, 17861.0, 18397.0, 18933.0, 19470.0, 20006.0, 20543.0, 21078.0, 21667.0, 22203.0, 22739.0, 23274.0, 23864.0, 24401.0, 24937.0, 25474.0, 26063.0, 26600.0, 27136.0, 27673.0, 28259.0, 28795.0, 29330.0, 29866.0, 30455.0, 30991.0, 31528.0, 32064.0, 32651.0, 33188.0, 33724.0, 34260.0, 34848.0, 35384.0, 35919.0, 36456.0, 37043.0, 37579.0, 38115.0, 38651.0, 39240.0, 39777.0, 40313.0, 40849.0, 41439.0, 41975.0, 42512.0, 43048.0, 43634.0, 44171.0, 44707.0, 45243.0, 45830.0, 46365.0, 46900.0, 47436.0, 48023.0, 48559.0, 49095.0, 49631.0, 50218.0, 50754.0, 51290.0, 51827.0, 52415.0, 52951.0, 53487.0, 54023.0, 54611.0, 55148.0, 55684.0, 56219.0, 56808.0, 57344.0, 57880.0, 58416.0, 59001.0, 59537.0, 60073.0, 60608.0, 61195.0, 61732.0, 62268.0, 62803.0, 63390.0, 63925.0, 64461.0, 64996.0, 65583.0, 66120.0, 66655.0, 67191.0, 67778.0, 68315.0, 68850.0, 69385.0, 69972.0, 70508.0, 71044.0, 71580.0, 72168.0, 72703.0, 73238.0, 73773.0, 74357.0, 74893.0, 75429.0, 75965.0, 76550.0, 77085.0, 77620.0, 78157.0, 78742.0, 79278.0, 79813.0, 80350.0, 80935.0, 81471.0, 82006.0, 82541.0, 83127.0, 83663.0, 84198.0, 84734.0, 85320.0, 85856.0, 86392.0, 86928.0, 87515.0, 88051.0, 88587.0, 89123.0, 89706.0, 90240.0, 90777.0, 91312.0, 91895.0, 92431.0, 92967.0, 93503.0, 94086.0, 94622.0, 95157.0, 95692.0, 96277.0, 96813.0, 97349.0, 97884.0, 98468.0, 99003.0, 99539.0, 100074.0, 100658.0, 101193.0, 101728.0, 102263.0, 102848.0, 103384.0, 103919.0, 104455.0]
source_col: d(18_16)Mean
PROC_DELTA_DH_CALIBRATE | -40.1228987629
actual_array: [ 18.1 -119. ]
varray: [-75.854, -71.414, -69.657, -68.615, -68.058, -67.881, -67.55, -67.602, -67.401, -67.286, 8.037, 14.966, 17.033, 17.922, 18.799, 18.668, 18.973, 19.294, 19.536, 19.849, -95.286, -105.733, -108.436, -109.698, -110.396, -110.983, -111.057, -111.389, -111.584, -111.753, -72.664, -69.402, -68.538, -68.205, -67.911, -67.542, -67.633, -67.748, -67.048, -67.336, -35.132, -31.895, -31.224, -30.85, -24.608, -23.275, -23.296, -23.214, -24.281, -24.484, -24.56, -24.096, -28.14, -28.398, -28.15, -28.15, -26.917, -26.457, -26.371, -26.224, -27.658, -28.082, -28.019, -28.271, -27.684, -27.875, -28.132, -27.821, -26.273, -26.293, -25.82, -26.199, -61.346, -64.711, -65.395, -65.921, -26.886, -23.126, -22.384, -22.012, -31.749, -32.496, -32.867, -32.952, -21.258, -20.949, -20.989, -20.635, -20.935, -20.477, -20.37, -20.374, -24.766, -25.228, -25.018, -24.881, -18.509, -17.941, -18.212, -17.729, -24.748, -25.528, -25.359, -25.314, -22.755, -22.317, -21.985, -22.066, -61.056, -64.329, -65.337, -65.786, -31.836, -29.189, -28.152, -28.03, -21.838, -21.43, -21.44, -21.754, -23.266, -22.682, -22.893, -22.922, -25.57, -25.987, -26.411, -26.143, -25.278, -25.197, -25.138, -25.409, -26.595, -26.696, -26.823, -27.3, -28.194, -27.731, -27.899, -28.119, -24.89, -25.035, -24.98, -25.136, -60.941, -64.283, -65.325, -65.764, -32.785, -29.046, -28.353, -28.149, -25.141, -24.344, -24.352, -23.802, -27.704, -27.453, -27.705, -27.758, -24.961, -24.751, -24.813, -24.328, -23.717, -23.984, -23.83, -23.636, -25.372, -25.751, -25.621, -25.395, -24.57, -24.689, -24.311, -24.663, -25.313, -25.401, -25.278, -25.699, -61.022, -64.822, -65.485, -65.639, -29.773, -27.029, -25.633, -26.606, -61.561, -64.595, -65.44, -66.01]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -36.1754895833
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3758.0, 4295.0, 4831.0, 5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0, 10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0, 16252.0, 16789.0, 17324.0, 17861.0, 18397.0, 18933.0, 19470.0, 20006.0, 20543.0, 21078.0, 21667.0, 22203.0, 22739.0, 23274.0, 23864.0, 24401.0, 24937.0, 25474.0, 26063.0, 26600.0, 27136.0, 27673.0, 28259.0, 28795.0, 29330.0, 29866.0, 30455.0, 30991.0, 31528.0, 32064.0, 32651.0, 33188.0, 33724.0, 34260.0, 34848.0, 35384.0, 35919.0, 36456.0, 37043.0, 37579.0, 38115.0, 38651.0, 39240.0, 39777.0, 40313.0, 40849.0, 41439.0, 41975.0, 42512.0, 43048.0, 43634.0, 44171.0, 44707.0, 45243.0, 45830.0, 46365.0, 46900.0, 47436.0, 48023.0, 48559.0, 49095.0, 49631.0, 50218.0, 50754.0, 51290.0, 51827.0, 52415.0, 52951.0, 53487.0, 54023.0, 54611.0, 55148.0, 55684.0, 56219.0, 56808.0, 57344.0, 57880.0, 58416.0, 59001.0, 59537.0, 60073.0, 60608.0, 61195.0, 61732.0, 62268.0, 62803.0, 63390.0, 63925.0, 64461.0, 64996.0, 65583.0, 66120.0, 66655.0, 67191.0, 67778.0, 68315.0, 68850.0, 69385.0, 69972.0, 70508.0, 71044.0, 71580.0, 72168.0, 72703.0, 73238.0, 73773.0, 74357.0, 74893.0, 75429.0, 75965.0, 76550.0, 77085.0, 77620.0, 78157.0, 78742.0, 79278.0, 79813.0, 80350.0, 80935.0, 81471.0, 82006.0, 82541.0, 83127.0, 83663.0, 84198.0, 84734.0, 85320.0, 85856.0, 86392.0, 86928.0, 87515.0, 88051.0, 88587.0, 89123.0, 89706.0, 90240.0, 90777.0, 91312.0, 91895.0, 92431.0, 92967.0, 93503.0, 94086.0, 94622.0, 95157.0, 95692.0, 96277.0, 96813.0, 97349.0, 97884.0, 98468.0, 99003.0, 99539.0, 100074.0, 100658.0, 101193.0, 101728.0, 102263.0, 102848.0, 103384.0, 103919.0, 104455.0]
source_col: d(D_H)Mean

## 1 - Sample 1 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 1
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1466
varray: [2.368, 2.208, 2.165, 2.088, 2.103, 2.112, 2.089, 2.11, 2.1, 2.123]
inst: mean_all
operation: mean_all(residuals)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3758.0, 4295.0, 4831.0]
source_col: residuals
SAMP_residuals_STD | 0.0863097779963
varray: [2.368, 2.208, 2.165, 2.088, 2.103, 2.112, 2.089, 2.11, 2.1, 2.123]
inst: std_all
operation: std_all(residuals)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3758.0, 4295.0, 4831.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0544
varray: [-1.267, -1.116, -1.035, -0.923, -1.093, -1.111, -1.04, -0.854, -1.086, -1.019]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3758.0, 4295.0, 4831.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.11252081289
varray: [-1.267, -1.116, -1.035, -0.923, -1.093, -1.111, -1.04, -0.854, -1.086, -1.019]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3758.0, 4295.0, 4831.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 68.4048
varray: [85.273, 127.087, 103.398, 55.828, 53.453, 52.949, 51.784, 51.415, 51.447, 51.414]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3758.0, 4295.0, 4831.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 27.3138820049
varray: [85.273, 127.087, 103.398, 55.828, 53.453, 52.949, 51.784, 51.415, 51.447, 51.414]
inst: std_all
operation: std_all(baseline_shift)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3758.0, 4295.0, 4831.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.4505
varray: [31.819, 27.747, 27.691, 27.761, 28.146, 28.427, 27.656, 28.131, 28.326, 28.801]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3758.0, 4295.0, 4831.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.24091760403
varray: [31.819, 27.747, 27.691, 27.761, 28.146, 28.427, 27.656, 28.131, 28.326, 28.801]
inst: std_all
operation: std_all(slope_shift)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3758.0, 4295.0, 4831.0]
source_col: slope_shift
SAMP_d18O_STD | 0.240814820511
varray: [-9.604, -9.214, -9.081, -8.931, -8.891, -8.877, -8.858, -8.891, -8.864, -8.849]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3758.0, 4295.0, 4831.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.6853905075
varray: [-75.854, -71.414, -69.657, -68.615, -68.058, -67.881, -67.55, -67.602, -67.401, -67.286]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3758.0, 4295.0, 4831.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.8724
varray: [67.539, 62.203, 61.666, 60.598, 60.773, 60.933, 60.738, 62.076, 60.956, 61.242]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3758.0, 4295.0, 4831.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.06943391079
varray: [67.539, 62.203, 61.666, 60.598, 60.773, 60.933, 60.738, 62.076, 60.956, 61.242]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3758.0, 4295.0, 4831.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1466
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0544
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 68.4048
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.4505
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.01725
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.01725
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.797190193405
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.9844
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 0
varray: [-9.604, -9.214, -9.081, -8.931, -8.891, -8.877, -8.858, -8.891, -8.864, -8.849]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3758.0, 4295.0, 4831.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 0
varray: [-75.854, -71.414, -69.657, -68.615, -68.058, -67.881, -67.55, -67.602, -67.401, -67.286]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3758.0, 4295.0, 4831.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.006
varray: [-9.604, -9.214, -9.081, -8.931, -8.891, -8.877, -8.858, -8.891, -8.864, -8.849]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3758.0, 4295.0, 4831.0]
source_col: d(18_16)Mean
dD_mean | -69.1318
varray: [-75.854, -71.414, -69.657, -68.615, -68.058, -67.881, -67.55, -67.602, -67.401, -67.286]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3758.0, 4295.0, 4831.0]
source_col: d(D_H)Mean
d18O_stdev | 0.240814820511
varray: [-9.604, -9.214, -9.081, -8.931, -8.891, -8.877, -8.858, -8.891, -8.864, -8.849]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3758.0, 4295.0, 4831.0]
source_col: d(18_16)Mean
dD_stdev | 2.6853905075
varray: [-75.854, -71.414, -69.657, -68.615, -68.058, -67.881, -67.55, -67.602, -67.401, -67.286]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3758.0, 4295.0, 4831.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.006
varray: [-9.604, -9.214, -9.081, -8.931, -8.891, -8.877, -8.858, -8.891, -8.864, -8.849]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3758.0, 4295.0, 4831.0]
source_col: d(18_16)Mean
measured_dD_mean | -69.1318
varray: [-75.854, -71.414, -69.657, -68.615, -68.058, -67.881, -67.55, -67.602, -67.401, -67.286]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3758.0, 4295.0, 4831.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.240814820511
varray: [-9.604, -9.214, -9.081, -8.931, -8.891, -8.877, -8.858, -8.891, -8.864, -8.849]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3758.0, 4295.0, 4831.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.6853905075
varray: [-75.854, -71.414, -69.657, -68.615, -68.058, -67.881, -67.55, -67.602, -67.401, -67.286]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3758.0, 4295.0, 4831.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.2469942398
actual_array: [ 1.93 -15.84]
varray: [-9.604, -9.214, -9.081, -8.931, -8.891, -8.877, -8.858, -8.891, -8.864, -8.849]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -9.006
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3758.0, 4295.0, 4831.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -75.9998152283
actual_array: [ 18.1 -119. ]
varray: [-75.854, -71.414, -69.657, -68.615, -68.058, -67.881, -67.55, -67.602, -67.401, -67.286]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -69.1318
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [1.0, 538.0, 1074.0, 1611.0, 2148.0, 2684.0, 3221.0, 3758.0, 4295.0, 4831.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.1835
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 20.0747
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 2 - Sample 2 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 2
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.2616
varray: [2.274, 2.244, 2.274, 2.261, 2.261, 2.231, 2.257, 2.258, 2.275, 2.281]
inst: mean_all
operation: mean_all(residuals)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0]
source_col: residuals
SAMP_residuals_STD | 0.0154358889173
varray: [2.274, 2.244, 2.274, 2.261, 2.261, 2.231, 2.257, 2.258, 2.275, 2.281]
inst: std_all
operation: std_all(residuals)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0843
varray: [-1.084, -1.014, -1.06, -1.173, -1.091, -0.918, -1.16, -1.257, -1.104, -0.982]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0986070202594
varray: [-1.084, -1.014, -1.06, -1.173, -1.091, -0.918, -1.16, -1.257, -1.104, -0.982]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 49.0937
varray: [47.956, 49.524, 49.904, 49.833, 49.505, 48.915, 48.893, 49.086, 48.77, 48.551]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.607674629679
varray: [47.956, 49.524, 49.904, 49.833, 49.505, 48.915, 48.893, 49.086, 48.77, 48.551]
inst: std_all
operation: std_all(baseline_shift)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.6967
varray: [29.464, 28.407, 28.442, 28.707, 28.683, 27.986, 28.938, 28.879, 28.805, 28.656]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.386979772655
varray: [29.464, 28.407, 28.442, 28.707, 28.683, 27.986, 28.938, 28.879, 28.805, 28.656]
inst: std_all
operation: std_all(slope_shift)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0]
source_col: slope_shift
SAMP_d18O_STD | 0.243298993011
varray: [2.157, 2.644, 2.812, 2.813, 2.909, 2.857, 2.916, 2.932, 2.946, 2.966]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.56056787137
varray: [8.037, 14.966, 17.033, 17.922, 18.799, 18.668, 18.973, 19.294, 19.536, 19.849]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.4339
varray: [64.517, 64.052, 65.537, 63.601, 64.274, 64.034, 64.273, 64.044, 65.086, 64.921]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.58587113676
varray: [64.517, 64.052, 65.537, 63.601, 64.274, 64.034, 64.273, 64.044, 65.086, 64.921]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.2616
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0843
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 49.0937
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.6967
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.09775
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.09775
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 4.51741109596
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.5771
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 2.7952
varray: [2.157, 2.644, 2.812, 2.813, 2.909, 2.857, 2.916, 2.932, 2.946, 2.966]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 17.3077
varray: [8.037, 14.966, 17.033, 17.922, 18.799, 18.668, 18.973, 19.294, 19.536, 19.849]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0]
source_col: d(D_H)Mean
known_dD | 18.1
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 1.93
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.7952
varray: [2.157, 2.644, 2.812, 2.813, 2.909, 2.857, 2.916, 2.932, 2.946, 2.966]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0]
source_col: d(18_16)Mean
dD_mean | 17.3077
varray: [8.037, 14.966, 17.033, 17.922, 18.799, 18.668, 18.973, 19.294, 19.536, 19.849]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0]
source_col: d(D_H)Mean
d18O_stdev | 0.243298993011
varray: [2.157, 2.644, 2.812, 2.813, 2.909, 2.857, 2.916, 2.932, 2.946, 2.966]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0]
source_col: d(18_16)Mean
dD_stdev | 3.56056787137
varray: [8.037, 14.966, 17.033, 17.922, 18.799, 18.668, 18.973, 19.294, 19.536, 19.849]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.7952
varray: [2.157, 2.644, 2.812, 2.813, 2.909, 2.857, 2.916, 2.932, 2.946, 2.966]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0]
source_col: d(18_16)Mean
measured_dD_mean | 17.3077
varray: [8.037, 14.966, 17.033, 17.922, 18.799, 18.668, 18.973, 19.294, 19.536, 19.849]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.243298993011
varray: [2.157, 2.644, 2.812, 2.813, 2.909, 2.857, 2.916, 2.932, 2.946, 2.966]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.56056787137
varray: [8.037, 14.966, 17.033, 17.922, 18.799, 18.668, 18.973, 19.294, 19.536, 19.849]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.93
actual_array: [ 1.93 -15.84]
varray: [2.157, 2.644, 2.812, 2.813, 2.909, 2.857, 2.916, 2.932, 2.946, 2.966]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 2.7952
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 18.1
actual_array: [ 18.1 -119. ]
varray: [8.037, 14.966, 17.033, 17.922, 18.799, 18.668, 18.973, 19.294, 19.536, 19.849]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: 17.3077
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [5420.0, 5956.0, 6492.0, 7028.0, 7565.0, 8102.0, 8639.0, 9176.0, 9712.0, 10248.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.0685
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.7636
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 3 - Sample 3 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 3
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0661
varray: [2.134, 2.043, 2.062, 2.09, 2.064, 2.055, 2.033, 2.066, 2.057, 2.057]
inst: mean_all
operation: mean_all(residuals)
tarray: [10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: residuals
SAMP_residuals_STD | 0.028128475094
varray: [2.134, 2.043, 2.062, 2.09, 2.064, 2.055, 2.033, 2.066, 2.057, 2.057]
inst: std_all
operation: std_all(residuals)
tarray: [10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9036
varray: [-1.173, -0.825, -0.782, -0.95, -0.947, -0.878, -0.832, -0.906, -0.827, -0.916]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.110245382871
varray: [-1.173, -0.825, -0.782, -0.95, -0.947, -0.878, -0.832, -0.906, -0.827, -0.916]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 47.5665
varray: [47.968, 47.642, 47.529, 48.072, 47.744, 47.762, 47.219, 47.43, 47.04, 47.259]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.334117826855
varray: [47.968, 47.642, 47.529, 48.072, 47.744, 47.762, 47.219, 47.43, 47.04, 47.259]
inst: std_all
operation: std_all(baseline_shift)
tarray: [10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.5904
varray: [29.719, 28.336, 28.402, 28.66, 28.535, 28.241, 28.474, 28.744, 28.064, 28.729]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.451835073155
varray: [29.719, 28.336, 28.402, 28.66, 28.535, 28.241, 28.474, 28.744, 28.064, 28.729]
inst: std_all
operation: std_all(slope_shift)
tarray: [10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: slope_shift
SAMP_d18O_STD | 0.336270459931
varray: [-13.531, -14.234, -14.445, -14.5, -14.536, -14.597, -14.551, -14.615, -14.621, -14.634]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 5.0362045067
varray: [-95.286, -105.733, -108.436, -109.698, -110.396, -110.983, -111.057, -111.389, -111.584, -111.753]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.2797
varray: [62.049, 60.708, 61.468, 61.372, 61.322, 61.571, 60.287, 61.086, 61.354, 61.58]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.49079958571
varray: [62.049, 60.708, 61.468, 61.372, 61.322, 61.571, 60.287, 61.086, 61.354, 61.58]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0661
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9036
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 47.5665
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.5904
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.09775
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.09775
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.51741109596
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.5771
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-13.531, -14.234, -14.445, -14.5, -14.536, -14.597, -14.551, -14.615, -14.621, -14.634]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-95.286, -105.733, -108.436, -109.698, -110.396, -110.983, -111.057, -111.389, -111.584, -111.753]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: d(D_H)Mean
known_dD | -119.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | -15.84
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -14.4264
varray: [-13.531, -14.234, -14.445, -14.5, -14.536, -14.597, -14.551, -14.615, -14.621, -14.634]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: d(18_16)Mean
dD_mean | -108.6315
varray: [-95.286, -105.733, -108.436, -109.698, -110.396, -110.983, -111.057, -111.389, -111.584, -111.753]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: d(D_H)Mean
d18O_stdev | 0.336270459931
varray: [-13.531, -14.234, -14.445, -14.5, -14.536, -14.597, -14.551, -14.615, -14.621, -14.634]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: d(18_16)Mean
dD_stdev | 5.0362045067
varray: [-95.286, -105.733, -108.436, -109.698, -110.396, -110.983, -111.057, -111.389, -111.584, -111.753]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: d(D_H)Mean
measured_d18O_mean | -14.4264
varray: [-13.531, -14.234, -14.445, -14.5, -14.536, -14.597, -14.551, -14.615, -14.621, -14.634]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: d(18_16)Mean
measured_dD_mean | -108.6315
varray: [-95.286, -105.733, -108.436, -109.698, -110.396, -110.983, -111.057, -111.389, -111.584, -111.753]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.336270459931
varray: [-13.531, -14.234, -14.445, -14.5, -14.536, -14.597, -14.551, -14.615, -14.621, -14.634]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: d(18_16)Mean
measured_dD_stdev | 5.0362045067
varray: [-95.286, -105.733, -108.436, -109.698, -110.396, -110.983, -111.057, -111.389, -111.584, -111.753]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -15.84
actual_array: [ 1.93 -15.84]
varray: [-13.531, -14.234, -14.445, -14.5, -14.536, -14.597, -14.551, -14.615, -14.621, -14.634]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -14.4264
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -119.0
actual_array: [ 18.1 -119. ]
varray: [-95.286, -105.733, -108.436, -109.698, -110.396, -110.983, -111.057, -111.389, -111.584, -111.753]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -108.6315
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [10836.0, 11372.0, 11909.0, 12444.0, 12981.0, 13517.0, 14053.0, 14589.0, 15126.0, 15663.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.264
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.7636
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 4 - Sample 4 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 4
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1119
varray: [2.144, 2.125, 2.112, 2.09, 2.108, 2.097, 2.099, 2.125, 2.123, 2.096]
inst: mean_all
operation: mean_all(residuals)
tarray: [16252.0, 16789.0, 17324.0, 17861.0, 18397.0, 18933.0, 19470.0, 20006.0, 20543.0, 21078.0]
source_col: residuals
SAMP_residuals_STD | 0.0171039306723
varray: [2.144, 2.125, 2.112, 2.09, 2.108, 2.097, 2.099, 2.125, 2.123, 2.096]
inst: std_all
operation: std_all(residuals)
tarray: [16252.0, 16789.0, 17324.0, 17861.0, 18397.0, 18933.0, 19470.0, 20006.0, 20543.0, 21078.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.8716
varray: [-0.895, -0.725, -0.691, -0.86, -0.93, -0.962, -0.991, -0.865, -0.926, -0.871]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [16252.0, 16789.0, 17324.0, 17861.0, 18397.0, 18933.0, 19470.0, 20006.0, 20543.0, 21078.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0963975564467
varray: [-0.895, -0.725, -0.691, -0.86, -0.93, -0.962, -0.991, -0.865, -0.926, -0.871]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [16252.0, 16789.0, 17324.0, 17861.0, 18397.0, 18933.0, 19470.0, 20006.0, 20543.0, 21078.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 46.7816
varray: [47.074, 46.884, 46.598, 46.482, 46.793, 46.957, 46.793, 46.962, 46.844, 46.429]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [16252.0, 16789.0, 17324.0, 17861.0, 18397.0, 18933.0, 19470.0, 20006.0, 20543.0, 21078.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.213760613772
varray: [47.074, 46.884, 46.598, 46.482, 46.793, 46.957, 46.793, 46.962, 46.844, 46.429]
inst: std_all
operation: std_all(baseline_shift)
tarray: [16252.0, 16789.0, 17324.0, 17861.0, 18397.0, 18933.0, 19470.0, 20006.0, 20543.0, 21078.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.4406
varray: [29.174, 28.13, 28.218, 28.188, 28.437, 28.318, 28.687, 28.545, 28.454, 28.255]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [16252.0, 16789.0, 17324.0, 17861.0, 18397.0, 18933.0, 19470.0, 20006.0, 20543.0, 21078.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.310731824354
varray: [29.174, 28.13, 28.218, 28.188, 28.437, 28.318, 28.687, 28.545, 28.454, 28.255]
inst: std_all
operation: std_all(slope_shift)
tarray: [16252.0, 16789.0, 17324.0, 17861.0, 18397.0, 18933.0, 19470.0, 20006.0, 20543.0, 21078.0]
source_col: slope_shift
SAMP_d18O_STD | 0.107427546643
varray: [-9.367, -9.2, -9.154, -9.139, -9.09, -9.019, -9.064, -9.032, -9.03, -9.048]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16252.0, 16789.0, 17324.0, 17861.0, 18397.0, 18933.0, 19470.0, 20006.0, 20543.0, 21078.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.63949091353
varray: [-72.664, -69.402, -68.538, -68.205, -67.911, -67.542, -67.633, -67.748, -67.048, -67.336]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16252.0, 16789.0, 17324.0, 17861.0, 18397.0, 18933.0, 19470.0, 20006.0, 20543.0, 21078.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.1482
varray: [62.728, 63.382, 63.052, 61.231, 61.181, 62.028, 60.758, 63.178, 62.074, 61.87]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [16252.0, 16789.0, 17324.0, 17861.0, 18397.0, 18933.0, 19470.0, 20006.0, 20543.0, 21078.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.91545432558
varray: [62.728, 63.382, 63.052, 61.231, 61.181, 62.028, 60.758, 63.178, 62.074, 61.87]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [16252.0, 16789.0, 17324.0, 17861.0, 18397.0, 18933.0, 19470.0, 20006.0, 20543.0, 21078.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1119
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.8716
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 46.7816
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.4406
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.05195
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.05195
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.40081336507
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.7086
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-9.367, -9.2, -9.154, -9.139, -9.09, -9.019, -9.064, -9.032, -9.03, -9.048]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [16252.0, 16789.0, 17324.0, 17861.0, 18397.0, 18933.0, 19470.0, 20006.0, 20543.0, 21078.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-72.664, -69.402, -68.538, -68.205, -67.911, -67.542, -67.633, -67.748, -67.048, -67.336]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [16252.0, 16789.0, 17324.0, 17861.0, 18397.0, 18933.0, 19470.0, 20006.0, 20543.0, 21078.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.1143
varray: [-9.367, -9.2, -9.154, -9.139, -9.09, -9.019, -9.064, -9.032, -9.03, -9.048]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16252.0, 16789.0, 17324.0, 17861.0, 18397.0, 18933.0, 19470.0, 20006.0, 20543.0, 21078.0]
source_col: d(18_16)Mean
dD_mean | -68.4027
varray: [-72.664, -69.402, -68.538, -68.205, -67.911, -67.542, -67.633, -67.748, -67.048, -67.336]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16252.0, 16789.0, 17324.0, 17861.0, 18397.0, 18933.0, 19470.0, 20006.0, 20543.0, 21078.0]
source_col: d(D_H)Mean
d18O_stdev | 0.107427546643
varray: [-9.367, -9.2, -9.154, -9.139, -9.09, -9.019, -9.064, -9.032, -9.03, -9.048]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16252.0, 16789.0, 17324.0, 17861.0, 18397.0, 18933.0, 19470.0, 20006.0, 20543.0, 21078.0]
source_col: d(18_16)Mean
dD_stdev | 1.63949091353
varray: [-72.664, -69.402, -68.538, -68.205, -67.911, -67.542, -67.633, -67.748, -67.048, -67.336]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16252.0, 16789.0, 17324.0, 17861.0, 18397.0, 18933.0, 19470.0, 20006.0, 20543.0, 21078.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.1143
varray: [-9.367, -9.2, -9.154, -9.139, -9.09, -9.019, -9.064, -9.032, -9.03, -9.048]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16252.0, 16789.0, 17324.0, 17861.0, 18397.0, 18933.0, 19470.0, 20006.0, 20543.0, 21078.0]
source_col: d(18_16)Mean
measured_dD_mean | -68.4027
varray: [-72.664, -69.402, -68.538, -68.205, -67.911, -67.542, -67.633, -67.748, -67.048, -67.336]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16252.0, 16789.0, 17324.0, 17861.0, 18397.0, 18933.0, 19470.0, 20006.0, 20543.0, 21078.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.107427546643
varray: [-9.367, -9.2, -9.154, -9.139, -9.09, -9.019, -9.064, -9.032, -9.03, -9.048]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16252.0, 16789.0, 17324.0, 17861.0, 18397.0, 18933.0, 19470.0, 20006.0, 20543.0, 21078.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.63949091353
varray: [-72.664, -69.402, -68.538, -68.205, -67.911, -67.542, -67.633, -67.748, -67.048, -67.336]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16252.0, 16789.0, 17324.0, 17861.0, 18397.0, 18933.0, 19470.0, 20006.0, 20543.0, 21078.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.3587429159
actual_array: [ 1.93 -15.84]
varray: [-9.367, -9.2, -9.154, -9.139, -9.09, -9.019, -9.064, -9.032, -9.03, -9.048]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -9.1143
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [16252.0, 16789.0, 17324.0, 17861.0, 18397.0, 18933.0, 19470.0, 20006.0, 20543.0, 21078.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -75.2061019921
actual_array: [ 18.1 -119. ]
varray: [-72.664, -69.402, -68.538, -68.205, -67.911, -67.542, -67.633, -67.748, -67.048, -67.336]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -68.4027
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [16252.0, 16789.0, 17324.0, 17861.0, 18397.0, 18933.0, 19470.0, 20006.0, 20543.0, 21078.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.2182
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.5485
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 5 - Sample 5 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 5
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.14525
varray: [2.147, 2.127, 2.131, 2.176]
inst: mean_all
operation: mean_all(residuals)
tarray: [21667.0, 22203.0, 22739.0, 23274.0]
source_col: residuals
SAMP_residuals_STD | 0.0222467226051
varray: [2.147, 2.127, 2.131, 2.176]
inst: std_all
operation: std_all(residuals)
tarray: [21667.0, 22203.0, 22739.0, 23274.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.074
varray: [-1.029, -1.099, -0.991, -1.177]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [21667.0, 22203.0, 22739.0, 23274.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0819512049942
varray: [-1.029, -1.099, -0.991, -1.177]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [21667.0, 22203.0, 22739.0, 23274.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 50.06675
varray: [49.994, 49.843, 49.864, 50.566]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [21667.0, 22203.0, 22739.0, 23274.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.339467597865
varray: [49.994, 49.843, 49.864, 50.566]
inst: std_all
operation: std_all(baseline_shift)
tarray: [21667.0, 22203.0, 22739.0, 23274.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.62825
varray: [27.979, 27.174, 27.201, 28.159]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [21667.0, 22203.0, 22739.0, 23274.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.514330227902
varray: [27.979, 27.174, 27.201, 28.159]
inst: std_all
operation: std_all(slope_shift)
tarray: [21667.0, 22203.0, 22739.0, 23274.0]
source_col: slope_shift
SAMP_d18O_STD | 0.152836241339
varray: [-3.936, -3.704, -3.658, -3.581]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21667.0, 22203.0, 22739.0, 23274.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.95295287108
varray: [-35.132, -31.895, -31.224, -30.85]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21667.0, 22203.0, 22739.0, 23274.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.0085
varray: [61.848, 61.779, 62.378, 62.029]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [21667.0, 22203.0, 22739.0, 23274.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.267943401984
varray: [61.848, 61.779, 62.378, 62.029]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [21667.0, 22203.0, 22739.0, 23274.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.14525
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.074
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 50.06675
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.62825
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0186
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0186
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.85957899115
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.8483
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-3.936, -3.704, -3.658, -3.581]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [21667.0, 22203.0, 22739.0, 23274.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-35.132, -31.895, -31.224, -30.85]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [21667.0, 22203.0, 22739.0, 23274.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.71975
varray: [-3.936, -3.704, -3.658, -3.581]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21667.0, 22203.0, 22739.0, 23274.0]
source_col: d(18_16)Mean
dD_mean | -32.27525
varray: [-35.132, -31.895, -31.224, -30.85]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21667.0, 22203.0, 22739.0, 23274.0]
source_col: d(D_H)Mean
d18O_stdev | 0.152836241339
varray: [-3.936, -3.704, -3.658, -3.581]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21667.0, 22203.0, 22739.0, 23274.0]
source_col: d(18_16)Mean
dD_stdev | 1.95295287108
varray: [-35.132, -31.895, -31.224, -30.85]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21667.0, 22203.0, 22739.0, 23274.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.71975
varray: [-3.936, -3.704, -3.658, -3.581]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21667.0, 22203.0, 22739.0, 23274.0]
source_col: d(18_16)Mean
measured_dD_mean | -32.27525
varray: [-35.132, -31.895, -31.224, -30.85]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21667.0, 22203.0, 22739.0, 23274.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.152836241339
varray: [-3.936, -3.704, -3.658, -3.581]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21667.0, 22203.0, 22739.0, 23274.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.95295287108
varray: [-35.132, -31.895, -31.224, -30.85]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21667.0, 22203.0, 22739.0, 23274.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.79241031611
actual_array: [ 1.93 -15.84]
varray: [-3.936, -3.704, -3.658, -3.581]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -3.71975
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [21667.0, 22203.0, 22739.0, 23274.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -35.8770178388
actual_array: [ 18.1 -119. ]
varray: [-35.132, -31.895, -31.224, -30.85]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -32.27525
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [21667.0, 22203.0, 22739.0, 23274.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.18485
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1.73665
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 6 - Sample 6 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 6
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.16325
varray: [2.185, 2.15, 2.155, 2.163]
inst: mean_all
operation: mean_all(residuals)
tarray: [23864.0, 24401.0, 24937.0, 25474.0]
source_col: residuals
SAMP_residuals_STD | 0.0154569294061
varray: [2.185, 2.15, 2.155, 2.163]
inst: std_all
operation: std_all(residuals)
tarray: [23864.0, 24401.0, 24937.0, 25474.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.1755
varray: [-1.262, -1.222, -1.087, -1.131]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [23864.0, 24401.0, 24937.0, 25474.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0805336368648
varray: [-1.262, -1.222, -1.087, -1.131]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [23864.0, 24401.0, 24937.0, 25474.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 71.13675
varray: [71.545, 70.986, 70.862, 71.154]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [23864.0, 24401.0, 24937.0, 25474.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.297309462794
varray: [71.545, 70.986, 70.862, 71.154]
inst: std_all
operation: std_all(baseline_shift)
tarray: [23864.0, 24401.0, 24937.0, 25474.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.585
varray: [26.951, 26.36, 26.5, 26.529]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [23864.0, 24401.0, 24937.0, 25474.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.254913057074
varray: [26.951, 26.36, 26.5, 26.529]
inst: std_all
operation: std_all(slope_shift)
tarray: [23864.0, 24401.0, 24937.0, 25474.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0670721005088
varray: [-2.132, -1.99, -2.008, -1.998]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23864.0, 24401.0, 24937.0, 25474.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.674064475156
varray: [-24.608, -23.275, -23.296, -23.214]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23864.0, 24401.0, 24937.0, 25474.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.5075
varray: [61.704, 61.321, 60.884, 62.121]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [23864.0, 24401.0, 24937.0, 25474.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.528686737744
varray: [61.704, 61.321, 60.884, 62.121]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [23864.0, 24401.0, 24937.0, 25474.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.16325
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.1755
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 71.13675
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.585
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0006
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0006
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.0277283545532
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.3493
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-2.132, -1.99, -2.008, -1.998]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [23864.0, 24401.0, 24937.0, 25474.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-24.608, -23.275, -23.296, -23.214]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [23864.0, 24401.0, 24937.0, 25474.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.032
varray: [-2.132, -1.99, -2.008, -1.998]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23864.0, 24401.0, 24937.0, 25474.0]
source_col: d(18_16)Mean
dD_mean | -23.59825
varray: [-24.608, -23.275, -23.296, -23.214]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23864.0, 24401.0, 24937.0, 25474.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0670721005088
varray: [-2.132, -1.99, -2.008, -1.998]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23864.0, 24401.0, 24937.0, 25474.0]
source_col: d(18_16)Mean
dD_stdev | 0.674064475156
varray: [-24.608, -23.275, -23.296, -23.214]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23864.0, 24401.0, 24937.0, 25474.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.032
varray: [-2.132, -1.99, -2.008, -1.998]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23864.0, 24401.0, 24937.0, 25474.0]
source_col: d(18_16)Mean
measured_dD_mean | -23.59825
varray: [-24.608, -23.275, -23.296, -23.214]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23864.0, 24401.0, 24937.0, 25474.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0670721005088
varray: [-2.132, -1.99, -2.008, -1.998]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23864.0, 24401.0, 24937.0, 25474.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.674064475156
varray: [-24.608, -23.275, -23.296, -23.214]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23864.0, 24401.0, 24937.0, 25474.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.0509160589
actual_array: [ 1.93 -15.84]
varray: [-2.132, -1.99, -2.008, -1.998]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.032
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [23864.0, 24401.0, 24937.0, 25474.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -26.4310574071
actual_array: [ 18.1 -119. ]
varray: [-24.608, -23.275, -23.296, -23.214]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -23.59825
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [23864.0, 24401.0, 24937.0, 25474.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.16685
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 22.80665
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 7 - Sample 7 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 7
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.16425
varray: [2.184, 2.169, 2.13, 2.174]
inst: mean_all
operation: mean_all(residuals)
tarray: [26063.0, 26600.0, 27136.0, 27673.0]
source_col: residuals
SAMP_residuals_STD | 0.0236696007571
varray: [2.184, 2.169, 2.13, 2.174]
inst: std_all
operation: std_all(residuals)
tarray: [26063.0, 26600.0, 27136.0, 27673.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.13075
varray: [-1.055, -1.09, -1.288, -1.09]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [26063.0, 26600.0, 27136.0, 27673.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.10612374852
varray: [-1.055, -1.09, -1.288, -1.09]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [26063.0, 26600.0, 27136.0, 27673.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 60.5535
varray: [60.881, 60.541, 60.174, 60.618]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [26063.0, 26600.0, 27136.0, 27673.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.291886393425
varray: [60.881, 60.541, 60.174, 60.618]
inst: std_all
operation: std_all(baseline_shift)
tarray: [26063.0, 26600.0, 27136.0, 27673.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.3165
varray: [27.7, 27.063, 26.915, 27.588]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [26063.0, 26600.0, 27136.0, 27673.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.385680783377
varray: [27.7, 27.063, 26.915, 27.588]
inst: std_all
operation: std_all(slope_shift)
tarray: [26063.0, 26600.0, 27136.0, 27673.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0338082830088
varray: [-2.404, -2.485, -2.458, -2.443]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26063.0, 26600.0, 27136.0, 27673.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.209143292187
varray: [-24.281, -24.484, -24.56, -24.096]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26063.0, 26600.0, 27136.0, 27673.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.62225
varray: [61.866, 61.857, 60.756, 62.01]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [26063.0, 26600.0, 27136.0, 27673.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.581738987863
varray: [61.866, 61.857, 60.756, 62.01]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [26063.0, 26600.0, 27136.0, 27673.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.16425
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.13075
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 60.5535
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.3165
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0004
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0004
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.0184855697021
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.23455
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-2.404, -2.485, -2.458, -2.443]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [26063.0, 26600.0, 27136.0, 27673.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-24.281, -24.484, -24.56, -24.096]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [26063.0, 26600.0, 27136.0, 27673.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.4475
varray: [-2.404, -2.485, -2.458, -2.443]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26063.0, 26600.0, 27136.0, 27673.0]
source_col: d(18_16)Mean
dD_mean | -24.35525
varray: [-24.281, -24.484, -24.56, -24.096]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26063.0, 26600.0, 27136.0, 27673.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0338082830088
varray: [-2.404, -2.485, -2.458, -2.443]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26063.0, 26600.0, 27136.0, 27673.0]
source_col: d(18_16)Mean
dD_stdev | 0.209143292187
varray: [-24.281, -24.484, -24.56, -24.096]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26063.0, 26600.0, 27136.0, 27673.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.4475
varray: [-2.404, -2.485, -2.458, -2.443]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26063.0, 26600.0, 27136.0, 27673.0]
source_col: d(18_16)Mean
measured_dD_mean | -24.35525
varray: [-24.281, -24.484, -24.56, -24.096]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26063.0, 26600.0, 27136.0, 27673.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0338082830088
varray: [-2.404, -2.485, -2.458, -2.443]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26063.0, 26600.0, 27136.0, 27673.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.209143292187
varray: [-24.281, -24.484, -24.56, -24.096]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26063.0, 26600.0, 27136.0, 27673.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.47964712919
actual_array: [ 1.93 -15.84]
varray: [-2.404, -2.485, -2.458, -2.443]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.4475
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [26063.0, 26600.0, 27136.0, 27673.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -27.2551431564
actual_array: [ 18.1 -119. ]
varray: [-24.281, -24.484, -24.56, -24.096]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -24.35525
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [26063.0, 26600.0, 27136.0, 27673.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.16585
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 12.2234
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 8 - Sample 8 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 8
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.16675
varray: [2.197, 2.156, 2.169, 2.145]
inst: mean_all
operation: mean_all(residuals)
tarray: [28259.0, 28795.0, 29330.0, 29866.0]
source_col: residuals
SAMP_residuals_STD | 0.0224258035902
varray: [2.197, 2.156, 2.169, 2.145]
inst: std_all
operation: std_all(residuals)
tarray: [28259.0, 28795.0, 29330.0, 29866.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.188
varray: [-1.204, -1.318, -1.116, -1.114]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [28259.0, 28795.0, 29330.0, 29866.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0962912249377
varray: [-1.204, -1.318, -1.116, -1.114]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [28259.0, 28795.0, 29330.0, 29866.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 69.369
varray: [70.025, 69.281, 69.194, 68.976]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [28259.0, 28795.0, 29330.0, 29866.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.455760902228
varray: [70.025, 69.281, 69.194, 68.976]
inst: std_all
operation: std_all(baseline_shift)
tarray: [28259.0, 28795.0, 29330.0, 29866.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.92475
varray: [27.45, 27.032, 26.741, 26.476]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [28259.0, 28795.0, 29330.0, 29866.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.417345081038
varray: [27.45, 27.032, 26.741, 26.476]
inst: std_all
operation: std_all(slope_shift)
tarray: [28259.0, 28795.0, 29330.0, 29866.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0160208197876
varray: [-3.089, -3.12, -3.122, -3.121]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28259.0, 28795.0, 29330.0, 29866.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.125755052914
varray: [-28.14, -28.398, -28.15, -28.15]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28259.0, 28795.0, 29330.0, 29866.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.553
varray: [61.817, 61.05, 62.259, 61.086]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [28259.0, 28795.0, 29330.0, 29866.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.588566054067
varray: [61.817, 61.05, 62.259, 61.086]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [28259.0, 28795.0, 29330.0, 29866.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.16675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.188
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 69.369
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.92475
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0029
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0029
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.134020380341
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.3038
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-3.089, -3.12, -3.122, -3.121]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [28259.0, 28795.0, 29330.0, 29866.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-28.14, -28.398, -28.15, -28.15]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [28259.0, 28795.0, 29330.0, 29866.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.113
varray: [-3.089, -3.12, -3.122, -3.121]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28259.0, 28795.0, 29330.0, 29866.0]
source_col: d(18_16)Mean
dD_mean | -28.2095
varray: [-28.14, -28.398, -28.15, -28.15]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28259.0, 28795.0, 29330.0, 29866.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0160208197876
varray: [-3.089, -3.12, -3.122, -3.121]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28259.0, 28795.0, 29330.0, 29866.0]
source_col: d(18_16)Mean
dD_stdev | 0.125755052914
varray: [-28.14, -28.398, -28.15, -28.15]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28259.0, 28795.0, 29330.0, 29866.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.113
varray: [-3.089, -3.12, -3.122, -3.121]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28259.0, 28795.0, 29330.0, 29866.0]
source_col: d(18_16)Mean
measured_dD_mean | -28.2095
varray: [-28.14, -28.398, -28.15, -28.15]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28259.0, 28795.0, 29330.0, 29866.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0160208197876
varray: [-3.089, -3.12, -3.122, -3.121]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28259.0, 28795.0, 29330.0, 29866.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.125755052914
varray: [-28.14, -28.398, -28.15, -28.15]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28259.0, 28795.0, 29330.0, 29866.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.16633913225
actual_array: [ 1.93 -15.84]
varray: [-3.089, -3.12, -3.122, -3.121]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -3.113
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [28259.0, 28795.0, 29330.0, 29866.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -31.4509588754
actual_array: [ 18.1 -119. ]
varray: [-28.14, -28.398, -28.15, -28.15]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -28.2095
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [28259.0, 28795.0, 29330.0, 29866.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.16335
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 21.0389
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 9 - Sample 9 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 9
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.12425
varray: [2.147, 2.114, 2.118, 2.118]
inst: mean_all
operation: mean_all(residuals)
tarray: [30455.0, 30991.0, 31528.0, 32064.0]
source_col: residuals
SAMP_residuals_STD | 0.0152834332967
varray: [2.147, 2.114, 2.118, 2.118]
inst: std_all
operation: std_all(residuals)
tarray: [30455.0, 30991.0, 31528.0, 32064.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9625
varray: [-0.949, -0.883, -0.922, -1.096]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [30455.0, 30991.0, 31528.0, 32064.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0930322524719
varray: [-0.949, -0.883, -0.922, -1.096]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [30455.0, 30991.0, 31528.0, 32064.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 46.57325
varray: [47.076, 46.25, 46.446, 46.521]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [30455.0, 30991.0, 31528.0, 32064.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.354104857352
varray: [47.076, 46.25, 46.446, 46.521]
inst: std_all
operation: std_all(baseline_shift)
tarray: [30455.0, 30991.0, 31528.0, 32064.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.50225
varray: [27.752, 27.317, 27.223, 27.717]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [30455.0, 30991.0, 31528.0, 32064.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.271287516607
varray: [27.752, 27.317, 27.223, 27.717]
inst: std_all
operation: std_all(slope_shift)
tarray: [30455.0, 30991.0, 31528.0, 32064.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0175973482851
varray: [-3.719, -3.712, -3.747, -3.708]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30455.0, 30991.0, 31528.0, 32064.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.299062284483
varray: [-26.917, -26.457, -26.371, -26.224]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30455.0, 30991.0, 31528.0, 32064.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.70975
varray: [62.71, 61.768, 61.878, 60.483]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [30455.0, 30991.0, 31528.0, 32064.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.919622522198
varray: [62.71, 61.768, 61.878, 60.483]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [30455.0, 30991.0, 31528.0, 32064.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.12425
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9625
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 46.57325
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.50225
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0396
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0396
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.83007140051
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.14705
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-3.719, -3.712, -3.747, -3.708]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [30455.0, 30991.0, 31528.0, 32064.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-26.917, -26.457, -26.371, -26.224]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [30455.0, 30991.0, 31528.0, 32064.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.7215
varray: [-3.719, -3.712, -3.747, -3.708]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30455.0, 30991.0, 31528.0, 32064.0]
source_col: d(18_16)Mean
dD_mean | -26.49225
varray: [-26.917, -26.457, -26.371, -26.224]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30455.0, 30991.0, 31528.0, 32064.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0175973482851
varray: [-3.719, -3.712, -3.747, -3.708]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30455.0, 30991.0, 31528.0, 32064.0]
source_col: d(18_16)Mean
dD_stdev | 0.299062284483
varray: [-26.917, -26.457, -26.371, -26.224]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30455.0, 30991.0, 31528.0, 32064.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.7215
varray: [-3.719, -3.712, -3.747, -3.708]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30455.0, 30991.0, 31528.0, 32064.0]
source_col: d(18_16)Mean
measured_dD_mean | -26.49225
varray: [-26.917, -26.457, -26.371, -26.224]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30455.0, 30991.0, 31528.0, 32064.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0175973482851
varray: [-3.719, -3.712, -3.747, -3.708]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30455.0, 30991.0, 31528.0, 32064.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.299062284483
varray: [-26.917, -26.457, -26.371, -26.224]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30455.0, 30991.0, 31528.0, 32064.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.79421604264
actual_array: [ 1.93 -15.84]
varray: [-3.719, -3.712, -3.747, -3.708]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -3.7215
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [30455.0, 30991.0, 31528.0, 32064.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -29.5815252519
actual_array: [ 18.1 -119. ]
varray: [-26.917, -26.457, -26.371, -26.224]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -26.49225
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [30455.0, 30991.0, 31528.0, 32064.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.20585
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.75685
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 10 - Sample 10 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 10
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.098
varray: [2.116, 2.108, 2.076, 2.092]
inst: mean_all
operation: mean_all(residuals)
tarray: [32651.0, 33188.0, 33724.0, 34260.0]
source_col: residuals
SAMP_residuals_STD | 0.0177388462609
varray: [2.116, 2.108, 2.076, 2.092]
inst: std_all
operation: std_all(residuals)
tarray: [32651.0, 33188.0, 33724.0, 34260.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.17175
varray: [-1.09, -1.12, -1.229, -1.248]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [32651.0, 33188.0, 33724.0, 34260.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0784277799422
varray: [-1.09, -1.12, -1.229, -1.248]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [32651.0, 33188.0, 33724.0, 34260.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 56.312
varray: [56.374, 56.45, 56.05, 56.374]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [32651.0, 33188.0, 33724.0, 34260.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.178303112704
varray: [56.374, 56.45, 56.05, 56.374]
inst: std_all
operation: std_all(baseline_shift)
tarray: [32651.0, 33188.0, 33724.0, 34260.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.20775
varray: [26.301, 26.398, 26.054, 26.078]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [32651.0, 33188.0, 33724.0, 34260.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.168685852005
varray: [26.301, 26.398, 26.054, 26.078]
inst: std_all
operation: std_all(slope_shift)
tarray: [32651.0, 33188.0, 33724.0, 34260.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0281350907824
varray: [-2.803, -2.779, -2.784, -2.841]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32651.0, 33188.0, 33724.0, 34260.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.256427377634
varray: [-27.658, -28.082, -28.019, -28.271]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32651.0, 33188.0, 33724.0, 34260.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 60.16825
varray: [61.065, 60.354, 59.054, 60.2]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [32651.0, 33188.0, 33724.0, 34260.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.832911509906
varray: [61.065, 60.354, 59.054, 60.2]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [32651.0, 33188.0, 33724.0, 34260.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.098
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.17175
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 56.312
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.20775
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.06585
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.06585
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.04318691222
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.68855
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-2.803, -2.779, -2.784, -2.841]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [32651.0, 33188.0, 33724.0, 34260.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-27.658, -28.082, -28.019, -28.271]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [32651.0, 33188.0, 33724.0, 34260.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.80175
varray: [-2.803, -2.779, -2.784, -2.841]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32651.0, 33188.0, 33724.0, 34260.0]
source_col: d(18_16)Mean
dD_mean | -28.0075
varray: [-27.658, -28.082, -28.019, -28.271]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32651.0, 33188.0, 33724.0, 34260.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0281350907824
varray: [-2.803, -2.779, -2.784, -2.841]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32651.0, 33188.0, 33724.0, 34260.0]
source_col: d(18_16)Mean
dD_stdev | 0.256427377634
varray: [-27.658, -28.082, -28.019, -28.271]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32651.0, 33188.0, 33724.0, 34260.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.80175
varray: [-2.803, -2.779, -2.784, -2.841]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32651.0, 33188.0, 33724.0, 34260.0]
source_col: d(18_16)Mean
measured_dD_mean | -28.0075
varray: [-27.658, -28.082, -28.019, -28.271]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32651.0, 33188.0, 33724.0, 34260.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0281350907824
varray: [-2.803, -2.779, -2.784, -2.841]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32651.0, 33188.0, 33724.0, 34260.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.256427377634
varray: [-27.658, -28.082, -28.019, -28.271]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32651.0, 33188.0, 33724.0, 34260.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.84517777094
actual_array: [ 1.93 -15.84]
varray: [-2.803, -2.779, -2.784, -2.841]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.80175
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [32651.0, 33188.0, 33724.0, 34260.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -31.2310575262
actual_array: [ 18.1 -119. ]
varray: [-27.658, -28.082, -28.019, -28.271]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -28.0075
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [32651.0, 33188.0, 33724.0, 34260.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.2321
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 7.9819
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 11 - Sample 11 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 11
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.14975
varray: [2.184, 2.129, 2.139, 2.147]
inst: mean_all
operation: mean_all(residuals)
tarray: [34848.0, 35384.0, 35919.0, 36456.0]
source_col: residuals
SAMP_residuals_STD | 0.023991317874
varray: [2.184, 2.129, 2.139, 2.147]
inst: std_all
operation: std_all(residuals)
tarray: [34848.0, 35384.0, 35919.0, 36456.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.1145
varray: [-1.018, -1.186, -1.014, -1.24]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [34848.0, 35384.0, 35919.0, 36456.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.115866302263
varray: [-1.018, -1.186, -1.014, -1.24]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [34848.0, 35384.0, 35919.0, 36456.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 63.25925
varray: [63.916, 63.03, 62.983, 63.108]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [34848.0, 35384.0, 35919.0, 36456.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.440857781301
varray: [63.916, 63.03, 62.983, 63.108]
inst: std_all
operation: std_all(baseline_shift)
tarray: [34848.0, 35384.0, 35919.0, 36456.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.77275
varray: [27.379, 26.557, 26.43, 26.725]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [34848.0, 35384.0, 35919.0, 36456.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.421839128105
varray: [27.379, 26.557, 26.43, 26.725]
inst: std_all
operation: std_all(slope_shift)
tarray: [34848.0, 35384.0, 35919.0, 36456.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0374076640989
varray: [-2.56, -2.585, -2.635, -2.552]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34848.0, 35384.0, 35919.0, 36456.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.187447770503
varray: [-27.684, -27.875, -28.132, -27.821]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34848.0, 35384.0, 35919.0, 36456.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.73625
varray: [62.837, 60.912, 62.409, 60.787]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [34848.0, 35384.0, 35919.0, 36456.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.03998505598
varray: [62.837, 60.912, 62.409, 60.787]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [34848.0, 35384.0, 35919.0, 36456.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.14975
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.1145
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 63.25925
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.77275
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0141
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0141
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.651616332001
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.12055
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-2.56, -2.585, -2.635, -2.552]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [34848.0, 35384.0, 35919.0, 36456.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-27.684, -27.875, -28.132, -27.821]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [34848.0, 35384.0, 35919.0, 36456.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.583
varray: [-2.56, -2.585, -2.635, -2.552]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [34848.0, 35384.0, 35919.0, 36456.0]
source_col: d(18_16)Mean
dD_mean | -27.878
varray: [-27.684, -27.875, -28.132, -27.821]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [34848.0, 35384.0, 35919.0, 36456.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0374076640989
varray: [-2.56, -2.585, -2.635, -2.552]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34848.0, 35384.0, 35919.0, 36456.0]
source_col: d(18_16)Mean
dD_stdev | 0.187447770503
varray: [-27.684, -27.875, -28.132, -27.821]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34848.0, 35384.0, 35919.0, 36456.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.583
varray: [-2.56, -2.585, -2.635, -2.552]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [34848.0, 35384.0, 35919.0, 36456.0]
source_col: d(18_16)Mean
measured_dD_mean | -27.878
varray: [-27.684, -27.875, -28.132, -27.821]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [34848.0, 35384.0, 35919.0, 36456.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0374076640989
varray: [-2.56, -2.585, -2.635, -2.552]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34848.0, 35384.0, 35919.0, 36456.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.187447770503
varray: [-27.684, -27.875, -28.132, -27.821]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34848.0, 35384.0, 35919.0, 36456.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.61946195475
actual_array: [ 1.93 -15.84]
varray: [-2.56, -2.585, -2.635, -2.552]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.583
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [34848.0, 35384.0, 35919.0, 36456.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -31.0900811662
actual_array: [ 18.1 -119. ]
varray: [-27.684, -27.875, -28.132, -27.821]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -27.878
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [34848.0, 35384.0, 35919.0, 36456.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.18035
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 14.92915
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 12 - Sample 12 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 12
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1545
varray: [2.169, 2.149, 2.134, 2.166]
inst: mean_all
operation: mean_all(residuals)
tarray: [37043.0, 37579.0, 38115.0, 38651.0]
source_col: residuals
SAMP_residuals_STD | 0.0162583311977
varray: [2.169, 2.149, 2.134, 2.166]
inst: std_all
operation: std_all(residuals)
tarray: [37043.0, 37579.0, 38115.0, 38651.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.02925
varray: [-1.026, -0.971, -1.023, -1.097]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [37043.0, 37579.0, 38115.0, 38651.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0517453701633
varray: [-1.026, -0.971, -1.023, -1.097]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [37043.0, 37579.0, 38115.0, 38651.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 47.00825
varray: [47.26, 46.901, 46.721, 47.151]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [37043.0, 37579.0, 38115.0, 38651.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.243427436142
varray: [47.26, 46.901, 46.721, 47.151]
inst: std_all
operation: std_all(baseline_shift)
tarray: [37043.0, 37579.0, 38115.0, 38651.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.84275
varray: [28.164, 27.625, 27.84, 27.742]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [37043.0, 37579.0, 38115.0, 38651.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.231498560111
varray: [28.164, 27.625, 27.84, 27.742]
inst: std_all
operation: std_all(slope_shift)
tarray: [37043.0, 37579.0, 38115.0, 38651.0]
source_col: slope_shift
SAMP_d18O_STD | 0.029268868558
varray: [-3.102, -3.131, -3.13, -3.173]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37043.0, 37579.0, 38115.0, 38651.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.221225940311
varray: [-26.273, -26.293, -25.82, -26.199]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37043.0, 37579.0, 38115.0, 38651.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.3395
varray: [62.915, 62.779, 61.786, 61.878]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [37043.0, 37579.0, 38115.0, 38651.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.589831896956
varray: [62.915, 62.779, 61.786, 61.878]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [37043.0, 37579.0, 38115.0, 38651.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1545
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.02925
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 47.00825
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.84275
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.00935
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.00935
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.432100191788
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.5173
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-3.102, -3.131, -3.13, -3.173]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [37043.0, 37579.0, 38115.0, 38651.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-26.273, -26.293, -25.82, -26.199]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [37043.0, 37579.0, 38115.0, 38651.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.134
varray: [-3.102, -3.131, -3.13, -3.173]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37043.0, 37579.0, 38115.0, 38651.0]
source_col: d(18_16)Mean
dD_mean | -26.14625
varray: [-26.273, -26.293, -25.82, -26.199]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37043.0, 37579.0, 38115.0, 38651.0]
source_col: d(D_H)Mean
d18O_stdev | 0.029268868558
varray: [-3.102, -3.131, -3.13, -3.173]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37043.0, 37579.0, 38115.0, 38651.0]
source_col: d(18_16)Mean
dD_stdev | 0.221225940311
varray: [-26.273, -26.293, -25.82, -26.199]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37043.0, 37579.0, 38115.0, 38651.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.134
varray: [-3.102, -3.131, -3.13, -3.173]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37043.0, 37579.0, 38115.0, 38651.0]
source_col: d(18_16)Mean
measured_dD_mean | -26.14625
varray: [-26.273, -26.293, -25.82, -26.199]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37043.0, 37579.0, 38115.0, 38651.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.029268868558
varray: [-3.102, -3.131, -3.13, -3.173]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37043.0, 37579.0, 38115.0, 38651.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.221225940311
varray: [-26.273, -26.293, -25.82, -26.199]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37043.0, 37579.0, 38115.0, 38651.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.18800785061
actual_array: [ 1.93 -15.84]
varray: [-3.102, -3.131, -3.13, -3.173]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -3.134
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [37043.0, 37579.0, 38115.0, 38651.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -29.2048625448
actual_array: [ 18.1 -119. ]
varray: [-26.273, -26.293, -25.82, -26.199]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -26.14625
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [37043.0, 37579.0, 38115.0, 38651.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.1756
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.32185
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 13 - Sample 13 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 13
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.103
varray: [2.15, 2.098, 2.079, 2.085]
inst: mean_all
operation: mean_all(residuals)
tarray: [39240.0, 39777.0, 40313.0, 40849.0]
source_col: residuals
SAMP_residuals_STD | 0.0323213036041
varray: [2.15, 2.098, 2.079, 2.085]
inst: std_all
operation: std_all(residuals)
tarray: [39240.0, 39777.0, 40313.0, 40849.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.02175
varray: [-1.112, -1.029, -0.913, -1.033]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [39240.0, 39777.0, 40313.0, 40849.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0819567975623
varray: [-1.112, -1.029, -0.913, -1.033]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [39240.0, 39777.0, 40313.0, 40849.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 46.542
varray: [47.199, 46.435, 46.26, 46.274]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [39240.0, 39777.0, 40313.0, 40849.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.445138929624
varray: [47.199, 46.435, 46.26, 46.274]
inst: std_all
operation: std_all(baseline_shift)
tarray: [39240.0, 39777.0, 40313.0, 40849.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.31225
varray: [28.927, 28.246, 28.001, 28.075]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [39240.0, 39777.0, 40313.0, 40849.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.422481064664
varray: [28.927, 28.246, 28.001, 28.075]
inst: std_all
operation: std_all(slope_shift)
tarray: [39240.0, 39777.0, 40313.0, 40849.0]
source_col: slope_shift
SAMP_d18O_STD | 0.144937917744
varray: [-8.643, -8.905, -8.886, -8.976]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39240.0, 39777.0, 40313.0, 40849.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.05865819973
varray: [-61.346, -64.711, -65.395, -65.921]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39240.0, 39777.0, 40313.0, 40849.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.32275
varray: [62.119, 61.414, 61.365, 60.393]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [39240.0, 39777.0, 40313.0, 40849.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.709121698535
varray: [62.119, 61.414, 61.365, 60.393]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [39240.0, 39777.0, 40313.0, 40849.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.103
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.02175
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 46.542
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.31225
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.06085
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.06085
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.81211729094
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.53405
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-8.643, -8.905, -8.886, -8.976]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [39240.0, 39777.0, 40313.0, 40849.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-61.346, -64.711, -65.395, -65.921]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [39240.0, 39777.0, 40313.0, 40849.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.8525
varray: [-8.643, -8.905, -8.886, -8.976]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39240.0, 39777.0, 40313.0, 40849.0]
source_col: d(18_16)Mean
dD_mean | -64.34325
varray: [-61.346, -64.711, -65.395, -65.921]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39240.0, 39777.0, 40313.0, 40849.0]
source_col: d(D_H)Mean
d18O_stdev | 0.144937917744
varray: [-8.643, -8.905, -8.886, -8.976]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39240.0, 39777.0, 40313.0, 40849.0]
source_col: d(18_16)Mean
dD_stdev | 2.05865819973
varray: [-61.346, -64.711, -65.395, -65.921]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39240.0, 39777.0, 40313.0, 40849.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.8525
varray: [-8.643, -8.905, -8.886, -8.976]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39240.0, 39777.0, 40313.0, 40849.0]
source_col: d(18_16)Mean
measured_dD_mean | -64.34325
varray: [-61.346, -64.711, -65.395, -65.921]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39240.0, 39777.0, 40313.0, 40849.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.144937917744
varray: [-8.643, -8.905, -8.886, -8.976]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39240.0, 39777.0, 40313.0, 40849.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.05865819973
varray: [-61.346, -64.711, -65.395, -65.921]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39240.0, 39777.0, 40313.0, 40849.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.0886062271
actual_array: [ 1.93 -15.84]
varray: [-8.643, -8.905, -8.886, -8.976]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.8525
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [39240.0, 39777.0, 40313.0, 40849.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -70.7869013381
actual_array: [ 18.1 -119. ]
varray: [-61.346, -64.711, -65.395, -65.921]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -64.34325
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [39240.0, 39777.0, 40313.0, 40849.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.2271
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.7881
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 14 - Sample 14 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 14
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.19775
varray: [2.234, 2.18, 2.171, 2.206]
inst: mean_all
operation: mean_all(residuals)
tarray: [41439.0, 41975.0, 42512.0, 43048.0]
source_col: residuals
SAMP_residuals_STD | 0.028359301825
varray: [2.234, 2.18, 2.171, 2.206]
inst: std_all
operation: std_all(residuals)
tarray: [41439.0, 41975.0, 42512.0, 43048.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.995
varray: [-0.951, -1.012, -0.906, -1.111]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [41439.0, 41975.0, 42512.0, 43048.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0886979894549
varray: [-0.951, -1.012, -0.906, -1.111]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [41439.0, 41975.0, 42512.0, 43048.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 47.50725
varray: [48.018, 47.238, 47.134, 47.639]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [41439.0, 41975.0, 42512.0, 43048.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.404159518837
varray: [48.018, 47.238, 47.134, 47.639]
inst: std_all
operation: std_all(baseline_shift)
tarray: [41439.0, 41975.0, 42512.0, 43048.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.4575
varray: [28.991, 28.324, 27.888, 28.627]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [41439.0, 41975.0, 42512.0, 43048.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.467441618458
varray: [28.991, 28.324, 27.888, 28.627]
inst: std_all
operation: std_all(slope_shift)
tarray: [41439.0, 41975.0, 42512.0, 43048.0]
source_col: slope_shift
SAMP_d18O_STD | 0.180737378536
varray: [-2.565, -2.228, -2.195, -2.192]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41439.0, 41975.0, 42512.0, 43048.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.2377709147
varray: [-26.886, -23.126, -22.384, -22.012]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41439.0, 41975.0, 42512.0, 43048.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.42875
varray: [64.85, 62.863, 62.663, 63.339]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [41439.0, 41975.0, 42512.0, 43048.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.989015124589
varray: [64.85, 62.863, 62.663, 63.339]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [41439.0, 41975.0, 42512.0, 43048.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.19775
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.995
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 47.50725
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.4575
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0339
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0339
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.56665203226
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.57195
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-2.565, -2.228, -2.195, -2.192]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [41439.0, 41975.0, 42512.0, 43048.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-26.886, -23.126, -22.384, -22.012]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [41439.0, 41975.0, 42512.0, 43048.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.295
varray: [-2.565, -2.228, -2.195, -2.192]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [41439.0, 41975.0, 42512.0, 43048.0]
source_col: d(18_16)Mean
dD_mean | -23.602
varray: [-26.886, -23.126, -22.384, -22.012]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [41439.0, 41975.0, 42512.0, 43048.0]
source_col: d(D_H)Mean
d18O_stdev | 0.180737378536
varray: [-2.565, -2.228, -2.195, -2.192]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41439.0, 41975.0, 42512.0, 43048.0]
source_col: d(18_16)Mean
dD_stdev | 2.2377709147
varray: [-26.886, -23.126, -22.384, -22.012]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41439.0, 41975.0, 42512.0, 43048.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.295
varray: [-2.565, -2.228, -2.195, -2.192]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [41439.0, 41975.0, 42512.0, 43048.0]
source_col: d(18_16)Mean
measured_dD_mean | -23.602
varray: [-26.886, -23.126, -22.384, -22.012]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [41439.0, 41975.0, 42512.0, 43048.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.180737378536
varray: [-2.565, -2.228, -2.195, -2.192]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41439.0, 41975.0, 42512.0, 43048.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.2377709147
varray: [-26.886, -23.126, -22.384, -22.012]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41439.0, 41975.0, 42512.0, 43048.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.32229096019
actual_array: [ 1.93 -15.84]
varray: [-2.565, -2.228, -2.195, -2.192]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.295
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [41439.0, 41975.0, 42512.0, 43048.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -26.4351397341
actual_array: [ 18.1 -119. ]
varray: [-26.886, -23.126, -22.384, -22.012]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -23.602
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [41439.0, 41975.0, 42512.0, 43048.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.13235
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.82285
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 15 - Sample 15 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 15
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.15425
varray: [2.192, 2.154, 2.125, 2.146]
inst: mean_all
operation: mean_all(residuals)
tarray: [43634.0, 44171.0, 44707.0, 45243.0]
source_col: residuals
SAMP_residuals_STD | 0.0279806480745
varray: [2.192, 2.154, 2.125, 2.146]
inst: std_all
operation: std_all(residuals)
tarray: [43634.0, 44171.0, 44707.0, 45243.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.133
varray: [-1.065, -0.963, -1.262, -1.242]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [43634.0, 44171.0, 44707.0, 45243.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.14381237777
varray: [-1.065, -0.963, -1.262, -1.242]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [43634.0, 44171.0, 44707.0, 45243.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 56.76475
varray: [57.06, 56.477, 56.536, 56.986]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [43634.0, 44171.0, 44707.0, 45243.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.30069405825
varray: [57.06, 56.477, 56.536, 56.986]
inst: std_all
operation: std_all(baseline_shift)
tarray: [43634.0, 44171.0, 44707.0, 45243.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.44925
varray: [27.568, 27.32, 27.203, 27.706]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [43634.0, 44171.0, 44707.0, 45243.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.229031839126
varray: [27.568, 27.32, 27.203, 27.706]
inst: std_all
operation: std_all(slope_shift)
tarray: [43634.0, 44171.0, 44707.0, 45243.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0584180337453
varray: [-3.979, -4.016, -4.086, -4.103]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43634.0, 44171.0, 44707.0, 45243.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.548326545044
varray: [-31.749, -32.496, -32.867, -32.952]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43634.0, 44171.0, 44707.0, 45243.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.7745
varray: [62.934, 62.869, 60.553, 60.742]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [43634.0, 44171.0, 44707.0, 45243.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.30390298719
varray: [62.934, 62.869, 60.553, 60.742]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [43634.0, 44171.0, 44707.0, 45243.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.15425
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.133
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 56.76475
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.44925
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0096
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0096
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.443653672852
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.0823
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-3.979, -4.016, -4.086, -4.103]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [43634.0, 44171.0, 44707.0, 45243.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-31.749, -32.496, -32.867, -32.952]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [43634.0, 44171.0, 44707.0, 45243.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.046
varray: [-3.979, -4.016, -4.086, -4.103]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [43634.0, 44171.0, 44707.0, 45243.0]
source_col: d(18_16)Mean
dD_mean | -32.516
varray: [-31.749, -32.496, -32.867, -32.952]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [43634.0, 44171.0, 44707.0, 45243.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0584180337453
varray: [-3.979, -4.016, -4.086, -4.103]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43634.0, 44171.0, 44707.0, 45243.0]
source_col: d(18_16)Mean
dD_stdev | 0.548326545044
varray: [-31.749, -32.496, -32.867, -32.952]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43634.0, 44171.0, 44707.0, 45243.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.046
varray: [-3.979, -4.016, -4.086, -4.103]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [43634.0, 44171.0, 44707.0, 45243.0]
source_col: d(18_16)Mean
measured_dD_mean | -32.516
varray: [-31.749, -32.496, -32.867, -32.952]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [43634.0, 44171.0, 44707.0, 45243.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0584180337453
varray: [-3.979, -4.016, -4.086, -4.103]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43634.0, 44171.0, 44707.0, 45243.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.548326545044
varray: [-31.749, -32.496, -32.867, -32.952]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43634.0, 44171.0, 44707.0, 45243.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.1290493334
actual_array: [ 1.93 -15.84]
varray: [-3.979, -4.016, -4.086, -4.103]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -4.046
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [43634.0, 44171.0, 44707.0, 45243.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -36.1391032339
actual_array: [ 18.1 -119. ]
varray: [-31.749, -32.496, -32.867, -32.952]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -32.516
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [43634.0, 44171.0, 44707.0, 45243.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.17585
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 8.43465
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 16 - Sample 16 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 16
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.56225
varray: [2.63, 2.525, 2.551, 2.543]
inst: mean_all
operation: mean_all(residuals)
tarray: [45830.0, 46365.0, 46900.0, 47436.0]
source_col: residuals
SAMP_residuals_STD | 0.0464569693372
varray: [2.63, 2.525, 2.551, 2.543]
inst: std_all
operation: std_all(residuals)
tarray: [45830.0, 46365.0, 46900.0, 47436.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -3.02675
varray: [-3.017, -2.946, -3.112, -3.032]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [45830.0, 46365.0, 46900.0, 47436.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0680949092566
varray: [-3.017, -2.946, -3.112, -3.032]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [45830.0, 46365.0, 46900.0, 47436.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 338.69625
varray: [345.064, 333.42, 338.415, 337.886]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [45830.0, 46365.0, 46900.0, 47436.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 4.80009238626
varray: [345.064, 333.42, 338.415, 337.886]
inst: std_all
operation: std_all(baseline_shift)
tarray: [45830.0, 46365.0, 46900.0, 47436.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 15.71975
varray: [16.2, 15.455, 15.816, 15.408]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [45830.0, 46365.0, 46900.0, 47436.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.368413151231
varray: [16.2, 15.455, 15.816, 15.408]
inst: std_all
operation: std_all(slope_shift)
tarray: [45830.0, 46365.0, 46900.0, 47436.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0357304725223
varray: [0.236, 0.3, 0.249, 0.307]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45830.0, 46365.0, 46900.0, 47436.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.255193227444
varray: [-21.258, -20.949, -20.989, -20.635]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45830.0, 46365.0, 46900.0, 47436.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 58.9045
varray: [60.437, 57.641, 58.748, 58.792]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [45830.0, 46365.0, 46900.0, 47436.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.15211935146
varray: [60.437, 57.641, 58.748, 58.792]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [45830.0, 46365.0, 46900.0, 47436.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.56225
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 3.02675
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 338.69625
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 15.71975
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.3984
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.3984
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 18.4116274233
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.9523
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [0.236, 0.3, 0.249, 0.307]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [45830.0, 46365.0, 46900.0, 47436.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-21.258, -20.949, -20.989, -20.635]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [45830.0, 46365.0, 46900.0, 47436.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.273
varray: [0.236, 0.3, 0.249, 0.307]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45830.0, 46365.0, 46900.0, 47436.0]
source_col: d(18_16)Mean
dD_mean | -20.95775
varray: [-21.258, -20.949, -20.989, -20.635]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45830.0, 46365.0, 46900.0, 47436.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0357304725223
varray: [0.236, 0.3, 0.249, 0.307]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45830.0, 46365.0, 46900.0, 47436.0]
source_col: d(18_16)Mean
dD_stdev | 0.255193227444
varray: [-21.258, -20.949, -20.989, -20.635]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45830.0, 46365.0, 46900.0, 47436.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.273
varray: [0.236, 0.3, 0.249, 0.307]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45830.0, 46365.0, 46900.0, 47436.0]
source_col: d(18_16)Mean
measured_dD_mean | -20.95775
varray: [-21.258, -20.949, -20.989, -20.635]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45830.0, 46365.0, 46900.0, 47436.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0357304725223
varray: [0.236, 0.3, 0.249, 0.307]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45830.0, 46365.0, 46900.0, 47436.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.255193227444
varray: [-21.258, -20.949, -20.989, -20.635]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45830.0, 46365.0, 46900.0, 47436.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.672516258652
actual_array: [ 1.93 -15.84]
varray: [0.236, 0.3, 0.249, 0.307]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 0.273
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [45830.0, 46365.0, 46900.0, 47436.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -23.5565548693
actual_array: [ 18.1 -119. ]
varray: [-21.258, -20.949, -20.989, -20.635]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -20.95775
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [45830.0, 46365.0, 46900.0, 47436.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 45.76785
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 290.36615
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 17 - Sample 17 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 17
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.33175
varray: [2.396, 2.294, 2.329, 2.308]
inst: mean_all
operation: mean_all(residuals)
tarray: [48023.0, 48559.0, 49095.0, 49631.0]
source_col: residuals
SAMP_residuals_STD | 0.0451838835575
varray: [2.396, 2.294, 2.329, 2.308]
inst: std_all
operation: std_all(residuals)
tarray: [48023.0, 48559.0, 49095.0, 49631.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.92075
varray: [-2.009, -1.941, -1.966, -1.767]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [48023.0, 48559.0, 49095.0, 49631.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.106277545449
varray: [-2.009, -1.941, -1.966, -1.767]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [48023.0, 48559.0, 49095.0, 49631.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 237.8395
varray: [243.492, 235.504, 236.679, 235.683]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [48023.0, 48559.0, 49095.0, 49631.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.80361972337
varray: [243.492, 235.504, 236.679, 235.683]
inst: std_all
operation: std_all(baseline_shift)
tarray: [48023.0, 48559.0, 49095.0, 49631.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 20.345
varray: [21.107, 19.862, 20.194, 20.217]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [48023.0, 48559.0, 49095.0, 49631.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.533266037421
varray: [21.107, 19.862, 20.194, 20.217]
inst: std_all
operation: std_all(slope_shift)
tarray: [48023.0, 48559.0, 49095.0, 49631.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0264496376787
varray: [-0.932, -0.873, -0.885, -0.911]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48023.0, 48559.0, 49095.0, 49631.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.268605038424
varray: [-20.935, -20.477, -20.37, -20.374]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48023.0, 48559.0, 49095.0, 49631.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 60.23
varray: [61.106, 59.863, 60.778, 59.173]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [48023.0, 48559.0, 49095.0, 49631.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.879321329208
varray: [61.106, 59.863, 60.778, 59.173]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [48023.0, 48559.0, 49095.0, 49631.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.33175
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.92075
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 237.8395
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 20.345
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.1679
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1679
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 7.75931788248
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.6268
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-0.932, -0.873, -0.885, -0.911]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [48023.0, 48559.0, 49095.0, 49631.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-20.935, -20.477, -20.37, -20.374]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [48023.0, 48559.0, 49095.0, 49631.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.90025
varray: [-0.932, -0.873, -0.885, -0.911]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48023.0, 48559.0, 49095.0, 49631.0]
source_col: d(18_16)Mean
dD_mean | -20.539
varray: [-20.935, -20.477, -20.37, -20.374]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48023.0, 48559.0, 49095.0, 49631.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0264496376787
varray: [-0.932, -0.873, -0.885, -0.911]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48023.0, 48559.0, 49095.0, 49631.0]
source_col: d(18_16)Mean
dD_stdev | 0.268605038424
varray: [-20.935, -20.477, -20.37, -20.374]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48023.0, 48559.0, 49095.0, 49631.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.90025
varray: [-0.932, -0.873, -0.885, -0.911]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48023.0, 48559.0, 49095.0, 49631.0]
source_col: d(18_16)Mean
measured_dD_mean | -20.539
varray: [-20.935, -20.477, -20.37, -20.374]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48023.0, 48559.0, 49095.0, 49631.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0264496376787
varray: [-0.932, -0.873, -0.885, -0.911]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48023.0, 48559.0, 49095.0, 49631.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.268605038424
varray: [-20.935, -20.477, -20.37, -20.374]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48023.0, 48559.0, 49095.0, 49631.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.8831269162
actual_array: [ 1.93 -15.84]
varray: [-0.932, -0.873, -0.885, -0.911]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -0.90025
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [48023.0, 48559.0, 49095.0, 49631.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -23.1006950179
actual_array: [ 18.1 -119. ]
varray: [-20.935, -20.477, -20.37, -20.374]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -20.539
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [48023.0, 48559.0, 49095.0, 49631.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 45.99835
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 189.5094
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 18 - Sample 18 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 18
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.239
varray: [2.302, 2.232, 2.211, 2.211]
inst: mean_all
operation: mean_all(residuals)
tarray: [50218.0, 50754.0, 51290.0, 51827.0]
source_col: residuals
SAMP_residuals_STD | 0.0431508980208
varray: [2.302, 2.232, 2.211, 2.211]
inst: std_all
operation: std_all(residuals)
tarray: [50218.0, 50754.0, 51290.0, 51827.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.62875
varray: [-1.702, -1.513, -1.834, -1.466]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [50218.0, 50754.0, 51290.0, 51827.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.170664143862
varray: [-1.702, -1.513, -1.834, -1.466]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [50218.0, 50754.0, 51290.0, 51827.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 174.38725
varray: [179.191, 173.146, 172.933, 172.279]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [50218.0, 50754.0, 51290.0, 51827.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.22367682158
varray: [179.191, 173.146, 172.933, 172.279]
inst: std_all
operation: std_all(baseline_shift)
tarray: [50218.0, 50754.0, 51290.0, 51827.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 23.762
varray: [24.696, 23.557, 23.627, 23.168]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [50218.0, 50754.0, 51290.0, 51827.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.654584346488
varray: [24.696, 23.557, 23.627, 23.168]
inst: std_all
operation: std_all(slope_shift)
tarray: [50218.0, 50754.0, 51290.0, 51827.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0414085337421
varray: [-2.015, -2.097, -2.005, -2.047]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50218.0, 50754.0, 51290.0, 51827.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.198630939181
varray: [-24.766, -25.228, -25.018, -24.881]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50218.0, 50754.0, 51290.0, 51827.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 59.67475
varray: [60.652, 59.439, 58.662, 59.946]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [50218.0, 50754.0, 51290.0, 51827.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.838616867228
varray: [60.652, 59.439, 58.662, 59.946]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [50218.0, 50754.0, 51290.0, 51827.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.239
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.62875
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 174.38725
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 23.762
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.07515
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.07515
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 3.47297640779
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.18205
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-2.015, -2.097, -2.005, -2.047]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [50218.0, 50754.0, 51290.0, 51827.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-24.766, -25.228, -25.018, -24.881]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [50218.0, 50754.0, 51290.0, 51827.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.041
varray: [-2.015, -2.097, -2.005, -2.047]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50218.0, 50754.0, 51290.0, 51827.0]
source_col: d(18_16)Mean
dD_mean | -24.97325
varray: [-24.766, -25.228, -25.018, -24.881]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50218.0, 50754.0, 51290.0, 51827.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0414085337421
varray: [-2.015, -2.097, -2.005, -2.047]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50218.0, 50754.0, 51290.0, 51827.0]
source_col: d(18_16)Mean
dD_stdev | 0.198630939181
varray: [-24.766, -25.228, -25.018, -24.881]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50218.0, 50754.0, 51290.0, 51827.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.041
varray: [-2.015, -2.097, -2.005, -2.047]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50218.0, 50754.0, 51290.0, 51827.0]
source_col: d(18_16)Mean
measured_dD_mean | -24.97325
varray: [-24.766, -25.228, -25.018, -24.881]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50218.0, 50754.0, 51290.0, 51827.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0414085337421
varray: [-2.015, -2.097, -2.005, -2.047]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50218.0, 50754.0, 51290.0, 51827.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.198630939181
varray: [-24.766, -25.228, -25.018, -24.881]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50218.0, 50754.0, 51290.0, 51827.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.06020265248
actual_array: [ 1.93 -15.84]
varray: [-2.015, -2.097, -2.005, -2.047]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.041
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [50218.0, 50754.0, 51290.0, 51827.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -27.9279106505
actual_array: [ 18.1 -119. ]
varray: [-24.766, -25.228, -25.018, -24.881]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -24.97325
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [50218.0, 50754.0, 51290.0, 51827.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.0911
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 126.05715
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 19 - Sample 19 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 19
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8785
varray: [2.938, 2.858, 2.866, 2.852]
inst: mean_all
operation: mean_all(residuals)
tarray: [52415.0, 52951.0, 53487.0, 54023.0]
source_col: residuals
SAMP_residuals_STD | 0.0400790884793
varray: [2.938, 2.858, 2.866, 2.852]
inst: std_all
operation: std_all(residuals)
tarray: [52415.0, 52951.0, 53487.0, 54023.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -4.0755
varray: [-4.12, -4.089, -3.964, -4.129]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [52415.0, 52951.0, 53487.0, 54023.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0762823701782
varray: [-4.12, -4.089, -3.964, -4.129]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [52415.0, 52951.0, 53487.0, 54023.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 212.29325
varray: [216.487, 210.478, 211.101, 211.107]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [52415.0, 52951.0, 53487.0, 54023.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.81136507472
varray: [216.487, 210.478, 211.101, 211.107]
inst: std_all
operation: std_all(baseline_shift)
tarray: [52415.0, 52951.0, 53487.0, 54023.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 16.9155
varray: [17.424, 16.807, 16.403, 17.028]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [52415.0, 52951.0, 53487.0, 54023.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.426480558369
varray: [17.424, 16.807, 16.403, 17.028]
inst: std_all
operation: std_all(slope_shift)
tarray: [52415.0, 52951.0, 53487.0, 54023.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0688621812027
varray: [2.784, 2.914, 2.893, 2.941]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52415.0, 52951.0, 53487.0, 54023.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.337997411233
varray: [-18.509, -17.941, -18.212, -17.729]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52415.0, 52951.0, 53487.0, 54023.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.51975
varray: [63.876, 62.28, 62.894, 61.029]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [52415.0, 52951.0, 53487.0, 54023.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.19153860617
varray: [63.876, 62.28, 62.894, 61.029]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [52415.0, 52951.0, 53487.0, 54023.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8785
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 4.0755
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 212.29325
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 16.9155
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.71465
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.71465
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 33.0267809691
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.33705
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [2.784, 2.914, 2.893, 2.941]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [52415.0, 52951.0, 53487.0, 54023.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-18.509, -17.941, -18.212, -17.729]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [52415.0, 52951.0, 53487.0, 54023.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.883
varray: [2.784, 2.914, 2.893, 2.941]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [52415.0, 52951.0, 53487.0, 54023.0]
source_col: d(18_16)Mean
dD_mean | -18.09775
varray: [-18.509, -17.941, -18.212, -17.729]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [52415.0, 52951.0, 53487.0, 54023.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0688621812027
varray: [2.784, 2.914, 2.893, 2.941]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52415.0, 52951.0, 53487.0, 54023.0]
source_col: d(18_16)Mean
dD_stdev | 0.337997411233
varray: [-18.509, -17.941, -18.212, -17.729]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52415.0, 52951.0, 53487.0, 54023.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.883
varray: [2.784, 2.914, 2.893, 2.941]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [52415.0, 52951.0, 53487.0, 54023.0]
source_col: d(18_16)Mean
measured_dD_mean | -18.09775
varray: [-18.509, -17.941, -18.212, -17.729]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [52415.0, 52951.0, 53487.0, 54023.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0688621812027
varray: [2.784, 2.914, 2.893, 2.941]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52415.0, 52951.0, 53487.0, 54023.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.337997411233
varray: [-18.509, -17.941, -18.212, -17.729]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52415.0, 52951.0, 53487.0, 54023.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 2.02059587959
actual_array: [ 1.93 -15.84]
varray: [2.784, 2.914, 2.893, 2.941]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 2.883
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [52415.0, 52951.0, 53487.0, 54023.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -20.4431001229
actual_array: [ 18.1 -119. ]
varray: [-18.509, -17.941, -18.212, -17.729]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -18.09775
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [52415.0, 52951.0, 53487.0, 54023.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 45.4516
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 163.96315
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 20 - Sample 20 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 20
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.5015
varray: [2.557, 2.48, 2.439, 2.53]
inst: mean_all
operation: mean_all(residuals)
tarray: [54611.0, 55148.0, 55684.0, 56219.0]
source_col: residuals
SAMP_residuals_STD | 0.0524753910578
varray: [2.557, 2.48, 2.439, 2.53]
inst: std_all
operation: std_all(residuals)
tarray: [54611.0, 55148.0, 55684.0, 56219.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.592
varray: [-2.74, -2.3, -2.69, -2.638]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [54611.0, 55148.0, 55684.0, 56219.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.199071176551
varray: [-2.74, -2.3, -2.69, -2.638]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [54611.0, 55148.0, 55684.0, 56219.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 53.9305
varray: [55.8, 53.468, 52.823, 53.631]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [54611.0, 55148.0, 55684.0, 56219.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.29424276445
varray: [55.8, 53.468, 52.823, 53.631]
inst: std_all
operation: std_all(baseline_shift)
tarray: [54611.0, 55148.0, 55684.0, 56219.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 25.14
varray: [25.832, 24.627, 24.845, 25.256]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [54611.0, 55148.0, 55684.0, 56219.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.529941506206
varray: [25.832, 24.627, 24.845, 25.256]
inst: std_all
operation: std_all(slope_shift)
tarray: [54611.0, 55148.0, 55684.0, 56219.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0278373250631
varray: [0.609, 0.544, 0.592, 0.59]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54611.0, 55148.0, 55684.0, 56219.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.338927106224
varray: [-24.748, -25.528, -25.359, -25.314]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54611.0, 55148.0, 55684.0, 56219.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.171
varray: [65.547, 64.594, 61.804, 64.739]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [54611.0, 55148.0, 55684.0, 56219.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.6327500319
varray: [65.547, 64.594, 61.804, 64.739]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [54611.0, 55148.0, 55684.0, 56219.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.5015
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.592
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 53.9305
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 25.14
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.33765
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.33765
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 15.6041315248
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.3142
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [0.609, 0.544, 0.592, 0.59]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [54611.0, 55148.0, 55684.0, 56219.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-24.748, -25.528, -25.359, -25.314]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [54611.0, 55148.0, 55684.0, 56219.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.58375
varray: [0.609, 0.544, 0.592, 0.59]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [54611.0, 55148.0, 55684.0, 56219.0]
source_col: d(18_16)Mean
dD_mean | -25.23725
varray: [-24.748, -25.528, -25.359, -25.314]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [54611.0, 55148.0, 55684.0, 56219.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0278373250631
varray: [0.609, 0.544, 0.592, 0.59]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54611.0, 55148.0, 55684.0, 56219.0]
source_col: d(18_16)Mean
dD_stdev | 0.338927106224
varray: [-24.748, -25.528, -25.359, -25.314]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54611.0, 55148.0, 55684.0, 56219.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.58375
varray: [0.609, 0.544, 0.592, 0.59]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [54611.0, 55148.0, 55684.0, 56219.0]
source_col: d(18_16)Mean
measured_dD_mean | -25.23725
varray: [-24.748, -25.528, -25.359, -25.314]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [54611.0, 55148.0, 55684.0, 56219.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0278373250631
varray: [0.609, 0.544, 0.592, 0.59]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54611.0, 55148.0, 55684.0, 56219.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.338927106224
varray: [-24.748, -25.528, -25.359, -25.314]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54611.0, 55148.0, 55684.0, 56219.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.351870819204
actual_array: [ 1.93 -15.84]
varray: [0.609, 0.544, 0.592, 0.59]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 0.58375
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [54611.0, 55148.0, 55684.0, 56219.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -28.2153064733
actual_array: [ 18.1 -119. ]
varray: [-24.748, -25.528, -25.359, -25.314]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -25.23725
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [54611.0, 55148.0, 55684.0, 56219.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 45.8286
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 5.6004
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 21 - Sample 21 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 21
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.228
varray: [2.274, 2.219, 2.202, 2.217]
inst: mean_all
operation: mean_all(residuals)
tarray: [56808.0, 57344.0, 57880.0, 58416.0]
source_col: residuals
SAMP_residuals_STD | 0.0315911379979
varray: [2.274, 2.219, 2.202, 2.217]
inst: std_all
operation: std_all(residuals)
tarray: [56808.0, 57344.0, 57880.0, 58416.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.58525
varray: [-1.511, -1.614, -1.66, -1.556]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [56808.0, 57344.0, 57880.0, 58416.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0652756973664
varray: [-1.511, -1.614, -1.66, -1.556]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [56808.0, 57344.0, 57880.0, 58416.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 288.49275
varray: [292.732, 287.143, 286.463, 287.633]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [56808.0, 57344.0, 57880.0, 58416.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.86659616677
varray: [292.732, 287.143, 286.463, 287.633]
inst: std_all
operation: std_all(baseline_shift)
tarray: [56808.0, 57344.0, 57880.0, 58416.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 19.542
varray: [19.826, 19.404, 19.624, 19.314]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [56808.0, 57344.0, 57880.0, 58416.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.229788308377
varray: [19.826, 19.404, 19.624, 19.314]
inst: std_all
operation: std_all(slope_shift)
tarray: [56808.0, 57344.0, 57880.0, 58416.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0224851803047
varray: [-1.705, -1.655, -1.673, -1.66]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56808.0, 57344.0, 57880.0, 58416.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.346320059098
varray: [-22.755, -22.317, -21.985, -22.066]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56808.0, 57344.0, 57880.0, 58416.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 57.629
varray: [59.433, 57.586, 56.391, 57.106]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [56808.0, 57344.0, 57880.0, 58416.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.29903015105
varray: [59.433, 57.586, 56.391, 57.106]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [56808.0, 57344.0, 57880.0, 58416.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.228
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.58525
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 288.49275
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 19.542
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.06415
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.06415
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.96462324098
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.2278
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-1.705, -1.655, -1.673, -1.66]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [56808.0, 57344.0, 57880.0, 58416.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-22.755, -22.317, -21.985, -22.066]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [56808.0, 57344.0, 57880.0, 58416.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.67325
varray: [-1.705, -1.655, -1.673, -1.66]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56808.0, 57344.0, 57880.0, 58416.0]
source_col: d(18_16)Mean
dD_mean | -22.28075
varray: [-22.755, -22.317, -21.985, -22.066]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56808.0, 57344.0, 57880.0, 58416.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0224851803047
varray: [-1.705, -1.655, -1.673, -1.66]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56808.0, 57344.0, 57880.0, 58416.0]
source_col: d(18_16)Mean
dD_stdev | 0.346320059098
varray: [-22.755, -22.317, -21.985, -22.066]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56808.0, 57344.0, 57880.0, 58416.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.67325
varray: [-1.705, -1.655, -1.673, -1.66]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56808.0, 57344.0, 57880.0, 58416.0]
source_col: d(18_16)Mean
measured_dD_mean | -22.28075
varray: [-22.755, -22.317, -21.985, -22.066]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56808.0, 57344.0, 57880.0, 58416.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0224851803047
varray: [-1.705, -1.655, -1.673, -1.66]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56808.0, 57344.0, 57880.0, 58416.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.346320059098
varray: [-22.755, -22.317, -21.985, -22.066]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56808.0, 57344.0, 57880.0, 58416.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.68074212036
actual_array: [ 1.93 -15.84]
varray: [-1.705, -1.655, -1.673, -1.66]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.67325
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [56808.0, 57344.0, 57880.0, 58416.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -24.9967998447
actual_array: [ 18.1 -119. ]
varray: [-22.755, -22.317, -21.985, -22.066]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -22.28075
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [56808.0, 57344.0, 57880.0, 58416.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.1021
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 240.16265
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 22 - Sample 22 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 22
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.113
varray: [2.165, 2.104, 2.098, 2.085]
inst: mean_all
operation: mean_all(residuals)
tarray: [59001.0, 59537.0, 60073.0, 60608.0]
source_col: residuals
SAMP_residuals_STD | 0.0355621521658
varray: [2.165, 2.104, 2.098, 2.085]
inst: std_all
operation: std_all(residuals)
tarray: [59001.0, 59537.0, 60073.0, 60608.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0945
varray: [-1.086, -1.147, -1.14, -1.005]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [59001.0, 59537.0, 60073.0, 60608.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0655972560402
varray: [-1.086, -1.147, -1.14, -1.005]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [59001.0, 59537.0, 60073.0, 60608.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 48.315
varray: [50.546, 48.064, 47.628, 47.022]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [59001.0, 59537.0, 60073.0, 60608.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.54749044154
varray: [50.546, 48.064, 47.628, 47.022]
inst: std_all
operation: std_all(baseline_shift)
tarray: [59001.0, 59537.0, 60073.0, 60608.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.415
varray: [29.021, 28.375, 28.398, 27.866]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [59001.0, 59537.0, 60073.0, 60608.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.472767032128
varray: [29.021, 28.375, 28.398, 27.866]
inst: std_all
operation: std_all(slope_shift)
tarray: [59001.0, 59537.0, 60073.0, 60608.0]
source_col: slope_shift
SAMP_d18O_STD | 0.164526340343
varray: [-8.562, -8.842, -8.877, -8.93]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59001.0, 59537.0, 60073.0, 60608.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.13605758349
varray: [-61.056, -64.329, -65.337, -65.786]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59001.0, 59537.0, 60073.0, 60608.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.468
varray: [62.703, 61.038, 60.781, 61.35]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [59001.0, 59537.0, 60073.0, 60608.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.855573491875
varray: [62.703, 61.038, 60.781, 61.35]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [59001.0, 59537.0, 60073.0, 60608.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.113
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0945
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 48.315
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.415
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.05085
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.05085
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.34997804839
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.3888
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-8.562, -8.842, -8.877, -8.93]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [59001.0, 59537.0, 60073.0, 60608.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-61.056, -64.329, -65.337, -65.786]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [59001.0, 59537.0, 60073.0, 60608.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.80275
varray: [-8.562, -8.842, -8.877, -8.93]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [59001.0, 59537.0, 60073.0, 60608.0]
source_col: d(18_16)Mean
dD_mean | -64.127
varray: [-61.056, -64.329, -65.337, -65.786]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [59001.0, 59537.0, 60073.0, 60608.0]
source_col: d(D_H)Mean
d18O_stdev | 0.164526340343
varray: [-8.562, -8.842, -8.877, -8.93]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59001.0, 59537.0, 60073.0, 60608.0]
source_col: d(18_16)Mean
dD_stdev | 2.13605758349
varray: [-61.056, -64.329, -65.337, -65.786]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59001.0, 59537.0, 60073.0, 60608.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.80275
varray: [-8.562, -8.842, -8.877, -8.93]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [59001.0, 59537.0, 60073.0, 60608.0]
source_col: d(18_16)Mean
measured_dD_mean | -64.127
varray: [-61.056, -64.329, -65.337, -65.786]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [59001.0, 59537.0, 60073.0, 60608.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.164526340343
varray: [-8.562, -8.842, -8.877, -8.93]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59001.0, 59537.0, 60073.0, 60608.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.13605758349
varray: [-61.056, -64.329, -65.337, -65.786]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59001.0, 59537.0, 60073.0, 60608.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.0372720014
actual_array: [ 1.93 -15.84]
varray: [-8.562, -8.842, -8.877, -8.93]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.80275
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [59001.0, 59537.0, 60073.0, 60608.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -70.5514871462
actual_array: [ 18.1 -119. ]
varray: [-61.056, -64.329, -65.337, -65.786]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -64.127
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [59001.0, 59537.0, 60073.0, 60608.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.2171
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.0151
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 23 - Sample 23 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 23
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.19675
varray: [2.24, 2.174, 2.188, 2.185]
inst: mean_all
operation: mean_all(residuals)
tarray: [61195.0, 61732.0, 62268.0, 62803.0]
source_col: residuals
SAMP_residuals_STD | 0.0294547675824
varray: [2.24, 2.174, 2.188, 2.185]
inst: std_all
operation: std_all(residuals)
tarray: [61195.0, 61732.0, 62268.0, 62803.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.31075
varray: [-1.318, -1.309, -1.285, -1.331]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [61195.0, 61732.0, 62268.0, 62803.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0193971647413
varray: [-1.318, -1.309, -1.285, -1.331]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [61195.0, 61732.0, 62268.0, 62803.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 48.498
varray: [48.989, 48.414, 48.389, 48.2]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [61195.0, 61732.0, 62268.0, 62803.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.340989735916
varray: [48.989, 48.414, 48.389, 48.2]
inst: std_all
operation: std_all(baseline_shift)
tarray: [61195.0, 61732.0, 62268.0, 62803.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.691
varray: [28.392, 27.332, 27.473, 27.567]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [61195.0, 61732.0, 62268.0, 62803.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.477207851291
varray: [28.392, 27.332, 27.473, 27.567]
inst: std_all
operation: std_all(slope_shift)
tarray: [61195.0, 61732.0, 62268.0, 62803.0]
source_col: slope_shift
SAMP_d18O_STD | 0.179652210303
varray: [-1.316, -1.061, -0.93, -0.94]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61195.0, 61732.0, 62268.0, 62803.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.7677112085
varray: [-31.836, -29.189, -28.152, -28.03]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61195.0, 61732.0, 62268.0, 62803.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.9645
varray: [64.206, 62.232, 63.122, 62.298]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [61195.0, 61732.0, 62268.0, 62803.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.921395137821
varray: [64.206, 62.232, 63.122, 62.298]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [61195.0, 61732.0, 62268.0, 62803.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.19675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.31075
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 48.498
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.691
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0329
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0329
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.520438108
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.1077
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-1.316, -1.061, -0.93, -0.94]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [61195.0, 61732.0, 62268.0, 62803.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-31.836, -29.189, -28.152, -28.03]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [61195.0, 61732.0, 62268.0, 62803.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.06175
varray: [-1.316, -1.061, -0.93, -0.94]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [61195.0, 61732.0, 62268.0, 62803.0]
source_col: d(18_16)Mean
dD_mean | -29.30175
varray: [-31.836, -29.189, -28.152, -28.03]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [61195.0, 61732.0, 62268.0, 62803.0]
source_col: d(D_H)Mean
d18O_stdev | 0.179652210303
varray: [-1.316, -1.061, -0.93, -0.94]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61195.0, 61732.0, 62268.0, 62803.0]
source_col: d(18_16)Mean
dD_stdev | 1.7677112085
varray: [-31.836, -29.189, -28.152, -28.03]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61195.0, 61732.0, 62268.0, 62803.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.06175
varray: [-1.316, -1.061, -0.93, -0.94]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [61195.0, 61732.0, 62268.0, 62803.0]
source_col: d(18_16)Mean
measured_dD_mean | -29.30175
varray: [-31.836, -29.189, -28.152, -28.03]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [61195.0, 61732.0, 62268.0, 62803.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.179652210303
varray: [-1.316, -1.061, -0.93, -0.94]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61195.0, 61732.0, 62268.0, 62803.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.7677112085
varray: [-31.836, -29.189, -28.152, -28.03]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61195.0, 61732.0, 62268.0, 62803.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.04976967878
actual_array: [ 1.93 -15.84]
varray: [-1.316, -1.061, -0.93, -0.94]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.06175
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [61195.0, 61732.0, 62268.0, 62803.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -32.640004661
actual_array: [ 18.1 -119. ]
varray: [-31.836, -29.189, -28.152, -28.03]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -29.30175
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [61195.0, 61732.0, 62268.0, 62803.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.13335
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.1679
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 24 - Sample 24 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 24
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.59625
varray: [2.633, 2.572, 2.593, 2.587]
inst: mean_all
operation: mean_all(residuals)
tarray: [63390.0, 63925.0, 64461.0, 64996.0]
source_col: residuals
SAMP_residuals_STD | 0.0260432332862
varray: [2.633, 2.572, 2.593, 2.587]
inst: std_all
operation: std_all(residuals)
tarray: [63390.0, 63925.0, 64461.0, 64996.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.935
varray: [-3.173, -2.845, -2.798, -2.924]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [63390.0, 63925.0, 64461.0, 64996.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.16696706262
varray: [-3.173, -2.845, -2.798, -2.924]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [63390.0, 63925.0, 64461.0, 64996.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 55.14375
varray: [55.756, 54.865, 54.838, 55.116]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [63390.0, 63925.0, 64461.0, 64996.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.426928858242
varray: [55.756, 54.865, 54.838, 55.116]
inst: std_all
operation: std_all(baseline_shift)
tarray: [63390.0, 63925.0, 64461.0, 64996.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 23.96375
varray: [24.699, 23.642, 23.598, 23.916]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [63390.0, 63925.0, 64461.0, 64996.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.509957106039
varray: [24.699, 23.642, 23.598, 23.916]
inst: std_all
operation: std_all(slope_shift)
tarray: [63390.0, 63925.0, 64461.0, 64996.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0247924719757
varray: [1.821, 1.843, 1.847, 1.793]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63390.0, 63925.0, 64461.0, 64996.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.211265236137
varray: [-21.838, -21.43, -21.44, -21.754]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63390.0, 63925.0, 64461.0, 64996.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.9335
varray: [65.61, 64.654, 64.892, 64.578]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [63390.0, 63925.0, 64461.0, 64996.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.470416482138
varray: [65.61, 64.654, 64.892, 64.578]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [63390.0, 63925.0, 64461.0, 64996.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.59625
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.935
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 55.14375
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 23.96375
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.4324
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.4324
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 19.982900848
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 2.0767
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [1.821, 1.843, 1.847, 1.793]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [63390.0, 63925.0, 64461.0, 64996.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-21.838, -21.43, -21.44, -21.754]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [63390.0, 63925.0, 64461.0, 64996.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.826
varray: [1.821, 1.843, 1.847, 1.793]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63390.0, 63925.0, 64461.0, 64996.0]
source_col: d(18_16)Mean
dD_mean | -21.6155
varray: [-21.838, -21.43, -21.44, -21.754]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63390.0, 63925.0, 64461.0, 64996.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0247924719757
varray: [1.821, 1.843, 1.847, 1.793]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63390.0, 63925.0, 64461.0, 64996.0]
source_col: d(18_16)Mean
dD_stdev | 0.211265236137
varray: [-21.838, -21.43, -21.44, -21.754]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63390.0, 63925.0, 64461.0, 64996.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.826
varray: [1.821, 1.843, 1.847, 1.793]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63390.0, 63925.0, 64461.0, 64996.0]
source_col: d(18_16)Mean
measured_dD_mean | -21.6155
varray: [-21.838, -21.43, -21.44, -21.754]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63390.0, 63925.0, 64461.0, 64996.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0247924719757
varray: [1.821, 1.843, 1.847, 1.793]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63390.0, 63925.0, 64461.0, 64996.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.211265236137
varray: [-21.838, -21.43, -21.44, -21.754]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63390.0, 63925.0, 64461.0, 64996.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 0.92993705579
actual_array: [ 1.93 -15.84]
varray: [1.821, 1.843, 1.847, 1.793]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 1.826
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [63390.0, 63925.0, 64461.0, 64996.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -24.27259503
actual_array: [ 18.1 -119. ]
varray: [-21.838, -21.43, -21.44, -21.754]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -21.6155
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [63390.0, 63925.0, 64461.0, 64996.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 45.73385
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 6.81365
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 25 - Sample 25 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 25
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.18375
varray: [2.183, 2.182, 2.178, 2.192]
inst: mean_all
operation: mean_all(residuals)
tarray: [65583.0, 66120.0, 66655.0, 67191.0]
source_col: residuals
SAMP_residuals_STD | 0.00590903263375
varray: [2.183, 2.182, 2.178, 2.192]
inst: std_all
operation: std_all(residuals)
tarray: [65583.0, 66120.0, 66655.0, 67191.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.23375
varray: [-1.373, -1.235, -1.189, -1.138]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [65583.0, 66120.0, 66655.0, 67191.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.100933559005
varray: [-1.373, -1.235, -1.189, -1.138]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [65583.0, 66120.0, 66655.0, 67191.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 51.207
varray: [51.335, 51.228, 51.168, 51.097]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [65583.0, 66120.0, 66655.0, 67191.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.100740590958
varray: [51.335, 51.228, 51.168, 51.097]
inst: std_all
operation: std_all(baseline_shift)
tarray: [65583.0, 66120.0, 66655.0, 67191.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.15
varray: [28.22, 28.014, 27.991, 28.375]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [65583.0, 66120.0, 66655.0, 67191.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.181935886143
varray: [28.22, 28.014, 27.991, 28.375]
inst: std_all
operation: std_all(slope_shift)
tarray: [65583.0, 66120.0, 66655.0, 67191.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0341894233158
varray: [-2.52, -2.519, -2.587, -2.567]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65583.0, 66120.0, 66655.0, 67191.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.241778652215
varray: [-23.266, -22.682, -22.893, -22.922]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65583.0, 66120.0, 66655.0, 67191.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.468
varray: [61.946, 62.817, 62.638, 62.471]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [65583.0, 66120.0, 66655.0, 67191.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.375585764728
varray: [61.946, 62.817, 62.638, 62.471]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [65583.0, 66120.0, 66655.0, 67191.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.18375
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.23375
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 51.207
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.15
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0199
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0199
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.919657092682
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.3888
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-2.52, -2.519, -2.587, -2.567]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [65583.0, 66120.0, 66655.0, 67191.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-23.266, -22.682, -22.893, -22.922]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [65583.0, 66120.0, 66655.0, 67191.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.54825
varray: [-2.52, -2.519, -2.587, -2.567]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65583.0, 66120.0, 66655.0, 67191.0]
source_col: d(18_16)Mean
dD_mean | -22.94075
varray: [-23.266, -22.682, -22.893, -22.922]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65583.0, 66120.0, 66655.0, 67191.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0341894233158
varray: [-2.52, -2.519, -2.587, -2.567]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65583.0, 66120.0, 66655.0, 67191.0]
source_col: d(18_16)Mean
dD_stdev | 0.241778652215
varray: [-23.266, -22.682, -22.893, -22.922]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65583.0, 66120.0, 66655.0, 67191.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.54825
varray: [-2.52, -2.519, -2.587, -2.567]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65583.0, 66120.0, 66655.0, 67191.0]
source_col: d(18_16)Mean
measured_dD_mean | -22.94075
varray: [-23.266, -22.682, -22.893, -22.922]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65583.0, 66120.0, 66655.0, 67191.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0341894233158
varray: [-2.52, -2.519, -2.587, -2.567]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65583.0, 66120.0, 66655.0, 67191.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.241778652215
varray: [-23.266, -22.682, -22.893, -22.922]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65583.0, 66120.0, 66655.0, 67191.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.5836053851
actual_array: [ 1.93 -15.84]
varray: [-2.52, -2.519, -2.587, -2.567]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.54825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [65583.0, 66120.0, 66655.0, 67191.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -25.7152894016
actual_array: [ 18.1 -119. ]
varray: [-23.266, -22.682, -22.893, -22.922]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -22.94075
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [65583.0, 66120.0, 66655.0, 67191.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.14635
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 2.8769
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 26 - Sample 26 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 26
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.213
varray: [2.265, 2.192, 2.187, 2.208]
inst: mean_all
operation: mean_all(residuals)
tarray: [67778.0, 68315.0, 68850.0, 69385.0]
source_col: residuals
SAMP_residuals_STD | 0.0358050275799
varray: [2.265, 2.192, 2.187, 2.208]
inst: std_all
operation: std_all(residuals)
tarray: [67778.0, 68315.0, 68850.0, 69385.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.51225
varray: [-1.443, -1.446, -1.554, -1.606]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [67778.0, 68315.0, 68850.0, 69385.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0810694147012
varray: [-1.443, -1.446, -1.554, -1.606]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [67778.0, 68315.0, 68850.0, 69385.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 319.0735
varray: [324.478, 316.745, 316.088, 318.983]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [67778.0, 68315.0, 68850.0, 69385.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.81015892057
varray: [324.478, 316.745, 316.088, 318.983]
inst: std_all
operation: std_all(baseline_shift)
tarray: [67778.0, 68315.0, 68850.0, 69385.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 19.12225
varray: [19.611, 19.007, 18.551, 19.32]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [67778.0, 68315.0, 68850.0, 69385.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.453721922915
varray: [19.611, 19.007, 18.551, 19.32]
inst: std_all
operation: std_all(slope_shift)
tarray: [67778.0, 68315.0, 68850.0, 69385.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0325883414736
varray: [-2.683, -2.758, -2.742, -2.737]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67778.0, 68315.0, 68850.0, 69385.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.351832701721
varray: [-25.57, -25.987, -26.411, -26.143]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67778.0, 68315.0, 68850.0, 69385.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 56.75475
varray: [58.502, 56.556, 56.103, 55.858]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [67778.0, 68315.0, 68850.0, 69385.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.20018370122
varray: [58.502, 56.556, 56.103, 55.858]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [67778.0, 68315.0, 68850.0, 69385.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.213
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.51225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 319.0735
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 19.12225
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.04915
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.04915
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.27141437715
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -6.10205
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-2.683, -2.758, -2.742, -2.737]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [67778.0, 68315.0, 68850.0, 69385.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-25.57, -25.987, -26.411, -26.143]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [67778.0, 68315.0, 68850.0, 69385.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.73
varray: [-2.683, -2.758, -2.742, -2.737]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [67778.0, 68315.0, 68850.0, 69385.0]
source_col: d(18_16)Mean
dD_mean | -26.02775
varray: [-25.57, -25.987, -26.411, -26.143]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [67778.0, 68315.0, 68850.0, 69385.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0325883414736
varray: [-2.683, -2.758, -2.742, -2.737]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67778.0, 68315.0, 68850.0, 69385.0]
source_col: d(18_16)Mean
dD_stdev | 0.351832701721
varray: [-25.57, -25.987, -26.411, -26.143]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67778.0, 68315.0, 68850.0, 69385.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.73
varray: [-2.683, -2.758, -2.742, -2.737]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [67778.0, 68315.0, 68850.0, 69385.0]
source_col: d(18_16)Mean
measured_dD_mean | -26.02775
varray: [-25.57, -25.987, -26.411, -26.143]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [67778.0, 68315.0, 68850.0, 69385.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0325883414736
varray: [-2.683, -2.758, -2.742, -2.737]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67778.0, 68315.0, 68850.0, 69385.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.351832701721
varray: [-25.57, -25.987, -26.411, -26.143]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67778.0, 68315.0, 68850.0, 69385.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.77114298323
actual_array: [ 1.93 -15.84]
varray: [-2.683, -2.758, -2.742, -2.737]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.73
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [67778.0, 68315.0, 68850.0, 69385.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -29.0758610107
actual_array: [ 18.1 -119. ]
varray: [-25.57, -25.987, -26.411, -26.143]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -26.02775
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [67778.0, 68315.0, 68850.0, 69385.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.1171
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 270.7434
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 27 - Sample 27 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 27
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.24175
varray: [2.298, 2.22, 2.211, 2.238]
inst: mean_all
operation: mean_all(residuals)
tarray: [69972.0, 70508.0, 71044.0, 71580.0]
source_col: residuals
SAMP_residuals_STD | 0.0391439650521
varray: [2.298, 2.22, 2.211, 2.238]
inst: std_all
operation: std_all(residuals)
tarray: [69972.0, 70508.0, 71044.0, 71580.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.6965
varray: [-1.795, -1.704, -1.655, -1.632]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [69972.0, 70508.0, 71044.0, 71580.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0722057246115
varray: [-1.795, -1.704, -1.655, -1.632]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [69972.0, 70508.0, 71044.0, 71580.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 210.123
varray: [215.408, 208.802, 208.13, 208.152]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [69972.0, 70508.0, 71044.0, 71580.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.53709654943
varray: [215.408, 208.802, 208.13, 208.152]
inst: std_all
operation: std_all(baseline_shift)
tarray: [69972.0, 70508.0, 71044.0, 71580.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 22.4365
varray: [23.001, 22.213, 22.227, 22.305]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [69972.0, 70508.0, 71044.0, 71580.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.378503632743
varray: [23.001, 22.213, 22.227, 22.305]
inst: std_all
operation: std_all(slope_shift)
tarray: [69972.0, 70508.0, 71044.0, 71580.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0420347078813
varray: [-1.536, -1.485, -1.435, -1.469]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [69972.0, 70508.0, 71044.0, 71580.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.117327177869
varray: [-25.278, -25.197, -25.138, -25.409]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [69972.0, 70508.0, 71044.0, 71580.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 59.11525
varray: [60.259, 58.315, 58.778, 59.109]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [69972.0, 70508.0, 71044.0, 71580.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.829124186517
varray: [60.259, 58.315, 58.778, 59.109]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [69972.0, 70508.0, 71044.0, 71580.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.24175
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.6965
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 210.123
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 22.4365
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0779
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0779
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 3.60006469949
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.74155
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-1.536, -1.485, -1.435, -1.469]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [69972.0, 70508.0, 71044.0, 71580.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-25.278, -25.197, -25.138, -25.409]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [69972.0, 70508.0, 71044.0, 71580.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.48125
varray: [-1.536, -1.485, -1.435, -1.469]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [69972.0, 70508.0, 71044.0, 71580.0]
source_col: d(18_16)Mean
dD_mean | -25.2555
varray: [-25.278, -25.197, -25.138, -25.409]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [69972.0, 70508.0, 71044.0, 71580.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0420347078813
varray: [-1.536, -1.485, -1.435, -1.469]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [69972.0, 70508.0, 71044.0, 71580.0]
source_col: d(18_16)Mean
dD_stdev | 0.117327177869
varray: [-25.278, -25.197, -25.138, -25.409]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [69972.0, 70508.0, 71044.0, 71580.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.48125
varray: [-1.536, -1.485, -1.435, -1.469]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [69972.0, 70508.0, 71044.0, 71580.0]
source_col: d(18_16)Mean
measured_dD_mean | -25.2555
varray: [-25.278, -25.197, -25.138, -25.409]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [69972.0, 70508.0, 71044.0, 71580.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0420347078813
varray: [-1.536, -1.485, -1.435, -1.469]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [69972.0, 70508.0, 71044.0, 71580.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.117327177869
varray: [-25.278, -25.197, -25.138, -25.409]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [69972.0, 70508.0, 71044.0, 71580.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.48262812398
actual_array: [ 1.93 -15.84]
varray: [-1.536, -1.485, -1.435, -1.469]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.48125
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [69972.0, 70508.0, 71044.0, 71580.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -28.2351737982
actual_array: [ 18.1 -119. ]
varray: [-25.278, -25.197, -25.138, -25.409]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -25.2555
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [69972.0, 70508.0, 71044.0, 71580.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.08835
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 161.7929
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 28 - Sample 28 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 28
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.231
varray: [2.268, 2.223, 2.215, 2.218]
inst: mean_all
operation: mean_all(residuals)
tarray: [72168.0, 72703.0, 73238.0, 73773.0]
source_col: residuals
SAMP_residuals_STD | 0.0248864086066
varray: [2.268, 2.223, 2.215, 2.218]
inst: std_all
operation: std_all(residuals)
tarray: [72168.0, 72703.0, 73238.0, 73773.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.3685
varray: [-1.356, -1.363, -1.445, -1.31]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [72168.0, 72703.0, 73238.0, 73773.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0561575165643
varray: [-1.356, -1.363, -1.445, -1.31]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [72168.0, 72703.0, 73238.0, 73773.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 52.71125
varray: [54.372, 52.521, 52.15, 51.802]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [72168.0, 72703.0, 73238.0, 73773.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.14542899533
varray: [54.372, 52.521, 52.15, 51.802]
inst: std_all
operation: std_all(baseline_shift)
tarray: [72168.0, 72703.0, 73238.0, 73773.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.17975
varray: [28.653, 28.012, 28.195, 27.859]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [72168.0, 72703.0, 73238.0, 73773.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.344102092409
varray: [28.653, 28.012, 28.195, 27.859]
inst: std_all
operation: std_all(slope_shift)
tarray: [72168.0, 72703.0, 73238.0, 73773.0]
source_col: slope_shift
SAMP_d18O_STD | 0.023514180119
varray: [-0.945, -0.913, -0.912, -0.959]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72168.0, 72703.0, 73238.0, 73773.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.311940699493
varray: [-26.595, -26.696, -26.823, -27.3]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72168.0, 72703.0, 73238.0, 73773.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.6685
varray: [64.122, 62.425, 61.788, 62.339]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [72168.0, 72703.0, 73238.0, 73773.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.00925797825
varray: [64.122, 62.425, 61.788, 62.339]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [72168.0, 72703.0, 73238.0, 73773.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.231
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.3685
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 52.71125
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.17975
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.06715
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.06715
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 3.10326501375
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.1883
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-0.945, -0.913, -0.912, -0.959]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [72168.0, 72703.0, 73238.0, 73773.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-26.595, -26.696, -26.823, -27.3]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [72168.0, 72703.0, 73238.0, 73773.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.93225
varray: [-0.945, -0.913, -0.912, -0.959]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72168.0, 72703.0, 73238.0, 73773.0]
source_col: d(18_16)Mean
dD_mean | -26.8535
varray: [-26.595, -26.696, -26.823, -27.3]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72168.0, 72703.0, 73238.0, 73773.0]
source_col: d(D_H)Mean
d18O_stdev | 0.023514180119
varray: [-0.945, -0.913, -0.912, -0.959]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72168.0, 72703.0, 73238.0, 73773.0]
source_col: d(18_16)Mean
dD_stdev | 0.311940699493
varray: [-26.595, -26.696, -26.823, -27.3]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72168.0, 72703.0, 73238.0, 73773.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.93225
varray: [-0.945, -0.913, -0.912, -0.959]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72168.0, 72703.0, 73238.0, 73773.0]
source_col: d(18_16)Mean
measured_dD_mean | -26.8535
varray: [-26.595, -26.696, -26.823, -27.3]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72168.0, 72703.0, 73238.0, 73773.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.023514180119
varray: [-0.945, -0.913, -0.912, -0.959]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72168.0, 72703.0, 73238.0, 73773.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.311940699493
varray: [-26.595, -26.696, -26.823, -27.3]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72168.0, 72703.0, 73238.0, 73773.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.91614591559
actual_array: [ 1.93 -15.84]
varray: [-0.945, -0.913, -0.912, -0.959]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -0.93225
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [72168.0, 72703.0, 73238.0, 73773.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -29.9747894222
actual_array: [ 18.1 -119. ]
varray: [-26.595, -26.696, -26.823, -27.3]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -26.8535
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [72168.0, 72703.0, 73238.0, 73773.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.0991
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 4.38115
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 29 - Sample 29 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 29
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.20575
varray: [2.261, 2.191, 2.192, 2.179]
inst: mean_all
operation: mean_all(residuals)
tarray: [74357.0, 74893.0, 75429.0, 75965.0]
source_col: residuals
SAMP_residuals_STD | 0.037303931875
varray: [2.261, 2.191, 2.192, 2.179]
inst: std_all
operation: std_all(residuals)
tarray: [74357.0, 74893.0, 75429.0, 75965.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.32475
varray: [-1.27, -1.411, -1.241, -1.377]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [74357.0, 74893.0, 75429.0, 75965.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0820177826247
varray: [-1.27, -1.411, -1.241, -1.377]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [74357.0, 74893.0, 75429.0, 75965.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 51.73025
varray: [52.379, 51.63, 51.368, 51.544]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [74357.0, 74893.0, 75429.0, 75965.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.446034658594
varray: [52.379, 51.63, 51.368, 51.544]
inst: std_all
operation: std_all(baseline_shift)
tarray: [74357.0, 74893.0, 75429.0, 75965.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.9495
varray: [28.99, 27.509, 27.611, 27.688]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [74357.0, 74893.0, 75429.0, 75965.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.697530166612
varray: [28.99, 27.509, 27.611, 27.688]
inst: std_all
operation: std_all(slope_shift)
tarray: [74357.0, 74893.0, 75429.0, 75965.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0288487434735
varray: [-0.86, -0.797, -0.813, -0.845]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74357.0, 74893.0, 75429.0, 75965.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.210987163907
varray: [-28.194, -27.731, -27.899, -28.119]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74357.0, 74893.0, 75429.0, 75965.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.23775
varray: [63.836, 61.053, 62.269, 61.793]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [74357.0, 74893.0, 75429.0, 75965.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.177117206
varray: [63.836, 61.053, 62.269, 61.793]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [74357.0, 74893.0, 75429.0, 75965.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.20575
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.32475
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 51.73025
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.9495
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0419
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0419
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.9363634263
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.61905
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-0.86, -0.797, -0.813, -0.845]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [74357.0, 74893.0, 75429.0, 75965.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-28.194, -27.731, -27.899, -28.119]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [74357.0, 74893.0, 75429.0, 75965.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.82875
varray: [-0.86, -0.797, -0.813, -0.845]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [74357.0, 74893.0, 75429.0, 75965.0]
source_col: d(18_16)Mean
dD_mean | -27.98575
varray: [-28.194, -27.731, -27.899, -28.119]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [74357.0, 74893.0, 75429.0, 75965.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0288487434735
varray: [-0.86, -0.797, -0.813, -0.845]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74357.0, 74893.0, 75429.0, 75965.0]
source_col: d(18_16)Mean
dD_stdev | 0.210987163907
varray: [-28.194, -27.731, -27.899, -28.119]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74357.0, 74893.0, 75429.0, 75965.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.82875
varray: [-0.86, -0.797, -0.813, -0.845]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [74357.0, 74893.0, 75429.0, 75965.0]
source_col: d(18_16)Mean
measured_dD_mean | -27.98575
varray: [-28.194, -27.731, -27.899, -28.119]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [74357.0, 74893.0, 75429.0, 75965.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0288487434735
varray: [-0.86, -0.797, -0.813, -0.845]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74357.0, 74893.0, 75429.0, 75965.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.210987163907
varray: [-28.194, -27.731, -27.899, -28.119]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74357.0, 74893.0, 75429.0, 75965.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.80935008942
actual_array: [ 1.93 -15.84]
varray: [-0.86, -0.797, -0.813, -0.845]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -0.82875
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [74357.0, 74893.0, 75429.0, 75965.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -31.2073800294
actual_array: [ 18.1 -119. ]
varray: [-28.194, -27.731, -27.899, -28.119]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -27.98575
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [74357.0, 74893.0, 75429.0, 75965.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.12435
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 3.40015
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 30 - Sample 30 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 30
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.28075
varray: [2.306, 2.248, 2.291, 2.278]
inst: mean_all
operation: mean_all(residuals)
tarray: [76550.0, 77085.0, 77620.0, 78157.0]
source_col: residuals
SAMP_residuals_STD | 0.0246492055315
varray: [2.306, 2.248, 2.291, 2.278]
inst: std_all
operation: std_all(residuals)
tarray: [76550.0, 77085.0, 77620.0, 78157.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.6365
varray: [-1.666, -1.626, -1.522, -1.732]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [76550.0, 77085.0, 77620.0, 78157.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0879602182808
varray: [-1.666, -1.626, -1.522, -1.732]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [76550.0, 77085.0, 77620.0, 78157.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 51.54325
varray: [51.932, 51.154, 51.627, 51.46]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [76550.0, 77085.0, 77620.0, 78157.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.324852145855
varray: [51.932, 51.154, 51.627, 51.46]
inst: std_all
operation: std_all(baseline_shift)
tarray: [76550.0, 77085.0, 77620.0, 78157.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.786
varray: [28.451, 27.1, 27.724, 27.869]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [76550.0, 77085.0, 77620.0, 78157.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.554843521485
varray: [28.451, 27.1, 27.724, 27.869]
inst: std_all
operation: std_all(slope_shift)
tarray: [76550.0, 77085.0, 77620.0, 78157.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0588444276603
varray: [-1.309, -1.391, -1.401, -1.451]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76550.0, 77085.0, 77620.0, 78157.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.10295751551
varray: [-24.89, -25.035, -24.98, -25.136]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76550.0, 77085.0, 77620.0, 78157.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.2355
varray: [63.341, 62.122, 64.239, 63.24]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [76550.0, 77085.0, 77620.0, 78157.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.867572283252
varray: [63.341, 62.122, 64.239, 63.24]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [76550.0, 77085.0, 77620.0, 78157.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.28075
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.6365
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 51.54325
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.786
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.1169
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1169
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 5.40240774545
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.3787
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-1.309, -1.391, -1.401, -1.451]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [76550.0, 77085.0, 77620.0, 78157.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-24.89, -25.035, -24.98, -25.136]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [76550.0, 77085.0, 77620.0, 78157.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.388
varray: [-1.309, -1.391, -1.401, -1.451]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [76550.0, 77085.0, 77620.0, 78157.0]
source_col: d(18_16)Mean
dD_mean | -25.01025
varray: [-24.89, -25.035, -24.98, -25.136]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [76550.0, 77085.0, 77620.0, 78157.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0588444276603
varray: [-1.309, -1.391, -1.401, -1.451]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76550.0, 77085.0, 77620.0, 78157.0]
source_col: d(18_16)Mean
dD_stdev | 0.10295751551
varray: [-24.89, -25.035, -24.98, -25.136]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76550.0, 77085.0, 77620.0, 78157.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.388
varray: [-1.309, -1.391, -1.401, -1.451]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [76550.0, 77085.0, 77620.0, 78157.0]
source_col: d(18_16)Mean
measured_dD_mean | -25.01025
varray: [-24.89, -25.035, -24.98, -25.136]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [76550.0, 77085.0, 77620.0, 78157.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0588444276603
varray: [-1.309, -1.391, -1.401, -1.451]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76550.0, 77085.0, 77620.0, 78157.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.10295751551
varray: [-24.89, -25.035, -24.98, -25.136]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76550.0, 77085.0, 77620.0, 78157.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.38640869606
actual_array: [ 1.93 -15.84]
varray: [-1.309, -1.391, -1.401, -1.451]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.388
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [76550.0, 77085.0, 77620.0, 78157.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -27.9681896105
actual_array: [ 18.1 -119. ]
varray: [-24.89, -25.035, -24.98, -25.136]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -25.01025
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [76550.0, 77085.0, 77620.0, 78157.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.04935
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 3.21315
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 31 - Sample 31 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 31
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.11075
varray: [2.168, 2.1, 2.081, 2.094]
inst: mean_all
operation: mean_all(residuals)
tarray: [78742.0, 79278.0, 79813.0, 80350.0]
source_col: residuals
SAMP_residuals_STD | 0.0389818333757
varray: [2.168, 2.1, 2.081, 2.094]
inst: std_all
operation: std_all(residuals)
tarray: [78742.0, 79278.0, 79813.0, 80350.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.07975
varray: [-1.17, -1.167, -1.09, -0.892]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [78742.0, 79278.0, 79813.0, 80350.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.130528094039
varray: [-1.17, -1.167, -1.09, -0.892]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [78742.0, 79278.0, 79813.0, 80350.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 46.76675
varray: [47.446, 46.696, 46.523, 46.402]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [78742.0, 79278.0, 79813.0, 80350.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.468630184687
varray: [47.446, 46.696, 46.523, 46.402]
inst: std_all
operation: std_all(baseline_shift)
tarray: [78742.0, 79278.0, 79813.0, 80350.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.3085
varray: [29.091, 28.2, 27.982, 27.961]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [78742.0, 79278.0, 79813.0, 80350.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.532740399569
varray: [29.091, 28.2, 27.982, 27.961]
inst: std_all
operation: std_all(slope_shift)
tarray: [78742.0, 79278.0, 79813.0, 80350.0]
source_col: slope_shift
SAMP_d18O_STD | 0.188333923303
varray: [-8.522, -8.803, -8.907, -8.934]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78742.0, 79278.0, 79813.0, 80350.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.18177288384
varray: [-60.941, -64.283, -65.325, -65.764]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78742.0, 79278.0, 79813.0, 80350.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.31325
varray: [62.21, 60.771, 60.795, 61.477]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [78742.0, 79278.0, 79813.0, 80350.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.681565049476
varray: [62.21, 60.771, 60.795, 61.477]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [78742.0, 79278.0, 79813.0, 80350.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.11075
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.07975
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 46.76675
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.3085
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0531
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0531
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.45395937796
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.54355
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-8.522, -8.803, -8.907, -8.934]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [78742.0, 79278.0, 79813.0, 80350.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-60.941, -64.283, -65.325, -65.764]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [78742.0, 79278.0, 79813.0, 80350.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.7915
varray: [-8.522, -8.803, -8.907, -8.934]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [78742.0, 79278.0, 79813.0, 80350.0]
source_col: d(18_16)Mean
dD_mean | -64.07825
varray: [-60.941, -64.283, -65.325, -65.764]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [78742.0, 79278.0, 79813.0, 80350.0]
source_col: d(D_H)Mean
d18O_stdev | 0.188333923303
varray: [-8.522, -8.803, -8.907, -8.934]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78742.0, 79278.0, 79813.0, 80350.0]
source_col: d(18_16)Mean
dD_stdev | 2.18177288384
varray: [-60.941, -64.283, -65.325, -65.764]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78742.0, 79278.0, 79813.0, 80350.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.7915
varray: [-8.522, -8.803, -8.907, -8.934]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [78742.0, 79278.0, 79813.0, 80350.0]
source_col: d(18_16)Mean
measured_dD_mean | -64.07825
varray: [-60.941, -64.283, -65.325, -65.764]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [78742.0, 79278.0, 79813.0, 80350.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.188333923303
varray: [-8.522, -8.803, -8.907, -8.934]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78742.0, 79278.0, 79813.0, 80350.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.18177288384
varray: [-60.941, -64.283, -65.325, -65.764]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78742.0, 79278.0, 79813.0, 80350.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.0256637595
actual_array: [ 1.93 -15.84]
varray: [-8.522, -8.803, -8.907, -8.934]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.7915
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [78742.0, 79278.0, 79813.0, 80350.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -70.4984168948
actual_array: [ 18.1 -119. ]
varray: [-60.941, -64.283, -65.325, -65.764]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -64.07825
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [78742.0, 79278.0, 79813.0, 80350.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.21935
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.56335
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 32 - Sample 32 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 32
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.21525
varray: [2.241, 2.215, 2.201, 2.204]
inst: mean_all
operation: mean_all(residuals)
tarray: [80935.0, 81471.0, 82006.0, 82541.0]
source_col: residuals
SAMP_residuals_STD | 0.0181911150474
varray: [2.241, 2.215, 2.201, 2.204]
inst: std_all
operation: std_all(residuals)
tarray: [80935.0, 81471.0, 82006.0, 82541.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.446
varray: [-1.713, -1.335, -1.43, -1.306]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [80935.0, 81471.0, 82006.0, 82541.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.185711245396
varray: [-1.713, -1.335, -1.43, -1.306]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [80935.0, 81471.0, 82006.0, 82541.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 306.5315
varray: [310.038, 305.673, 305.965, 304.45]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [80935.0, 81471.0, 82006.0, 82541.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.42803960704
varray: [310.038, 305.673, 305.965, 304.45]
inst: std_all
operation: std_all(baseline_shift)
tarray: [80935.0, 81471.0, 82006.0, 82541.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 19.51775
varray: [20.438, 19.193, 19.587, 18.853]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [80935.0, 81471.0, 82006.0, 82541.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.682888656127
varray: [20.438, 19.193, 19.587, 18.853]
inst: std_all
operation: std_all(slope_shift)
tarray: [80935.0, 81471.0, 82006.0, 82541.0]
source_col: slope_shift
SAMP_d18O_STD | 0.164036581286
varray: [-2.996, -2.69, -2.664, -2.654]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [80935.0, 81471.0, 82006.0, 82541.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.16874992795
varray: [-32.785, -29.046, -28.353, -28.149]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [80935.0, 81471.0, 82006.0, 82541.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 57.5495
varray: [56.968, 58.288, 56.735, 58.207]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [80935.0, 81471.0, 82006.0, 82541.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.812248114802
varray: [56.968, 58.288, 56.735, 58.207]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [80935.0, 81471.0, 82006.0, 82541.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.21525
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.446
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 306.5315
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 19.51775
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0514
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0514
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.37539570673
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.3073
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-2.996, -2.69, -2.664, -2.654]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [80935.0, 81471.0, 82006.0, 82541.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-32.785, -29.046, -28.353, -28.149]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [80935.0, 81471.0, 82006.0, 82541.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.751
varray: [-2.996, -2.69, -2.664, -2.654]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [80935.0, 81471.0, 82006.0, 82541.0]
source_col: d(18_16)Mean
dD_mean | -29.58325
varray: [-32.785, -29.046, -28.353, -28.149]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [80935.0, 81471.0, 82006.0, 82541.0]
source_col: d(D_H)Mean
d18O_stdev | 0.164036581286
varray: [-2.996, -2.69, -2.664, -2.654]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [80935.0, 81471.0, 82006.0, 82541.0]
source_col: d(18_16)Mean
dD_stdev | 2.16874992795
varray: [-32.785, -29.046, -28.353, -28.149]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [80935.0, 81471.0, 82006.0, 82541.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.751
varray: [-2.996, -2.69, -2.664, -2.654]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [80935.0, 81471.0, 82006.0, 82541.0]
source_col: d(18_16)Mean
measured_dD_mean | -29.58325
varray: [-32.785, -29.046, -28.353, -28.149]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [80935.0, 81471.0, 82006.0, 82541.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.164036581286
varray: [-2.996, -2.69, -2.664, -2.654]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [80935.0, 81471.0, 82006.0, 82541.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.16874992795
varray: [-32.785, -29.046, -28.353, -28.149]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [80935.0, 81471.0, 82006.0, 82541.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.79281170158
actual_array: [ 1.93 -15.84]
varray: [-2.996, -2.69, -2.664, -2.654]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.751
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [80935.0, 81471.0, 82006.0, 82541.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -32.9464513432
actual_array: [ 18.1 -119. ]
varray: [-32.785, -29.046, -28.353, -28.149]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -29.58325
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [80935.0, 81471.0, 82006.0, 82541.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.11485
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 258.2014
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 33 - Sample 33 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 33
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.199
varray: [2.226, 2.16, 2.213, 2.197]
inst: mean_all
operation: mean_all(residuals)
tarray: [83127.0, 83663.0, 84198.0, 84734.0]
source_col: residuals
SAMP_residuals_STD | 0.0285773803325
varray: [2.226, 2.16, 2.213, 2.197]
inst: std_all
operation: std_all(residuals)
tarray: [83127.0, 83663.0, 84198.0, 84734.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.61025
varray: [-1.632, -1.601, -1.497, -1.711]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [83127.0, 83663.0, 84198.0, 84734.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0885715341781
varray: [-1.632, -1.601, -1.497, -1.711]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [83127.0, 83663.0, 84198.0, 84734.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 142.9165
varray: [145.549, 140.235, 143.022, 142.86]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [83127.0, 83663.0, 84198.0, 84734.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.1706236431
varray: [145.549, 140.235, 143.022, 142.86]
inst: std_all
operation: std_all(baseline_shift)
tarray: [83127.0, 83663.0, 84198.0, 84734.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 22.2325
varray: [22.89, 21.524, 22.352, 22.164]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [83127.0, 83663.0, 84198.0, 84734.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.563693474624
varray: [22.89, 21.524, 22.352, 22.164]
inst: std_all
operation: std_all(slope_shift)
tarray: [83127.0, 83663.0, 84198.0, 84734.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0630997886948
varray: [-1.063, -0.957, -0.958, -0.915]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83127.0, 83663.0, 84198.0, 84734.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.551284787262
varray: [-25.141, -24.344, -24.352, -23.802]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83127.0, 83663.0, 84198.0, 84734.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 59.58375
varray: [60.082, 59.164, 59.766, 59.323]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [83127.0, 83663.0, 84198.0, 84734.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.418588401655
varray: [60.082, 59.164, 59.766, 59.323]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [83127.0, 83663.0, 84198.0, 84734.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.199
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.61025
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 142.9165
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 22.2325
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.03515
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.03515
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.62441943758
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.27305
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-1.063, -0.957, -0.958, -0.915]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [83127.0, 83663.0, 84198.0, 84734.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-25.141, -24.344, -24.352, -23.802]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [83127.0, 83663.0, 84198.0, 84734.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.97325
varray: [-1.063, -0.957, -0.958, -0.915]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [83127.0, 83663.0, 84198.0, 84734.0]
source_col: d(18_16)Mean
dD_mean | -24.40975
varray: [-25.141, -24.344, -24.352, -23.802]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [83127.0, 83663.0, 84198.0, 84734.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0630997886948
varray: [-1.063, -0.957, -0.958, -0.915]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83127.0, 83663.0, 84198.0, 84734.0]
source_col: d(18_16)Mean
dD_stdev | 0.551284787262
varray: [-25.141, -24.344, -24.352, -23.802]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83127.0, 83663.0, 84198.0, 84734.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.97325
varray: [-1.063, -0.957, -0.958, -0.915]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [83127.0, 83663.0, 84198.0, 84734.0]
source_col: d(18_16)Mean
measured_dD_mean | -24.40975
varray: [-25.141, -24.344, -24.352, -23.802]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [83127.0, 83663.0, 84198.0, 84734.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0630997886948
varray: [-1.063, -0.957, -0.958, -0.915]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83127.0, 83663.0, 84198.0, 84734.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.551284787262
varray: [-25.141, -24.344, -24.352, -23.802]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83127.0, 83663.0, 84198.0, 84734.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.95845150857
actual_array: [ 1.93 -15.84]
varray: [-1.063, -0.957, -0.958, -0.915]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -0.97325
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [83127.0, 83663.0, 84198.0, 84734.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -27.3144729758
actual_array: [ 18.1 -119. ]
varray: [-25.141, -24.344, -24.352, -23.802]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -24.40975
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [83127.0, 83663.0, 84198.0, 84734.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.1311
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 94.5864
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 34 - Sample 34 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 34
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.694
varray: [1.48, 1.636, 1.8, 1.86]
inst: mean_all
operation: mean_all(residuals)
tarray: [85320.0, 85856.0, 86392.0, 86928.0]
source_col: residuals
SAMP_residuals_STD | 0.171223051408
varray: [1.48, 1.636, 1.8, 1.86]
inst: std_all
operation: std_all(residuals)
tarray: [85320.0, 85856.0, 86392.0, 86928.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.2855
varray: [-1.11, -1.364, -1.338, -1.33]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [85320.0, 85856.0, 86392.0, 86928.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.117896847569
varray: [-1.11, -1.364, -1.338, -1.33]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [85320.0, 85856.0, 86392.0, 86928.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 274.30625
varray: [238.825, 266.803, 291.136, 300.461]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [85320.0, 85856.0, 86392.0, 86928.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 27.5833844597
varray: [238.825, 266.803, 291.136, 300.461]
inst: std_all
operation: std_all(baseline_shift)
tarray: [85320.0, 85856.0, 86392.0, 86928.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 11.025
varray: [8.536, 10.517, 12.173, 12.874]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [85320.0, 85856.0, 86392.0, 86928.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.93131129892
varray: [8.536, 10.517, 12.173, 12.874]
inst: std_all
operation: std_all(slope_shift)
tarray: [85320.0, 85856.0, 86392.0, 86928.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0095
varray: [-2.037, -2.022, -2.016, -2.018]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85320.0, 85856.0, 86392.0, 86928.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.13700851555
varray: [-27.704, -27.453, -27.705, -27.758]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85320.0, 85856.0, 86392.0, 86928.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 45.46175
varray: [41.157, 44.081, 47.531, 49.078]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [85320.0, 85856.0, 86392.0, 86928.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.54947896411
varray: [41.157, 44.081, 47.531, 49.078]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [85320.0, 85856.0, 86392.0, 86928.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.694
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.2855
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 274.30625
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 11.025
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.46985
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.46985
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -21.7136123114
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -17.39505
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-2.037, -2.022, -2.016, -2.018]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [85320.0, 85856.0, 86392.0, 86928.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-27.704, -27.453, -27.705, -27.758]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [85320.0, 85856.0, 86392.0, 86928.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.02325
varray: [-2.037, -2.022, -2.016, -2.018]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [85320.0, 85856.0, 86392.0, 86928.0]
source_col: d(18_16)Mean
dD_mean | -27.655
varray: [-27.704, -27.453, -27.705, -27.758]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [85320.0, 85856.0, 86392.0, 86928.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0095
varray: [-2.037, -2.022, -2.016, -2.018]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85320.0, 85856.0, 86392.0, 86928.0]
source_col: d(18_16)Mean
dD_stdev | 0.13700851555
varray: [-27.704, -27.453, -27.705, -27.758]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85320.0, 85856.0, 86392.0, 86928.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.02325
varray: [-2.037, -2.022, -2.016, -2.018]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [85320.0, 85856.0, 86392.0, 86928.0]
source_col: d(18_16)Mean
measured_dD_mean | -27.655
varray: [-27.704, -27.453, -27.705, -27.758]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [85320.0, 85856.0, 86392.0, 86928.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0095
varray: [-2.037, -2.022, -2.016, -2.018]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85320.0, 85856.0, 86392.0, 86928.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.13700851555
varray: [-27.704, -27.453, -27.705, -27.758]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85320.0, 85856.0, 86392.0, 86928.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.04188742626
actual_array: [ 1.93 -15.84]
varray: [-2.037, -2.022, -2.016, -2.018]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.02325
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [85320.0, 85856.0, 86392.0, 86928.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -30.8473187856
actual_array: [ 18.1 -119. ]
varray: [-27.704, -27.453, -27.705, -27.758]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -27.655
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [85320.0, 85856.0, 86392.0, 86928.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.6361
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 225.97615
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 35 - Sample 35 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 35
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.12625
varray: [2.073, 2.126, 2.144, 2.162]
inst: mean_all
operation: mean_all(residuals)
tarray: [87515.0, 88051.0, 88587.0, 89123.0]
source_col: residuals
SAMP_residuals_STD | 0.0384219989069
varray: [2.073, 2.126, 2.144, 2.162]
inst: std_all
operation: std_all(residuals)
tarray: [87515.0, 88051.0, 88587.0, 89123.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.17375
varray: [-1.166, -1.233, -1.203, -1.093]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [87515.0, 88051.0, 88587.0, 89123.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0604062634721
varray: [-1.166, -1.233, -1.203, -1.093]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [87515.0, 88051.0, 88587.0, 89123.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 64.05425
varray: [64.987, 63.926, 63.839, 63.465]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [87515.0, 88051.0, 88587.0, 89123.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.653202048884
varray: [64.987, 63.926, 63.839, 63.465]
inst: std_all
operation: std_all(baseline_shift)
tarray: [87515.0, 88051.0, 88587.0, 89123.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.57625
varray: [25.741, 26.487, 26.82, 27.257]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [87515.0, 88051.0, 88587.0, 89123.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.639906959383
varray: [25.741, 26.487, 26.82, 27.257]
inst: std_all
operation: std_all(slope_shift)
tarray: [87515.0, 88051.0, 88587.0, 89123.0]
source_col: slope_shift
SAMP_d18O_STD | 0.019663841605
varray: [-2.594, -2.608, -2.626, -2.58]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87515.0, 88051.0, 88587.0, 89123.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.271522098548
varray: [-24.961, -24.751, -24.813, -24.328]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87515.0, 88051.0, 88587.0, 89123.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.0755
varray: [59.814, 61.077, 61.709, 61.702]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [87515.0, 88051.0, 88587.0, 89123.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.891666791277
varray: [59.814, 61.077, 61.709, 61.702]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [87515.0, 88051.0, 88587.0, 89123.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.12625
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.17375
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 64.05425
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.57625
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0376
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0376
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.737643552
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.7813
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-2.594, -2.608, -2.626, -2.58]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [87515.0, 88051.0, 88587.0, 89123.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-24.961, -24.751, -24.813, -24.328]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [87515.0, 88051.0, 88587.0, 89123.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.602
varray: [-2.594, -2.608, -2.626, -2.58]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [87515.0, 88051.0, 88587.0, 89123.0]
source_col: d(18_16)Mean
dD_mean | -24.71325
varray: [-24.961, -24.751, -24.813, -24.328]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [87515.0, 88051.0, 88587.0, 89123.0]
source_col: d(D_H)Mean
d18O_stdev | 0.019663841605
varray: [-2.594, -2.608, -2.626, -2.58]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87515.0, 88051.0, 88587.0, 89123.0]
source_col: d(18_16)Mean
dD_stdev | 0.271522098548
varray: [-24.961, -24.751, -24.813, -24.328]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87515.0, 88051.0, 88587.0, 89123.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.602
varray: [-2.594, -2.608, -2.626, -2.58]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [87515.0, 88051.0, 88587.0, 89123.0]
source_col: d(18_16)Mean
measured_dD_mean | -24.71325
varray: [-24.961, -24.751, -24.813, -24.328]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [87515.0, 88051.0, 88587.0, 89123.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.019663841605
varray: [-2.594, -2.608, -2.626, -2.58]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87515.0, 88051.0, 88587.0, 89123.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.271522098548
varray: [-24.961, -24.751, -24.813, -24.328]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87515.0, 88051.0, 88587.0, 89123.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.63906698565
actual_array: [ 1.93 -15.84]
varray: [-2.594, -2.608, -2.626, -2.58]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.602
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [87515.0, 88051.0, 88587.0, 89123.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -27.64486931
actual_array: [ 18.1 -119. ]
varray: [-24.961, -24.751, -24.813, -24.328]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -24.71325
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [87515.0, 88051.0, 88587.0, 89123.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.20385
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 15.72415
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 36 - Sample 36 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 36
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.119
varray: [2.105, 2.115, 2.123, 2.133]
inst: mean_all
operation: mean_all(residuals)
tarray: [89706.0, 90240.0, 90777.0, 91312.0]
source_col: residuals
SAMP_residuals_STD | 0.0118883696668
varray: [2.105, 2.115, 2.123, 2.133]
inst: std_all
operation: std_all(residuals)
tarray: [89706.0, 90240.0, 90777.0, 91312.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.19775
varray: [-1.144, -1.152, -1.331, -1.164]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [89706.0, 90240.0, 90777.0, 91312.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0892127606717
varray: [-1.144, -1.152, -1.331, -1.164]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [89706.0, 90240.0, 90777.0, 91312.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 60.77275
varray: [60.483, 60.712, 60.947, 60.949]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [89706.0, 90240.0, 90777.0, 91312.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.222914595006
varray: [60.483, 60.712, 60.947, 60.949]
inst: std_all
operation: std_all(baseline_shift)
tarray: [89706.0, 90240.0, 90777.0, 91312.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.481
varray: [26.254, 26.46, 26.737, 26.473]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [89706.0, 90240.0, 90777.0, 91312.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.197964643308
varray: [26.254, 26.46, 26.737, 26.473]
inst: std_all
operation: std_all(slope_shift)
tarray: [89706.0, 90240.0, 90777.0, 91312.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0227064601674
varray: [-2.574, -2.544, -2.566, -2.599]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89706.0, 90240.0, 90777.0, 91312.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.150851748415
varray: [-23.717, -23.984, -23.83, -23.636]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89706.0, 90240.0, 90777.0, 91312.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 60.2405
varray: [59.638, 60.336, 59.739, 61.249]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [89706.0, 90240.0, 90777.0, 91312.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.739527101149
varray: [59.638, 60.336, 59.739, 61.249]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [89706.0, 90240.0, 90777.0, 91312.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.119
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.19775
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 60.77275
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.481
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.04485
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.04485
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.07269450285
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.6163
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-2.574, -2.544, -2.566, -2.599]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [89706.0, 90240.0, 90777.0, 91312.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-23.717, -23.984, -23.83, -23.636]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [89706.0, 90240.0, 90777.0, 91312.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.57075
varray: [-2.574, -2.544, -2.566, -2.599]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [89706.0, 90240.0, 90777.0, 91312.0]
source_col: d(18_16)Mean
dD_mean | -23.79175
varray: [-23.717, -23.984, -23.83, -23.636]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [89706.0, 90240.0, 90777.0, 91312.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0227064601674
varray: [-2.574, -2.544, -2.566, -2.599]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89706.0, 90240.0, 90777.0, 91312.0]
source_col: d(18_16)Mean
dD_stdev | 0.150851748415
varray: [-23.717, -23.984, -23.83, -23.636]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89706.0, 90240.0, 90777.0, 91312.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.57075
varray: [-2.574, -2.544, -2.566, -2.599]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [89706.0, 90240.0, 90777.0, 91312.0]
source_col: d(18_16)Mean
measured_dD_mean | -23.79175
varray: [-23.717, -23.984, -23.83, -23.636]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [89706.0, 90240.0, 90777.0, 91312.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0227064601674
varray: [-2.574, -2.544, -2.566, -2.599]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89706.0, 90240.0, 90777.0, 91312.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.150851748415
varray: [-23.717, -23.984, -23.83, -23.636]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89706.0, 90240.0, 90777.0, 91312.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.60682186905
actual_array: [ 1.93 -15.84]
varray: [-2.574, -2.544, -2.566, -2.599]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.57075
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [89706.0, 90240.0, 90777.0, 91312.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -26.6417054817
actual_array: [ 18.1 -119. ]
varray: [-23.717, -23.984, -23.83, -23.636]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -23.79175
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [89706.0, 90240.0, 90777.0, 91312.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.2111
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 12.44265
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 37 - Sample 37 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 37
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.002
varray: [2.043, 1.989, 1.984, 1.992]
inst: mean_all
operation: mean_all(residuals)
tarray: [91895.0, 92431.0, 92967.0, 93503.0]
source_col: residuals
SAMP_residuals_STD | 0.0275317997959
varray: [2.043, 1.989, 1.984, 1.992]
inst: std_all
operation: std_all(residuals)
tarray: [91895.0, 92431.0, 92967.0, 93503.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.99625
varray: [-0.997, -1.075, -0.96, -0.953]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [91895.0, 92431.0, 92967.0, 93503.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0559367201994
varray: [-0.997, -1.075, -0.96, -0.953]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [91895.0, 92431.0, 92967.0, 93503.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 55.34525
varray: [55.868, 55.333, 55.074, 55.106]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [91895.0, 92431.0, 92967.0, 93503.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.36707617284
varray: [55.868, 55.333, 55.074, 55.106]
inst: std_all
operation: std_all(baseline_shift)
tarray: [91895.0, 92431.0, 92967.0, 93503.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 24.57825
varray: [25.02, 24.501, 24.153, 24.639]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [91895.0, 92431.0, 92967.0, 93503.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.358533471241
varray: [25.02, 24.501, 24.153, 24.639]
inst: std_all
operation: std_all(slope_shift)
tarray: [91895.0, 92431.0, 92967.0, 93503.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0310094071748
varray: [-1.849, -1.784, -1.811, -1.783]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [91895.0, 92431.0, 92967.0, 93503.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.182775591003
varray: [-25.372, -25.751, -25.621, -25.395]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [91895.0, 92431.0, 92967.0, 93503.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 58.677
varray: [59.908, 58.429, 58.085, 58.286]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [91895.0, 92431.0, 92967.0, 93503.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.832708432366
varray: [59.908, 58.429, 58.085, 58.286]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [91895.0, 92431.0, 92967.0, 93503.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.002
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.99625
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 55.34525
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 24.57825
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.16185
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.16185
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -7.47972364073
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.1798
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-1.849, -1.784, -1.811, -1.783]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [91895.0, 92431.0, 92967.0, 93503.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-25.372, -25.751, -25.621, -25.395]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [91895.0, 92431.0, 92967.0, 93503.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.80675
varray: [-1.849, -1.784, -1.811, -1.783]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [91895.0, 92431.0, 92967.0, 93503.0]
source_col: d(18_16)Mean
dD_mean | -25.53475
varray: [-25.372, -25.751, -25.621, -25.395]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [91895.0, 92431.0, 92967.0, 93503.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0310094071748
varray: [-1.849, -1.784, -1.811, -1.783]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [91895.0, 92431.0, 92967.0, 93503.0]
source_col: d(18_16)Mean
dD_stdev | 0.182775591003
varray: [-25.372, -25.751, -25.621, -25.395]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [91895.0, 92431.0, 92967.0, 93503.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.80675
varray: [-1.849, -1.784, -1.811, -1.783]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [91895.0, 92431.0, 92967.0, 93503.0]
source_col: d(18_16)Mean
measured_dD_mean | -25.53475
varray: [-25.372, -25.751, -25.621, -25.395]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [91895.0, 92431.0, 92967.0, 93503.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0310094071748
varray: [-1.849, -1.784, -1.811, -1.783]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [91895.0, 92431.0, 92967.0, 93503.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.182775591003
varray: [-25.372, -25.751, -25.621, -25.395]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [91895.0, 92431.0, 92967.0, 93503.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.81849325847
actual_array: [ 1.93 -15.84]
varray: [-1.849, -1.784, -1.811, -1.783]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.80675
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [91895.0, 92431.0, 92967.0, 93503.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -28.5391710841
actual_array: [ 18.1 -119. ]
varray: [-25.372, -25.751, -25.621, -25.395]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -25.53475
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [91895.0, 92431.0, 92967.0, 93503.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.3281
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 7.01515
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 38 - Sample 38 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 38
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.3275
varray: [2.349, 2.325, 2.325, 2.311]
inst: mean_all
operation: mean_all(residuals)
tarray: [94086.0, 94622.0, 95157.0, 95692.0]
source_col: residuals
SAMP_residuals_STD | 0.0157797338381
varray: [2.349, 2.325, 2.325, 2.311]
inst: std_all
operation: std_all(residuals)
tarray: [94086.0, 94622.0, 95157.0, 95692.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.07275
varray: [-1.964, -2.17, -2.178, -1.979]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [94086.0, 94622.0, 95157.0, 95692.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.117119241232
varray: [-1.964, -2.17, -2.178, -1.979]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [94086.0, 94622.0, 95157.0, 95692.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 264.5335
varray: [265.782, 264.16, 264.768, 263.424]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [94086.0, 94622.0, 95157.0, 95692.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.997369039022
varray: [265.782, 264.16, 264.768, 263.424]
inst: std_all
operation: std_all(baseline_shift)
tarray: [94086.0, 94622.0, 95157.0, 95692.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 19.7105
varray: [20.289, 19.449, 19.723, 19.381]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [94086.0, 94622.0, 95157.0, 95692.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.413025826473
varray: [20.289, 19.449, 19.723, 19.381]
inst: std_all
operation: std_all(slope_shift)
tarray: [94086.0, 94622.0, 95157.0, 95692.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0373217988134
varray: [-1.024, -1.086, -1.007, -1.07]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94086.0, 94622.0, 95157.0, 95692.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.172567619597
varray: [-24.57, -24.689, -24.311, -24.663]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94086.0, 94622.0, 95157.0, 95692.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 59.2995
varray: [59.031, 59.367, 59.55, 59.25]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [94086.0, 94622.0, 95157.0, 95692.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.217446545155
varray: [59.031, 59.367, 59.55, 59.25]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [94086.0, 94622.0, 95157.0, 95692.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.3275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.07275
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 264.5335
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 19.7105
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.16365
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.16365
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 7.56290870439
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.5573
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-1.024, -1.086, -1.007, -1.07]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [94086.0, 94622.0, 95157.0, 95692.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-24.57, -24.689, -24.311, -24.663]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [94086.0, 94622.0, 95157.0, 95692.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.04675
varray: [-1.024, -1.086, -1.007, -1.07]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [94086.0, 94622.0, 95157.0, 95692.0]
source_col: d(18_16)Mean
dD_mean | -24.55825
varray: [-24.57, -24.689, -24.311, -24.663]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [94086.0, 94622.0, 95157.0, 95692.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0373217988134
varray: [-1.024, -1.086, -1.007, -1.07]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94086.0, 94622.0, 95157.0, 95692.0]
source_col: d(18_16)Mean
dD_stdev | 0.172567619597
varray: [-24.57, -24.689, -24.311, -24.663]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94086.0, 94622.0, 95157.0, 95692.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.04675
varray: [-1.024, -1.086, -1.007, -1.07]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [94086.0, 94622.0, 95157.0, 95692.0]
source_col: d(18_16)Mean
measured_dD_mean | -24.55825
varray: [-24.57, -24.689, -24.311, -24.663]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [94086.0, 94622.0, 95157.0, 95692.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0373217988134
varray: [-1.024, -1.086, -1.007, -1.07]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94086.0, 94622.0, 95157.0, 95692.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.172567619597
varray: [-24.57, -24.689, -24.311, -24.663]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94086.0, 94622.0, 95157.0, 95692.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -2.03429202281
actual_array: [ 1.93 -15.84]
varray: [-1.024, -1.086, -1.007, -1.07]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.04675
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [94086.0, 94622.0, 95157.0, 95692.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -27.4761331261
actual_array: [ 18.1 -119. ]
varray: [-24.57, -24.689, -24.311, -24.663]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -24.55825
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [94086.0, 94622.0, 95157.0, 95692.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.0026
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 216.2034
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 39 - Sample 39 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 39
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.303
varray: [2.326, 2.315, 2.302, 2.269]
inst: mean_all
operation: mean_all(residuals)
tarray: [96277.0, 96813.0, 97349.0, 97884.0]
source_col: residuals
SAMP_residuals_STD | 0.0246981780705
varray: [2.326, 2.315, 2.302, 2.269]
inst: std_all
operation: std_all(residuals)
tarray: [96277.0, 96813.0, 97349.0, 97884.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.75325
varray: [-1.714, -1.674, -1.797, -1.828]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [96277.0, 96813.0, 97349.0, 97884.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0714673585166
varray: [-1.714, -1.674, -1.797, -1.828]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [96277.0, 96813.0, 97349.0, 97884.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 154.64375
varray: [157.127, 154.189, 154.451, 152.808]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [96277.0, 96813.0, 97349.0, 97884.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.80558935438
varray: [157.127, 154.189, 154.451, 152.808]
inst: std_all
operation: std_all(baseline_shift)
tarray: [96277.0, 96813.0, 97349.0, 97884.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 23.80675
varray: [24.352, 23.481, 23.909, 23.485]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [96277.0, 96813.0, 97349.0, 97884.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.415286547659
varray: [24.352, 23.481, 23.909, 23.485]
inst: std_all
operation: std_all(slope_shift)
tarray: [96277.0, 96813.0, 97349.0, 97884.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0328772971314
varray: [-2.032, -2.091, -2.068, -2.108]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96277.0, 96813.0, 97349.0, 97884.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.191297978731
varray: [-25.313, -25.401, -25.278, -25.699]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96277.0, 96813.0, 97349.0, 97884.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.501
varray: [62.226, 62.265, 61.298, 60.215]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [96277.0, 96813.0, 97349.0, 97884.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.966837800944
varray: [62.226, 62.265, 61.298, 60.215]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [96277.0, 96813.0, 97349.0, 97884.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.303
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.75325
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 154.64375
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 23.80675
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.13915
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.13915
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 6.43066756014
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.3558
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-2.032, -2.091, -2.068, -2.108]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [96277.0, 96813.0, 97349.0, 97884.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-25.313, -25.401, -25.278, -25.699]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [96277.0, 96813.0, 97349.0, 97884.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.07475
varray: [-2.032, -2.091, -2.068, -2.108]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [96277.0, 96813.0, 97349.0, 97884.0]
source_col: d(18_16)Mean
dD_mean | -25.42275
varray: [-25.313, -25.401, -25.278, -25.699]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [96277.0, 96813.0, 97349.0, 97884.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0328772971314
varray: [-2.032, -2.091, -2.068, -2.108]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96277.0, 96813.0, 97349.0, 97884.0]
source_col: d(18_16)Mean
dD_stdev | 0.191297978731
varray: [-25.313, -25.401, -25.278, -25.699]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96277.0, 96813.0, 97349.0, 97884.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.07475
varray: [-2.032, -2.091, -2.068, -2.108]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [96277.0, 96813.0, 97349.0, 97884.0]
source_col: d(18_16)Mean
measured_dD_mean | -25.42275
varray: [-25.313, -25.401, -25.278, -25.699]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [96277.0, 96813.0, 97349.0, 97884.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0328772971314
varray: [-2.032, -2.091, -2.068, -2.108]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96277.0, 96813.0, 97349.0, 97884.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.191297978731
varray: [-25.313, -25.401, -25.278, -25.699]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96277.0, 96813.0, 97349.0, 97884.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.09502737841
actual_array: [ 1.93 -15.84]
varray: [-2.032, -2.091, -2.068, -2.108]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.07475
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [96277.0, 96813.0, 97349.0, 97884.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -28.4172455836
actual_array: [ 18.1 -119. ]
varray: [-25.313, -25.401, -25.278, -25.699]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -25.42275
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [96277.0, 96813.0, 97349.0, 97884.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.0271
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 106.31365
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 40 - Sample 40 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 40
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.12275
varray: [2.175, 2.087, 2.119, 2.11]
inst: mean_all
operation: mean_all(residuals)
tarray: [98468.0, 99003.0, 99539.0, 100074.0]
source_col: residuals
SAMP_residuals_STD | 0.0373485831949
varray: [2.175, 2.087, 2.119, 2.11]
inst: std_all
operation: std_all(residuals)
tarray: [98468.0, 99003.0, 99539.0, 100074.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.96625
varray: [-1.097, -0.953, -0.9, -0.915]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [98468.0, 99003.0, 99539.0, 100074.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0899754596172
varray: [-1.097, -0.953, -0.9, -0.915]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [98468.0, 99003.0, 99539.0, 100074.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 47.60275
varray: [49.11, 47.13, 47.164, 47.007]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [98468.0, 99003.0, 99539.0, 100074.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.00709396285
varray: [49.11, 47.13, 47.164, 47.007]
inst: std_all
operation: std_all(baseline_shift)
tarray: [98468.0, 99003.0, 99539.0, 100074.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.53875
varray: [29.386, 28.192, 28.327, 28.25]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [98468.0, 99003.0, 99539.0, 100074.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.567533479189
varray: [29.386, 28.192, 28.327, 28.25]
inst: std_all
operation: std_all(slope_shift)
tarray: [98468.0, 99003.0, 99539.0, 100074.0]
source_col: slope_shift
SAMP_d18O_STD | 0.171134450068
varray: [-8.581, -8.915, -8.95, -8.896]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [98468.0, 99003.0, 99539.0, 100074.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.17573420558
varray: [-61.022, -64.822, -65.485, -65.639]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [98468.0, 99003.0, 99539.0, 100074.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.24275
varray: [62.922, 61.26, 62.802, 61.987]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [98468.0, 99003.0, 99539.0, 100074.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.775746253616
varray: [62.922, 61.26, 62.802, 61.987]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [98468.0, 99003.0, 99539.0, 100074.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.12275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.96625
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 47.60275
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.53875
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0411
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0411
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.8993922869
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.61405
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-8.581, -8.915, -8.95, -8.896]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [98468.0, 99003.0, 99539.0, 100074.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-61.022, -64.822, -65.485, -65.639]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [98468.0, 99003.0, 99539.0, 100074.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.8355
varray: [-8.581, -8.915, -8.95, -8.896]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [98468.0, 99003.0, 99539.0, 100074.0]
source_col: d(18_16)Mean
dD_mean | -64.242
varray: [-61.022, -64.822, -65.485, -65.639]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [98468.0, 99003.0, 99539.0, 100074.0]
source_col: d(D_H)Mean
d18O_stdev | 0.171134450068
varray: [-8.581, -8.915, -8.95, -8.896]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [98468.0, 99003.0, 99539.0, 100074.0]
source_col: d(18_16)Mean
dD_stdev | 2.17573420558
varray: [-61.022, -64.822, -65.485, -65.639]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [98468.0, 99003.0, 99539.0, 100074.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.8355
varray: [-8.581, -8.915, -8.95, -8.896]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [98468.0, 99003.0, 99539.0, 100074.0]
source_col: d(18_16)Mean
measured_dD_mean | -64.242
varray: [-61.022, -64.822, -65.485, -65.639]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [98468.0, 99003.0, 99539.0, 100074.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.171134450068
varray: [-8.581, -8.915, -8.95, -8.896]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [98468.0, 99003.0, 99539.0, 100074.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.17573420558
varray: [-61.022, -64.822, -65.485, -65.639]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [98468.0, 99003.0, 99539.0, 100074.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.0710648836
actual_array: [ 1.93 -15.84]
varray: [-8.581, -8.915, -8.95, -8.896]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.8355
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [98468.0, 99003.0, 99539.0, 100074.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -70.6766785084
actual_array: [ 18.1 -119. ]
varray: [-61.022, -64.822, -65.485, -65.639]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -64.242
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [98468.0, 99003.0, 99539.0, 100074.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.20735
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.72735
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 41 - Sample 41 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 41
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.2565
varray: [2.295, 2.263, 2.222, 2.246]
inst: mean_all
operation: mean_all(residuals)
tarray: [100658.0, 101193.0, 101728.0, 102263.0]
source_col: residuals
SAMP_residuals_STD | 0.0306865877325
varray: [2.295, 2.263, 2.222, 2.246]
inst: std_all
operation: std_all(residuals)
tarray: [100658.0, 101193.0, 101728.0, 102263.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.60075
varray: [-1.503, -1.566, -1.752, -1.582]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [100658.0, 101193.0, 101728.0, 102263.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.106443647063
varray: [-1.503, -1.566, -1.752, -1.582]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [100658.0, 101193.0, 101728.0, 102263.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 331.29225
varray: [333.933, 330.394, 328.157, 332.685]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [100658.0, 101193.0, 101728.0, 102263.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.55277357333
varray: [333.933, 330.394, 328.157, 332.685]
inst: std_all
operation: std_all(baseline_shift)
tarray: [100658.0, 101193.0, 101728.0, 102263.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 18.15325
varray: [18.269, 18.325, 18.057, 17.962]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [100658.0, 101193.0, 101728.0, 102263.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.171985222233
varray: [18.269, 18.325, 18.057, 17.962]
inst: std_all
operation: std_all(slope_shift)
tarray: [100658.0, 101193.0, 101728.0, 102263.0]
source_col: slope_shift
SAMP_d18O_STD | 0.135023454753
varray: [-2.281, -2.082, -1.953, -2.116]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100658.0, 101193.0, 101728.0, 102263.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.77420167493
varray: [-29.773, -27.029, -25.633, -26.606]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100658.0, 101193.0, 101728.0, 102263.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 58.0
varray: [59.642, 57.838, 56.9, 57.62]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [100658.0, 101193.0, 101728.0, 102263.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.16573410347
varray: [59.642, 57.838, 56.9, 57.62]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [100658.0, 101193.0, 101728.0, 102263.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.2565
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.60075
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 331.29225
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 18.15325
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.09265
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.09265
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 4.28172008226
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.8568
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-2.281, -2.082, -1.953, -2.116]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [100658.0, 101193.0, 101728.0, 102263.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-29.773, -27.029, -25.633, -26.606]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [100658.0, 101193.0, 101728.0, 102263.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.108
varray: [-2.281, -2.082, -1.953, -2.116]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [100658.0, 101193.0, 101728.0, 102263.0]
source_col: d(18_16)Mean
dD_mean | -27.26025
varray: [-29.773, -27.029, -25.633, -26.606]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [100658.0, 101193.0, 101728.0, 102263.0]
source_col: d(D_H)Mean
d18O_stdev | 0.135023454753
varray: [-2.281, -2.082, -1.953, -2.116]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100658.0, 101193.0, 101728.0, 102263.0]
source_col: d(18_16)Mean
dD_stdev | 1.77420167493
varray: [-29.773, -27.029, -25.633, -26.606]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100658.0, 101193.0, 101728.0, 102263.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.108
varray: [-2.281, -2.082, -1.953, -2.116]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [100658.0, 101193.0, 101728.0, 102263.0]
source_col: d(18_16)Mean
measured_dD_mean | -27.26025
varray: [-29.773, -27.029, -25.633, -26.606]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [100658.0, 101193.0, 101728.0, 102263.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.135023454753
varray: [-2.281, -2.082, -1.953, -2.116]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100658.0, 101193.0, 101728.0, 102263.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.77420167493
varray: [-29.773, -27.029, -25.633, -26.606]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100658.0, 101193.0, 101728.0, 102263.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.12933618247
actual_array: [ 1.93 -15.84]
varray: [-2.281, -2.082, -1.953, -2.116]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.108
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [100658.0, 101193.0, 101728.0, 102263.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -30.4175858271
actual_array: [ 18.1 -119. ]
varray: [-29.773, -27.029, -25.633, -26.606]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -27.26025
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [100658.0, 101193.0, 101728.0, 102263.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.0736
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 282.96215
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 42 - Sample 42 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Sample Variables
Sample 42
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.124
varray: [2.181, 2.11, 2.112, 2.093]
inst: mean_all
operation: mean_all(residuals)
tarray: [102848.0, 103384.0, 103919.0, 104455.0]
source_col: residuals
SAMP_residuals_STD | 0.0389444048185
varray: [2.181, 2.11, 2.112, 2.093]
inst: std_all
operation: std_all(residuals)
tarray: [102848.0, 103384.0, 103919.0, 104455.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.90025
varray: [-0.993, -0.937, -0.891, -0.78]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [102848.0, 103384.0, 103919.0, 104455.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0903673060349
varray: [-0.993, -0.937, -0.891, -0.78]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [102848.0, 103384.0, 103919.0, 104455.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 48.397
varray: [50.977, 48.161, 47.573, 46.877]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [102848.0, 103384.0, 103919.0, 104455.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.79828362613
varray: [50.977, 48.161, 47.573, 46.877]
inst: std_all
operation: std_all(baseline_shift)
tarray: [102848.0, 103384.0, 103919.0, 104455.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.2415
varray: [29.113, 28.214, 27.978, 27.661]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [102848.0, 103384.0, 103919.0, 104455.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.623613395195
varray: [29.113, 28.214, 27.978, 27.661]
inst: std_all
operation: std_all(slope_shift)
tarray: [102848.0, 103384.0, 103919.0, 104455.0]
source_col: slope_shift
SAMP_d18O_STD | 0.174289749173
varray: [-8.607, -8.825, -8.926, -9.011]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [102848.0, 103384.0, 103919.0, 104455.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.98087834053
varray: [-61.561, -64.595, -65.44, -66.01]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [102848.0, 103384.0, 103919.0, 104455.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.68625
varray: [63.596, 61.801, 62.973, 62.375]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [102848.0, 103384.0, 103919.0, 104455.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.772531498818
varray: [63.596, 61.801, 62.973, 62.375]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [102848.0, 103384.0, 103919.0, 104455.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.124
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.90025
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 48.397
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.2415
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.03985
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.03985
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.84162488158
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.17055
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.4264
varray: [-8.607, -8.825, -8.926, -9.011]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [102848.0, 103384.0, 103919.0, 104455.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -108.6315
varray: [-61.561, -64.595, -65.44, -66.01]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [102848.0, 103384.0, 103919.0, 104455.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.84225
varray: [-8.607, -8.825, -8.926, -9.011]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [102848.0, 103384.0, 103919.0, 104455.0]
source_col: d(18_16)Mean
dD_mean | -64.4015
varray: [-61.561, -64.595, -65.44, -66.01]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [102848.0, 103384.0, 103919.0, 104455.0]
source_col: d(D_H)Mean
d18O_stdev | 0.174289749173
varray: [-8.607, -8.825, -8.926, -9.011]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [102848.0, 103384.0, 103919.0, 104455.0]
source_col: d(18_16)Mean
dD_stdev | 1.98087834053
varray: [-61.561, -64.595, -65.44, -66.01]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [102848.0, 103384.0, 103919.0, 104455.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.84225
varray: [-8.607, -8.825, -8.926, -9.011]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [102848.0, 103384.0, 103919.0, 104455.0]
source_col: d(18_16)Mean
measured_dD_mean | -64.4015
varray: [-61.561, -64.595, -65.44, -66.01]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [102848.0, 103384.0, 103919.0, 104455.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.174289749173
varray: [-8.607, -8.825, -8.926, -9.011]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [102848.0, 103384.0, 103919.0, 104455.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.98087834053
varray: [-61.561, -64.595, -65.44, -66.01]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [102848.0, 103384.0, 103919.0, 104455.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.0780298288
actual_array: [ 1.93 -15.84]
varray: [-8.607, -8.825, -8.926, -9.011]
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.84225
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: [102848.0, 103384.0, 103919.0, 104455.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -70.8503134846
actual_array: [ 18.1 -119. ]
varray: [-61.561, -64.595, -65.44, -66.01]
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -64.4015
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: [102848.0, 103384.0, 103919.0, 104455.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 46.2061
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.0669
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 1 - Sample 1 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 1 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.604
operation: col(d(18_16)Mean)
inst: col
dD_actual | -75.854
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.864036791
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -9.604
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.317740227
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -75.854
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 1 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.214
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.414
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4616177359
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -9.214
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.4842650263
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -71.414
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 1 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.081
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.657
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3243825196
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -9.081
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.5715587363
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -69.657
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 1 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.931
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.615
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.16960596
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.931
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.4372161329
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -68.615
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 1 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.891
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.058
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1283322107
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.891
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.8308544917
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -68.058
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 1 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.877
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.881
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1138863985
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.877
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.638168656
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -67.881
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 1 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.858
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.55
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0942813676
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.858
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.277835257
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -67.55
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 1 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.891
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.602
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1283322107
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.891
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3344435251
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -67.602
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 1 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.864
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.401
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.10047243
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.864
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.1156307964
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -67.401
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 1 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.849
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.286
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.084994774
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.849
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.9904394343
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -67.286
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 2 - Sample 2 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 2 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.966
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7952
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3077
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.157
operation: col(d(18_16)Mean)
inst: col
dD_actual | 8.037
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.27147733079
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 2.157
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 8.00772555328
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: 8.037
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 2 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.966
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7952
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3077
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.644
operation: col(d(18_16)Mean)
inst: col
dD_actual | 14.966
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.77398522785
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 2.644
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 15.5507772798
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: 14.966
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 2 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.966
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7952
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3077
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.812
operation: col(d(18_16)Mean)
inst: col
dD_actual | 17.033
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.94733497468
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 2.812
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 17.8009559375
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: 17.033
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 2 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.966
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7952
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3077
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.813
operation: col(d(18_16)Mean)
inst: col
dD_actual | 17.922
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.94836681841
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 2.813
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 18.7687395982
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: 17.922
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 2 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.966
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7952
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3077
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.909
operation: col(d(18_16)Mean)
inst: col
dD_actual | 18.799
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.0474238166
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 2.909
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.7234598124
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: 18.799
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 2 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.966
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7952
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3077
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.857
operation: col(d(18_16)Mean)
inst: col
dD_actual | 18.668
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.99376794258
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 2.857
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.5808505215
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: 18.668
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 2 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.966
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7952
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3077
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.916
operation: col(d(18_16)Mean)
inst: col
dD_actual | 18.973
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.05464672272
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 2.916
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.9128797864
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: 18.973
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 2 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.966
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7952
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3077
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.932
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.294
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.07115622242
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 2.932
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.26232698
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: 19.294
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 2 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.966
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7952
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3077
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.946
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.536
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.08560203465
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 2.946
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.5257731509
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: 19.536
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 2 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.966
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.849
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7952
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3077
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.966
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.849
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.10623890928
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 2.966
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.8665113801
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: 19.849
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 3 - Sample 3 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 3 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.531
operation: col(d(18_16)Mean)
inst: col
dD_actual | -95.286
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.9160871231
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -13.531
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -104.471814574
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -95.286
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 3 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.234
operation: col(d(18_16)Mean)
inst: col
dD_actual | -105.733
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.6414732661
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -14.234
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -115.844633363
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -105.733
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 3 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.445
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.436
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.8591922934
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -14.445
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -118.787174684
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -108.436
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 3 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.5
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.698
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9159436986
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -14.5
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.161013807
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -109.698
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 3 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.536
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.396
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9530900729
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -14.536
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.920870944
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -110.396
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 3 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.597
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.983
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0160325405
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -14.597
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.559891201
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -110.983
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 3 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.551
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.057
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9685677289
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -14.551
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.640449121
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -111.057
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 3 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.615
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.389
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0346057277
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -14.615
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.001871141
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -111.389
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 3 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.621
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.584
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0407967901
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -14.621
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.214152146
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -111.584
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 3 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.634
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.753
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0542107586
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -14.634
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.398129018
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -111.753
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 4 - Sample 4 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 4 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.367
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.664
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.6194898267
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -9.367
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.8450407022
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -72.664
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 4 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.2
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.402
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4471719236
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -9.2
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.2939604984
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -69.402
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 4 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.154
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.538
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.399707112
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -9.154
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.3533923512
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -68.538
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 4 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.139
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.205
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.384229456
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -9.139
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.9908817112
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -68.205
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 4 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.09
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.911
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3336691132
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -9.09
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.6708272722
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -67.911
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 4 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.019
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.542
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2604082083
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -9.019
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.2691262927
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -67.542
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 4 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.064
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.633
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3068411762
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -9.064
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3681907619
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -67.633
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 4 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.032
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.748
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2738221768
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -9.032
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.4933821241
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -67.748
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 4 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.03
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.048
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2717584893
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -9.03
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.7313477456
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -67.048
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 4 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.048
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.336
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2903316765
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -9.048
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.0448704613
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -67.336
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 5 - Sample 5 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 5 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.936
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.132
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.01554652297
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -3.936
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -38.9869345684
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -35.132
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 5 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.704
operation: col(d(18_16)Mean)
inst: col
dD_actual | -31.895
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.77615877735
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -3.704
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -35.4630698782
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -31.895
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 5 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.658
operation: col(d(18_16)Mean)
inst: col
dD_actual | -31.224
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.72869396572
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -3.658
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -34.7326054954
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -31.224
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 5 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.581
operation: col(d(18_16)Mean)
inst: col
dD_actual | -30.85
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.64924199842
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -3.581
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -34.3254614131
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -30.85
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 6 - Sample 6 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 6 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.132
operation: col(d(18_16)Mean)
inst: col
dD_actual | -24.608
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.15410043202
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.132
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.530291998
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -24.608
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 6 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.99
operation: col(d(18_16)Mean)
inst: col
dD_actual | -23.275
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.0075786222
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.99
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -26.0791608173
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -23.275
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 6 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.008
operation: col(d(18_16)Mean)
inst: col
dD_actual | -23.296
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.02615180936
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.008
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -26.1020218486
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -23.296
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 6 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.998
operation: col(d(18_16)Mean)
inst: col
dD_actual | -23.214
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.01583337204
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.998
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -26.0127549643
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -23.214
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 7 - Sample 7 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 7 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.404
operation: col(d(18_16)Mean)
inst: col
dD_actual | -24.281
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.43476192688
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.404
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.1743130812
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -24.281
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 7 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.485
operation: col(d(18_16)Mean)
inst: col
dD_actual | -24.484
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.5183412691
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.485
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.395303051
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -24.484
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 7 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.458
operation: col(d(18_16)Mean)
inst: col
dD_actual | -24.56
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.49048148836
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.458
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.4780382121
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -24.56
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 7 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.443
operation: col(d(18_16)Mean)
inst: col
dD_actual | -24.096
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.4750038324
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.443
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -26.9729182812
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -24.096
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 8 - Sample 8 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 8 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.089
operation: col(d(18_16)Mean)
inst: col
dD_actual | -28.14
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.14157488271
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -3.089
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -31.3752997478
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -28.14
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 8 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.12
operation: col(d(18_16)Mean)
inst: col
dD_actual | -28.398
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.17356203837
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -3.12
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -31.6561638473
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -28.398
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 8 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.122
operation: col(d(18_16)Mean)
inst: col
dD_actual | -28.15
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.17562572583
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -3.122
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -31.3861859532
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -28.15
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 8 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.121
operation: col(d(18_16)Mean)
inst: col
dD_actual | -28.15
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.1745938821
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -3.121
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -31.3861859532
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -28.15
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 9 - Sample 9 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 9 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.719
operation: col(d(18_16)Mean)
inst: col
dD_actual | -26.917
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.79163643332
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -3.719
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -30.0439168265
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -26.917
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 9 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.712
operation: col(d(18_16)Mean)
inst: col
dD_actual | -26.457
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.7844135272
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -3.712
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -29.5431513778
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -26.457
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 9 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.747
operation: col(d(18_16)Mean)
inst: col
dD_actual | -26.371
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.82052805779
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -3.747
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -29.4495300113
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -26.371
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 9 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.708
operation: col(d(18_16)Mean)
inst: col
dD_actual | -26.224
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.78028615227
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -3.708
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -29.2895027918
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -26.224
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 10 - Sample 10 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 10 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.803
operation: col(d(18_16)Mean)
inst: col
dD_actual | -27.658
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.8464675756
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.803
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -30.8505846472
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -27.658
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 10 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.779
operation: col(d(18_16)Mean)
inst: col
dD_actual | -28.082
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.82170332606
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.779
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -31.3121597565
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -28.082
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 10 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.784
operation: col(d(18_16)Mean)
inst: col
dD_actual | -28.019
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.82686254471
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.784
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -31.2435766624
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -28.019
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 10 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.841
operation: col(d(18_16)Mean)
inst: col
dD_actual | -28.271
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.88567763739
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.841
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -31.5179090386
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -28.271
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 11 - Sample 11 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 11 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.56
operation: col(d(18_16)Mean)
inst: col
dD_actual | -27.684
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.59572954894
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.56
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -30.8788887813
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -27.684
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 11 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.585
operation: col(d(18_16)Mean)
inst: col
dD_actual | -27.875
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.62152564222
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.585
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -31.0868153045
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -27.875
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 11 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.635
operation: col(d(18_16)Mean)
inst: col
dD_actual | -28.132
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.67311782877
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.635
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -31.3665907835
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -28.132
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 11 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.552
operation: col(d(18_16)Mean)
inst: col
dD_actual | -27.821
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.58747479909
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.552
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -31.0280297953
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -27.821
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 12 - Sample 12 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 12 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.102
operation: col(d(18_16)Mean)
inst: col
dD_actual | -26.273
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.15498885121
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -3.102
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -29.3428451983
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -26.273
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 12 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.131
operation: col(d(18_16)Mean)
inst: col
dD_actual | -26.293
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.18491231941
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -3.131
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -29.3646176091
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -26.293
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 12 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.13
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.82
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.18388047568
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -3.13
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.8497000934
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -25.82
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 12 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.173
operation: col(d(18_16)Mean)
inst: col
dD_actual | -26.199
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.22824975612
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -3.173
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -29.2622872783
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -26.199
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 13 - Sample 13 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 13 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.643
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.346
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.87243496539
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.643
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.5240334225
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -61.346
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 13 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.905
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.711
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1427780229
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.905
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.187241542
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -64.711
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 13 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.886
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.395
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1231729921
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.886
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.9318579918
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -65.395
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 13 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.976
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.921
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2160389279
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.976
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.5044723962
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -65.921
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 14 - Sample 14 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 14 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.565
operation: col(d(18_16)Mean)
inst: col
dD_actual | -26.886
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.60088876759
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.565
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -30.0101695898
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -26.886
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 14 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.228
operation: col(d(18_16)Mean)
inst: col
dD_actual | -23.126
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.2531574302
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.228
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.9169563567
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -23.126
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 14 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.195
operation: col(d(18_16)Mean)
inst: col
dD_actual | -22.384
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.21910658708
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.195
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.1091999155
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -22.384
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 14 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.192
operation: col(d(18_16)Mean)
inst: col
dD_actual | -22.012
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.21601105588
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.192
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -24.7042330744
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -22.012
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 15 - Sample 15 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 15 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.979
operation: col(d(18_16)Mean)
inst: col
dD_actual | -31.749
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.05991580341
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -3.979
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -35.3041312792
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -31.749
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 15 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.016
operation: col(d(18_16)Mean)
inst: col
dD_actual | -32.496
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.09809402146
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -4.016
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -36.1173308231
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -32.496
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 15 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.086
operation: col(d(18_16)Mean)
inst: col
dD_actual | -32.867
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.17032308264
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -4.086
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -36.5212090437
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -32.867
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 15 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.103
operation: col(d(18_16)Mean)
inst: col
dD_actual | -32.952
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.18786442607
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -4.103
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -36.6137417897
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -32.952
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 16 - Sample 16 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 16 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.236
operation: col(d(18_16)Mean)
inst: col
dD_actual | -21.258
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.710694476704
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 0.236
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -23.8834131867
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -21.258
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 16 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.3
operation: col(d(18_16)Mean)
inst: col
dD_actual | -20.949
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.644656477911
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 0.3
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -23.5470294396
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -20.949
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 16 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.249
operation: col(d(18_16)Mean)
inst: col
dD_actual | -20.989
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.697280508199
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 0.249
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -23.5905742612
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -20.989
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 16 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.307
operation: col(d(18_16)Mean)
inst: col
dD_actual | -20.635
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.637433571794
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 0.307
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -23.2052025898
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -20.635
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 17 - Sample 17 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 17 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.932
operation: col(d(18_16)Mean)
inst: col
dD_actual | -20.935
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.91588795466
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -0.932
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -23.531788752
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -20.935
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 17 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.873
operation: col(d(18_16)Mean)
inst: col
dD_actual | -20.477
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.85500917453
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -0.873
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -23.0332005444
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -20.477
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 17 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.885
operation: col(d(18_16)Mean)
inst: col
dD_actual | -20.37
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.8673912993
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -0.885
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -22.9167181465
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -20.37
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 17 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.911
operation: col(d(18_16)Mean)
inst: col
dD_actual | -20.374
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.89421923631
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -0.911
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -22.9210726287
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -20.374
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 18 - Sample 18 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 18 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.015
operation: col(d(18_16)Mean)
inst: col
dD_actual | -24.766
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.03337471547
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.015
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.7022940435
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -24.766
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 18 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.097
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.228
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.11798590143
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.097
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.2052367333
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -25.228
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 18 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.005
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.018
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.02305627816
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.005
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.9766264197
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -25.018
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 18 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.047
operation: col(d(18_16)Mean)
inst: col
dD_actual | -24.881
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.06639371487
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.047
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.8274854057
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -24.881
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 19 - Sample 19 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 19 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.784
operation: col(d(18_16)Mean)
inst: col
dD_actual | -18.509
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.91844335021
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 2.784
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -20.8907953203
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -18.509
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 19 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.914
operation: col(d(18_16)Mean)
inst: col
dD_actual | -17.941
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.05258303526
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 2.914
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -20.2724588532
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -17.941
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 19 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.893
operation: col(d(18_16)Mean)
inst: col
dD_actual | -18.212
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.0309143169
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 2.893
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -20.5674750197
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -18.212
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 19 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.941
operation: col(d(18_16)Mean)
inst: col
dD_actual | -17.729
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.080442816
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 2.941
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -20.0416712985
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -17.729
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 20 - Sample 20 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 20 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.609
operation: col(d(18_16)Mean)
inst: col
dD_actual | -24.748
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.325816764993
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 0.609
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.6826988737
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -24.748
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 20 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.544
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.528
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.392886607516
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 0.544
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.5318228955
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -25.528
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 20 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.592
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.359
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.343358108422
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 0.592
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.3478460241
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -25.359
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 20 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.59
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.314
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.345421795884
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 0.59
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.2988580998
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -25.314
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 21 - Sample 21 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 21 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.705
operation: col(d(18_16)Mean)
inst: col
dD_actual | -22.755
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.71350315882
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.705
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.5130781361
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -22.755
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 21 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.655
operation: col(d(18_16)Mean)
inst: col
dD_actual | -22.317
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.66191097227
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.655
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.0362623393
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -22.317
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 21 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.673
operation: col(d(18_16)Mean)
inst: col
dD_actual | -21.985
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.68048415943
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.673
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -24.6748403198
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -21.985
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 21 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.66
operation: col(d(18_16)Mean)
inst: col
dD_actual | -22.066
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.66707019092
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.66
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -24.7630185836
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -22.066
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 22 - Sample 22 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 22 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.562
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.056
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.78885562317
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.562
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.2083334657
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -61.056
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 22 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.842
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.329
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0777718679
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.842
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.7713884954
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -64.329
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 22 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.877
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.337
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1138863985
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.877
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.8687180004
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -65.337
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 22 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.93
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.786
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1685741162
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.93
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.3575086232
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -65.786
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 23 - Sample 23 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 23 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.316
operation: col(d(18_16)Mean)
inst: col
dD_actual | -31.836
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.31211594741
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.316
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -35.3988412663
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -31.836
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 23 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.061
operation: col(d(18_16)Mean)
inst: col
dD_actual | -29.189
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.04899579598
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.061
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -32.517262695
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -29.189
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 23 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.93
operation: col(d(18_16)Mean)
inst: col
dD_actual | -28.152
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.9138242672
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -0.93
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -31.3883631943
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -28.152
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 23 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.94
operation: col(d(18_16)Mean)
inst: col
dD_actual | -28.03
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.92414270451
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -0.94
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -31.2555514883
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -28.03
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 24 - Sample 24 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 24 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.821
operation: col(d(18_16)Mean)
inst: col
dD_actual | -21.838
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.924777837135
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 1.821
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -24.5148131003
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -21.838
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 24 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.843
operation: col(d(18_16)Mean)
inst: col
dD_actual | -21.43
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.94747839922
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 1.843
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -24.0706559197
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -21.43
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 24 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.847
operation: col(d(18_16)Mean)
inst: col
dD_actual | -21.44
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.951605774144
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 1.847
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -24.0815421251
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -21.44
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 24 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.793
operation: col(d(18_16)Mean)
inst: col
dD_actual | -21.754
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.895886212663
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: 1.793
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -24.4233689749
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -21.754
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 25 - Sample 25 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 25 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.52
operation: col(d(18_16)Mean)
inst: col
dD_actual | -23.266
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.55445579969
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.52
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -26.0693632324
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -23.266
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 25 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.519
operation: col(d(18_16)Mean)
inst: col
dD_actual | -22.682
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.55342395596
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.519
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.4336088366
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -22.682
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 25 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.587
operation: col(d(18_16)Mean)
inst: col
dD_actual | -22.893
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.62358932968
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.587
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.6633077707
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -22.893
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 25 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.567
operation: col(d(18_16)Mean)
inst: col
dD_actual | -22.922
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.60295245506
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.567
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.6948777664
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -22.922
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 26 - Sample 26 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 26 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.683
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.57
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.72264632787
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.683
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.5775449582
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -25.57
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 26 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.758
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.987
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.8000346077
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.758
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -29.0314997237
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -25.987
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 26 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.742
operation: col(d(18_16)Mean)
inst: col
dD_actual | -26.411
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.783525108
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.742
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -29.4930748329
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -26.411
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 26 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.737
operation: col(d(18_16)Mean)
inst: col
dD_actual | -26.143
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.77836588935
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.737
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -29.201324528
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -26.143
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 27 - Sample 27 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 27 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.536
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.278
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.53912156826
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.536
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.2596677603
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -25.278
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 27 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.485
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.197
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.48649753798
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.485
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.1714894965
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -25.197
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 27 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.435
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.138
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.43490535142
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.435
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.1072608846
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -25.138
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 27 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.469
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.409
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.46998803828
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.469
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.4022770511
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -25.409
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 28 - Sample 28 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 28 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.945
operation: col(d(18_16)Mean)
inst: col
dD_actual | -26.595
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.92930192317
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -0.945
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -29.6933810124
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -26.595
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 28 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.913
operation: col(d(18_16)Mean)
inst: col
dD_actual | -26.696
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.89628292377
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -0.913
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -29.803331687
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -26.696
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 28 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.912
operation: col(d(18_16)Mean)
inst: col
dD_actual | -26.823
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.89525108004
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -0.912
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -29.9415864957
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -26.823
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 28 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.959
operation: col(d(18_16)Mean)
inst: col
dD_actual | -27.3
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.9437477354
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -0.959
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -30.4608584936
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -27.3
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 29 - Sample 29 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 29 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.86
operation: col(d(18_16)Mean)
inst: col
dD_actual | -28.194
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.84159520602
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -0.86
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -31.434085257
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -28.194
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 29 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.797
operation: col(d(18_16)Mean)
inst: col
dD_actual | -27.731
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.77658905096
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -0.797
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -30.9300539467
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -27.731
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 29 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.813
operation: col(d(18_16)Mean)
inst: col
dD_actual | -27.899
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.79309855066
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -0.813
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -31.1129421975
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -27.899
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 29 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.845
operation: col(d(18_16)Mean)
inst: col
dD_actual | -28.119
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.82611755005
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -0.845
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -31.3524387165
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -28.119
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 30 - Sample 30 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 30 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.309
operation: col(d(18_16)Mean)
inst: col
dD_actual | -24.89
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.3048930413
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.309
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.8372829905
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -24.89
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 30 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.391
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.035
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.38950422725
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.391
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.9951329689
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -25.035
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 30 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.401
operation: col(d(18_16)Mean)
inst: col
dD_actual | -24.98
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.39982266456
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.401
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.9352588392
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -24.98
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 30 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.451
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.136
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.45141485112
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.451
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.1050836435
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -25.136
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 31 - Sample 31 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 31 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.522
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.941
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.74758187393
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.522
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.0831421035
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -60.941
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 31 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.803
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.283
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0375299624
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.803
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.7213119505
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -64.283
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 31 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.907
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.325
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1448417104
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.907
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.8556545539
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -65.325
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 31 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.934
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.764
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1727014912
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.934
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.3335589713
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -65.764
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 32 - Sample 32 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 32 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.996
operation: col(d(18_16)Mean)
inst: col
dD_actual | -32.785
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.04561341571
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.996
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -36.4319421594
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -32.785
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 32 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.69
operation: col(d(18_16)Mean)
inst: col
dD_actual | -29.046
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.72986923399
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.69
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -32.3615899577
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -29.046
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 32 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.664
operation: col(d(18_16)Mean)
inst: col
dD_actual | -28.353
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.70304129698
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.664
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -31.607175923
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -28.353
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 32 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.654
operation: col(d(18_16)Mean)
inst: col
dD_actual | -28.149
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.69272285966
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.654
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -31.3850973327
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -28.149
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 33 - Sample 33 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 33 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.063
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.141
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.05105948344
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.063
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.1105267462
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -25.141
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 33 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.957
operation: col(d(18_16)Mean)
inst: col
dD_actual | -24.344
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.94168404794
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -0.957
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.2428961753
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -24.344
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 33 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.958
operation: col(d(18_16)Mean)
inst: col
dD_actual | -24.352
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.94271589167
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -0.958
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.2516051396
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -24.352
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 33 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.915
operation: col(d(18_16)Mean)
inst: col
dD_actual | -23.802
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.89834661123
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -0.915
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -26.6528638422
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -23.802
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 34 - Sample 34 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 34 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.037
operation: col(d(18_16)Mean)
inst: col
dD_actual | -27.704
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.05607527756
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.037
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -30.9006611921
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -27.704
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 34 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.022
operation: col(d(18_16)Mean)
inst: col
dD_actual | -27.453
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.04059762159
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.022
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -30.6274174364
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -27.453
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 34 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.016
operation: col(d(18_16)Mean)
inst: col
dD_actual | -27.705
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.0344065592
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.016
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -30.9017498126
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -27.705
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 34 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.018
operation: col(d(18_16)Mean)
inst: col
dD_actual | -27.758
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.03647024667
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.018
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -30.9594467013
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -27.758
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 35 - Sample 35 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 35 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.594
operation: col(d(18_16)Mean)
inst: col
dD_actual | -24.961
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.6308122358
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.594
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.9145750489
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -24.961
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 35 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.608
operation: col(d(18_16)Mean)
inst: col
dD_actual | -24.751
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.64525804803
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.608
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.6859647354
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -24.751
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 35 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.626
operation: col(d(18_16)Mean)
inst: col
dD_actual | -24.813
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.66383123519
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.626
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.7534592089
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -24.813
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 35 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.58
operation: col(d(18_16)Mean)
inst: col
dD_actual | -24.328
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.61636642356
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.58
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.2254782466
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -24.328
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 36 - Sample 36 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 36 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.574
operation: col(d(18_16)Mean)
inst: col
dD_actual | -23.717
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.61017536117
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.574
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -26.5603310963
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -23.717
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 36 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.544
operation: col(d(18_16)Mean)
inst: col
dD_actual | -23.984
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.57922004924
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.544
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -26.8509927806
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -23.984
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 36 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.566
operation: col(d(18_16)Mean)
inst: col
dD_actual | -23.83
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.60192061133
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.566
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -26.6833452174
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -23.83
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 36 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.599
operation: col(d(18_16)Mean)
inst: col
dD_actual | -23.636
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.63597145445
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.599
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -26.4721528325
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -23.636
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 37 - Sample 37 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 37 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.849
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.372
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.86208865611
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.849
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.3619980911
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -25.372
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 37 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.784
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.751
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.79501881358
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.784
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.7745852761
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -25.751
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 37 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.811
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.621
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.82287859432
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.811
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.6330646058
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -25.621
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 37 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.783
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.395
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.79398696985
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.783
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.3870363636
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -25.395
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 38 - Sample 38 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 38 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.024
operation: col(d(18_16)Mean)
inst: col
dD_actual | -24.57
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.01081757793
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.024
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.4889244175
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -24.57
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 38 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.086
operation: col(d(18_16)Mean)
inst: col
dD_actual | -24.689
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.07479188926
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.086
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.6184702618
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -24.689
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 38 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.007
operation: col(d(18_16)Mean)
inst: col
dD_actual | -24.311
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.9932762345
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.007
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.2069716975
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -24.311
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 38 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.07
operation: col(d(18_16)Mean)
inst: col
dD_actual | -24.663
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.05828238956
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.07
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.5901661278
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -24.663
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 39 - Sample 39 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 39 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.032
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.313
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.0509160589
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.032
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.2977694792
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -25.313
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 39 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.091
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.401
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.11179483904
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.091
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.3935680868
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -25.401
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 39 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.068
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.278
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.08806243322
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.068
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.2596677603
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -25.278
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 39 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.108
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.699
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.12933618247
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.108
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.717977008
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -25.699
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 40 - Sample 40 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 40 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.581
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.022
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.80846065406
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.581
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.1713203673
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -61.022
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 40 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.915
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.822
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1530964603
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.915
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.308078422
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -64.822
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 40 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.95
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.485
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1892109908
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.95
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.0298338405
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -65.485
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 40 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.896
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.639
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1334914294
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.896
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.1974814037
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -65.639
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 41 - Sample 41 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 41 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.281
operation: col(d(18_16)Mean)
inst: col
dD_actual | -29.773
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.30784514795
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.281
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -33.1530170908
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -29.773
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 41 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.082
operation: col(d(18_16)Mean)
inst: col
dD_actual | -27.029
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.10250824546
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.082
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -30.1658423271
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -27.029
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 41 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.953
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.633
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -2.96940040414
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -1.953
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.6461280523
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -25.633
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 41 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.116
operation: col(d(18_16)Mean)
inst: col
dD_actual | -26.606
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.13759093232
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -2.116
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -29.7053558384
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -26.606
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean

## 42 - Sample 42 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 13:48:09
Detail Variables
Detail 42 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.607
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.561
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.83528859107
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.607
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.7580868387
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -61.561
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 42 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.825
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.595
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0602305245
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.0609615592
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -64.595
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 42 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.926
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.44
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1644467413
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -8.926
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.9808459161
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -65.44
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean
Detail 42 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7952
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3077
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.634
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.753
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.4264
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -108.6315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.011
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.01
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2521534584
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.924761733258
measure_array: [ 2.7952 -14.4264]
inst: fit_type1
actual_mean: -9.011
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.969138998312
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.6013596243
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.681153537564
measure_array: [ 17.3077 -108.6315]
inst: fit_type1
actual_mean: -66.01
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.918593727206
tarray: []
source_col: d(D_H)Mean