## Summary
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
230425_23-074_23-096_HIDS2052
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 | * | EV | -9.4564 | -69.7837 |  |  |  | 2.3495 | 19.416
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.28695 | 4.1906
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.39107 | 6.2768
4 |  | EV | -10.4225 | -75.4514 |  |  |  | 0.12039 | 1.9423
5 |  | WOOD4-010-1 | -11.8153 | -96.6541 |  |  |  | 0.059701 | 1.4135
6 |  | WOOD4-B-4 | -9.9763 | -91.7307 |  |  |  | 0.05945 | 0.24577
7 |  | WOOD4-D-4 | 13.0158 | -43.2243 | -13.6635 | 368.7114 | 113.437 | 0.15471 | 1.0461 | -51.8222 | 22.7982
8 |  | WOOD4-B-3 | 5.1895 | -57.4481 | -9.4623 | 929.7646 | 51.1678 | 0.091994 | 0.41605 | -57.3105 | 4.4127
9 |  | WOOD4-010-2 | -8.7138 | -86.6567 |  |  |  | 0.094165 | 0.69957
10 |  | WOOD4-D-1 | 10.3559 | -52.085 | -12.729 | 726.4534 | 98.4441 | 0.12157 | 0.5126 | -61.5538 | 16.0062
11 |  | WOOD4-2030-2 | -11.2978 | -104.0423 |  |  |  | 0.20074 | 1.6951
12 |  | WOOD4-deep-3 | -15.1272 | -119.6614 |  |  |  | 0.11165 | 0.99066
13 | * | EV | -8.5499 | -61.7877 |  |  |  | 3.8282 | 30.8638
14 | * | WOOD4-deep-1 | -11.524 | -90.0091 |  |  |  | 5.8033 | 50.8743
15 |  | WOOD4-3040-1 | -14.2877 | -114.2969 |  |  |  | 0.01919 | 0.064813
16 |  | WOOD4-d-5 | 9.2863 | -54.0858 | -11.2287 | 614.9281 | 79.1101 | 0.29393 | 2.4341 | -53.088 | 11.9246
17 |  | WOOD4-1020-1 | -10.8858 | -92.8191 |  |  |  | 0.15158 | 0.79858
18 |  | WOOD4-2030-1 | -12.5453 | -108.824 |  |  |  | 0.080914 | 0.99233
19 |  | WOOD4-010-3 | -12.1631 | -100.5998 |  |  |  | 0.039911 | 0.43666
20 |  | WOOD4-1020-3 | -11.1857 | -91.5004 |  |  |  | 0.024125 | 0.55009
21 | * | WOOD4-3040-3 | -11.7157 | -91.6341 |  |  |  | 5.9307 | 52.0258
22 | * | EV | -8.5587 | -61.6514 |  |  |  | 3.8341 | 30.7679
23 |  | WOOD4-B-1 | -4.4957 | -99.2893 | -6.371 | 437.6934 | 22.0519 | 0.037652 | 2.2816 | -32.336 | 0.7564
24 |  | WOOD4-2030-3 | -13.8787 | -113.447 |  |  |  | 0.03353 | 0.068103
25 |  | WOOD4-1020-2 | -7.6627 | -84.0206 |  |  |  | 0.2055 | 2.0011
26 |  | WOOD4-C-5 | -6.9194 | -83.5907 |  |  |  | 0.042458 | 0.18133
27 |  | WOOD4-E-1 | -9.6556 | -86.4804 |  | 407.9931 |  | 0.10323 | 0.30066 | -14.0718 | -16.5088
28 |  | WOOD4-A-1 | -8.1692 | -76.4512 | -1.2208 | 1025.2574 |  | 0.029799 | 0.58917 | -38.824 | -20.4453
29 | * | WOOD4-D-2 | -8.5624 | -77.9056 |  |  |  | 3.8365 | 42.2898
30 | * | WOOD4-C-2 | -4.5283 | -58.1866 |  |  |  | 1.1584 | 28.3112
31 |  | EV | -10.2953 | -75.0274 |  |  |  | 0.17919 | 0.294
32 |  | WOOD4-C-4 | -6.0423 | -76.977 |  |  |  | 0.14288 | 0.15218
33 |  | WOOD4-A-4 | -7.9825 | -85.3003 |  |  |  | 0.12206 | 0.8879
34 |  | WOOD4-A-2 | -7.3864 | -77.1342 |  |  |  | 0.042201 | 0.55824
35 |  | WOOD4-E-3 | -8.9055 | -86.7401 |  |  |  | 0.053544 | 0.7275
36 |  | WOOD4-C-1 | -7.3255 | -79.9597 |  |  |  | 0.073469 | 0.61531
37 | * | WOOD4-A-3 | 10.4377 | -32.7234 | -9.5613 |  | 49.9532 | 8.783 | 10.2794 | -35.9308 | -24.1666
38 |  | WOOD4-C-3 | -9.505 | -83.7536 |  | 114.0151 |  | 0.20033 | 1.022 |  | -10.3051
39 |  | WOOD4-D-3 | -8.6412 | -86.9314 |  |  |  | 0.096927 | 0.057703
40 | * | EV | -8.4457 | -60.9263 |  |  |  | 3.7592 | 30.2483
41 |  | WOOD4-E-2 | -7.5531 | -81.9636 |  |  |  | 0.09456 | 0.4883
42 |  | WOOD4-B-5 | 2.5474 | -69.4229 | -7.1145 | 228.0131 | 24.9241 | 0.1029 | 0.45901 | -27.7547 | -6.0013
43 |  | WOOD4-deep-2 | -12.2889 | -110.2642 |  |  |  | 0.17607 | 1.6946
44 |  | WOOD4-B-2 | 11.1979 | -52.5717 | -10.9565 | 670.6504 | 66.7104 | 0.25544 | 2.499 | -54.7747 | -12.4551
45 |  | WOOD4-3040-2 | -12.198 | -109.3417 |  |  |  | 0.20966 | 1.9247
46 | * | WOOD4-E-5 | -6.0403 | -62.1475 |  |  |  | 2.1615 | 31.1174
47 |  | A00000277285 | -4.9241 | -31.0606 |  |  |  | 0.062222 | 3.0076
48 | * | EV | -8.3571 | -59.6691 |  |  |  | 3.7006 | 29.364

## Detail
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
1 |  | EV |  |  | -2.7857 | -18.2712 |  |  | 2023/04/25 12:34:11
1 |  | EV | -7.973 | -66.916 | -10.7892 | -81.1802 | 0.644 | 1328.557 | 2023/04/25 12:44:13
1 |  | EV | -7.588 | -62.272 | -10.4027 | -76.8143 | 0.348 | 1346.381 | 2023/04/25 12:53:09
1 |  | EV | -7.563 | -62.339 | -10.3776 | -76.8773 | 0.413 | 1379.411 | 2023/04/25 13:02:06
1 |  | EV | -7.361 | -60.824 | -10.1748 | -75.453 | 0.302 | 1324.937 | 2023/04/25 13:11:03
1 |  | EV | -7.255 | -59.622 | -10.0684 | -74.323 | 0.446 | 1325.198 | 2023/04/25 13:20:00
1 |  | EV | -7.238 | -59.615 | -10.0513 | -74.3164 | 0.512 | 1310.335 | 2023/04/25 13:28:56
1 |  | EV | -7.171 | -58.914 | -9.9841 | -73.6574 | 0.324 | 1307.313 | 2023/04/25 13:37:54
1 |  | EV | -7.121 | -58.741 | -9.9339 | -73.4948 | 0.487 | 1305.189 | 2023/04/25 13:46:51
1 |  | EV | -7.183 | -58.693 | -9.9961 | -73.4496 | 0.543 | 1303.666 | 2023/04/25 13:55:48
Sample 1 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 | * | EV | -9.4564 | -69.7837 |  |  |  | 2.3495 | 19.416
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
2 |  | PZ | 3.937 | 27.51 | 1.1664 | 7.5914 | 0.36 | 1301.698 | 2023/04/25 14:05:36
2 |  | PZ | 4.541 | 36.325 | 1.7727 | 15.8786 | 0.182 | 1302.669 | 2023/04/25 14:14:33
2 |  | PZ | 4.7 | 38.866 | 1.9323 | 18.2674 | 0.258 | 1304.042 | 2023/04/25 14:23:30
2 |  | PZ | 4.777 | 39.653 | 2.0096 | 19.0073 | 0.352 | 1303.083 | 2023/04/25 14:32:27
2 |  | PZ | 4.769 | 40.114 | 2.0016 | 19.4407 | 0.808 | 1302.412 | 2023/04/25 14:41:24
2 |  | PZ | 4.784 | 40.396 | 2.0166 | 19.7058 | 0.314 | 1303.232 | 2023/04/25 14:50:22
2 |  | PZ | 4.832 | 40.695 | 2.0648 | 19.9869 | 0.628 | 1302.083 | 2023/04/25 14:59:19
2 |  | PZ | 4.852 | 41.027 | 2.0849 | 20.299 | 0.379 | 1302.591 | 2023/04/25 15:08:16
2 |  | PZ | 4.918 | 41.052 | 2.1511 | 20.3225 | 0.495 | 1302.382 | 2023/04/25 15:17:13
2 |  | PZ | 4.867 | 41.241 | 2.0999 | 20.5002 | 0.516 | 1302.557 | 2023/04/25 15:26:11
Sample 2 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.28695 | 4.1906
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
3 |  | UT2 | -11.957 | -90.283 | -14.7884 | -103.148 | 0.675 | 1299.562 | 2023/04/25 15:36:00
3 |  | UT2 | -12.806 | -104.086 | -15.6406 | -116.1245 | 0.414 | 1301.321 | 2023/04/25 15:44:57
3 |  | UT2 | -13.051 | -107.145 | -15.8866 | -119.0003 | 0.522 | 1301.135 | 2023/04/25 15:53:54
3 |  | UT2 | -13.053 | -108.517 | -15.8886 | -120.2901 | 0.345 | 1301.284 | 2023/04/25 16:02:52
3 |  | UT2 | -13.144 | -109.271 | -15.9799 | -120.999 | 0.508 | 1301.216 | 2023/04/25 16:11:48
3 |  | UT2 | -13.151 | -109.877 | -15.987 | -121.5687 | 0.474 | 1301.536 | 2023/04/25 16:20:46
3 |  | UT2 | -13.224 | -110.253 | -16.0602 | -121.9222 | 0.407 | 1301.078 | 2023/04/25 16:29:44
3 |  | UT2 | -13.199 | -110.5 | -16.0351 | -122.1544 | 0.425 | 1301.359 | 2023/04/25 16:38:41
3 |  | UT2 | -13.175 | -110.672 | -16.0111 | -122.3161 | 0.333 | 1301.224 | 2023/04/25 16:47:38
3 |  | UT2 | -13.286 | -110.843 | -16.1225 | -122.4768 | 0.404 | 1300.274 | 2023/04/25 16:56:35
Sample 3 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.39107 | 6.2768
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
4 |  | EV | -7.945 | -66.123 | -10.7611 | -80.4347 | 0.23 | 1300.303 | 2023/04/25 17:06:24
4 |  | EV | -7.578 | -61.595 | -10.3926 | -76.1779 | 0.603 | 1300.315 | 2023/04/25 17:15:21
4 |  | EV | -7.585 | -60.611 | -10.3997 | -75.2528 | 0.274 | 1298.504 | 2023/04/25 17:24:19
4 |  | EV | -7.595 | -60.407 | -10.4097 | -75.061 | 0.198 | 1298.801 | 2023/04/25 17:33:16
4 |  | EV | -7.565 | -60.11 | -10.3796 | -74.7818 | 0.038 | 1297.719 | 2023/04/25 17:42:12
4 |  | EV | -7.536 | -59.737 | -10.3505 | -74.4311 | 0.362 | 1297.981 | 2023/04/25 17:51:09
4 |  | EV | -7.573 | -59.848 | -10.3876 | -74.5355 | 0.298 | 1296.921 | 2023/04/25 18:00:06
4 |  | EV | -7.532 | -59.885 | -10.3465 | -74.5703 | 0.319 | 1296.548 | 2023/04/25 18:09:03
4 |  | EV | -7.577 | -60.05 | -10.3916 | -74.7254 | 0.258 | 1297.248 | 2023/04/25 18:18:00
4 |  | EV | -7.591 | -59.857 | -10.4057 | -74.5439 | 0.484 | 1296.973 | 2023/04/25 18:26:58
Sample 4 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
4 |  | EV | -10.4225 | -75.4514 |  |  |  | 0.12039 | 1.9423
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
5 |  | WOOD4-010-1 | -8.926 | -81.38 | -11.7458 | -94.7781 | 0.382 | 1297.568 | 2023/04/25 18:36:48
5 |  | WOOD4-010-1 | -8.969 | -83.399 | -11.789 | -96.6762 | 0.277 | 1299.544 | 2023/04/25 18:45:45
5 |  | WOOD4-010-1 | -9.025 | -84.178 | -11.8452 | -97.4086 | 0.589 | 1297.768 | 2023/04/25 18:54:42
5 |  | WOOD4-010-1 | -9.061 | -84.545 | -11.8813 | -97.7536 | 0.401 | 1297.128 | 2023/04/25 19:03:40
Sample 5 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
5 |  | WOOD4-010-1 | -11.8153 | -96.6541 |  |  |  | 0.059701 | 1.4135
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
6 |  | WOOD4-B-4 | -7.238 | -78.453 | -10.0513 | -92.0264 | 0.392 | 1300.071 | 2023/04/25 19:13:30
6 |  | WOOD4-B-4 | -7.184 | -78.214 | -9.9971 | -91.8017 | -0.092 | 1298.901 | 2023/04/25 19:22:26
6 |  | WOOD4-B-4 | -7.119 | -77.927 | -9.9319 | -91.5319 | 0.143 | 1299.358 | 2023/04/25 19:31:24
6 |  | WOOD4-B-4 | -7.112 | -77.96 | -9.9249 | -91.5629 | 0.337 | 1298.44 | 2023/04/25 19:40:21
Sample 6 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
6 |  | WOOD4-B-4 | -9.9763 | -91.7307 |  |  |  | 0.05945 | 0.24577
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
7 |  | WOOD4-D-4 | 15.521 | -28.048 | 12.7947 | -44.6397 | -13.878 | 1680.041 | 2023/04/25 19:50:12
7 |  | WOOD4-D-4 | 15.746 | -26.456 | 13.0206 | -43.143 | -13.437 | 1667.745 | 2023/04/25 19:59:09
7 |  | WOOD4-D-4 | 15.854 | -25.865 | 13.129 | -42.5874 | -13.825 | 1668.101 | 2023/04/25 20:08:06
7 |  | WOOD4-D-4 | 15.844 | -25.801 | 13.1189 | -42.5272 | -13.514 | 1666.306 | 2023/04/25 20:17:04
Sample 7 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
7 |  | WOOD4-D-4 | 13.0158 | -43.2243 | -13.6635 | 368.7114 | 113.437 | 0.15471 | 1.0461 | -51.8222 | 22.7982
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
8 |  | WOOD4-B-3 | 8.079 | -41.103 | 5.3242 | -56.9129 | -9.664 | 2255.738 | 2023/04/25 20:26:50
8 |  | WOOD4-B-3 | 7.93 | -41.734 | 5.1747 | -57.5062 | -9.259 | 2218.743 | 2023/04/25 20:35:47
8 |  | WOOD4-B-3 | 7.883 | -42.103 | 5.1275 | -57.8531 | -9.279 | 2225.417 | 2023/04/25 20:44:44
8 |  | WOOD4-B-3 | 7.887 | -41.749 | 5.1315 | -57.5203 | -9.647 | 2226.508 | 2023/04/25 20:53:41
Sample 8 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
8 |  | WOOD4-B-3 | 5.1895 | -57.4481 | -9.4623 | 929.7646 | 51.1678 | 0.091994 | 0.41605 | -57.3105 | 4.4127
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
9 |  | WOOD4-010-2 | -5.765 | -71.707 | -8.5727 | -85.6843 | 0.148 | 1311.727 | 2023/04/25 21:03:30
9 |  | WOOD4-010-2 | -5.945 | -72.919 | -8.7534 | -86.8238 | 0.313 | 1301.922 | 2023/04/25 21:12:27
9 |  | WOOD4-010-2 | -5.966 | -73.162 | -8.7745 | -87.0522 | 0.218 | 1300.054 | 2023/04/25 21:21:24
9 |  | WOOD4-010-2 | -5.946 | -73.177 | -8.7544 | -87.0663 | 0.484 | 1298.796 | 2023/04/25 21:30:22
Sample 9 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
9 |  | WOOD4-010-2 | -8.7138 | -86.6567 |  |  |  | 0.094165 | 0.69957
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
10 |  | WOOD4-D-1 | 12.915 | -36.736 | 10.1787 | -52.8074 | -12.55 | 2028.928 | 2023/04/25 21:40:10
10 |  | WOOD4-D-1 | 13.11 | -35.73 | 10.3745 | -51.8617 | -12.938 | 2032.711 | 2023/04/25 21:49:07
10 |  | WOOD4-D-1 | 13.157 | -35.714 | 10.4217 | -51.8466 | -12.942 | 2025.968 | 2023/04/25 21:58:05
10 |  | WOOD4-D-1 | 13.184 | -35.69 | 10.4488 | -51.8241 | -12.486 | 2025.554 | 2023/04/25 22:07:02
Sample 10 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
10 |  | WOOD4-D-1 | 10.3559 | -52.085 | -12.729 | 726.4534 | 98.4441 | 0.12157 | 0.5126 | -61.5538 | 16.0062
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
11 |  | WOOD4-2030-2 | -8.186 | -88.77 | -11.003 | -101.7256 | 0.114 | 1296.462 | 2023/04/25 22:16:51
11 |  | WOOD4-2030-2 | -8.545 | -91.492 | -11.3633 | -104.2846 | 0.228 | 1290.26 | 2023/04/25 22:25:48
11 |  | WOOD4-2030-2 | -8.548 | -92.185 | -11.3664 | -104.9361 | 0.404 | 1289.341 | 2023/04/25 22:34:44
11 |  | WOOD4-2030-2 | -8.64 | -92.49 | -11.4587 | -105.2228 | 0.395 | 1288.27 | 2023/04/25 22:43:42
Sample 11 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
11 |  | WOOD4-2030-2 | -11.2978 | -104.0423 |  |  |  | 0.20074 | 1.6951
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
12 |  | WOOD4-deep-3 | -12.128 | -106.41 | -14.96 | -118.3093 | 0.585 | 1291.536 | 2023/04/25 22:53:32
12 |  | WOOD4-deep-3 | -12.334 | -108.026 | -15.1668 | -119.8285 | 0.266 | 1292.706 | 2023/04/25 23:02:29
12 |  | WOOD4-deep-3 | -12.363 | -108.321 | -15.1959 | -120.1059 | -0.118 | 1295.176 | 2023/04/25 23:11:29
12 |  | WOOD4-deep-3 | -12.353 | -108.636 | -15.1859 | -120.402 | 0.036 | 1291.82 | 2023/04/25 23:20:26
Sample 12 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
12 |  | WOOD4-deep-3 | -15.1272 | -119.6614 |  |  |  | 0.11165 | 0.99066
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
13 |  | EV |  |  | -2.7857 | -18.2712 |  |  | 2023/04/25 23:30:16
13 |  | EV | -7.685 | -62.481 | -10.5001 | -77.0108 | 0.197 | 1292.921 | 2023/04/25 23:39:13
13 |  | EV | -7.651 | -61.489 | -10.4659 | -76.0782 | 0.354 | 1291.144 | 2023/04/25 23:48:10
13 |  | EV | -7.633 | -61.183 | -10.4479 | -75.7905 | 0.465 | 1291.326 | 2023/04/25 23:57:07
Sample 13 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
13 | * | EV | -8.5499 | -61.7877 |  |  |  | 3.8282 | 30.8638
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
14 |  | WOOD4-deep-1 |  |  | -2.7857 | -18.2712 |  |  | 2023/04/26 00:06:57
14 |  | WOOD4-deep-1 | -11.595 | -101.044 | -14.425 | -113.2646 | 0.419 | 1293.668 | 2023/04/26 00:15:54
14 |  | WOOD4-deep-1 | -11.609 | -101.844 | -14.4391 | -114.0167 | 0.232 | 1293.349 | 2023/04/26 00:24:51
14 |  | WOOD4-deep-1 | -11.616 | -102.341 | -14.4461 | -114.484 | 0.201 | 1293.105 | 2023/04/26 00:33:48
Sample 14 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
14 | * | WOOD4-deep-1 | -11.524 | -90.0091 |  |  |  | 5.8033 | 50.8743
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
15 |  | WOOD4-3040-1 | -11.48 | -102.117 | -14.3096 | -114.2734 | 0.552 | 1291.803 | 2023/04/26 00:43:36
15 |  | WOOD4-3040-1 | -11.438 | -102.127 | -14.2674 | -114.2828 | 0.37 | 1292.287 | 2023/04/26 00:52:32
15 |  | WOOD4-3040-1 | -11.447 | -102.088 | -14.2764 | -114.2461 | 0.168 | 1292.406 | 2023/04/26 01:01:29
15 |  | WOOD4-3040-1 | -11.468 | -102.236 | -14.2975 | -114.3852 | 0.105 | 1292.662 | 2023/04/26 01:10:27
Sample 15 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
15 |  | WOOD4-3040-1 | -14.2877 | -114.2969 |  |  |  | 0.01919 | 0.064813
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
16 |  | WOOD4-d-5 | 11.604 | -41.673 | 8.8627 | -57.4488 | -11.292 | 1909.159 | 2023/04/26 01:20:15
16 |  | WOOD4-d-5 | 12.056 | -37.586 | 9.3164 | -53.6065 | -11.297 | 1920.665 | 2023/04/26 01:29:12
16 |  | WOOD4-d-5 | 12.181 | -36.643 | 9.4419 | -52.72 | -11.315 | 1919.074 | 2023/04/26 01:38:09
16 |  | WOOD4-d-5 | 12.263 | -36.481 | 9.5242 | -52.5677 | -11.011 | 1918.162 | 2023/04/26 01:47:05
Sample 16 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
16 |  | WOOD4-d-5 | 9.2863 | -54.0858 | -11.2287 | 614.9281 | 79.1101 | 0.29393 | 2.4341 | -53.088 | 11.9246
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
17 |  | WOOD4-1020-1 | -7.862 | -78.137 | -10.6777 | -91.7293 | 0.376 | 1299.306 | 2023/04/26 01:56:53
17 |  | WOOD4-1020-1 | -8.052 | -79.414 | -10.8685 | -92.9298 | 0.11 | 1297.846 | 2023/04/26 02:05:51
17 |  | WOOD4-1020-1 | -8.166 | -79.736 | -10.9829 | -93.2326 | 0.056 | 1295.553 | 2023/04/26 02:14:48
17 |  | WOOD4-1020-1 | -8.197 | -79.898 | -11.014 | -93.3849 | 0.096 | 1295.086 | 2023/04/26 02:23:45
Sample 17 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
17 |  | WOOD4-1020-1 | -10.8858 | -92.8191 |  |  |  | 0.15158 | 0.79858
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
18 |  | WOOD4-2030-1 | -9.635 | -94.839 | -12.4575 | -107.4312 | 0.381 | 1292.692 | 2023/04/26 02:33:33
18 |  | WOOD4-2030-1 | -9.823 | -96.679 | -12.6462 | -109.161 | 0.384 | 1292.989 | 2023/04/26 02:42:30
18 |  | WOOD4-2030-1 | -9.686 | -96.893 | -12.5087 | -109.3622 | 0.323 | 1293.314 | 2023/04/26 02:51:27
18 |  | WOOD4-2030-1 | -9.746 | -96.871 | -12.5689 | -109.3415 | 0.314 | 1293.979 | 2023/04/26 03:00:24
Sample 18 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
18 |  | WOOD4-2030-1 | -12.5453 | -108.824 |  |  |  | 0.080914 | 0.99233
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
19 |  | WOOD4-010-3 | -9.309 | -88.153 | -12.1303 | -101.1455 | 0.358 | 1287.857 | 2023/04/26 03:10:12
19 |  | WOOD4-010-3 | -9.386 | -87.665 | -12.2076 | -100.6868 | 0.214 | 1284.385 | 2023/04/26 03:19:09
19 |  | WOOD4-010-3 | -9.307 | -87.237 | -12.1283 | -100.2844 | -0.127 | 1284.442 | 2023/04/26 03:28:05
19 |  | WOOD4-010-3 | -9.365 | -87.235 | -12.1865 | -100.2825 | 0.189 | 1286.919 | 2023/04/26 03:37:02
Sample 19 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
19 |  | WOOD4-010-3 | -12.1631 | -100.5998 |  |  |  | 0.039911 | 0.43666
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
20 |  | WOOD4-1020-3 | -8.392 | -78.632 | -11.2098 | -92.1947 | 0.049 | 1291.86 | 2023/04/26 03:46:51
20 |  | WOOD4-1020-3 | -8.377 | -77.981 | -11.1947 | -91.5827 | 0.153 | 1290.954 | 2023/04/26 03:55:47
20 |  | WOOD4-1020-3 | -8.335 | -77.563 | -11.1525 | -91.1897 | 0.158 | 1291.883 | 2023/04/26 04:04:45
20 |  | WOOD4-1020-3 | -8.368 | -77.398 | -11.1857 | -91.0346 | 0.344 | 1292.631 | 2023/04/26 04:13:42
Sample 20 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
20 |  | WOOD4-1020-3 | -11.1857 | -91.5004 |  |  |  | 0.024125 | 0.55009
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
21 |  | WOOD4-3040-3 |  |  | -2.7857 | -18.2712 |  |  | 2023/04/26 04:23:32
21 |  | WOOD4-3040-3 | -11.857 | -103.508 | -14.688 | -115.5811 | 0.298 | 1292.165 | 2023/04/26 04:32:29
21 |  | WOOD4-3040-3 | -11.832 | -104.014 | -14.6629 | -116.0568 | 0.395 | 1292.525 | 2023/04/26 04:41:26
21 |  | WOOD4-3040-3 | -11.895 | -104.621 | -14.7262 | -116.6274 | 0.361 | 1292.56 | 2023/04/26 04:50:23
Sample 21 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
21 | * | WOOD4-3040-3 | -11.7157 | -91.6341 |  |  |  | 5.9307 | 52.0258
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
22 |  | EV |  |  | -2.7857 | -18.2712 |  |  | 2023/04/26 05:00:10
22 |  | EV | -7.699 | -62.364 | -10.5141 | -76.9008 | 0.485 | 1291.487 | 2023/04/26 05:09:06
22 |  | EV | -7.653 | -61.144 | -10.4679 | -75.7539 | 0.076 | 1292.429 | 2023/04/26 05:18:03
22 |  | EV | -7.652 | -61.065 | -10.4669 | -75.6796 | 0.674 | 1291.395 | 2023/04/26 05:27:00
Sample 22 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
22 | * | EV | -8.5587 | -61.6514 |  |  |  | 3.8341 | 30.7679
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
23 |  | WOOD4-B-1 | -1.648 | -82.813 | -4.44 | -96.1253 | -6.003 | 1732.762 | 2023/04/26 05:36:46
23 |  | WOOD4-B-1 | -1.723 | -86.743 | -4.5153 | -99.82 | -6.532 | 1740.72 | 2023/04/26 05:45:44
23 |  | WOOD4-B-1 | -1.713 | -87.423 | -4.5052 | -100.4593 | -6.377 | 1740.737 | 2023/04/26 05:54:41
23 |  | WOOD4-B-1 | -1.73 | -87.735 | -4.5223 | -100.7526 | -6.572 | 1743.902 | 2023/04/26 06:03:37
Sample 23 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
23 |  | WOOD4-B-1 | -4.4957 | -99.2893 | -6.371 | 437.6934 | 22.0519 | 0.037652 | 2.2816 | -32.336 | 0.7564
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
24 |  | WOOD4-2030-3 | -11.005 | -101.333 | -13.8328 | -113.5363 | 0.238 | 1298.73 | 2023/04/26 06:13:25
24 |  | WOOD4-2030-3 | -11.052 | -101.226 | -13.8799 | -113.4357 | 0.445 | 1293.879 | 2023/04/26 06:22:22
24 |  | WOOD4-2030-3 | -11.085 | -101.171 | -13.9131 | -113.384 | 0.464 | 1292.894 | 2023/04/26 06:31:18
24 |  | WOOD4-2030-3 | -11.061 | -101.222 | -13.889 | -113.432 | 0.335 | 1291.752 | 2023/04/26 06:40:16
Sample 24 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
24 |  | WOOD4-2030-3 | -13.8787 | -113.447 |  |  |  | 0.03353 | 0.068103
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
25 |  | WOOD4-1020-2 | -5.162 | -72.855 | -7.9674 | -86.7636 | 0.236 | 1291.449 | 2023/04/26 06:50:03
25 |  | WOOD4-1020-2 | -4.801 | -69.605 | -7.605 | -83.7082 | 0.35 | 1291.315 | 2023/04/26 06:59:00
25 |  | WOOD4-1020-2 | -4.758 | -68.799 | -7.5619 | -82.9505 | 0.391 | 1291.613 | 2023/04/26 07:07:58
25 |  | WOOD4-1020-2 | -4.713 | -68.49 | -7.5167 | -82.66 | 0.357 | 1291.642 | 2023/04/26 07:16:55
Sample 25 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
25 |  | WOOD4-1020-2 | -7.6627 | -84.0206 |  |  |  | 0.2055 | 2.0011
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
26 |  | WOOD4-C-5 | -4.102 | -69.415 | -6.9034 | -83.5296 | 0.475 | 1289.687 | 2023/04/26 07:26:43
26 |  | WOOD4-C-5 | -4.13 | -69.476 | -6.9315 | -83.5869 | 0.118 | 1290.347 | 2023/04/26 07:35:41
26 |  | WOOD4-C-5 | -4.17 | -69.729 | -6.9716 | -83.8248 | 0.122 | 1289.831 | 2023/04/26 07:44:38
26 |  | WOOD4-C-5 | -4.07 | -69.3 | -6.8712 | -83.4215 | 0.094 | 1290.203 | 2023/04/26 07:53:34
Sample 26 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
26 |  | WOOD4-C-5 | -6.9194 | -83.5907 |  |  |  | 0.042458 | 0.18133
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
27 |  | WOOD4-E-1 | -6.706 | -72.508 | -9.5173 | -86.4374 | -0.291 | 1697.542 | 2023/04/26 08:03:23
27 |  | WOOD4-E-1 | -6.923 | -72.783 | -9.7351 | -86.6959 | -0.36 | 1706.821 | 2023/04/26 08:12:20
27 |  | WOOD4-E-1 | -6.923 | -72.778 | -9.7351 | -86.6912 | -0.408 | 1706.552 | 2023/04/26 08:21:17
27 |  | WOOD4-E-1 | -6.823 | -72.146 | -9.6348 | -86.0971 | -0.647 | 1728.405 | 2023/04/26 08:30:13
Sample 27 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
27 |  | WOOD4-E-1 | -9.6556 | -86.4804 |  | 407.9931 |  | 0.10323 | 0.30066 | -14.0718 | -16.5088
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
28 |  | WOOD4-A-1 | -5.339 | -62.766 | -8.1451 | -77.2787 | -0.889 | 2317.416 | 2023/04/26 08:40:02
28 |  | WOOD4-A-1 | -5.363 | -61.666 | -8.1692 | -76.2446 | -1.312 | 2334.196 | 2023/04/26 08:48:58
28 |  | WOOD4-A-1 | -5.345 | -61.563 | -8.1511 | -76.1478 | -1.398 | 2321.801 | 2023/04/26 08:57:54
28 |  | WOOD4-A-1 | -5.405 | -61.548 | -8.2113 | -76.1337 | -1.284 | 2334.964 | 2023/04/26 09:06:52
Sample 28 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
28 |  | WOOD4-A-1 | -8.1692 | -76.4512 | -1.2208 | 1025.2574 |  | 0.029799 | 0.58917 | -38.824 | -20.4453
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
29 |  | WOOD4-D-2 |  |  | -2.7857 | -18.2712 |  |  | 2023/04/26 09:16:37
29 |  | WOOD4-D-2 | -7.661 | -84.119 | -10.476 | -97.3531 | 0.032 | 1296.241 | 2023/04/26 09:25:34
29 |  | WOOD4-D-2 | -7.673 | -84.7 | -10.488 | -97.8993 | 0.082 | 1293.246 | 2023/04/26 09:34:31
29 |  | WOOD4-D-2 | -7.685 | -84.912 | -10.5001 | -98.0986 | 0.21 | 1291.564 | 2023/04/26 09:43:27
Sample 29 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
29 | * | WOOD4-D-2 | -8.5624 | -77.9056 |  |  |  | 3.8365 | 42.2898
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
30 |  | WOOD4-C-2 |  |  | -2.7857 | -18.2712 |  |  | 2023/04/26 09:53:13
30 |  | WOOD4-C-2 | -2.381 | -57.412 | -5.1758 | -72.2453 | 0.232 | 1291.865 | 2023/04/26 10:02:09
30 |  | WOOD4-C-2 | -2.302 | -56.387 | -5.0965 | -71.2817 | 0.144 | 1295.162 | 2023/04/26 10:11:06
30 |  | WOOD4-C-2 | -2.261 | -56.032 | -5.0553 | -70.948 | 0.441 | 1295.981 | 2023/04/26 10:20:02
Sample 30 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
30 | * | WOOD4-C-2 | -4.5283 | -58.1866 |  |  |  | 1.1584 | 28.3112
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
31 |  | EV | -7.224 | -60.002 | -10.0373 | -74.6803 | -0.008 | 1291.901 | 2023/04/26 10:29:48
31 |  | EV | -7.506 | -60.366 | -10.3204 | -75.0225 | 0.094 | 1292.306 | 2023/04/26 10:38:45
31 |  | EV | -7.56 | -60.396 | -10.3746 | -75.0507 | 0.425 | 1291.95 | 2023/04/26 10:47:42
31 |  | EV | -7.634 | -60.721 | -10.4489 | -75.3562 | 0.305 | 1291.564 | 2023/04/26 10:56:39
Sample 31 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
31 |  | EV | -10.2953 | -75.0274 |  |  |  | 0.17919 | 0.294
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
32 |  | WOOD4-C-4 | -3.454 | -62.305 | -6.2529 | -76.8453 | 0.44 | 1289.445 | 2023/04/26 11:06:25
32 |  | WOOD4-C-4 | -3.214 | -62.332 | -6.012 | -76.8707 | -0.255 | 1290.805 | 2023/04/26 11:15:21
32 |  | WOOD4-C-4 | -3.165 | -62.521 | -5.9628 | -77.0484 | -0.179 | 1290.309 | 2023/04/26 11:24:18
32 |  | WOOD4-C-4 | -3.144 | -62.622 | -5.9417 | -77.1434 | 0.086 | 1290.696 | 2023/04/26 11:33:15
Sample 32 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
32 |  | WOOD4-C-4 | -6.0423 | -76.977 |  |  |  | 0.14288 | 0.15218
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
33 |  | WOOD4-A-4 | -5.071 | -70.368 | -7.8761 | -84.4255 | -0.203 | 1289.577 | 2023/04/26 11:43:01
33 |  | WOOD4-A-4 | -5.146 | -71.151 | -7.9513 | -85.1616 | -0.077 | 1291.153 | 2023/04/26 11:51:58
33 |  | WOOD4-A-4 | -5.138 | -71.168 | -7.9433 | -85.1776 | 0.284 | 1292.376 | 2023/04/26 12:00:54
33 |  | WOOD4-A-4 | -5.353 | -72.507 | -8.1591 | -86.4364 | 0.085 | 1275.635 | 2023/04/26 12:09:51
Sample 33 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
33 |  | WOOD4-A-4 | -7.9825 | -85.3003 |  |  |  | 0.12206 | 0.8879
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
34 |  | WOOD4-A-2 | -4.612 | -63.442 | -7.4153 | -77.9143 | 0.468 | 1288.312 | 2023/04/26 12:19:38
34 |  | WOOD4-A-2 | -4.56 | -62.422 | -7.3631 | -76.9553 | -0.016 | 1289.242 | 2023/04/26 12:28:35
34 |  | WOOD4-A-2 | -4.625 | -62.346 | -7.4284 | -76.8839 | 0.188 | 1288.536 | 2023/04/26 12:37:31
34 |  | WOOD4-A-2 | -4.536 | -62.239 | -7.339 | -76.7833 | -0.124 | 1289.06 | 2023/04/26 12:46:28
Sample 34 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
34 |  | WOOD4-A-2 | -7.3864 | -77.1342 |  |  |  | 0.042201 | 0.55824
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
35 |  | WOOD4-E-3 | -6.029 | -71.772 | -8.8377 | -85.7455 | 0.197 | 1290.538 | 2023/04/26 12:56:15
35 |  | WOOD4-E-3 | -6.141 | -73.01 | -8.9502 | -86.9093 | 0.204 | 1290.459 | 2023/04/26 13:05:12
35 |  | WOOD4-E-3 | -6.138 | -73.11 | -8.9471 | -87.0033 | -0.017 | 1291.023 | 2023/04/26 13:14:08
35 |  | WOOD4-E-3 | -6.078 | -73.428 | -8.8869 | -87.3023 | 0.648 | 1290.382 | 2023/04/26 13:23:05
Sample 35 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
35 |  | WOOD4-E-3 | -8.9055 | -86.7401 |  |  |  | 0.053544 | 0.7275
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
36 |  | WOOD4-C-1 | -4.628 | -66.473 | -7.4314 | -80.7638 | 0.239 | 1293.618 | 2023/04/26 13:32:48
36 |  | WOOD4-C-1 | -4.488 | -65.608 | -7.2908 | -79.9506 | -0.059 | 1293.241 | 2023/04/26 13:41:45
36 |  | WOOD4-C-1 | -4.513 | -65.348 | -7.3159 | -79.7061 | 0.116 | 1293.202 | 2023/04/26 13:50:41
36 |  | WOOD4-C-1 | -4.461 | -65.042 | -7.2637 | -79.4185 | 0.243 | 1293.119 | 2023/04/26 13:59:37
Sample 36 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
36 |  | WOOD4-C-1 | -7.3255 | -79.9597 |  |  |  | 0.073469 | 0.61531
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
37 |  | WOOD4-A-3 | 17.411 | -21.328 | 14.6919 | -38.3221 | -13.055 | 1618.085 | 2023/04/26 14:09:21
37 |  | WOOD4-A-3 |  |  | -2.7857 | -18.2712 |  |  | 2023/04/26 14:18:18
37 |  | WOOD4-A-3 | 17.541 | -20.739 | 14.8224 | -37.7684 | -11.911 | 1579.545 | 2023/04/26 14:27:14
37 |  | WOOD4-A-3 | 17.74 | -19.424 | 15.0222 | -36.5321 | -13.279 | 1619.673 | 2023/04/26 14:36:10
Sample 37 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
37 | * | WOOD4-A-3 | 10.4377 | -32.7234 | -9.5613 |  | 49.9532 | 8.783 | 10.2794 | -35.9308 | -24.1666
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
38 |  | WOOD4-C-3 | -6.406 | -68.159 | -9.2162 | -82.3488 | -0.273 | 1413.403 | 2023/04/26 14:45:55
38 |  | WOOD4-C-3 | -6.714 | -69.882 | -9.5253 | -83.9686 | 0.172 | 1416.409 | 2023/04/26 14:54:51
38 |  | WOOD4-C-3 | -6.801 | -70.132 | -9.6127 | -84.2037 | -0.485 | 1416.382 | 2023/04/26 15:03:48
38 |  | WOOD4-C-3 | -6.854 | -70.44 | -9.6659 | -84.4932 | 0.082 | 1417.214 | 2023/04/26 15:12:44
Sample 38 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
38 |  | WOOD4-C-3 | -9.505 | -83.7536 |  | 114.0151 |  | 0.20033 | 1.022 |  | -10.3051
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
39 |  | WOOD4-D-3 | -5.978 | -72.968 | -8.7865 | -86.8698 | 0.22 | 1291.466 | 2023/04/26 15:22:28
39 |  | WOOD4-D-3 | -5.787 | -73.018 | -8.5948 | -86.9168 | 0.138 | 1292.496 | 2023/04/26 15:31:25
39 |  | WOOD4-D-3 | -5.795 | -73.041 | -8.6028 | -86.9385 | 0.189 | 1291.945 | 2023/04/26 15:40:21
39 |  | WOOD4-D-3 | -5.773 | -73.107 | -8.5807 | -87.0005 | 0.079 | 1291.974 | 2023/04/26 15:49:18
Sample 39 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
39 |  | WOOD4-D-3 | -8.6412 | -86.9314 |  |  |  | 0.096927 | 0.057703
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
40 |  | EV |  |  | -2.7857 | -18.2712 |  |  | 2023/04/26 15:59:03
40 |  | EV | -7.469 | -60.451 | -10.2832 | -75.1024 | 0.374 | 1293.237 | 2023/04/26 16:08:00
40 |  | EV | -7.547 | -60.68 | -10.3615 | -75.3177 | 0.788 | 1290.023 | 2023/04/26 16:16:56
40 |  | EV | -7.538 | -60.357 | -10.3525 | -75.014 | 0.246 | 1289.37 | 2023/04/26 16:25:52
Sample 40 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
40 | * | EV | -8.4457 | -60.9263 |  |  |  | 3.7592 | 30.2483
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
41 |  | WOOD4-E-2 | -4.887 | -67.081 | -7.6914 | -81.3354 | 0.06 | 1292.016 | 2023/04/26 16:35:37
41 |  | WOOD4-E-2 | -4.71 | -67.706 | -7.5137 | -81.9229 | 0.165 | 1290.937 | 2023/04/26 16:44:33
41 |  | WOOD4-E-2 | -4.727 | -68.188 | -7.5307 | -82.3761 | 0.51 | 1288.852 | 2023/04/26 16:53:30
41 |  | WOOD4-E-2 | -4.673 | -68.022 | -7.4765 | -82.22 | 0.32 | 1290.654 | 2023/04/26 17:02:27
Sample 41 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
41 |  | WOOD4-E-2 | -7.5531 | -81.9636 |  |  |  | 0.09456 | 0.4883
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
42 |  | WOOD4-B-5 | 5.169 | -54.857 | 2.4031 | -69.8433 | -6.383 | 1497.647 | 2023/04/26 17:12:12
42 |  | WOOD4-B-5 | 5.308 | -54.754 | 2.5426 | -69.7465 | -6.808 | 1518.962 | 2023/04/26 17:21:09
42 |  | WOOD4-B-5 | 5.382 | -54.002 | 2.6169 | -69.0395 | -7.518 | 1553.769 | 2023/04/26 17:30:05
42 |  | WOOD4-B-5 | 5.392 | -54.026 | 2.627 | -69.0621 | -7.749 | 1549.022 | 2023/04/26 17:39:02
Sample 42 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
42 |  | WOOD4-B-5 | 2.5474 | -69.4229 | -7.1145 | 228.0131 | 24.9241 | 0.1029 | 0.45901 | -27.7547 | -6.0013
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
43 |  | WOOD4-deep-2 | -9.22 | -95.403 | -12.0409 | -107.9614 | 0.11 | 1292.784 | 2023/04/26 17:48:47
43 |  | WOOD4-deep-2 | -9.492 | -98.079 | -12.314 | -110.4772 | 0.127 | 1290.059 | 2023/04/26 17:57:43
43 |  | WOOD4-deep-2 | -9.52 | -98.749 | -12.3421 | -111.107 | -0.124 | 1289.33 | 2023/04/26 18:06:40
43 |  | WOOD4-deep-2 | -9.636 | -99.179 | -12.4585 | -111.5113 | 0.466 | 1288.152 | 2023/04/26 18:15:36
Sample 43 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
43 |  | WOOD4-deep-2 | -12.2889 | -110.2642 |  |  |  | 0.17607 | 1.6946
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
44 |  | WOOD4-B-2 | 13.562 | -40.006 | 10.8282 | -55.8816 | -10.827 | 1962.888 | 2023/04/26 18:25:21
44 |  | WOOD4-B-2 | 14.035 | -36.46 | 11.303 | -52.548 | -10.659 | 1924.072 | 2023/04/26 18:34:17
44 |  | WOOD4-B-2 | 13.977 | -35.076 | 11.2448 | -51.2468 | -11.395 | 2009.786 | 2023/04/26 18:43:13
44 |  | WOOD4-B-2 | 14.147 | -34.399 | 11.4154 | -50.6104 | -10.945 | 1993.203 | 2023/04/26 18:52:10
Sample 44 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
44 |  | WOOD4-B-2 | 11.1979 | -52.5717 | -10.9565 | 670.6504 | 66.7104 | 0.25544 | 2.499 | -54.7747 | -12.4551
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
45 |  | WOOD4-3040-2 | -9.084 | -94.109 | -11.9044 | -106.7449 | -0.263 | 1290.305 | 2023/04/26 19:01:55
45 |  | WOOD4-3040-2 | -9.37 | -97.015 | -12.1915 | -109.4769 | 0.103 | 1291.347 | 2023/04/26 19:10:51
45 |  | WOOD4-3040-2 | -9.498 | -98.05 | -12.32 | -110.4499 | 0.024 | 1286.487 | 2023/04/26 19:19:47
45 |  | WOOD4-3040-2 | -9.554 | -98.311 | -12.3762 | -110.6953 | 0.147 | 1287.874 | 2023/04/26 19:28:44
Sample 45 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
45 |  | WOOD4-3040-2 | -12.198 | -109.3417 |  |  |  | 0.20966 | 1.9247
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
46 |  | WOOD4-E-5 |  |  | -2.7857 | -18.2712 |  |  | 2023/04/26 19:38:30
46 |  | WOOD4-E-5 | -4.316 | -62.77 | -7.1182 | -77.2825 | 0.225 | 1295.948 | 2023/04/26 19:47:26
46 |  | WOOD4-E-5 | -4.317 | -62.267 | -7.1192 | -76.8096 | 0.3 | 1294.829 | 2023/04/26 19:56:22
46 |  | WOOD4-E-5 | -4.336 | -61.647 | -7.1382 | -76.2267 | -0.045 | 1294.717 | 2023/04/26 20:05:19
Sample 46 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
46 | * | WOOD4-E-5 | -6.0403 | -62.1475 |  |  |  | 2.1615 | 31.1174
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
47 |  | A00000277285 | -2.223 | -17.985 | -5.0172 | -35.1793 | 0.173 | 1292.128 | 2023/04/26 20:15:05
47 |  | A00000277285 | -2.109 | -13.103 | -4.9027 | -30.5896 | 0.419 | 1292.662 | 2023/04/26 20:24:02
47 |  | A00000277285 | -2.096 | -11.951 | -4.8897 | -29.5066 | -0.015 | 1291.404 | 2023/04/26 20:32:58
47 |  | A00000277285 | -2.093 | -11.377 | -4.8867 | -28.967 | -0.213 | 1293.255 | 2023/04/26 20:41:55
Sample 47 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
47 |  | A00000277285 | -4.9241 | -31.0606 |  |  |  | 0.062222 | 3.0076
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
48 |  | EV |  |  | -2.7857 | -18.2712 |  |  | 2023/04/26 20:51:41
48 |  | EV | -7.329 | -57.801 | -10.1427 | -72.611 | 0.411 | 1291.48 | 2023/04/26 21:00:38
48 |  | EV | -7.405 | -58.97 | -10.219 | -73.71 | 0.087 | 1291.777 | 2023/04/26 21:09:34
48 |  | EV | -7.467 | -59.368 | -10.2812 | -74.0842 | 0.11 | 1290.968 | 2023/04/26 21:18:30
Sample 48 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
48 | * | EV | -8.3571 | -59.6691 |  |  |  | 3.7006 | 29.364

## Source
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Line | Analysis | Time Code | Port | Inj Nr | d(18_16)Mean | d(D_H)Mean | H2O_Mean | Ignore | Good | Identifier 1 | Identifier 2 | Gas Configuration | Timestamp Mean | d(18_16)_SD | d(D_H)_SD | H2O_SD | d(18_16)_Sl | d(D_H)_Sl | H2O_Sl | baseline_shift | slope_shift | residuals | baseline_curvature | interval | ch4_ppm | h16od_adjust | h16od_shift | n2_flag | DAS Temp | Tray | Sample | Job | Method | Error Code
1 | P-3859 | 2023/04/25 12:34:11 | 2-01 | 1 | [5 empty] | EV | standard | H2O | [17 empty] | 1 | 1 | 1 | GJB_Rinse | 0
2 | P-3859 | 2023/04/25 12:44:13 | 2-01 | 2 | -7.973 | -66.916 | 19547 | -1 | 1 | EV | standard | H2O | 1682426955 | 0.331 | 0.71 | 102.504 | 0.002 | 0.002 | 2.412 | 1328.557 | 0.608 | 3.159 | 121.892 | 0.864 | 0.644 | 0 | 0 | 1 | 41.562 | 1 | 1 | 1 | GJB_Rinse | 0
3 | P-3859 | 2023/04/25 12:53:09 | 2-01 | 3 | -7.588 | -62.272 | 20150 | -1 | 1 | EV | standard | H2O | 1682427491 | 0.327 | 0.713 | 104.645 | 0.002 | -0.001 | 2.456 | 1346.381 | -0.261 | 3.232 | 124.447 | 0.866 | 0.348 | 0 | 0 | 1 | 41.188 | 1 | 1 | 1 | GJB_Rinse | 0
4 | P-3859 | 2023/04/25 13:02:06 | 2-01 | 4 | -7.563 | -62.339 | 21376 | 0 | 1 | EV | standard | H2O | 1682428028 | 0.346 | 0.682 | 109.208 | 0.001 | 0.002 | 2.571 | 1379.411 | -0.716 | 3.495 | 131.309 | 0.869 | 0.413 | 0 | 0 | 1 | 41 | 1 | 1 | 1 | GJB_Rinse | 0
5 | P-3859 | 2023/04/25 13:11:03 | 2-01 | 5 | -7.361 | -60.824 | 19085 | 0 | 1 | EV | standard | H2O | 1682428564 | 0.322 | 0.701 | 106.801 | 0.001 | -0.001 | 2.499 | 1324.937 | 1.217 | 3.048 | 117.086 | 0.864 | 0.302 | 0 | 0 | 1 | 40.938 | 1 | 1 | 1 | GJB_Rinse | 0
6 | P-3859 | 2023/04/25 13:20:00 | 2-01 | 6 | -7.255 | -59.622 | 19655 | 0 | 1 | EV | standard | H2O | 1682429100 | 0.308 | 0.689 | 103.305 | 0.001 | 0 | 2.447 | 1325.198 | 0.995 | 3.18 | 121.274 | 0.864 | 0.446 | 0 | 0 | 1 | 41 | 1 | 1 | 1 | GJB_Rinse | 0
7 | P-3859 | 2023/04/25 13:28:56 | 2-01 | 7 | -7.238 | -59.615 | 18732 | 0 | 1 | EV | standard | H2O | 1682429638 | 0.293 | 0.694 | 104.628 | 0.001 | 0.003 | 2.454 | 1310.335 | 1.986 | 2.991 | 115.269 | 0.863 | 0.512 | 0 | 0 | 1 | 41.062 | 1 | 1 | 1 | GJB_Rinse | 0
8 | P-3859 | 2023/04/25 13:37:54 | 2-01 | 8 | -7.171 | -58.914 | 18928 | 0 | 1 | EV | standard | H2O | 1682430176 | 0.302 | 0.672 | 103.791 | 0 | 0.002 | 2.433 | 1307.313 | 1.63 | 3.051 | 116.445 | 0.867 | 0.324 | 0 | 0 | 1 | 41.125 | 1 | 1 | 1 | GJB_Rinse | 0
9 | P-3859 | 2023/04/25 13:46:51 | 2-01 | 9 | -7.121 | -58.741 | 18878 | 0 | 1 | EV | standard | H2O | 1682430712 | 0.31 | 0.732 | 103.453 | 0.001 | 0.001 | 2.446 | 1305.189 | 1.818 | 3.029 | 116.74 | 0.866 | 0.487 | 0 | 0 | 1 | 41.062 | 1 | 1 | 1 | GJB_Rinse | 0
10 | P-3859 | 2023/04/25 13:55:48 | 2-01 | 10 | -7.183 | -58.693 | 18811 | 0 | 1 | EV | standard | H2O | 1682431250 | 0.278 | 0.758 | 102.172 | 0.001 | 0.004 | 2.42 | 1303.666 | 1.757 | 3.047 | 117.038 | 0.864 | 0.543 | 0 | 0 | 1 | 41.125 | 1 | 1 | 1 | GJB_Rinse | 0
11 | P-3860 | 2023/04/25 14:05:36 | 2-02 | 1 | 3.937 | 27.51 | 18932 | -1 | 1 | PZ | standard | H2O | 1682431837 | 0.323 | 0.823 | 101.193 | 0 | -0.009 | 2.408 | 1301.698 | 1.05 | 3.072 | 117.406 | 0.865 | 0.36 | 0 | 0 | 1 | 41.188 | 1 | 2 | 1 | GJB_Rinse | 0
12 | P-3860 | 2023/04/25 14:14:33 | 2-02 | 2 | 4.541 | 36.325 | 18811 | -1 | 1 | PZ | standard | H2O | 1682432374 | 0.325 | 0.766 | 104.101 | 0 | -0.006 | 2.457 | 1302.669 | 0.88 | 3.036 | 116.372 | 0.866 | 0.182 | 0 | 0 | 1 | 41.188 | 1 | 2 | 1 | GJB_Rinse | 0
13 | P-3860 | 2023/04/25 14:23:30 | 2-02 | 3 | 4.7 | 38.866 | 19018 | -1 | 1 | PZ | standard | H2O | 1682432912 | 0.311 | 0.749 | 104.457 | 0 | -0.005 | 2.454 | 1304.042 | 1.132 | 3.097 | 117.625 | 0.866 | 0.258 | 0 | 0 | 1 | 41.062 | 1 | 2 | 1 | GJB_Rinse | 0
14 | P-3860 | 2023/04/25 14:32:27 | 2-02 | 4 | 4.777 | 39.653 | 18896 | 0 | 1 | PZ | standard | H2O | 1682433449 | 0.308 | 0.79 | 103.795 | 0.002 | 0 | 2.432 | 1303.083 | 1.026 | 3.079 | 117.268 | 0.864 | 0.352 | 0 | 0 | 1 | 41.062 | 1 | 2 | 1 | GJB_Rinse | 0
15 | P-3860 | 2023/04/25 14:41:24 | 2-02 | 5 | 4.769 | 40.114 | 18905 | 0 | 1 | PZ | standard | H2O | 1682433986 | 0.364 | 0.785 | 102.297 | 0.002 | -0.002 | 2.403 | 1302.412 | 1.362 | 3.15 | 118.966 | 0.865 | 0.808 | 0 | 0 | 1 | 40.812 | 1 | 2 | 1 | GJB_Rinse | 0
16 | P-3860 | 2023/04/25 14:50:22 | 2-02 | 6 | 4.784 | 40.396 | 18992 | 0 | 1 | PZ | standard | H2O | 1682434523 | 0.355 | 0.711 | 103.459 | 0 | -0.003 | 2.433 | 1303.232 | 1.102 | 3.067 | 117.206 | 0.865 | 0.314 | 0 | 0 | 1 | 40.875 | 1 | 2 | 1 | GJB_Rinse | 0
17 | P-3860 | 2023/04/25 14:59:19 | 2-02 | 7 | 4.832 | 40.695 | 18882 | 0 | 1 | PZ | standard | H2O | 1682435060 | 0.348 | 0.722 | 102.836 | 0.001 | -0.002 | 2.43 | 1302.083 | 1.135 | 3.079 | 117.939 | 0.866 | 0.628 | 0 | 0 | 1 | 40.875 | 1 | 2 | 1 | GJB_Rinse | 0
18 | P-3860 | 2023/04/25 15:08:16 | 2-02 | 8 | 4.852 | 41.027 | 18925 | 0 | 1 | PZ | standard | H2O | 1682435597 | 0.305 | 0.701 | 98.58 | 0.001 | -0.002 | 2.347 | 1302.591 | 0.739 | 3.064 | 117.682 | 0.864 | 0.379 | 0 | 0 | 1 | 40.812 | 1 | 2 | 1 | GJB_Rinse | 0
19 | P-3860 | 2023/04/25 15:17:13 | 2-02 | 9 | 4.918 | 41.052 | 18958 | 0 | 1 | PZ | standard | H2O | 1682436135 | 0.331 | 0.734 | 101.959 | 0 | -0.003 | 2.396 | 1302.382 | 1.003 | 3.116 | 118.904 | 0.866 | 0.495 | 0 | 0 | 1 | 40.812 | 1 | 2 | 1 | GJB_Rinse | 0
20 | P-3860 | 2023/04/25 15:26:11 | 2-02 | 10 | 4.867 | 41.241 | 18988 | 0 | 1 | PZ | standard | H2O | 1682436672 | 0.325 | 0.744 | 101.446 | 0.001 | -0.002 | 2.394 | 1302.557 | 0.918 | 3.131 | 118.658 | 0.867 | 0.516 | 0 | 0 | 1 | 40.75 | 1 | 2 | 1 | GJB_Rinse | 0
21 | P-3861 | 2023/04/25 15:36:00 | 2-03 | 1 | -11.957 | -90.283 | 18732 | -1 | 1 | UT2 | standard | H2O | 1682437262 | 0.306 | 1.015 | 97.213 | 0.002 | 0.017 | 2.297 | 1299.562 | 2.287 | 3.042 | 115.334 | 0.866 | 0.675 | 0 | 0 | 1 | 40.875 | 1 | 3 | 1 | GJB_Rinse | 0
22 | P-3861 | 2023/04/25 15:44:57 | 2-03 | 2 | -12.806 | -104.086 | 18879 | -1 | 1 | UT2 | standard | H2O | 1682437799 | 0.302 | 1.008 | 102.367 | 0.002 | 0.017 | 2.409 | 1301.321 | 2.107 | 2.983 | 116.053 | 0.864 | 0.414 | 0 | 0 | 1 | 40.812 | 1 | 3 | 1 | GJB_Rinse | 0
23 | P-3861 | 2023/04/25 15:53:54 | 2-03 | 3 | -13.051 | -107.145 | 18925 | -1 | 1 | UT2 | standard | H2O | 1682438336 | 0.33 | 0.937 | 101.327 | 0.001 | 0.015 | 2.384 | 1301.135 | 2.897 | 3.005 | 116.106 | 0.864 | 0.522 | 0 | 0 | 1 | 40.625 | 1 | 3 | 1 | GJB_Rinse | 0
24 | P-3861 | 2023/04/25 16:02:52 | 2-03 | 4 | -13.053 | -108.517 | 18870 | 0 | 1 | UT2 | standard | H2O | 1682438873 | 0.274 | 0.741 | 100.877 | 0.001 | 0.01 | 2.361 | 1301.284 | 1.963 | 2.982 | 115.447 | 0.864 | 0.345 | 0 | 0 | 1 | 40.812 | 1 | 3 | 1 | GJB_Rinse | 0
25 | P-3861 | 2023/04/25 16:11:48 | 2-03 | 5 | -13.144 | -109.271 | 18889 | 0 | 1 | UT2 | standard | H2O | 1682439410 | 0.3 | 0.835 | 101.321 | 0.003 | 0.011 | 2.384 | 1301.216 | 1.934 | 3.045 | 116.843 | 0.865 | 0.508 | 0 | 0 | 1 | 40.812 | 1 | 3 | 1 | GJB_Rinse | 0
26 | P-3861 | 2023/04/25 16:20:46 | 2-03 | 6 | -13.151 | -109.877 | 18917 | 0 | 1 | UT2 | standard | H2O | 1682439948 | 0.337 | 0.843 | 98.472 | 0.002 | 0.008 | 2.342 | 1301.536 | 1.556 | 3.003 | 117.281 | 0.865 | 0.474 | 0 | 0 | 1 | 40.875 | 1 | 3 | 1 | GJB_Rinse | 0
27 | P-3861 | 2023/04/25 16:29:44 | 2-03 | 7 | -13.224 | -110.253 | 18854 | 0 | 1 | UT2 | standard | H2O | 1682440485 | 0.335 | 0.709 | 100.195 | 0.002 | 0.006 | 2.359 | 1301.078 | 1.569 | 2.993 | 116.468 | 0.865 | 0.407 | 0 | 0 | 1 | 40.812 | 1 | 3 | 1 | GJB_Rinse | 0
28 | P-3861 | 2023/04/25 16:38:41 | 2-03 | 8 | -13.199 | -110.5 | 18920 | 0 | 1 | UT2 | standard | H2O | 1682441022 | 0.329 | 0.777 | 100.383 | 0.001 | 0.007 | 2.377 | 1301.359 | 2.156 | 3.001 | 116.446 | 0.869 | 0.425 | 0 | 0 | 1 | 40.75 | 1 | 3 | 1 | GJB_Rinse | 0
29 | P-3861 | 2023/04/25 16:47:38 | 2-03 | 9 | -13.175 | -110.672 | 18891 | 0 | 1 | UT2 | standard | H2O | 1682441559 | 0.329 | 0.758 | 97.681 | 0.002 | 0.008 | 2.319 | 1301.224 | 1.803 | 2.976 | 116.35 | 0.867 | 0.333 | 0 | 0 | 1 | 40.688 | 1 | 3 | 1 | GJB_Rinse | 0
30 | P-3861 | 2023/04/25 16:56:35 | 2-03 | 10 | -13.286 | -110.843 | 18794 | 0 | 1 | UT2 | standard | H2O | 1682442097 | 0.316 | 0.75 | 104.083 | 0.002 | 0.006 | 2.442 | 1300.274 | 1.966 | 3.008 | 115.768 | 0.867 | 0.404 | 0 | 0 | 1 | 40.5 | 1 | 3 | 1 | GJB_Rinse | 0
31 | P-3862 | 2023/04/25 17:06:24 | 2-04 | 1 | -7.945 | -66.123 | 18846 | -1 | 1 | EV | standard | H2O | 1682442685 | 0.311 | 0.717 | 97.545 | 0.002 | 0.001 | 2.317 | 1300.303 | 2.123 | 2.995 | 115.997 | 0.866 | 0.23 | 0 | 0 | 1 | 40.625 | 1 | 4 | 1 | GJB_Rinse | 0
32 | P-3862 | 2023/04/25 17:15:21 | 2-04 | 2 | -7.578 | -61.595 | 18816 | -1 | 1 | EV | standard | H2O | 1682443222 | 0.327 | 0.685 | 100.591 | 0.001 | 0.001 | 2.367 | 1300.315 | 1.512 | 3.046 | 117.612 | 0.865 | 0.603 | 0 | 0 | 1 | 40.625 | 1 | 4 | 1 | GJB_Rinse | 0
33 | P-3862 | 2023/04/25 17:24:19 | 2-04 | 3 | -7.585 | -60.611 | 18546 | -1 | 1 | EV | standard | H2O | 1682443760 | 0.282 | 0.731 | 95.983 | 0 | 0.001 | 2.27 | 1298.504 | 1.557 | 2.932 | 114.346 | 0.87 | 0.274 | 0 | 0 | 1 | 40.688 | 1 | 4 | 1 | GJB_Rinse | 0
34 | P-3862 | 2023/04/25 17:33:16 | 2-04 | 4 | -7.595 | -60.407 | 18596 | 0 | 1 | EV | standard | H2O | 1682444297 | 0.301 | 0.652 | 99.582 | 0.001 | 0.002 | 2.352 | 1298.801 | 2.132 | 2.922 | 113.473 | 0.867 | 0.198 | 0 | 0 | 1 | 40.75 | 1 | 4 | 1 | GJB_Rinse | 0
35 | P-3862 | 2023/04/25 17:42:12 | 2-04 | 5 | -7.565 | -60.11 | 18425 | 0 | 1 | EV | standard | H2O | 1682444834 | 0.296 | 0.737 | 99.909 | 0.002 | 0.004 | 2.371 | 1297.719 | 1.82 | 2.903 | 112.269 | 0.867 | 0.038 | 0 | 0 | 1 | 40.688 | 1 | 4 | 1 | GJB_Rinse | 0
36 | P-3862 | 2023/04/25 17:51:09 | 2-04 | 6 | -7.536 | -59.737 | 18492 | 0 | 1 | EV | standard | H2O | 1682445371 | 0.298 | 0.735 | 100.585 | 0.002 | 0.004 | 2.366 | 1297.981 | 1.864 | 2.979 | 114.202 | 0.871 | 0.362 | 0 | 0 | 1 | 41 | 1 | 4 | 1 | GJB_Rinse | 0
37 | P-3862 | 2023/04/25 18:00:06 | 2-04 | 7 | -7.573 | -59.848 | 18343 | 0 | 1 | EV | standard | H2O | 1682445907 | 0.287 | 0.742 | 99.405 | 0.001 | 0.002 | 2.359 | 1296.921 | 1.634 | 2.928 | 113.131 | 0.863 | 0.298 | 0 | 0 | 1 | 40.812 | 1 | 4 | 1 | GJB_Rinse | 0
38 | P-3862 | 2023/04/25 18:09:03 | 2-04 | 8 | -7.532 | -59.885 | 18269 | 0 | 1 | EV | standard | H2O | 1682446444 | 0.324 | 0.797 | 97.49 | 0.002 | 0.003 | 2.309 | 1296.548 | 1.401 | 2.867 | 112.919 | 0.864 | 0.319 | 0 | 0 | 1 | 40.938 | 1 | 4 | 1 | GJB_Rinse | 0
39 | P-3862 | 2023/04/25 18:18:00 | 2-04 | 9 | -7.577 | -60.05 | 18379 | 0 | 1 | EV | standard | H2O | 1682446982 | 0.311 | 0.726 | 97.72 | 0 | 0.001 | 2.319 | 1297.248 | 1.927 | 2.876 | 112.67 | 0.868 | 0.258 | 0 | 0 | 1 | 40.562 | 1 | 4 | 1 | GJB_Rinse | 0
40 | P-3862 | 2023/04/25 18:26:58 | 2-04 | 10 | -7.591 | -59.857 | 18386 | 0 | 1 | EV | standard | H2O | 1682447519 | 0.315 | 0.784 | 96.992 | 0.001 | 0.006 | 2.292 | 1296.973 | 1.617 | 2.973 | 114.746 | 0.871 | 0.484 | 0 | 0 | 1 | 40.625 | 1 | 4 | 1 | GJB_Rinse | 0
41 | P-3863 | 2023/04/25 18:36:48 | 2-05 | 1 | -8.926 | -81.38 | 18417 | -1 | 1 | WOOD4-010-1 | 23-074 | H2O | 1682448110 | 0.32 | 0.734 | 99.601 | 0.002 | 0.007 | 2.369 | 1297.568 | 1.864 | 2.957 | 113.931 | 0.87 | 0.382 | 0 | 0 | 1 | 40.688 | 1 | 5 | 1 | GJB_Rinse | 0
42 | P-3863 | 2023/04/25 18:45:45 | 2-05 | 2 | -8.969 | -83.399 | 18649 | -1 | 1 | WOOD4-010-1 | 23-074 | H2O | 1682448648 | 0.332 | 0.821 | 100.126 | 0.001 | 0.009 | 2.361 | 1299.544 | 1.928 | 2.951 | 114.198 | 0.868 | 0.277 | 0 | 0 | 1 | 40.625 | 1 | 5 | 1 | GJB_Rinse | 0
43 | P-3863 | 2023/04/25 18:54:42 | 2-05 | 3 | -9.025 | -84.178 | 18475 | -1 | 1 | WOOD4-010-1 | 23-074 | H2O | 1682449184 | 0.315 | 0.779 | 97.869 | 0.002 | 0.006 | 2.32 | 1297.768 | 2.43 | 2.937 | 114.056 | 0.868 | 0.589 | 0 | 0 | 1 | 40.562 | 1 | 5 | 1 | GJB_Rinse | 0
44 | P-3863 | 2023/04/25 19:03:40 | 2-05 | 4 | -9.061 | -84.545 | 18348 | 0 | 1 | WOOD4-010-1 | 23-074 | H2O | 1682449721 | 0.311 | 0.719 | 97.576 | 0.001 | 0.003 | 2.293 | 1297.128 | 1.65 | 2.93 | 113.784 | 0.866 | 0.401 | 0 | 0 | 1 | 40.625 | 1 | 5 | 1 | GJB_Rinse | 0
45 | P-3864 | 2023/04/25 19:13:30 | 2-06 | 1 | -7.238 | -78.453 | 18754 | -1 | 1 | WOOD4-B-4 | 23-074 | H2O | 1682450312 | 0.322 | 0.828 | 98.911 | 0.001 | 0.005 | 2.325 | 1300.071 | 1.302 | 3.021 | 116.261 | 0.866 | 0.392 | 0 | 0 | 1 | 40.375 | 1 | 6 | 1 | GJB_Rinse | 0
46 | P-3864 | 2023/04/25 19:22:26 | 2-06 | 2 | -7.184 | -78.214 | 18552 | -1 | 1 | WOOD4-B-4 | 23-074 | H2O | 1682450849 | 0.32 | 0.765 | 99.205 | 0.001 | 0.004 | 2.344 | 1298.901 | 1.538 | 2.883 | 113.007 | 0.871 | -0.092 | 0 | 0 | 1 | 40.562 | 1 | 6 | 1 | GJB_Rinse | 0
47 | P-3864 | 2023/04/25 19:31:24 | 2-06 | 3 | -7.119 | -77.927 | 18626 | -1 | 1 | WOOD4-B-4 | 23-074 | H2O | 1682451385 | 0.321 | 0.769 | 102.478 | 0.002 | 0.005 | 2.409 | 1299.358 | 1.501 | 2.954 | 114.118 | 0.865 | 0.143 | 0 | 0 | 1 | 40.625 | 1 | 6 | 1 | GJB_Rinse | 0
48 | P-3864 | 2023/04/25 19:40:21 | 2-06 | 4 | -7.112 | -77.96 | 18538 | 0 | 1 | WOOD4-B-4 | 23-074 | H2O | 1682451923 | 0.333 | 0.747 | 98.948 | 0 | 0.005 | 2.324 | 1298.44 | 1.665 | 2.925 | 114.444 | 0.866 | 0.337 | 0 | 0 | 1 | 40.438 | 1 | 6 | 1 | GJB_Rinse | 0
49 | P-3865 | 2023/04/25 19:50:12 | 2-07 | 1 | 15.521 | -28.048 | 18381 | -1 | 1 | WOOD4-D-4 | 23-074 | H2O | 1682452514 | 0.344 | 1.224 | 99.601 | -0.002 | -0.008 | 2.341 | 1680.041 | -52.729 | 6.673 | 143.299 | 0.88 | -13.878 | 0 | 0 | 1 | 40.312 | 1 | 7 | 1 | GJB_Rinse | 0
50 | P-3865 | 2023/04/25 19:59:09 | 2-07 | 2 | 15.746 | -26.456 | 17795 | -1 | 1 | WOOD4-D-4 | 23-074 | H2O | 1682453050 | 0.395 | 1.221 | 100.322 | -0.002 | 0.001 | 2.366 | 1667.745 | -51.8 | 6.435 | 139.445 | 0.867 | -13.437 | 0 | 0 | 1 | 40.5 | 1 | 7 | 1 | GJB_Rinse | 0
51 | P-3865 | 2023/04/25 20:08:06 | 2-07 | 3 | 15.854 | -25.865 | 17785 | -1 | 1 | WOOD4-D-4 | 23-074 | H2O | 1682453588 | 0.393 | 1.292 | 99.278 | -0.002 | -0.003 | 2.361 | 1668.101 | -51.45 | 6.421 | 138.14 | 0.87 | -13.825 | 0 | 0 | 1 | 40.5 | 1 | 7 | 1 | GJB_Rinse | 0
52 | P-3865 | 2023/04/25 20:17:04 | 2-07 | 4 | 15.844 | -25.801 | 17673 | 0 | 1 | WOOD4-D-4 | 23-074 | H2O | 1682454125 | 0.373 | 1.171 | 98.179 | -0.002 | -0.005 | 2.305 | 1666.306 | -51.31 | 6.48 | 138.333 | 0.871 | -13.514 | 0 | 0 | 1 | 40.562 | 1 | 7 | 1 | GJB_Rinse | 0
53 | P-3866 | 2023/04/25 20:26:50 | 2-08 | 1 | 8.079 | -41.103 | 17981 | -1 | 1 | WOOD4-B-3 | 23-074 | H2O | 1682454711 | 0.466 | 1.365 | 97.625 | 0 | -0.005 | 2.305 | 2255.738 | -58.201 | 4.752 | 123.722 | 0.875 | -9.664 | 0 | 0 | 1 | 40.562 | 1 | 8 | 1 | GJB_Rinse | 0
54 | P-3866 | 2023/04/25 20:35:47 | 2-08 | 2 | 7.93 | -41.734 | 17323 | -1 | 1 | WOOD4-B-3 | 23-074 | H2O | 1682455249 | 0.453 | 1.481 | 96.293 | 0.001 | -0.003 | 2.273 | 2218.743 | -56.811 | 4.513 | 119.968 | 0.875 | -9.259 | 0 | 0 | 1 | 40.625 | 1 | 8 | 1 | GJB_Rinse | 0
55 | P-3866 | 2023/04/25 20:44:44 | 2-08 | 3 | 7.883 | -42.103 | 17462 | -1 | 1 | WOOD4-B-3 | 23-074 | H2O | 1682455785 | 0.445 | 1.384 | 93.525 | -0.002 | 0.002 | 2.241 | 2225.417 | -56.968 | 4.57 | 121.432 | 0.874 | -9.279 | 0 | 0 | 1 | 40.562 | 1 | 8 | 1 | GJB_Rinse | 0
56 | P-3866 | 2023/04/25 20:53:41 | 2-08 | 4 | 7.887 | -41.749 | 17423 | 0 | 1 | WOOD4-B-3 | 23-074 | H2O | 1682456322 | 0.444 | 1.357 | 98.898 | 0 | 0.006 | 2.349 | 2226.508 | -57.262 | 4.586 | 120.553 | 0.876 | -9.647 | 0 | 0 | 1 | 40.625 | 1 | 8 | 1 | GJB_Rinse | 0
57 | P-3867 | 2023/04/25 21:03:30 | 2-09 | 1 | -5.765 | -71.707 | 18504 | -1 | 1 | WOOD4-010-2 | 23-074 | H2O | 1682456911 | 0.333 | 0.727 | 98.528 | 0.001 | 0.004 | 2.313 | 1311.727 | 0.489 | 2.928 | 115.519 | 0.866 | 0.148 | 0 | 0 | 1 | 40.688 | 1 | 9 | 1 | GJB_Rinse | 0
58 | P-3867 | 2023/04/25 21:12:27 | 2-09 | 2 | -5.945 | -72.919 | 18171 | -1 | 1 | WOOD4-010-2 | 23-074 | H2O | 1682457448 | 0.316 | 0.868 | 97.972 | 0.001 | 0.007 | 2.31 | 1301.922 | 1.2 | 2.893 | 113.444 | 0.863 | 0.313 | 0 | 0 | 1 | 40.75 | 1 | 9 | 1 | GJB_Rinse | 0
59 | P-3867 | 2023/04/25 21:21:24 | 2-09 | 3 | -5.966 | -73.162 | 18220 | -1 | 1 | WOOD4-010-2 | 23-074 | H2O | 1682457986 | 0.286 | 0.86 | 100.263 | 0.001 | 0.006 | 2.351 | 1300.054 | 1.782 | 2.855 | 111.557 | 0.868 | 0.218 | 0 | 0 | 1 | 40.75 | 1 | 9 | 1 | GJB_Rinse | 0
60 | P-3867 | 2023/04/25 21:30:22 | 2-09 | 4 | -5.946 | -73.177 | 18219 | 0 | 1 | WOOD4-010-2 | 23-074 | H2O | 1682458523 | 0.305 | 0.753 | 98.413 | 0 | 0.004 | 2.328 | 1298.796 | 1.661 | 2.923 | 112.744 | 0.865 | 0.484 | 0 | 0 | 1 | 40.812 | 1 | 9 | 1 | GJB_Rinse | 0
61 | P-3868 | 2023/04/25 21:40:10 | 2-10 | 1 | 12.915 | -36.736 | 17596 | -1 | 1 | WOOD4-D-1 | 23-074 | H2O | 1682459112 | 0.434 | 1.294 | 98.339 | -0.002 | 0.002 | 2.323 | 2028.928 | -60.902 | 5.996 | 132.745 | 0.876 | -12.55 | 0 | 0 | 1 | 40.688 | 1 | 10 | 1 | GJB_Rinse | 0
62 | P-3868 | 2023/04/25 21:49:07 | 2-10 | 2 | 13.11 | -35.73 | 17521 | -1 | 1 | WOOD4-D-1 | 23-074 | H2O | 1682459649 | 0.396 | 1.213 | 96.955 | -0.001 | 0.001 | 2.302 | 2032.711 | -61.693 | 6.083 | 133.165 | 0.871 | -12.938 | 0 | 0 | 1 | 40.625 | 1 | 10 | 1 | GJB_Rinse | 0
63 | P-3868 | 2023/04/25 21:58:05 | 2-10 | 3 | 13.157 | -35.714 | 17338 | -1 | 1 | WOOD4-D-1 | 23-074 | H2O | 1682460186 | 0.43 | 1.39 | 95.73 | -0.001 | 0.002 | 2.268 | 2025.968 | -62.185 | 6.049 | 134.06 | 0.873 | -12.942 | 0 | 0 | 1 | 40.625 | 1 | 10 | 1 | GJB_Rinse | 0
64 | P-3868 | 2023/04/25 22:07:02 | 2-10 | 4 | 13.184 | -35.69 | 17313 | 0 | 1 | WOOD4-D-1 | 23-074 | H2O | 1682460724 | 0.443 | 1.285 | 97.236 | -0.002 | -0.002 | 2.287 | 2025.554 | -61.435 | 6.054 | 132.079 | 0.874 | -12.486 | 0 | 0 | 1 | 40.5 | 1 | 10 | 1 | GJB_Rinse | 0
65 | P-3869 | 2023/04/25 22:16:51 | 2-11 | 1 | -8.186 | -88.77 | 16699 | -1 | 1 | WOOD4-2030-2 | 23-074 | H2O | 1682461312 | 0.307 | 0.771 | 94.393 | 0.002 | 0.009 | 2.28 | 1296.462 | 0.785 | 2.628 | 103.475 | 0.863 | 0.114 | 0 | 0 | 1 | 40.562 | 1 | 11 | 1 | GJB_Rinse | 0
66 | P-3869 | 2023/04/25 22:25:48 | 2-11 | 2 | -8.545 | -91.492 | 16651 | -1 | 1 | WOOD4-2030-2 | 23-074 | H2O | 1682461850 | 0.313 | 0.806 | 94.516 | 0.002 | 0.008 | 2.238 | 1290.26 | 1.305 | 2.64 | 102.943 | 0.864 | 0.228 | 0 | 0 | 1 | 40.625 | 1 | 11 | 1 | GJB_Rinse | 0
67 | P-3869 | 2023/04/25 22:34:44 | 2-11 | 3 | -8.548 | -92.185 | 16805 | -1 | 1 | WOOD4-2030-2 | 23-074 | H2O | 1682462386 | 0.299 | 0.786 | 96.28 | 0.001 | 0.007 | 2.291 | 1289.341 | 1.354 | 2.64 | 105.32 | 0.865 | 0.404 | 0 | 0 | 1 | 40.562 | 1 | 11 | 1 | GJB_Rinse | 0
68 | P-3869 | 2023/04/25 22:43:42 | 2-11 | 4 | -8.64 | -92.49 | 16784 | 0 | 1 | WOOD4-2030-2 | 23-074 | H2O | 1682462924 | 0.294 | 0.783 | 94.577 | 0.001 | 0.003 | 2.257 | 1288.27 | 1.381 | 2.647 | 104.604 | 0.863 | 0.395 | 0 | 0 | 1 | 40.625 | 1 | 11 | 1 | GJB_Rinse | 0
69 | P-3870 | 2023/04/25 22:53:32 | 2-12 | 1 | -12.128 | -106.41 | 17485 | -1 | 1 | WOOD4-deep-3 | 23-074 | H2O | 1682463513 | 0.324 | 0.787 | 96.661 | 0.001 | 0.008 | 2.295 | 1291.536 | 1.516 | 2.851 | 110.109 | 0.862 | 0.585 | 0 | 0 | 1 | 40.688 | 1 | 12 | 1 | GJB_Rinse | 0
70 | P-3870 | 2023/04/25 23:02:29 | 2-12 | 2 | -12.334 | -108.026 | 17553 | -1 | 1 | WOOD4-deep-3 | 23-074 | H2O | 1682464050 | 0.308 | 0.842 | 96.293 | 0.001 | 0.004 | 2.285 | 1292.706 | 2.06 | 2.741 | 107.535 | 0.863 | 0.266 | 0 | 0 | 1 | 40.562 | 1 | 12 | 1 | GJB_Rinse | 0
71 | P-3870 | 2023/04/25 23:11:29 | 2-12 | 3 | -12.363 | -108.321 | 17850 | -1 | 1 | WOOD4-deep-3 | 23-074 | H2O | 1682464591 | 0.291 | 0.79 | 97.057 | 0.001 | 0.005 | 2.296 | 1295.176 | 1.905 | 2.738 | 107.825 | 0.864 | -0.118 | 0 | 0 | 1 | 40.688 | 1 | 12 | 1 | GJB_Rinse | 0
72 | P-3870 | 2023/04/25 23:20:26 | 2-12 | 4 | -12.353 | -108.636 | 17416 | 0 | 1 | WOOD4-deep-3 | 23-074 | H2O | 1682465127 | 0.351 | 0.762 | 97.405 | 0.003 | 0.005 | 2.308 | 1291.82 | 1.878 | 2.691 | 105.599 | 0.864 | 0.036 | 0 | 0 | 1 | 40.688 | 1 | 12 | 1 | GJB_Rinse | 0
73 | P-3871 | 2023/04/25 23:30:16 | 2-13 | 1 | [5 empty] | EV | standard | [18 empty] | 1 | 13 | 1 | GJB_Rinse | 0
74 | P-3871 | 2023/04/25 23:39:13 | 2-13 | 2 | -7.685 | -62.481 | 17679 | -1 | 1 | EV | standard | H2O | 1682466255 | 0.293 | 0.727 | 96.786 | 0.001 | 0.001 | 2.278 | 1292.921 | 1.765 | 2.79 | 108.188 | 0.869 | 0.197 | 0 | 0 | 1 | 40.75 | 1 | 13 | 1 | GJB_Rinse | 0
75 | P-3871 | 2023/04/25 23:48:10 | 2-13 | 3 | -7.651 | -61.489 | 17449 | -1 | 1 | EV | standard | H2O | 1682466791 | 0.299 | 0.756 | 94.907 | 0 | 0.002 | 2.248 | 1291.144 | 1.647 | 2.75 | 108.261 | 0.864 | 0.354 | 0 | 0 | 1 | 40.688 | 1 | 13 | 1 | GJB_Rinse | 0
76 | P-3871 | 2023/04/25 23:57:07 | 2-13 | 4 | -7.633 | -61.183 | 17488 | 0 | 1 | EV | standard | H2O | 1682467329 | 0.319 | 0.731 | 95.928 | 0.002 | 0.002 | 2.274 | 1291.326 | 1.553 | 2.804 | 108.702 | 0.863 | 0.465 | 0 | 0 | 1 | 40.75 | 1 | 13 | 1 | GJB_Rinse | 0
77 | P-3872 | 2023/04/26 00:06:57 | 2-14 | 1 | [5 empty] | WOOD4-deep-1 | 23-074 | [18 empty] | 1 | 14 | 1 | GJB_Rinse | 0
78 | P-3872 | 2023/04/26 00:15:54 | 2-14 | 2 | -11.595 | -101.044 | 17810 | -1 | 1 | WOOD4-deep-1 | 23-074 | H2O | 1682468456 | 0.307 | 0.738 | 97.946 | 0.002 | 0.006 | 2.307 | 1293.668 | 1.775 | 2.816 | 110.024 | 0.864 | 0.419 | 0 | 0 | 1 | 40.75 | 1 | 14 | 1 | GJB_Rinse | 0
79 | P-3872 | 2023/04/26 00:24:51 | 2-14 | 3 | -11.609 | -101.844 | 17770 | -1 | 1 | WOOD4-deep-1 | 23-074 | H2O | 1682468992 | 0.304 | 0.745 | 96.236 | 0.001 | 0.006 | 2.269 | 1293.349 | 2.183 | 2.769 | 108.808 | 0.863 | 0.232 | 0 | 0 | 1 | 40.25 | 1 | 14 | 1 | GJB_Rinse | 0
80 | P-3872 | 2023/04/26 00:33:48 | 2-14 | 4 | -11.616 | -102.341 | 17726 | 0 | 1 | WOOD4-deep-1 | 23-074 | H2O | 1682469530 | 0.314 | 0.76 | 96.337 | 0.001 | 0.004 | 2.297 | 1293.105 | 2.175 | 2.765 | 107.729 | 0.864 | 0.201 | 0 | 0 | 1 | 40.438 | 1 | 14 | 1 | GJB_Rinse | 0
81 | P-3873 | 2023/04/26 00:43:36 | 2-15 | 1 | -11.48 | -102.117 | 17584 | -1 | 1 | WOOD4-3040-1 | 23-074 | H2O | 1682470117 | 0.288 | 0.732 | 95.19 | 0.002 | 0.004 | 2.249 | 1291.803 | 1.17 | 2.801 | 110.05 | 0.866 | 0.552 | 0 | 0 | 1 | 40.625 | 1 | 15 | 1 | GJB_Rinse | 0
82 | P-3873 | 2023/04/26 00:52:32 | 2-15 | 2 | -11.438 | -102.127 | 17670 | -1 | 1 | WOOD4-3040-1 | 23-074 | H2O | 1682470653 | 0.351 | 0.779 | 98.677 | 0.003 | 0.005 | 2.337 | 1292.287 | 2.292 | 2.755 | 108.008 | 0.864 | 0.37 | 0 | 0 | 1 | 40.625 | 1 | 15 | 1 | GJB_Rinse | 0
83 | P-3873 | 2023/04/26 01:01:29 | 2-15 | 3 | -11.447 | -102.088 | 17657 | -1 | 1 | WOOD4-3040-1 | 23-074 | H2O | 1682471190 | 0.319 | 0.639 | 96.921 | 0.001 | 0.004 | 2.3 | 1292.406 | 2.356 | 2.727 | 107.021 | 0.867 | 0.168 | 0 | 0 | 1 | 40.625 | 1 | 15 | 1 | GJB_Rinse | 0
84 | P-3873 | 2023/04/26 01:10:27 | 2-15 | 4 | -11.468 | -102.236 | 17651 | 0 | 1 | WOOD4-3040-1 | 23-074 | H2O | 1682471729 | 0.314 | 0.841 | 98.214 | 0 | 0.007 | 2.314 | 1292.662 | 1.503 | 2.757 | 108.27 | 0.868 | 0.105 | 0 | 0 | 1 | 40.625 | 1 | 15 | 1 | GJB_Rinse | 0
85 | P-3874 | 2023/04/26 01:20:15 | 2-16 | 1 | 11.604 | -41.673 | 17200 | -1 | 1 | WOOD4-d-5 | 23-074 | H2O | 1682472317 | 0.388 | 1.329 | 94.885 | -0.002 | -0.003 | 2.247 | 1909.159 | -53.029 | 5.421 | 128.561 | 0.87 | -11.292 | 0 | 0 | 1 | 40.688 | 1 | 16 | 1 | GJB_Rinse | 0
86 | P-3874 | 2023/04/26 01:29:12 | 2-16 | 2 | 12.056 | -37.586 | 17337 | -1 | 1 | WOOD4-d-5 | 23-074 | H2O | 1682472854 | 0.439 | 1.296 | 95.59 | -0.002 | -0.003 | 2.251 | 1920.665 | -52.852 | 5.492 | 128.857 | 0.868 | -11.297 | 0 | 0 | 1 | 40.688 | 1 | 16 | 1 | GJB_Rinse | 0
87 | P-3874 | 2023/04/26 01:38:09 | 2-16 | 3 | 12.181 | -36.643 | 17249 | -1 | 1 | WOOD4-d-5 | 23-074 | H2O | 1682473390 | 0.388 | 1.412 | 97.106 | -0.001 | -0.004 | 2.313 | 1919.074 | -53.525 | 5.436 | 129.401 | 0.869 | -11.315 | 0 | 0 | 1 | 40.688 | 1 | 16 | 1 | GJB_Rinse | 0
88 | P-3874 | 2023/04/26 01:47:05 | 2-16 | 4 | 12.263 | -36.481 | 17204 | 0 | 1 | WOOD4-d-5 | 23-074 | H2O | 1682473927 | 0.386 | 1.403 | 95.153 | -0.001 | 0.001 | 2.256 | 1918.162 | -52.946 | 5.477 | 128.904 | 0.873 | -11.011 | 0 | 0 | 1 | 40.688 | 1 | 16 | 1 | GJB_Rinse | 0
89 | P-3875 | 2023/04/26 01:56:53 | 2-17 | 1 | -7.862 | -78.137 | 17273 | -1 | 1 | WOOD4-1020-1 | 23-074 | H2O | 1682474515 | 0.327 | 0.802 | 96.665 | 0.002 | 0.005 | 2.291 | 1299.306 | 1.101 | 2.731 | 107.896 | 0.864 | 0.376 | 0 | 0 | 1 | 40.75 | 1 | 17 | 1 | GJB_Rinse | 0
90 | P-3875 | 2023/04/26 02:05:51 | 2-17 | 2 | -8.052 | -79.414 | 17777 | -1 | 1 | WOOD4-1020-1 | 23-074 | H2O | 1682475052 | 0.297 | 0.735 | 94.32 | 0.001 | 0.002 | 2.238 | 1297.846 | 1.32 | 2.818 | 109.348 | 0.863 | 0.11 | 0 | 0 | 1 | 40.688 | 1 | 17 | 1 | GJB_Rinse | 0
91 | P-3875 | 2023/04/26 02:14:48 | 2-17 | 3 | -8.166 | -79.736 | 17707 | -1 | 1 | WOOD4-1020-1 | 23-074 | H2O | 1682475589 | 0.33 | 0.636 | 95.983 | 0.002 | 0.003 | 2.276 | 1295.553 | 1.431 | 2.767 | 108.739 | 0.863 | 0.056 | 0 | 0 | 1 | 40.688 | 1 | 17 | 1 | GJB_Rinse | 0
92 | P-3875 | 2023/04/26 02:23:45 | 2-17 | 4 | -8.197 | -79.898 | 17738 | 0 | 1 | WOOD4-1020-1 | 23-074 | H2O | 1682476126 | 0.312 | 0.715 | 96.37 | 0.001 | 0.003 | 2.284 | 1295.086 | 1.248 | 2.783 | 109.115 | 0.864 | 0.096 | 0 | 0 | 1 | 40.75 | 1 | 17 | 1 | GJB_Rinse | 0
93 | P-3876 | 2023/04/26 02:33:33 | 2-18 | 1 | -9.635 | -94.839 | 17641 | -1 | 1 | WOOD4-2030-1 | 23-074 | H2O | 1682476714 | 0.312 | 0.785 | 94.497 | 0.001 | 0.007 | 2.266 | 1292.692 | 1.912 | 2.757 | 108.492 | 0.863 | 0.381 | 0 | 0 | 1 | 40.812 | 1 | 18 | 1 | GJB_Rinse | 0
94 | P-3876 | 2023/04/26 02:42:30 | 2-18 | 2 | -9.823 | -96.679 | 17640 | -1 | 1 | WOOD4-2030-1 | 23-074 | H2O | 1682477251 | 0.305 | 0.753 | 97.205 | 0.002 | 0.006 | 2.303 | 1292.989 | 1.91 | 2.755 | 108.238 | 0.864 | 0.384 | 0 | 0 | 1 | 40.812 | 1 | 18 | 1 | GJB_Rinse | 0
95 | P-3876 | 2023/04/26 02:51:27 | 2-18 | 3 | -9.686 | -96.893 | 17682 | -1 | 1 | WOOD4-2030-1 | 23-074 | H2O | 1682477788 | 0.329 | 0.707 | 98.543 | 0.002 | 0.003 | 2.321 | 1293.314 | 1.707 | 2.804 | 108.798 | 0.864 | 0.323 | 0 | 0 | 1 | 40.812 | 1 | 18 | 1 | GJB_Rinse | 0
96 | P-3876 | 2023/04/26 03:00:24 | 2-18 | 4 | -9.746 | -96.871 | 17800 | 0 | 1 | WOOD4-2030-1 | 23-074 | H2O | 1682478325 | 0.311 | 0.659 | 97.67 | 0.001 | 0.002 | 2.306 | 1293.979 | 1.672 | 2.775 | 109.581 | 0.862 | 0.314 | 0 | 0 | 1 | 40.812 | 1 | 18 | 1 | GJB_Rinse | 0
97 | P-3877 | 2023/04/26 03:10:12 | 2-19 | 1 | -9.309 | -88.153 | 16997 | -1 | 1 | WOOD4-010-3 | 23-074 | H2O | 1682478913 | 0.311 | 0.775 | 95.364 | 0.001 | 0.005 | 2.248 | 1287.857 | 1.277 | 2.671 | 105.56 | 0.863 | 0.358 | 0 | 0 | 1 | 40.688 | 1 | 19 | 1 | GJB_Rinse | 0
98 | P-3877 | 2023/04/26 03:19:09 | 2-19 | 2 | -9.386 | -87.665 | 16504 | -1 | 1 | WOOD4-010-3 | 23-074 | H2O | 1682479450 | 0.281 | 0.742 | 91.921 | 0.001 | 0.001 | 2.197 | 1284.385 | 1.384 | 2.633 | 102.483 | 0.862 | 0.214 | 0 | 0 | 1 | 40.75 | 1 | 19 | 1 | GJB_Rinse | 0
99 | P-3877 | 2023/04/26 03:28:05 | 2-19 | 3 | -9.307 | -87.237 | 16498 | -1 | 1 | WOOD4-010-3 | 23-074 | H2O | 1682479987 | 0.322 | 0.845 | 93.509 | 0.002 | 0.008 | 2.235 | 1284.442 | 1.859 | 2.555 | 100.437 | 0.862 | -0.127 | 0 | 0 | 1 | 40.75 | 1 | 19 | 1 | GJB_Rinse | 0
100 | P-3877 | 2023/04/26 03:37:02 | 2-19 | 4 | -9.365 | -87.235 | 16861 | 0 | 1 | WOOD4-010-3 | 23-074 | H2O | 1682480523 | 0.299 | 0.779 | 93.63 | 0.001 | 0.002 | 2.236 | 1286.919 | 1.432 | 2.684 | 104.148 | 0.862 | 0.189 | 0 | 0 | 1 | 40.75 | 1 | 19 | 1 | GJB_Rinse | 0
101 | P-3878 | 2023/04/26 03:46:51 | 2-20 | 1 | -8.392 | -78.632 | 17588 | -1 | 1 | WOOD4-1020-3 | 23-074 | H2O | 1682481112 | 0.318 | 0.829 | 96.919 | 0.002 | 0.003 | 2.299 | 1291.86 | 2.112 | 2.763 | 106.595 | 0.863 | 0.049 | 0 | 0 | 1 | 41.375 | 1 | 20 | 1 | GJB_Rinse | 0
102 | P-3878 | 2023/04/26 03:55:47 | 2-20 | 2 | -8.377 | -77.981 | 17439 | -1 | 1 | WOOD4-1020-3 | 23-074 | H2O | 1682481649 | 0.292 | 0.74 | 95.691 | 0 | 0 | 2.254 | 1290.954 | 1.759 | 2.706 | 106.454 | 0.864 | 0.153 | 0 | 0 | 1 | 40.812 | 1 | 20 | 1 | GJB_Rinse | 0
103 | P-3878 | 2023/04/26 04:04:45 | 2-20 | 3 | -8.335 | -77.563 | 17575 | -1 | 1 | WOOD4-1020-3 | 23-074 | H2O | 1682482186 | 0.314 | 0.722 | 93.082 | 0 | 0.002 | 2.229 | 1291.883 | 1.802 | 2.753 | 107.789 | 0.865 | 0.158 | 0 | 0 | 1 | 40.688 | 1 | 20 | 1 | GJB_Rinse | 0
104 | P-3878 | 2023/04/26 04:13:42 | 2-20 | 4 | -8.368 | -77.398 | 17693 | 0 | 1 | WOOD4-1020-3 | 23-074 | H2O | 1682482723 | 0.34 | 0.76 | 96.933 | 0.001 | 0.006 | 2.306 | 1292.631 | 1.53 | 2.82 | 109.423 | 0.865 | 0.344 | 0 | 0 | 1 | 40.688 | 1 | 20 | 1 | GJB_Rinse | 0
105 | P-3879 | 2023/04/26 04:23:32 | 2-21 | 1 | [5 empty] | WOOD4-3040-3 | 23-074 | [18 empty] | 1 | 21 | 1 | GJB_Rinse | 0
106 | P-3879 | 2023/04/26 04:32:29 | 2-21 | 2 | -11.857 | -103.508 | 17655 | -1 | 1 | WOOD4-3040-3 | 23-074 | H2O | 1682483851 | 0.281 | 0.812 | 95.15 | 0 | 0.007 | 2.258 | 1292.165 | 2.078 | 2.79 | 108.34 | 0.868 | 0.298 | 0 | 0 | 1 | 40.375 | 1 | 21 | 1 | GJB_Rinse | 0
107 | P-3879 | 2023/04/26 04:41:26 | 2-21 | 3 | -11.832 | -104.014 | 17680 | -1 | 1 | WOOD4-3040-3 | 23-074 | H2O | 1682484388 | 0.324 | 0.807 | 99.206 | 0.001 | 0.008 | 2.354 | 1292.525 | 1.485 | 2.77 | 109.757 | 0.867 | 0.395 | 0 | 0 | 1 | 40.562 | 1 | 21 | 1 | GJB_Rinse | 0
108 | P-3879 | 2023/04/26 04:50:23 | 2-21 | 4 | -11.895 | -104.621 | 17724 | 0 | 1 | WOOD4-3040-3 | 23-074 | H2O | 1682484925 | 0.313 | 0.786 | 95.162 | 0.001 | 0.007 | 2.254 | 1292.56 | 2.099 | 2.796 | 109.071 | 0.869 | 0.361 | 0 | 0 | 1 | 40.625 | 1 | 21 | 1 | GJB_Rinse | 0
109 | P-3880 | 2023/04/26 05:00:10 | 2-22 | 1 | [5 empty] | EV | standard | [18 empty] | 1 | 22 | 1 | GJB_Rinse | 0
110 | P-3880 | 2023/04/26 05:09:06 | 2-22 | 2 | -7.699 | -62.364 | 17612 | -1 | 1 | EV | standard | H2O | 1682486048 | 0.303 | 0.725 | 97.844 | 0.002 | 0 | 2.307 | 1291.487 | 2.023 | 2.78 | 109.013 | 0.863 | 0.485 | 0 | 0 | 1 | 40.5 | 1 | 22 | 1 | GJB_Rinse | 0
111 | P-3880 | 2023/04/26 05:18:03 | 2-22 | 3 | -7.653 | -61.144 | 17655 | -1 | 1 | EV | standard | H2O | 1682486585 | 0.304 | 0.762 | 96.684 | 0.001 | 0 | 2.307 | 1292.429 | 1.079 | 2.749 | 108.537 | 0.863 | 0.076 | 0 | 0 | 1 | 40.625 | 1 | 22 | 1 | GJB_Rinse | 0
112 | P-3880 | 2023/04/26 05:27:00 | 2-22 | 4 | -7.652 | -61.065 | 17623 | 0 | 1 | EV | standard | H2O | 1682487121 | 0.325 | 0.759 | 99.97 | 0.002 | 0.004 | 2.398 | 1291.395 | 2.151 | 2.816 | 109.531 | 0.864 | 0.674 | 0 | 0 | 1 | 40.688 | 1 | 22 | 1 | GJB_Rinse | 0
113 | P-3881 | 2023/04/26 05:36:46 | 2-23 | 1 | -1.648 | -82.813 | 17268 | -1 | 1 | WOOD4-B-1 | 23-074 | H2O | 1682487709 | 0.35 | 0.997 | 95.398 | 0 | 0.004 | 2.244 | 1732.762 | -32.07 | 3.671 | 117.479 | 0.87 | -6.003 | 0 | 0 | 1 | 40.75 | 1 | 23 | 1 | GJB_Rinse | 0
114 | P-3881 | 2023/04/26 05:45:44 | 2-23 | 2 | -1.723 | -86.743 | 17384 | -1 | 1 | WOOD4-B-1 | 23-074 | H2O | 1682488246 | 0.363 | 1.008 | 96.549 | 0 | 0.002 | 2.29 | 1740.72 | -32.226 | 3.713 | 117.136 | 0.867 | -6.532 | 0 | 0 | 1 | 40.75 | 1 | 23 | 1 | GJB_Rinse | 0
115 | P-3881 | 2023/04/26 05:54:41 | 2-23 | 3 | -1.713 | -87.423 | 17348 | -1 | 1 | WOOD4-B-1 | 23-074 | H2O | 1682488782 | 0.37 | 1.118 | 96.696 | 0 | 0.006 | 2.296 | 1740.737 | -32.393 | 3.793 | 118.809 | 0.867 | -6.377 | 0 | 0 | 1 | 40.688 | 1 | 23 | 1 | GJB_Rinse | 0
116 | P-3881 | 2023/04/26 06:03:37 | 2-23 | 4 | -1.73 | -87.735 | 17400 | 0 | 1 | WOOD4-B-1 | 23-074 | H2O | 1682489319 | 0.412 | 1.048 | 97.674 | 0 | 0.004 | 2.309 | 1743.902 | -32.655 | 3.696 | 117.626 | 0.881 | -6.572 | 0 | 0 | 1 | 40.688 | 1 | 23 | 1 | GJB_Rinse | 0
117 | P-3882 | 2023/04/26 06:13:25 | 2-24 | 1 | -11.005 | -101.333 | 17562 | -1 | 1 | WOOD4-2030-3 | 23-074 | H2O | 1682489906 | 0.31 | 0.765 | 95.987 | 0.001 | 0.006 | 2.259 | 1298.73 | 1.81 | 2.762 | 108.236 | 0.869 | 0.238 | 0 | 0 | 1 | 40.688 | 1 | 24 | 1 | GJB_Rinse | 0
118 | P-3882 | 2023/04/26 06:22:22 | 2-24 | 2 | -11.052 | -101.226 | 17575 | -1 | 1 | WOOD4-2030-3 | 23-074 | H2O | 1682490443 | 0.326 | 0.721 | 96.652 | 0 | 0.004 | 2.278 | 1293.879 | 1.981 | 2.78 | 108.649 | 0.863 | 0.445 | 0 | 0 | 1 | 40.688 | 1 | 24 | 1 | GJB_Rinse | 0
119 | P-3882 | 2023/04/26 06:31:18 | 2-24 | 3 | -11.085 | -101.171 | 17553 | -1 | 1 | WOOD4-2030-3 | 23-074 | H2O | 1682490980 | 0.322 | 0.729 | 97.441 | 0.002 | 0.007 | 2.31 | 1292.894 | 2.107 | 2.719 | 108.011 | 0.864 | 0.464 | 0 | 0 | 1 | 40.75 | 1 | 24 | 1 | GJB_Rinse | 0
120 | P-3882 | 2023/04/26 06:40:16 | 2-24 | 4 | -11.061 | -101.222 | 17507 | 0 | 1 | WOOD4-2030-3 | 23-074 | H2O | 1682491517 | 0.297 | 0.837 | 96.909 | 0.001 | 0.008 | 2.297 | 1291.752 | 1.745 | 2.768 | 108.777 | 0.863 | 0.335 | 0 | 0 | 1 | 40.562 | 1 | 24 | 1 | GJB_Rinse | 0
121 | P-3883 | 2023/04/26 06:50:03 | 2-25 | 1 | -5.162 | -72.855 | 17476 | -1 | 1 | WOOD4-1020-2 | 23-074 | H2O | 1682492105 | 0.28 | 0.796 | 95.022 | 0.001 | -0.001 | 2.255 | 1291.449 | 1.242 | 2.785 | 108.254 | 0.862 | 0.236 | 0 | 0 | 1 | 40.312 | 1 | 25 | 1 | GJB_Rinse | 0
122 | P-3883 | 2023/04/26 06:59:00 | 2-25 | 2 | -4.801 | -69.605 | 17439 | -1 | 1 | WOOD4-1020-2 | 23-074 | H2O | 1682492641 | 0.315 | 0.77 | 92.725 | 0 | -0.001 | 2.225 | 1291.315 | 1.55 | 2.757 | 107.883 | 0.864 | 0.35 | 0 | 0 | 1 | 40.5 | 1 | 25 | 1 | GJB_Rinse | 0
123 | P-3883 | 2023/04/26 07:07:58 | 2-25 | 3 | -4.758 | -68.799 | 17514 | -1 | 1 | WOOD4-1020-2 | 23-074 | H2O | 1682493180 | 0.272 | 0.662 | 96.752 | 0.001 | -0.001 | 2.282 | 1291.613 | 1.789 | 2.761 | 108.373 | 0.868 | 0.391 | 0 | 0 | 1 | 40.75 | 1 | 25 | 1 | GJB_Rinse | 0
124 | P-3883 | 2023/04/26 07:16:55 | 2-25 | 4 | -4.713 | -68.49 | 17527 | 0 | 1 | WOOD4-1020-2 | 23-074 | H2O | 1682493716 | 0.302 | 0.77 | 94.48 | 0.001 | -0.001 | 2.227 | 1291.642 | 1.708 | 2.83 | 108.266 | 0.863 | 0.357 | 0 | 0 | 1 | 40.812 | 1 | 25 | 1 | GJB_Rinse | 0
125 | P-3884 | 2023/04/26 07:26:43 | 2-26 | 1 | -4.102 | -69.415 | 17271 | -1 | 1 | WOOD4-C-5 | 23-074 | H2O | 1682494305 | 0.322 | 0.733 | 95.419 | 0.001 | 0.002 | 2.264 | 1289.687 | 1.402 | 2.76 | 107.637 | 0.863 | 0.475 | 0 | 0 | 1 | 40.938 | 1 | 26 | 1 | GJB_Rinse | 0
126 | P-3884 | 2023/04/26 07:35:41 | 2-26 | 2 | -4.13 | -69.476 | 17309 | -1 | 1 | WOOD4-C-5 | 23-074 | H2O | 1682494842 | 0.302 | 0.695 | 94.656 | 0 | 0.002 | 2.256 | 1290.347 | 1.287 | 2.737 | 106.454 | 0.863 | 0.118 | 0 | 0 | 1 | 41 | 1 | 26 | 1 | GJB_Rinse | 0
127 | P-3884 | 2023/04/26 07:44:38 | 2-26 | 3 | -4.17 | -69.729 | 17285 | -1 | 1 | WOOD4-C-5 | 23-074 | H2O | 1682495379 | 0.297 | 0.768 | 95.843 | 0.001 | 0.005 | 2.273 | 1289.831 | 1.619 | 2.703 | 106.892 | 0.863 | 0.122 | 0 | 0 | 1 | 40.75 | 1 | 26 | 1 | GJB_Rinse | 0
128 | P-3884 | 2023/04/26 07:53:34 | 2-26 | 4 | -4.07 | -69.3 | 17338 | 0 | 1 | WOOD4-C-5 | 23-074 | H2O | 1682495915 | 0.281 | 0.77 | 95.066 | 0.001 | 0.003 | 2.255 | 1290.203 | 1.778 | 2.741 | 106.804 | 0.863 | 0.094 | 0 | 0 | 1 | 40.625 | 1 | 26 | 1 | GJB_Rinse | 0
129 | P-3885 | 2023/04/26 08:03:23 | 2-27 | 1 | -6.706 | -72.508 | 16628 | -1 | 1 | WOOD4-E-1 | 23-074 | H2O | 1682496504 | 0.374 | 0.966 | 94.868 | 0.001 | 0.003 | 2.27 | 1697.542 | -13.641 | 2.59 | 100.055 | 0.866 | -0.291 | 0 | 0 | 1 | 40.812 | 1 | 27 | 1 | GJB_Rinse | 0
130 | P-3885 | 2023/04/26 08:12:20 | 2-27 | 2 | -6.923 | -72.783 | 16780 | -1 | 1 | WOOD4-E-1 | 23-074 | H2O | 1682497042 | 0.376 | 0.985 | 96.235 | 0 | -0.001 | 2.268 | 1706.821 | -13.652 | 2.596 | 99.374 | 0.867 | -0.36 | 0 | 0 | 1 | 40.812 | 1 | 27 | 1 | GJB_Rinse | 0
131 | P-3885 | 2023/04/26 08:21:17 | 2-27 | 3 | -6.923 | -72.778 | 16720 | -1 | 1 | WOOD4-E-1 | 23-074 | H2O | 1682497578 | 0.352 | 0.936 | 93.649 | 0 | 0.003 | 2.236 | 1706.552 | -14.035 | 2.581 | 99.466 | 0.868 | -0.408 | 0 | 0 | 1 | 40.812 | 1 | 27 | 1 | GJB_Rinse | 0
132 | P-3885 | 2023/04/26 08:30:13 | 2-27 | 4 | -6.823 | -72.146 | 17382 | 0 | 1 | WOOD4-E-1 | 23-074 | H2O | 1682498115 | 0.367 | 0.937 | 95.415 | 0 | 0.003 | 2.266 | 1728.405 | -14.959 | 2.667 | 103.094 | 0.867 | -0.647 | 0 | 0 | 1 | 40.812 | 1 | 27 | 1 | GJB_Rinse | 0
133 | P-3886 | 2023/04/26 08:40:02 | 2-28 | 1 | -5.339 | -62.766 | 16628 | -1 | 1 | WOOD4-A-1 | 23-074 | H2O | 1682498703 | 0.429 | 1.379 | 95.138 | 0.001 | 0.002 | 2.26 | 2317.416 | -38.544 | 2.691 | 98.556 | 0.871 | -0.889 | 0 | 0 | 1 | 41.312 | 1 | 28 | 1 | GJB_Rinse | 0
134 | P-3886 | 2023/04/26 08:48:58 | 2-28 | 2 | -5.363 | -61.666 | 16723 | -1 | 1 | WOOD4-A-1 | 23-074 | H2O | 1682499239 | 0.442 | 1.443 | 93.553 | 0.002 | 0.003 | 2.209 | 2334.196 | -39.573 | 2.586 | 96.828 | 0.871 | -1.312 | 0 | 0 | 1 | 41.125 | 1 | 28 | 1 | GJB_Rinse | 0
135 | P-3886 | 2023/04/26 08:57:54 | 2-28 | 3 | -5.345 | -61.563 | 16556 | -1 | 1 | WOOD4-A-1 | 23-074 | H2O | 1682499776 | 0.442 | 1.273 | 93.455 | 0 | 0 | 2.222 | 2321.801 | -38.511 | 2.583 | 94.857 | 0.87 | -1.398 | 0 | 0 | 1 | 40.938 | 1 | 28 | 1 | GJB_Rinse | 0
136 | P-3886 | 2023/04/26 09:06:52 | 2-28 | 4 | -5.405 | -61.548 | 16713 | 0 | 1 | WOOD4-A-1 | 23-074 | H2O | 1682500314 | 0.493 | 1.358 | 96.189 | 0 | 0.003 | 2.266 | 2334.964 | -38.668 | 2.653 | 96.002 | 0.873 | -1.284 | 0 | 0 | 1 | 40.875 | 1 | 28 | 1 | GJB_Rinse | 0
137 | P-3887 | 2023/04/26 09:16:37 | 2-29 | 1 | [5 empty] | WOOD4-D-2 | 23-074 | [18 empty] | 1 | 29 | 1 | GJB_Rinse | 0
138 | P-3887 | 2023/04/26 09:25:34 | 2-29 | 2 | -7.661 | -84.119 | 16983 | -1 | 1 | WOOD4-D-2 | 23-074 | H2O | 1682501436 | 0.321 | 0.789 | 95.688 | 0.001 | 0.004 | 2.272 | 1296.241 | 0.948 | 2.636 | 103.695 | 0.862 | 0.032 | 0 | 0 | 1 | 40.938 | 1 | 29 | 1 | GJB_Rinse | 0
139 | P-3887 | 2023/04/26 09:34:31 | 2-29 | 3 | -7.673 | -84.7 | 16987 | -1 | 1 | WOOD4-D-2 | 23-074 | H2O | 1682501972 | 0.3 | 0.724 | 96.01 | 0.002 | 0.003 | 2.274 | 1293.246 | 0.722 | 2.658 | 105.418 | 0.864 | 0.082 | 0 | 0 | 1 | 40.938 | 1 | 29 | 1 | GJB_Rinse | 0
140 | P-3887 | 2023/04/26 09:43:27 | 2-29 | 4 | -7.685 | -84.912 | 16980 | 0 | 1 | WOOD4-D-2 | 23-074 | H2O | 1682502509 | 0.283 | 0.735 | 95.568 | 0.001 | 0.007 | 2.277 | 1291.564 | 1.068 | 2.695 | 105.733 | 0.864 | 0.21 | 0 | 0 | 1 | 40.75 | 1 | 29 | 1 | GJB_Rinse | 0
141 | P-3888 | 2023/04/26 09:53:13 | 2-30 | 1 | [5 empty] | WOOD4-C-2 | 23-074 | [18 empty] | 1 | 30 | 1 | GJB_Rinse | 0
142 | P-3888 | 2023/04/26 10:02:09 | 2-30 | 2 | -2.381 | -57.412 | 16811 | -1 | 1 | WOOD4-C-2 | 23-074 | H2O | 1682503631 | 0.288 | 0.734 | 95.642 | 0.001 | 0.001 | 2.272 | 1291.865 | 1.076 | 2.639 | 104.812 | 0.862 | 0.232 | 0 | 0 | 1 | 41 | 1 | 30 | 1 | GJB_Rinse | 0
143 | P-3888 | 2023/04/26 10:11:06 | 2-30 | 3 | -2.302 | -56.387 | 17307 | -1 | 1 | WOOD4-C-2 | 23-074 | H2O | 1682504167 | 0.322 | 0.696 | 95.807 | 0.001 | 0.002 | 2.295 | 1295.162 | 1.711 | 2.706 | 106.173 | 0.865 | 0.144 | 0 | 0 | 1 | 41.062 | 1 | 30 | 1 | GJB_Rinse | 0
144 | P-3888 | 2023/04/26 10:20:02 | 2-30 | 4 | -2.261 | -56.032 | 17452 | 0 | 1 | WOOD4-C-2 | 23-074 | H2O | 1682504704 | 0.289 | 0.778 | 96.879 | 0.001 | 0.001 | 2.287 | 1295.981 | 1.446 | 2.798 | 108.343 | 0.862 | 0.441 | 0 | 0 | 1 | 41.125 | 1 | 30 | 1 | GJB_Rinse | 0
145 | P-3889 | 2023/04/26 10:29:48 | 2-31 | 1 | -7.224 | -60.002 | 17537 | -1 | 1 | EV | standard | H2O | 1682505290 | 0.322 | 0.781 | 97.694 | 0.001 | 0.005 | 2.319 | 1291.901 | 1.95 | 2.722 | 106.501 | 0.862 | -0.008 | 0 | 0 | 1 | 41.25 | 1 | 31 | 1 | GJB_Rinse | 0
146 | P-3889 | 2023/04/26 10:38:45 | 2-31 | 2 | -7.506 | -60.366 | 17549 | -1 | 1 | EV | standard | H2O | 1682505827 | 0.293 | 0.716 | 94.011 | 0 | 0.001 | 2.233 | 1292.306 | 2.02 | 2.744 | 106.394 | 0.861 | 0.094 | 0 | 0 | 1 | 41.25 | 1 | 31 | 1 | GJB_Rinse | 0
147 | P-3889 | 2023/04/26 10:47:42 | 2-31 | 3 | -7.56 | -60.396 | 17550 | -1 | 1 | EV | standard | H2O | 1682506364 | 0.282 | 0.812 | 96.466 | 0.001 | 0.003 | 2.287 | 1291.95 | 1.764 | 2.806 | 108.812 | 0.864 | 0.425 | 0 | 0 | 1 | 41.062 | 1 | 31 | 1 | GJB_Rinse | 0
148 | P-3889 | 2023/04/26 10:56:39 | 2-31 | 4 | -7.634 | -60.721 | 17497 | 0 | 1 | EV | standard | H2O | 1682506900 | 0.321 | 0.769 | 97.026 | 0.001 | 0.001 | 2.3 | 1291.564 | 1.44 | 2.772 | 108.5 | 0.864 | 0.305 | 0 | 0 | 1 | 41.125 | 1 | 31 | 1 | GJB_Rinse | 0
149 | P-3890 | 2023/04/26 11:06:25 | 2-32 | 1 | -3.454 | -62.305 | 16899 | -1 | 1 | WOOD4-C-4 | 23-074 | H2O | 1682507486 | 0.31 | 0.766 | 93.584 | 0 | 0.001 | 2.232 | 1289.445 | 0.909 | 2.738 | 106.23 | 0.863 | 0.44 | 0 | 0 | 1 | 40.75 | 1 | 32 | 1 | GJB_Rinse | 0
150 | P-3890 | 2023/04/26 11:15:21 | 2-32 | 2 | -3.214 | -62.332 | 16991 | -1 | 1 | WOOD4-C-4 | 23-074 | H2O | 1682508022 | 0.339 | 0.809 | 92.779 | 0.002 | 0.003 | 2.238 | 1290.805 | 0.811 | 2.669 | 104.245 | 0.864 | -0.255 | 0 | 0 | 1 | 40.812 | 1 | 32 | 1 | GJB_Rinse | 0
151 | P-3890 | 2023/04/26 11:24:18 | 2-32 | 3 | -3.165 | -62.521 | 16920 | -1 | 1 | WOOD4-C-4 | 23-074 | H2O | 1682508559 | 0.329 | 0.718 | 95.297 | 0.002 | 0.003 | 2.254 | 1290.309 | 1.07 | 2.64 | 103.145 | 0.865 | -0.179 | 0 | 0 | 1 | 40.812 | 1 | 32 | 1 | GJB_Rinse | 0
152 | P-3890 | 2023/04/26 11:33:15 | 2-32 | 4 | -3.144 | -62.622 | 17007 | 0 | 1 | WOOD4-C-4 | 23-074 | H2O | 1682509097 | 0.298 | 0.739 | 97.182 | 0.001 | 0.002 | 2.29 | 1290.696 | 0.942 | 2.672 | 105.306 | 0.864 | 0.086 | 0 | 0 | 1 | 40.688 | 1 | 32 | 1 | GJB_Rinse | 0
153 | P-3891 | 2023/04/26 11:43:01 | 2-33 | 1 | -5.071 | -70.368 | 16688 | -1 | 1 | WOOD4-A-4 | 23-074 | H2O | 1682509683 | 0.338 | 0.782 | 94.098 | 0.001 | 0.003 | 2.233 | 1289.577 | 0.606 | 2.59 | 102.371 | 0.867 | -0.203 | 0 | 0 | 1 | 40.562 | 1 | 33 | 1 | GJB_Rinse | 0
154 | P-3891 | 2023/04/26 11:51:58 | 2-33 | 2 | -5.146 | -71.151 | 16921 | -1 | 1 | WOOD4-A-4 | 23-074 | H2O | 1682510219 | 0.321 | 0.757 | 95.306 | 0.001 | 0.004 | 2.258 | 1291.153 | 1.272 | 2.631 | 103.318 | 0.873 | -0.077 | 0 | 0 | 1 | 40.25 | 1 | 33 | 1 | GJB_Rinse | 0
155 | P-3891 | 2023/04/26 12:00:54 | 2-33 | 3 | -5.138 | -71.168 | 17104 | -1 | 1 | WOOD4-A-4 | 23-074 | H2O | 1682510756 | 0.293 | 0.778 | 94.893 | 0.001 | 0.002 | 2.24 | 1292.376 | 0.858 | 2.714 | 105.929 | 0.872 | 0.284 | 0 | 0 | 1 | 40.062 | 1 | 33 | 1 | GJB_Rinse | 0
156 | P-3891 | 2023/04/26 12:09:51 | 2-33 | 4 | -5.353 | -72.507 | 14809 | 0 | 1 | WOOD4-A-4 | 23-074 | H2O | 1682511293 | 0.313 | 0.805 | 90.672 | 0.001 | 0.002 | 2.165 | 1275.635 | 1.245 | 2.26 | 90.823 | 0.874 | 0.085 | 0 | 0 | 1 | 39.938 | 1 | 33 | 1 | GJB_Rinse | 0
157 | P-3892 | 2023/04/26 12:19:38 | 2-34 | 1 | -4.612 | -63.442 | 17045 | -1 | 1 | WOOD4-A-2 | 23-074 | H2O | 1682511880 | 0.31 | 0.775 | 95.229 | 0 | 0.001 | 2.239 | 1288.312 | 1.047 | 2.731 | 106.328 | 0.875 | 0.468 | 0 | 0 | 1 | 39.938 | 1 | 34 | 1 | GJB_Rinse | 0
158 | P-3892 | 2023/04/26 12:28:35 | 2-34 | 2 | -4.56 | -62.422 | 17131 | -1 | 1 | WOOD4-A-2 | 23-074 | H2O | 1682512417 | 0.299 | 0.731 | 95.254 | 0.001 | 0.003 | 2.252 | 1289.242 | 1.491 | 2.687 | 104.263 | 0.876 | -0.016 | 0 | 0 | 1 | 40.125 | 1 | 34 | 1 | GJB_Rinse | 0
159 | P-3892 | 2023/04/26 12:37:31 | 2-34 | 3 | -4.625 | -62.346 | 17087 | -1 | 1 | WOOD4-A-2 | 23-074 | H2O | 1682512953 | 0.329 | 0.689 | 95.814 | 0 | 0 | 2.26 | 1288.536 | 1.672 | 2.684 | 104.808 | 0.877 | 0.188 | 0 | 0 | 1 | 40.062 | 1 | 34 | 1 | GJB_Rinse | 0
160 | P-3892 | 2023/04/26 12:46:28 | 2-34 | 4 | -4.536 | -62.239 | 17078 | 0 | 1 | WOOD4-A-2 | 23-074 | H2O | 1682513491 | 0.322 | 0.715 | 95.31 | 0.001 | 0 | 2.251 | 1289.06 | 0.91 | 2.624 | 104.465 | 0.877 | -0.124 | 0 | 0 | 1 | 40.062 | 1 | 34 | 1 | GJB_Rinse | 0
161 | P-3893 | 2023/04/26 12:56:15 | 2-35 | 1 | -6.029 | -71.772 | 17371 | -1 | 1 | WOOD4-E-3 | 23-074 | H2O | 1682514077 | 0.348 | 0.786 | 96.66 | 0.002 | 0.007 | 2.286 | 1290.538 | 1.773 | 2.71 | 107.108 | 0.876 | 0.197 | 0 | 0 | 1 | 40.062 | 1 | 35 | 1 | GJB_Rinse | 0
162 | P-3893 | 2023/04/26 13:05:12 | 2-35 | 2 | -6.141 | -73.01 | 17340 | -1 | 1 | WOOD4-E-3 | 23-074 | H2O | 1682514613 | 0.329 | 0.768 | 95.088 | 0.002 | 0.002 | 2.259 | 1290.459 | 1.463 | 2.743 | 106.902 | 0.877 | 0.204 | 0 | 0 | 1 | 40 | 1 | 35 | 1 | GJB_Rinse | 0
163 | P-3893 | 2023/04/26 13:14:08 | 2-35 | 3 | -6.138 | -73.11 | 17375 | -1 | 1 | WOOD4-E-3 | 23-074 | H2O | 1682515150 | 0.299 | 0.74 | 95.918 | 0.001 | 0.003 | 2.26 | 1291.023 | 1.28 | 2.718 | 106.547 | 0.874 | -0.017 | 0 | 0 | 1 | 40 | 1 | 35 | 1 | GJB_Rinse | 0
164 | P-3893 | 2023/04/26 13:23:05 | 2-35 | 4 | -6.078 | -73.428 | 17349 | 0 | 1 | WOOD4-E-3 | 23-074 | H2O | 1682515686 | 0.285 | 0.764 | 94.853 | 0.002 | 0.005 | 2.235 | 1290.382 | 1.319 | 2.774 | 108.647 | 0.879 | 0.648 | 0 | 0 | 1 | 40 | 1 | 35 | 1 | GJB_Rinse | 0
165 | P-3894 | 2023/04/26 13:32:48 | 2-36 | 1 | -4.628 | -66.473 | 17626 | -1 | 1 | WOOD4-C-1 | 23-074 | H2O | 1682516269 | 0.298 | 0.7 | 94.72 | 0.001 | -0.001 | 2.234 | 1293.618 | 1.286 | 2.798 | 109.521 | 0.873 | 0.239 | 0 | 0 | 1 | 39.938 | 1 | 36 | 1 | GJB_Rinse | 0
166 | P-3894 | 2023/04/26 13:41:45 | 2-36 | 2 | -4.488 | -65.608 | 17547 | -1 | 1 | WOOD4-C-1 | 23-074 | H2O | 1682516806 | 0.336 | 0.773 | 96.423 | 0.002 | 0.005 | 2.265 | 1293.241 | 1.817 | 2.729 | 106.065 | 0.881 | -0.059 | 0 | 0 | 1 | 39.812 | 1 | 36 | 1 | GJB_Rinse | 0
167 | P-3894 | 2023/04/26 13:50:41 | 2-36 | 3 | -4.513 | -65.348 | 17534 | -1 | 1 | WOOD4-C-1 | 23-074 | H2O | 1682517343 | 0.32 | 0.723 | 96.684 | 0.001 | 0.002 | 2.282 | 1293.202 | 1.326 | 2.734 | 107.229 | 0.877 | 0.116 | 0 | 0 | 1 | 39.75 | 1 | 36 | 1 | GJB_Rinse | 0
168 | P-3894 | 2023/04/26 13:59:37 | 2-36 | 4 | -4.461 | -65.042 | 17549 | 0 | 1 | WOOD4-C-1 | 23-074 | H2O | 1682517879 | 0.336 | 0.82 | 95.983 | 0.001 | 0.006 | 2.272 | 1293.119 | 1.806 | 2.737 | 106.821 | 0.88 | 0.243 | 0 | 0 | 1 | 39.562 | 1 | 36 | 1 | GJB_Rinse | 0
169 | P-3895 | 2023/04/26 14:09:21 | 2-37 | 1 | 17.411 | -21.328 | 15911 | -1 | 1 | WOOD4-A-3 | 23-074 | H2O | 1682518463 | 0.35 | 1.319 | 91.478 | -0.002 | -0.006 | 2.18 | 1618.085 | -49.476 | 6.273 | 128.907 | 0.874 | -13.055 | 0 | 0 | 1 | 39.688 | 1 | 37 | 1 | GJB_Rinse | 0
170 | P-3895 | 2023/04/26 14:18:18 | 2-37 | 2 | [5 empty] | WOOD4-A-3 | 23-074 | [18 empty] | 1 | 37 | 1 | GJB_Rinse | 0
171 | P-3895 | 2023/04/26 14:27:14 | 2-37 | 3 | 17.541 | -20.739 | 14417 | -1 | 1 | WOOD4-A-3 | 23-074 | H2O | 1682519536 | 0.383 | 1.334 | 90.348 | -0.004 | -0.005 | 2.158 | 1579.545 | -44.571 | 5.721 | 116.261 | 0.872 | -11.911 | 0 | 0 | 1 | 39.688 | 1 | 37 | 1 | GJB_Rinse | 0
172 | P-3895 | 2023/04/26 14:36:10 | 2-37 | 4 | 17.74 | -19.424 | 15719 | 0 | 1 | WOOD4-A-3 | 23-074 | H2O | 1682520072 | 0.355 | 1.314 | 91.612 | -0.002 | 0.002 | 2.167 | 1619.673 | -49.676 | 6.279 | 126.19 | 0.875 | -13.279 | 0 | 0 | 1 | 39.562 | 1 | 37 | 1 | GJB_Rinse | 0
173 | P-3896 | 2023/04/26 14:45:55 | 2-38 | 1 | -6.406 | -68.159 | 17023 | -1 | 1 | WOOD4-C-3 | 23-074 | H2O | 1682520657 | 0.365 | 0.899 | 95.183 | 0.002 | 0.004 | 2.25 | 1413.403 | -4.341 | 2.634 | 105.007 | 0.876 | -0.273 | 0 | 0 | 1 | 39.625 | 1 | 38 | 1 | GJB_Rinse | 0
174 | P-3896 | 2023/04/26 14:54:51 | 2-38 | 2 | -6.714 | -69.882 | 17390 | -1 | 1 | WOOD4-C-3 | 23-074 | H2O | 1682521193 | 0.346 | 0.831 | 95.315 | 0.001 | 0.005 | 2.261 | 1416.409 | -4.203 | 2.774 | 108.246 | 0.884 | 0.172 | 0 | 0 | 1 | 39.5 | 1 | 38 | 1 | GJB_Rinse | 0
175 | P-3896 | 2023/04/26 15:03:48 | 2-38 | 3 | -6.801 | -70.132 | 17397 | -1 | 1 | WOOD4-C-3 | 23-074 | H2O | 1682521729 | 0.353 | 0.806 | 95.915 | 0.001 | 0.004 | 2.263 | 1416.382 | -4.216 | 2.704 | 105.493 | 0.888 | -0.485 | 0 | 0 | 1 | 39.562 | 1 | 38 | 1 | GJB_Rinse | 0
176 | P-3896 | 2023/04/26 15:12:44 | 2-38 | 4 | -6.854 | -70.44 | 17524 | 0 | 1 | WOOD4-C-3 | 23-074 | H2O | 1682522266 | 0.336 | 0.803 | 93.046 | 0.002 | 0.001 | 2.209 | 1417.214 | -3.98 | 2.801 | 108.058 | 0.901 | 0.082 | 0 | 0 | 1 | 39.5 | 1 | 38 | 1 | GJB_Rinse | 0
177 | P-3897 | 2023/04/26 15:22:28 | 2-39 | 1 | -5.978 | -72.968 | 17057 | -1 | 1 | WOOD4-D-3 | 23-074 | H2O | 1682522850 | 0.311 | 0.796 | 94.674 | 0.001 | 0.004 | 2.251 | 1291.466 | 1.524 | 2.716 | 105.197 | 0.889 | 0.22 | 0 | 0 | 1 | 39.5 | 1 | 39 | 1 | GJB_Rinse | 0
178 | P-3897 | 2023/04/26 15:31:25 | 2-39 | 2 | -5.787 | -73.018 | 17334 | -1 | 1 | WOOD4-D-3 | 23-074 | H2O | 1682523386 | 0.318 | 0.749 | 96.487 | 0.001 | 0.003 | 2.265 | 1292.496 | 1.117 | 2.73 | 106.518 | 0.882 | 0.138 | 0 | 0 | 1 | 39.5 | 1 | 39 | 1 | GJB_Rinse | 0
179 | P-3897 | 2023/04/26 15:40:21 | 2-39 | 3 | -5.795 | -73.041 | 17353 | -1 | 1 | WOOD4-D-3 | 23-074 | H2O | 1682523923 | 0.31 | 0.804 | 94.003 | 0.001 | 0.002 | 2.221 | 1291.945 | 1.385 | 2.715 | 106.698 | 0.876 | 0.189 | 0 | 0 | 1 | 39.5 | 1 | 39 | 1 | GJB_Rinse | 0
180 | P-3897 | 2023/04/26 15:49:18 | 2-39 | 4 | -5.773 | -73.107 | 17354 | 0 | 1 | WOOD4-D-3 | 23-074 | H2O | 1682524460 | 0.357 | 0.752 | 97.859 | 0.001 | 0.003 | 2.309 | 1291.974 | 0.591 | 2.763 | 107.831 | 0.882 | 0.079 | 0 | 0 | 1 | 39.5 | 1 | 39 | 1 | GJB_Rinse | 0
181 | P-3898 | 2023/04/26 15:59:03 | 2-40 | 1 | [5 empty] | EV | standard | [18 empty] | 1 | 40 | 1 | GJB_Rinse | 0
182 | P-3898 | 2023/04/26 16:08:00 | 2-40 | 2 | -7.469 | -60.451 | 17728 | -1 | 1 | EV | standard | H2O | 1682525581 | 0.307 | 0.722 | 98.145 | 0.002 | 0.003 | 2.325 | 1293.237 | 1.213 | 2.808 | 109.959 | 0.875 | 0.374 | 0 | 0 | 1 | 39.562 | 1 | 40 | 1 | GJB_Rinse | 0
183 | P-3898 | 2023/04/26 16:16:56 | 2-40 | 3 | -7.547 | -60.68 | 17337 | -1 | 1 | EV | standard | H2O | 1682526117 | 0.309 | 0.759 | 94.372 | 0.002 | 0.002 | 2.247 | 1290.023 | 0.826 | 2.829 | 110.201 | 0.882 | 0.788 | 0 | 0 | 1 | 39.5 | 1 | 40 | 1 | GJB_Rinse | 0
184 | P-3898 | 2023/04/26 16:25:52 | 2-40 | 4 | -7.538 | -60.357 | 17272 | 0 | 1 | EV | standard | H2O | 1682526654 | 0.32 | 0.795 | 98.544 | 0.002 | 0.003 | 2.34 | 1289.37 | 2.32 | 2.717 | 105.541 | 0.872 | 0.246 | 0 | 0 | 1 | 39.812 | 1 | 40 | 1 | GJB_Rinse | 0
185 | P-3899 | 2023/04/26 16:35:37 | 2-41 | 1 | -4.887 | -67.081 | 17540 | -1 | 1 | WOOD4-E-2 | 23-074 | H2O | 1682527239 | 0.309 | 0.737 | 96.399 | 0.001 | 0.003 | 2.273 | 1292.016 | 1.711 | 2.705 | 106.96 | 0.868 | 0.06 | 0 | 0 | 1 | 40 | 1 | 41 | 1 | GJB_Rinse | 0
186 | P-3899 | 2023/04/26 16:44:33 | 2-41 | 2 | -4.71 | -67.706 | 17386 | -1 | 1 | WOOD4-E-2 | 23-074 | H2O | 1682527775 | 0.307 | 0.86 | 96.859 | 0.001 | 0.008 | 2.296 | 1290.937 | 1.285 | 2.751 | 107.519 | 0.873 | 0.165 | 0 | 0 | 1 | 39.688 | 1 | 41 | 1 | GJB_Rinse | 0
187 | P-3899 | 2023/04/26 16:53:30 | 2-41 | 3 | -4.727 | -68.188 | 17105 | -1 | 1 | WOOD4-E-2 | 23-074 | H2O | 1682528313 | 0.311 | 0.75 | 95.454 | 0 | 0.003 | 2.251 | 1288.852 | 1.327 | 2.718 | 105.941 | 0.887 | 0.51 | 0 | 0 | 1 | 39.312 | 1 | 41 | 1 | GJB_Rinse | 0
188 | P-3899 | 2023/04/26 17:02:27 | 2-41 | 4 | -4.673 | -68.022 | 17348 | 0 | 1 | WOOD4-E-2 | 23-074 | H2O | 1682528848 | 0.316 | 0.755 | 96.069 | 0.002 | 0.006 | 2.271 | 1290.654 | 1.211 | 2.776 | 107.521 | 0.887 | 0.32 | 0 | 0 | 1 | 39.312 | 1 | 41 | 1 | GJB_Rinse | 0
189 | P-3900 | 2023/04/26 17:12:12 | 2-42 | 1 | 5.169 | -54.857 | 14488 | -1 | 1 | WOOD4-B-5 | 23-074 | H2O | 1682529434 | 0.357 | 1.125 | 88.594 | -0.002 | 0.001 | 2.122 | 1497.647 | -25.127 | 3.48 | 102.747 | 0.879 | -6.383 | 0 | 0 | 1 | 39.375 | 1 | 42 | 1 | GJB_Rinse | 0
190 | P-3900 | 2023/04/26 17:21:09 | 2-42 | 2 | 5.308 | -54.754 | 15236 | -1 | 1 | WOOD4-B-5 | 23-074 | H2O | 1682529971 | 0.391 | 0.966 | 91.993 | -0.001 | 0.003 | 2.181 | 1518.962 | -26.371 | 3.7 | 107.032 | 0.884 | -6.808 | 0 | 0 | 1 | 39.312 | 1 | 42 | 1 | GJB_Rinse | 0
191 | P-3900 | 2023/04/26 17:30:05 | 2-42 | 3 | 5.382 | -54.002 | 16708 | -1 | 1 | WOOD4-B-5 | 23-074 | H2O | 1682530507 | 0.361 | 1.058 | 93.531 | -0.001 | -0.001 | 2.211 | 1553.769 | -29.839 | 4.082 | 118.888 | 0.909 | -7.518 | 0 | 0 | 1 | 39.312 | 1 | 42 | 1 | GJB_Rinse | 0
192 | P-3900 | 2023/04/26 17:39:02 | 2-42 | 4 | 5.392 | -54.026 | 16463 | 0 | 1 | WOOD4-B-5 | 23-074 | H2O | 1682531044 | 0.305 | 1.042 | 94.002 | -0.001 | 0.005 | 2.243 | 1549.022 | -29.682 | 3.961 | 115.352 | 0.885 | -7.749 | 0 | 0 | 1 | 39.25 | 1 | 42 | 1 | GJB_Rinse | 0
193 | P-3901 | 2023/04/26 17:48:47 | 2-43 | 1 | -9.22 | -95.403 | 17069 | -1 | 1 | WOOD4-deep-2 | 23-074 | H2O | 1682531628 | 0.308 | 0.774 | 94.575 | 0.002 | 0.007 | 2.241 | 1292.784 | 1.107 | 2.657 | 105.381 | 0.878 | 0.11 | 0 | 0 | 1 | 39.312 | 1 | 43 | 1 | GJB_Rinse | 0
194 | P-3901 | 2023/04/26 17:57:43 | 2-43 | 2 | -9.492 | -98.079 | 17049 | -1 | 1 | WOOD4-deep-2 | 23-074 | H2O | 1682532165 | 0.268 | 0.914 | 93.305 | 0.001 | 0.006 | 2.203 | 1290.059 | 1.305 | 2.655 | 104.719 | 0.878 | 0.127 | 0 | 0 | 1 | 39.25 | 1 | 43 | 1 | GJB_Rinse | 0
195 | P-3901 | 2023/04/26 18:06:40 | 2-43 | 3 | -9.52 | -98.749 | 17047 | -1 | 1 | WOOD4-deep-2 | 23-074 | H2O | 1682532702 | 0.321 | 0.801 | 92.668 | 0.001 | 0.005 | 2.192 | 1289.33 | 1.635 | 2.596 | 103.174 | 0.876 | -0.124 | 0 | 0 | 1 | 39.562 | 1 | 43 | 1 | GJB_Rinse | 0
196 | P-3901 | 2023/04/26 18:15:36 | 2-43 | 4 | -9.636 | -99.179 | 16996 | 0 | 1 | WOOD4-deep-2 | 23-074 | H2O | 1682533238 | 0.319 | 0.816 | 96.003 | 0.001 | 0.002 | 2.262 | 1288.152 | 1.573 | 2.699 | 106.018 | 0.879 | 0.466 | 0 | 0 | 1 | 39.5 | 1 | 43 | 1 | GJB_Rinse | 0
197 | P-3902 | 2023/04/26 18:25:21 | 2-44 | 1 | 13.562 | -40.006 | 13625 | -1 | 1 | WOOD4-B-2 | 23-074 | H2O | 1682533822 | 0.456 | 1.436 | 89.561 | -0.003 | -0.009 | 2.147 | 1962.888 | -54.186 | 5.008 | 103.402 | 0.884 | -10.827 | 0 | 0 | 1 | 39.438 | 1 | 44 | 1 | GJB_Rinse | 0
198 | P-3902 | 2023/04/26 18:34:17 | 2-44 | 2 | 14.035 | -36.46 | 12757 | -1 | 1 | WOOD4-B-2 | 23-074 | H2O | 1682534359 | 0.422 | 1.594 | 85.49 | -0.003 | -0.007 | 2.059 | 1924.072 | -51.517 | 4.762 | 97.67 | 0.876 | -10.659 | 0 | 0 | 1 | 39.5 | 1 | 44 | 1 | GJB_Rinse | 0
199 | P-3902 | 2023/04/26 18:43:13 | 2-44 | 3 | 13.977 | -35.076 | 14324 | -1 | 1 | WOOD4-B-2 | 23-074 | H2O | 1682534895 | 0.401 | 1.613 | 86.441 | -0.002 | -0.006 | 2.051 | 2009.786 | -56.485 | 5.321 | 108.736 | 0.876 | -11.395 | 0 | 0 | 1 | 39.5 | 1 | 44 | 1 | GJB_Rinse | 0
200 | P-3902 | 2023/04/26 18:52:10 | 2-44 | 4 | 14.147 | -34.399 | 13947 | 0 | 1 | WOOD4-B-2 | 23-074 | H2O | 1682535431 | 0.436 | 1.401 | 87.724 | -0.004 | -0.002 | 2.095 | 1993.203 | -56.911 | 5.224 | 108.396 | 0.877 | -10.945 | 0 | 0 | 1 | 39.5 | 1 | 44 | 1 | GJB_Rinse | 0
201 | P-3903 | 2023/04/26 19:01:55 | 2-45 | 1 | -9.084 | -94.109 | 15631 | -1 | 1 | WOOD4-3040-2 | 23-074 | H2O | 1682536017 | 0.319 | 0.767 | 89.209 | 0 | 0.003 | 2.118 | 1290.305 | 0.387 | 2.392 | 96.24 | 0.872 | -0.263 | 0 | 0 | 1 | 39.5 | 1 | 45 | 1 | GJB_Rinse | 0
202 | P-3903 | 2023/04/26 19:10:51 | 2-45 | 2 | -9.37 | -97.015 | 16959 | -1 | 1 | WOOD4-3040-2 | 23-074 | H2O | 1682536553 | 0.296 | 0.763 | 93.912 | 0.001 | 0.005 | 2.216 | 1291.347 | 1.295 | 2.661 | 104.41 | 0.873 | 0.103 | 0 | 0 | 1 | 39.438 | 1 | 45 | 1 | GJB_Rinse | 0
203 | P-3903 | 2023/04/26 19:19:47 | 2-45 | 3 | -9.498 | -98.05 | 16465 | -1 | 1 | WOOD4-3040-2 | 23-074 | H2O | 1682537089 | 0.299 | 0.727 | 93.906 | 0.001 | 0.003 | 2.224 | 1286.487 | 1.434 | 2.525 | 100.346 | 0.87 | 0.024 | 0 | 0 | 1 | 39.5 | 1 | 45 | 1 | GJB_Rinse | 0
204 | P-3903 | 2023/04/26 19:28:44 | 2-45 | 4 | -9.554 | -98.311 | 16770 | 0 | 1 | WOOD4-3040-2 | 23-074 | H2O | 1682537626 | 0.31 | 0.749 | 92.21 | 0.001 | 0.002 | 2.174 | 1287.874 | 1.373 | 2.585 | 103.121 | 0.872 | 0.147 | 0 | 0 | 1 | 39.5 | 1 | 45 | 1 | GJB_Rinse | 0
205 | P-3904 | 2023/04/26 19:38:30 | 2-46 | 1 | [5 empty] | WOOD4-E-5 | 23-074 | [18 empty] | 1 | 46 | 1 | GJB_Rinse | 0
206 | P-3904 | 2023/04/26 19:47:26 | 2-46 | 2 | -4.316 | -62.77 | 17707 | -1 | 1 | WOOD4-E-5 | 23-074 | H2O | 1682538748 | 0.305 | 0.746 | 97.575 | 0 | 0.002 | 2.3 | 1295.948 | 1.435 | 2.79 | 109.127 | 0.878 | 0.225 | 0 | 0 | 1 | 39.562 | 1 | 46 | 1 | GJB_Rinse | 0
207 | P-3904 | 2023/04/26 19:56:22 | 2-46 | 3 | -4.317 | -62.267 | 17600 | -1 | 1 | WOOD4-E-5 | 23-074 | H2O | 1682539284 | 0.3 | 0.719 | 94.408 | 0.001 | 0.002 | 2.237 | 1294.829 | 1.658 | 2.813 | 108.707 | 0.879 | 0.3 | 0 | 0 | 1 | 39.562 | 1 | 46 | 1 | GJB_Rinse | 0
208 | P-3904 | 2023/04/26 20:05:19 | 2-46 | 4 | -4.336 | -61.647 | 17559 | 0 | 1 | WOOD4-E-5 | 23-074 | H2O | 1682539820 | 0.325 | 0.72 | 95.936 | 0 | 0.002 | 2.268 | 1294.717 | 1.074 | 2.746 | 107.293 | 0.871 | -0.045 | 0 | 0 | 1 | 39.5 | 1 | 46 | 1 | GJB_Rinse | 0
209 | P-3905 | 2023/04/26 20:15:05 | 2-47 | 1 | -2.223 | -17.985 | 17581 | -1 | 1 | A00000277285 | 23-096 | H2O | 1682540407 | 0.315 | 0.748 | 97.233 | 0.001 | -0.002 | 2.298 | 1292.128 | 1.325 | 2.767 | 108.195 | 0.876 | 0.173 | 0 | 0 | 1 | 39.562 | 1 | 47 | 1 | GJB_Rinse | 0
210 | P-3905 | 2023/04/26 20:24:02 | 2-47 | 2 | -2.109 | -13.103 | 17640 | -1 | 1 | A00000277285 | 23-096 | H2O | 1682540943 | 0.322 | 0.73 | 97.542 | 0.001 | -0.001 | 2.3 | 1292.662 | 0.78 | 2.881 | 110.408 | 0.873 | 0.419 | 0 | 0 | 1 | 39.562 | 1 | 47 | 1 | GJB_Rinse | 0
211 | P-3905 | 2023/04/26 20:32:58 | 2-47 | 3 | -2.096 | -11.951 | 17433 | -1 | 1 | A00000277285 | 23-096 | H2O | 1682541480 | 0.291 | 0.795 | 94.993 | 0.002 | -0.002 | 2.249 | 1291.404 | 1.056 | 2.749 | 107.042 | 0.88 | -0.015 | 0 | 0 | 1 | 39.625 | 1 | 47 | 1 | GJB_Rinse | 0
212 | P-3905 | 2023/04/26 20:41:55 | 2-47 | 4 | -2.093 | -11.377 | 17681 | 0 | 1 | A00000277285 | 23-096 | H2O | 1682542017 | 0.301 | 0.713 | 94.293 | 0.001 | -0.002 | 2.237 | 1293.255 | 1.354 | 2.742 | 106.887 | 0.873 | -0.213 | 0 | 0 | 1 | 39.625 | 1 | 47 | 1 | GJB_Rinse | 0
213 | P-3906 | 2023/04/26 20:51:41 | 2-48 | 1 | [5 empty] | EV | standard | [18 empty] | 1 | 48 | 1 | GJB_Rinse | 0
214 | P-3906 | 2023/04/26 21:00:38 | 2-48 | 2 | -7.329 | -57.801 | 17517 | -1 | 1 | EV | standard | H2O | 1682543139 | 0.302 | 0.843 | 99.025 | 0.001 | 0.009 | 2.336 | 1291.48 | 1.541 | 2.812 | 108.506 | 0.877 | 0.411 | 0 | 0 | 1 | 39.688 | 1 | 48 | 1 | GJB_Rinse | 0
215 | P-3906 | 2023/04/26 21:09:34 | 2-48 | 3 | -7.405 | -58.97 | 17525 | -1 | 1 | EV | standard | H2O | 1682543675 | 0.318 | 0.761 | 96.446 | 0.003 | 0.006 | 2.295 | 1291.777 | 2.031 | 2.73 | 106.003 | 0.881 | 0.087 | 0 | 0 | 1 | 39.625 | 1 | 48 | 1 | GJB_Rinse | 0
216 | P-3906 | 2023/04/26 21:18:30 | 2-48 | 4 | -7.467 | -59.368 | 17439 | 0 | 1 | EV | standard | H2O | 1682544212 | 0.311 | 0.926 | 100.655 | 0.001 | 0.01 | 2.377 | 1290.968 | 1.612 | 2.716 | 106.73 | 0.887 | 0.11 | 0 | 0 | 1 | 39.438 | 1 | 48 | 1 | GJB_Rinse | 0

## Instructions
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Variable | Instruction | Comment
StandardsFile | standards file.csv
OutputBase | ChemCorrect
CorrectionType | M
timecode | Time Code
identifier | Identifier 1
sample | Sample
injection | Inj Nr | Inj Nr
line | Line
job | Job
Variable | Instruction | Comment
baseline_shift_MEAN | mean_standards(baseline_shift) | The mean of the organic baseline of the standards.
baseline_shift_STD | std_standards(baseline_shift) | The standard deviation of the organic baseline of the standards.
ch4_ppm_MEAN | mean_standards(ch4_ppm) | The mean methane concentration of the standards.
ch4_ppm_STD | std_standards(ch4_ppm) | The standard deviation of the methane concentrations of the standards.
residuals_MEAN | mean_standards(residuals) | The mean difference between the measured and modeled spectra of the standards.
residuals_STD | std_standards(residuals) | The standard deviation of the difference between measured and modeled spectra of the standards.
slope_shift_MEAN | mean_standards(slope_shift) | The mean organic slope of the standards.
slope_shift_STD | std_standards(slope_shift) | The standard deviation of the organic slope of the standards.
baseline_curvature_MEAN | mean_standards(baseline_curvature) | The mean of the baseline curvature of the standards.
baseline_curvature_STD | std_standards(baseline_curvature) | The standard deviation of the baseline curvature of the standards.
residuals_MEAN_ABS | abs(residuals_MEAN) | The absolute value of the mean difference between the measured and modeled spectra of the standards.
baseline_shift_MEAN_ABS | abs(baseline_shift_MEAN) | The absolute value of the mean of the organic baseline of the standards.
ch4_ppm_MEAN_ABS | abs(ch4_ppm_MEAN) | The absolute value of the mean methane concentration of the standards.
slope_shift_MEAN_ABS | abs(slope_shift_MEAN) | The absolute value of the mean organic slope of the standards.
residuals_15SIGMA | evaluation(residuals_STD*1.5) | One and half times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_15SIGMA | evaluation(residuals_MEAN+residuals_15SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_2SIGMA | evaluation(residuals_STD*2) | Twice the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_2SIGMA | evaluation(residuals_MEAN+residuals_2SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_3SIGMA | evaluation(residuals_STD*3) | Three times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_3SIGMA | evaluation(residuals_MEAN+residuals_3SIGMA) | The mean difference between measured and modeled spectra of the standards plus three standard deviation.
residuals_10PCT | evaluation(residuals_MEAN*0.1) | 30% of the difference between measured and modeled spectra of the standards.
ch4_ppm_3SIGMA | evaluation(ch4_ppm_STD*3) | Three times the standard deviation of the methanol amplitudes of the standards.
ch4_ppm_PLUS_3SIGMA | evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA) | The mean methanol amplitude of the standards plus three standard deviation.
HALF | float(0.5) | A fixed value of 0.5 or half.
baseline_shift_HALFSIGMA | evaluation(baseline_shift_STD*HALF) | Half the standard deviation of the organic baseline of the standards.
baseline_shift_PLUS_HALFSIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA) | The absolute value of the mean organic baseline of the standards plus one-half standard deviation.
baseline_shift_PLUS_1SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD) | The absolute value of the mean organic baseline of the standards plus one standard deviation.
baseline_shift_3SIGMA | evaluation(baseline_shift_STD*3) | Three times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_3SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA) | The absolute value of the mean organic baseline of the standards plus three standard deviation.
baseline_shift_18SIGMA | evaluation(baseline_shift_STD*18) | Eighteen times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_18SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA) | The absolute value of the mean organic baseline of the standards plus eighteen standard deviation.
slope_shift_5SIGMA | evaluation(slope_shift_STD*5) | Five times the standard deviation of the organic slope of the standards.
slope_shift_PLUS_5SIGMA | evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA) | The absolute value of the mean organic slope of the standards plus five standard deviation.
baseline_curvature_3SIGMA | evaluation(baseline_curvature_STD*3) | Three times the standard deviation of the baseline curvature of the standards.
baseline_curvature_PLUS_3SIGMA | evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
TRUE | evaluation(1>0)
FALSE | evaluation(1<0)
PROC_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the uncalibrated d18O value in order to calibrate it.
PROC_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the uncalibrated dD value in order to calibrate it.
Variable | Instruction | Comment
IS_STANDARD | is_standard_sample() | Standard recognition.
IS_NOT_STANDARD | bool_xor(IS_STANDARD,TRUE)
SAMP_residuals_MEAN | mean_all(residuals) | The mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_residuals_STD | std_all(residuals) | One standard deviation of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN | mean_all(ch4_ppm) | The mean methane concentration of a sample.
SAMP_ch4_ppm_STD | std_all(ch4_ppm) | The standard deviation of the methane concentrations of a sample.
SAMP_baseline_shift_MEAN | mean_all(baseline_shift) | The mean of the organic baseline of a sample.
SAMP_baseline_shift_STD | std_all(baseline_shift) | The standard deviation of the organic baseline of a sample.
SAMP_slope_shift_MEAN | mean_all(slope_shift) | The mean organic slope of a sample.
SAMP_slope_shift_STD | std_all(slope_shift) | The standard deviation of the organic slope of a sample.
SAMP_d18O_STD | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
SAMP_dD_STD | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_baseline_curvature_MEAN | mean_all(baseline_curvature) | The mean of the baseline curvature of the ssamples
SAMP_baseline_curvature_STD | std_all(baseline_curvature) | The standard deviation of the baseline curvature of the samples
SAMP_residuals_MEAN_ABS | abs(SAMP_residuals_MEAN) | Absolute value of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN_ABS | abs(SAMP_ch4_ppm_MEAN) | Absolute value of the mean methane concentration of a sample.
SAMP_baseline_shift_MEAN_ABS | abs(SAMP_baseline_shift_MEAN) | Absolute value of the mean organic baseline of a sample.
SAMP_slope_shift_MEAN_ABS | abs(SAMP_slope_shift_MEAN) | Absolute value of the mean organic slope of a sample.
SAMP_residuals_DIFFERENTIAL | evaluation(SAMP_residuals_MEAN-residuals_MEAN)
SAMP_residuals_DIFFERENTIAL_ABS | abs(SAMP_residuals_DIFFERENTIAL) | Absolute value of the difference between the mean of the sample residual and the mean of all standards residuals
SAMP_residuals_DIFFERENTIAL_PCT | evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
SAMP_baseline_curvature_DIFFERENTIAL | evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
SAMP_STD_DELTA_18_16O_CURR_MEAN | current_standards_mean(d(18_16)Mean) | The d18O mean of the last run standard.
SAMP_STD_DELTA_DH_CURR_MEAN | current_standards_mean(d(D_H)Mean) | The dD mean of the last run standard.
known_dD | known_standard(d(D_H)Mean) | The dD value of the known Standard.
known_d18O | known_standard(d(18_16)Mean) | The d18O value of the known Standard.
d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
measured_d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
measured_dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
measured_d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
measured_dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the mean uncalibrated d18O value in order to calibrate it.
SAMP_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the mean uncalibrated dD value in order to calibrate it.
SAMP_TEST_VALUE | evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
residuals_GT_3SIGMA | evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA) | Threshold of red trigger. If the mean difference between the measured spectra of the sample and the modeled spectra of the standard is greater than that of the mean difference of the standard plus three standard deviations, this will trigger.
residuals_GE_2SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GE_15SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GT_10PCT | evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
ch4_ppm_GE_3SIGMA | evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA) | True if the mean methane concentration of the sample is greater than or equal to that of the standard plus three standard deviations, otherwise false.
baseline_shift_GT_1SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus one standard deviation, otherwise false.
baseline_shift_GT_3SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus three standard deviation, otherwise false. Sets yellow trigger for "other", trigger threshold is approximately sample to sample dD precision specification
baseline_shift_GT_18SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus eighteen standard deviation, otherwise false. Sets red trigger for "other", trigger threshold is approximately 24 hour dD drift specification
slope_shift_GT_5SIGMA | evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA) | True if the mean organic slope of the sample is greater than that of the standard plus five standard deviations, otherwise false.
baseline_curvature_GT_3SIGMA | evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
RED | truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA)) | Red trigger. If the sample is not a standard and the threshold is triggered (measured-modeled > standard+3stdev's), then RED.
NOT_RED | bool_xor(RED,TRUE) | If the sample is determined to be "red" than it can not be "not red."
YELLOW | truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA))) | Truth evaluation used to determine whether or not a given sample is yellow.
NOT_YELLOW | bool_xor(YELLOW,TRUE) | If sample is determined to be "yellow" than it can not be "not yellow."
GREEN | truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW)) | This is the pre-requisite for a sample to be colored green. It must not be a standard, must not be yellow, and must not be red.
NOT_GREEN | bool_xor(GREEN,TRUE) | Pre-requisite for being "not-green."
CYAN | truth_evaluation((IS_STANDARD)) | This is the pre-requisite for a sample to be a standard and therefore cyan colored.
NOT_CYAN | bool_xor(CYAN,TRUE) | Pre-requisite for being not a standard.
METHANE | truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA)) | This is the threshold for amounts of methane being included in the main ChemCorrect spreadsheet. Methane will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured methane concentration must be greater than or equal to that of the standard plus three standard deviations.
OTHER1 | truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA)) | This is the threshold for amounts of "OTHER1" being included in the main ChemCorrect spreadsheet. This column (OTHER1) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic baseline must be greater than that of the standard plus one standard deviation.
OTHER2 | truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA)) | This is the threshold for amounts of "OTHER2" being included in the main ChemCorrect spreadsheet. This column (OTHER2) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic slope must be greater than that of the standard plus three standard deviations.
OTHER3 | truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | truth_evaluation(RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN)) | Organic baseline differential showing the difference between the measured baseline and the standard baseline.
Variable | Instruction | Comment
DTL_TEST | math_evaluation((1 + 2)) | Test evaluation of 1+2 being equal to 3.
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample | FALSE | FALSE | FALSE | FALSE | Testing Sample
B | identifier | TRUE | Name | RED | YELLOW | GREEN | CYAN | Testing Name
C | SAMP_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d<sup>18</sup>O Mean | RED | YELLOW | GREEN | CYAN | Testing18 16 Mean
D | SAMP_DELTA_DH_CALIBRATE | TRUE | Calibrated d<sup>2</sup>H Mean | RED | YELLOW | GREEN | CYAN | This is a test to see that I can use tool tips
E | SAMP_ch4_ppm_MEAN | METHANE | CH<sub>4</sub> | RED | YELLOW | FALSE | FALSE
F | baseline_shift_DIFF | OTHER1 | C<sub>2+</sub> alcohols | RED | YELLOW | FALSE | FALSE
G | SAMP_residuals_DIFFERENTIAL_PCT | OTHER3 | relative %deviation | RED | YELLOW | FALSE | FALSE
H | SAMP_d18O_STD | TRUE | Uncalibrated d<sup>18</sup>O precision | RED | YELLOW | GREEN | CYAN
I | SAMP_dD_STD | TRUE | Uncalibrated d<sup>2</sup>H precision | RED | YELLOW | GREEN | CYAN
J | SAMP_slope_shift_MEAN | OTHER2 | slope | RED | YELLOW | FALSE | FALSE
K | SAMP_baseline_curvature_DIFFERENTIAL | OTHER4 | curvature | RED | YELLOW | FALSE | FALSE
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample
B | identifier | TRUE | Name
C | col(d(18_16)Mean) | TRUE | Measured d(18_16) Mean
D | col(d(D_H)Mean) | TRUE | Measured d(D_H) Mean
E | DTL_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d(18_16) Mean
F | DTL_DELTA_DH_CALIBRATE | TRUE | Calibrated d(D_H) Mean
G | col(ch4_ppm) | METHANE | CH<sub>4</sub>
H | col(baseline_shift) | OTHER1 | Other
I | col(Time Code) | TRUE | Time Code
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | d18O_mean | TRUE | d(18_16) Mean | RED | YELLOW | GREEN | CYAN
B | dD_mean | TRUE | d(D_H) Mean | RED | YELLOW | GREEN | CYAN
C | XY_X(d18O_mean) XY_Y(dD_mean) | TRUE | d18O x dD | gmwl | best_fit

## Process Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Process Variables
Variable | Value | Operations
baseline_shift_MEAN | 1301.8369
varray: [1301.698, 1302.669, 1304.042, 1303.083, 1302.412, 1303.232, 1302.083, 1302.591, 1302.382, 1302.557, 1299.562, 1301.321, 1301.135, 1301.284, 1301.216, 1301.536, 1301.078, 1301.359, 1301.224, 1300.274]
inst: mean_standards
operation: mean_standards(baseline_shift)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0, 10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: baseline_shift
baseline_shift_STD | 1.05599122305
varray: [1301.698, 1302.669, 1304.042, 1303.083, 1302.412, 1303.232, 1302.083, 1302.591, 1302.382, 1302.557, 1299.562, 1301.321, 1301.135, 1301.284, 1301.216, 1301.536, 1301.078, 1301.359, 1301.224, 1300.274]
inst: std_standards
operation: std_standards(baseline_shift)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0, 10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: baseline_shift
ch4_ppm_MEAN | 0.43995
varray: [0.36, 0.182, 0.258, 0.352, 0.808, 0.314, 0.628, 0.379, 0.495, 0.516, 0.675, 0.414, 0.522, 0.345, 0.508, 0.474, 0.407, 0.425, 0.333, 0.404]
inst: mean_standards
operation: mean_standards(ch4_ppm)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0, 10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: ch4_ppm
ch4_ppm_STD | 0.14590677449
varray: [0.36, 0.182, 0.258, 0.352, 0.808, 0.314, 0.628, 0.379, 0.495, 0.516, 0.675, 0.414, 0.522, 0.345, 0.508, 0.474, 0.407, 0.425, 0.333, 0.404]
inst: std_standards
operation: std_standards(ch4_ppm)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0, 10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: ch4_ppm
residuals_MEAN | 3.04645
varray: [3.072, 3.036, 3.097, 3.079, 3.15, 3.067, 3.079, 3.064, 3.116, 3.131, 3.042, 2.983, 3.005, 2.982, 3.045, 3.003, 2.993, 3.001, 2.976, 3.008]
inst: mean_standards
operation: mean_standards(residuals)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0, 10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: residuals
residuals_STD | 0.05233844818
varray: [3.072, 3.036, 3.097, 3.079, 3.15, 3.067, 3.079, 3.064, 3.116, 3.131, 3.042, 2.983, 3.005, 2.982, 3.045, 3.003, 2.993, 3.001, 2.976, 3.008]
inst: std_standards
operation: std_standards(residuals)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0, 10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: residuals
slope_shift_MEAN | 1.52925
varray: [1.05, 0.88, 1.132, 1.026, 1.362, 1.102, 1.135, 0.739, 1.003, 0.918, 2.287, 2.107, 2.897, 1.963, 1.934, 1.556, 1.569, 2.156, 1.803, 1.966]
inst: mean_standards
operation: mean_standards(slope_shift)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0, 10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: slope_shift
slope_shift_STD | 0.584662104395
varray: [1.05, 0.88, 1.132, 1.026, 1.362, 1.102, 1.135, 0.739, 1.003, 0.918, 2.287, 2.107, 2.897, 1.963, 1.934, 1.556, 1.569, 2.156, 1.803, 1.966]
inst: std_standards
operation: std_standards(slope_shift)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0, 10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: slope_shift
baseline_curvature_MEAN | 117.0061
varray: [117.406, 116.372, 117.625, 117.268, 118.966, 117.206, 117.939, 117.682, 118.904, 118.658, 115.334, 116.053, 116.106, 115.447, 116.843, 117.281, 116.468, 116.446, 116.35, 115.768]
inst: mean_standards
operation: mean_standards(baseline_curvature)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0, 10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: baseline_curvature
baseline_curvature_STD | 1.0801273268
varray: [117.406, 116.372, 117.625, 117.268, 118.966, 117.206, 117.939, 117.682, 118.904, 118.658, 115.334, 116.053, 116.106, 115.447, 116.843, 117.281, 116.468, 116.446, 116.35, 115.768]
inst: std_standards
operation: std_standards(baseline_curvature)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0, 10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: baseline_curvature
residuals_MEAN_ABS | 3.04645
operation: abs(residuals_MEAN)
inst: abs
source_col: residuals_MEAN
baseline_shift_MEAN_ABS | 1301.8369
operation: abs(baseline_shift_MEAN)
inst: abs
source_col: baseline_shift_MEAN
ch4_ppm_MEAN_ABS | 0.43995
operation: abs(ch4_ppm_MEAN)
inst: abs
source_col: ch4_ppm_MEAN
slope_shift_MEAN_ABS | 1.52925
operation: abs(slope_shift_MEAN)
inst: abs
source_col: slope_shift_MEAN
residuals_15SIGMA | 0.0785076722701
operation: evaluation(residuals_STD*1.5)
inst: evaluation
factors: residuals_STD*1.5
residuals_PLUS_15SIGMA | 3.12495767227
operation: evaluation(residuals_MEAN+residuals_15SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_15SIGMA
residuals_2SIGMA | 0.10467689636
operation: evaluation(residuals_STD*2)
inst: evaluation
factors: residuals_STD*2
residuals_PLUS_2SIGMA | 3.15112689636
operation: evaluation(residuals_MEAN+residuals_2SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_2SIGMA
residuals_3SIGMA | 0.15701534454
operation: evaluation(residuals_STD*3)
inst: evaluation
factors: residuals_STD*3
residuals_PLUS_3SIGMA | 3.20346534454
operation: evaluation(residuals_MEAN+residuals_3SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_3SIGMA
residuals_10PCT | 0.304645
operation: evaluation(residuals_MEAN*0.1)
inst: evaluation
factors: residuals_MEAN*0.1
ch4_ppm_3SIGMA | 0.43772032347
operation: evaluation(ch4_ppm_STD*3)
inst: evaluation
factors: ch4_ppm_STD*3
ch4_ppm_PLUS_3SIGMA | 0.87767032347
operation: evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA)
inst: evaluation
factors: ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA
HALF | 0.5
operation: float(0.5)
inst: float
baseline_shift_HALFSIGMA | 0.527995611525
operation: evaluation(baseline_shift_STD*HALF)
inst: evaluation
factors: baseline_shift_STD*HALF
baseline_shift_PLUS_HALFSIGMA | 1302.36489561
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA
baseline_shift_PLUS_1SIGMA | 1302.89289122
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_STD
baseline_shift_3SIGMA | 3.16797366915
operation: evaluation(baseline_shift_STD*3)
inst: evaluation
factors: baseline_shift_STD*3
baseline_shift_PLUS_3SIGMA | 1305.00487367
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_3SIGMA
baseline_shift_18SIGMA | 19.0078420149
operation: evaluation(baseline_shift_STD*18)
inst: evaluation
factors: baseline_shift_STD*18
baseline_shift_PLUS_18SIGMA | 1320.84474201
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_18SIGMA
slope_shift_5SIGMA | 2.92331052198
operation: evaluation(slope_shift_STD*5)
inst: evaluation
factors: slope_shift_STD*5
slope_shift_PLUS_5SIGMA | 4.45256052198
operation: evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA)
inst: evaluation
factors: slope_shift_MEAN_ABS+slope_shift_5SIGMA
baseline_curvature_3SIGMA | 3.24038198041
operation: evaluation(baseline_curvature_STD*3)
inst: evaluation
factors: baseline_curvature_STD*3
baseline_curvature_PLUS_3SIGMA | 120.24648198
operation: evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
inst: evaluation
factors: baseline_curvature_3SIGMA+baseline_curvature_MEAN
TRUE | True
operation: evaluation(1>0)
inst: evaluation
factors: 1>0
FALSE | False
operation: evaluation(1<0)
inst: evaluation
factors: 1<0
PROC_DELTA_18_16O_CALIBRATE | -6.86205237037
actual_array: [ 1.93 -15.84]
varray: [0.0, -7.973, -7.588, -7.563, -7.361, -7.255, -7.238, -7.171, -7.121, -7.183, 3.937, 4.541, 4.7, 4.777, 4.769, 4.784, 4.832, 4.852, 4.918, 4.867, -11.957, -12.806, -13.051, -13.053, -13.144, -13.151, -13.224, -13.199, -13.175, -13.286, -7.945, -7.578, -7.585, -7.595, -7.565, -7.536, -7.573, -7.532, -7.577, -7.591, -8.926, -8.969, -9.025, -9.061, -7.238, -7.184, -7.119, -7.112, 15.521, 15.746, 15.854, 15.844, 8.079, 7.93, 7.883, 7.887, -5.765, -5.945, -5.966, -5.946, 12.915, 13.11, 13.157, 13.184, -8.186, -8.545, -8.548, -8.64, -12.128, -12.334, -12.363, -12.353, 0.0, -7.685, -7.651, -7.633, 0.0, -11.595, -11.609, -11.616, -11.48, -11.438, -11.447, -11.468, 11.604, 12.056, 12.181, 12.263, -7.862, -8.052, -8.166, -8.197, -9.635, -9.823, -9.686, -9.746, -9.309, -9.386, -9.307, -9.365, -8.392, -8.377, -8.335, -8.368, 0.0, -11.857, -11.832, -11.895, 0.0, -7.699, -7.653, -7.652, -1.648, -1.723, -1.713, -1.73, -11.005, -11.052, -11.085, -11.061, -5.162, -4.801, -4.758, -4.713, -4.102, -4.13, -4.17, -4.07, -6.706, -6.923, -6.923, -6.823, -5.339, -5.363, -5.345, -5.405, 0.0, -7.661, -7.673, -7.685, 0.0, -2.381, -2.302, -2.261, -7.224, -7.506, -7.56, -7.634, -3.454, -3.214, -3.165, -3.144, -5.071, -5.146, -5.138, -5.353, -4.612, -4.56, -4.625, -4.536, -6.029, -6.141, -6.138, -6.078, -4.628, -4.488, -4.513, -4.461, 17.411, 0.0, 17.541, 17.74, -6.406, -6.714, -6.801, -6.854, -5.978, -5.787, -5.795, -5.773, 0.0, -7.469, -7.547, -7.538, -4.887, -4.71, -4.727, -4.673, 5.169, 5.308, 5.382, 5.392, -9.22, -9.492, -9.52, -9.636, 13.562, 14.035, 13.977, 14.147, -9.084, -9.37, -9.498, -9.554, 0.0, -4.316, -4.317, -4.336, -2.223, -2.109, -2.096, -2.093, 0.0, -7.329, -7.405, -7.467]
fit_rsq: 1.0
fit_intercept: 2.77505289814
except: True
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -4.06085648148
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [1.0, 603.0, 1139.0, 1676.0, 2213.0, 2750.0, 3286.0, 3824.0, 4361.0, 4898.0, 5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0, 10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0, 16334.0, 16871.0, 17409.0, 17946.0, 18482.0, 19019.0, 19556.0, 20093.0, 20630.0, 21168.0, 21758.0, 22295.0, 22832.0, 23370.0, 23960.0, 24496.0, 25034.0, 25571.0, 26162.0, 26699.0, 27236.0, 27774.0, 28360.0, 28897.0, 29434.0, 29971.0, 30560.0, 31097.0, 31634.0, 32172.0, 32760.0, 33297.0, 33835.0, 34372.0, 34961.0, 35498.0, 36034.0, 36572.0, 37162.0, 37699.0, 38239.0, 38776.0, 39366.0, 39903.0, 40440.0, 40977.0, 41567.0, 42104.0, 42641.0, 43178.0, 43766.0, 44302.0, 44839.0, 45377.0, 45965.0, 46502.0, 47039.0, 47575.0, 48163.0, 48701.0, 49238.0, 49775.0, 50363.0, 50900.0, 51437.0, 51974.0, 52562.0, 53099.0, 53635.0, 54172.0, 54761.0, 55297.0, 55835.0, 56372.0, 56962.0, 57499.0, 58036.0, 58573.0, 59160.0, 59696.0, 60233.0, 60770.0, 61356.0, 61894.0, 62431.0, 62967.0, 63555.0, 64092.0, 64628.0, 65166.0, 65753.0, 66290.0, 66828.0, 67365.0, 67953.0, 68491.0, 69028.0, 69564.0, 70153.0, 70690.0, 71227.0, 71763.0, 72352.0, 72888.0, 73424.0, 73962.0, 74547.0, 75084.0, 75621.0, 76157.0, 76743.0, 77279.0, 77816.0, 78352.0, 78938.0, 79475.0, 80012.0, 80549.0, 81135.0, 81671.0, 82208.0, 82745.0, 83331.0, 83868.0, 84404.0, 84941.0, 85528.0, 86065.0, 86601.0, 87138.0, 87725.0, 88262.0, 88798.0, 89335.0, 89918.0, 90455.0, 90991.0, 91527.0, 92111.0, 92648.0, 93184.0, 93720.0, 94305.0, 94841.0, 95378.0, 95914.0, 96498.0, 97035.0, 97571.0, 98108.0, 98693.0, 99230.0, 99766.0, 100302.0, 100887.0, 101423.0, 101960.0, 102497.0, 103082.0, 103619.0, 104155.0, 104692.0, 105277.0, 105813.0, 106350.0, 106886.0, 107471.0, 108007.0, 108543.0, 109080.0, 109665.0, 110201.0, 110737.0, 111274.0, 111860.0, 112396.0, 112932.0, 113469.0, 114055.0, 114592.0, 115128.0, 115665.0, 116251.0, 116788.0, 117324.0, 117860.0]
source_col: d(18_16)Mean
PROC_DELTA_DH_CALIBRATE | -76.5802754475
actual_array: [ 18.1 -119. ]
varray: [0.0, -66.916, -62.272, -62.339, -60.824, -59.622, -59.615, -58.914, -58.741, -58.693, 27.51, 36.325, 38.866, 39.653, 40.114, 40.396, 40.695, 41.027, 41.052, 41.241, -90.283, -104.086, -107.145, -108.517, -109.271, -109.877, -110.253, -110.5, -110.672, -110.843, -66.123, -61.595, -60.611, -60.407, -60.11, -59.737, -59.848, -59.885, -60.05, -59.857, -81.38, -83.399, -84.178, -84.545, -78.453, -78.214, -77.927, -77.96, -28.048, -26.456, -25.865, -25.801, -41.103, -41.734, -42.103, -41.749, -71.707, -72.919, -73.162, -73.177, -36.736, -35.73, -35.714, -35.69, -88.77, -91.492, -92.185, -92.49, -106.41, -108.026, -108.321, -108.636, 0.0, -62.481, -61.489, -61.183, 0.0, -101.044, -101.844, -102.341, -102.117, -102.127, -102.088, -102.236, -41.673, -37.586, -36.643, -36.481, -78.137, -79.414, -79.736, -79.898, -94.839, -96.679, -96.893, -96.871, -88.153, -87.665, -87.237, -87.235, -78.632, -77.981, -77.563, -77.398, 0.0, -103.508, -104.014, -104.621, 0.0, -62.364, -61.144, -61.065, -82.813, -86.743, -87.423, -87.735, -101.333, -101.226, -101.171, -101.222, -72.855, -69.605, -68.799, -68.49, -69.415, -69.476, -69.729, -69.3, -72.508, -72.783, -72.778, -72.146, -62.766, -61.666, -61.563, -61.548, 0.0, -84.119, -84.7, -84.912, 0.0, -57.412, -56.387, -56.032, -60.002, -60.366, -60.396, -60.721, -62.305, -62.332, -62.521, -62.622, -70.368, -71.151, -71.168, -72.507, -63.442, -62.422, -62.346, -62.239, -71.772, -73.01, -73.11, -73.428, -66.473, -65.608, -65.348, -65.042, -21.328, 0.0, -20.739, -19.424, -68.159, -69.882, -70.132, -70.44, -72.968, -73.018, -73.041, -73.107, 0.0, -60.451, -60.68, -60.357, -67.081, -67.706, -68.188, -68.022, -54.857, -54.754, -54.002, -54.026, -95.403, -98.079, -98.749, -99.179, -40.006, -36.46, -35.076, -34.399, -94.109, -97.015, -98.05, -98.311, 0.0, -62.77, -62.267, -61.647, -17.985, -13.103, -11.951, -11.377, 0.0, -57.801, -58.97, -59.368]
fit_rsq: 1.0
fit_intercept: 19.4350184537
except: True
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -62.0230462963
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [1.0, 603.0, 1139.0, 1676.0, 2213.0, 2750.0, 3286.0, 3824.0, 4361.0, 4898.0, 5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0, 10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0, 16334.0, 16871.0, 17409.0, 17946.0, 18482.0, 19019.0, 19556.0, 20093.0, 20630.0, 21168.0, 21758.0, 22295.0, 22832.0, 23370.0, 23960.0, 24496.0, 25034.0, 25571.0, 26162.0, 26699.0, 27236.0, 27774.0, 28360.0, 28897.0, 29434.0, 29971.0, 30560.0, 31097.0, 31634.0, 32172.0, 32760.0, 33297.0, 33835.0, 34372.0, 34961.0, 35498.0, 36034.0, 36572.0, 37162.0, 37699.0, 38239.0, 38776.0, 39366.0, 39903.0, 40440.0, 40977.0, 41567.0, 42104.0, 42641.0, 43178.0, 43766.0, 44302.0, 44839.0, 45377.0, 45965.0, 46502.0, 47039.0, 47575.0, 48163.0, 48701.0, 49238.0, 49775.0, 50363.0, 50900.0, 51437.0, 51974.0, 52562.0, 53099.0, 53635.0, 54172.0, 54761.0, 55297.0, 55835.0, 56372.0, 56962.0, 57499.0, 58036.0, 58573.0, 59160.0, 59696.0, 60233.0, 60770.0, 61356.0, 61894.0, 62431.0, 62967.0, 63555.0, 64092.0, 64628.0, 65166.0, 65753.0, 66290.0, 66828.0, 67365.0, 67953.0, 68491.0, 69028.0, 69564.0, 70153.0, 70690.0, 71227.0, 71763.0, 72352.0, 72888.0, 73424.0, 73962.0, 74547.0, 75084.0, 75621.0, 76157.0, 76743.0, 77279.0, 77816.0, 78352.0, 78938.0, 79475.0, 80012.0, 80549.0, 81135.0, 81671.0, 82208.0, 82745.0, 83331.0, 83868.0, 84404.0, 84941.0, 85528.0, 86065.0, 86601.0, 87138.0, 87725.0, 88262.0, 88798.0, 89335.0, 89918.0, 90455.0, 90991.0, 91527.0, 92111.0, 92648.0, 93184.0, 93720.0, 94305.0, 94841.0, 95378.0, 95914.0, 96498.0, 97035.0, 97571.0, 98108.0, 98693.0, 99230.0, 99766.0, 100302.0, 100887.0, 101423.0, 101960.0, 102497.0, 103082.0, 103619.0, 104155.0, 104692.0, 105277.0, 105813.0, 106350.0, 106886.0, 107471.0, 108007.0, 108543.0, 109080.0, 109665.0, 110201.0, 110737.0, 111274.0, 111860.0, 112396.0, 112932.0, 113469.0, 114055.0, 114592.0, 115128.0, 115665.0, 116251.0, 116788.0, 117324.0, 117860.0]
source_col: d(D_H)Mean

## 1 - Sample 1 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 1
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.8232
varray: [0.0, 3.159, 3.232, 3.495, 3.048, 3.18, 2.991, 3.051, 3.029, 3.047]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [1.0, 603.0, 1139.0, 1676.0, 2213.0, 2750.0, 3286.0, 3824.0, 4361.0, 4898.0]
source_col: residuals
SAMP_residuals_STD | 1.00288381957
varray: [0.0, 3.159, 3.232, 3.495, 3.048, 3.18, 2.991, 3.051, 3.029, 3.047]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [1.0, 603.0, 1139.0, 1676.0, 2213.0, 2750.0, 3286.0, 3824.0, 4361.0, 4898.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4019
varray: [0.0, 0.644, 0.348, 0.413, 0.302, 0.446, 0.512, 0.324, 0.487, 0.543]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [1.0, 603.0, 1139.0, 1676.0, 2213.0, 2750.0, 3286.0, 3824.0, 4361.0, 4898.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.176519089304
varray: [0.0, 0.644, 0.348, 0.413, 0.302, 0.446, 0.512, 0.324, 0.487, 0.543]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [1.0, 603.0, 1139.0, 1676.0, 2213.0, 2750.0, 3286.0, 3824.0, 4361.0, 4898.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1193.0987
varray: [0.0, 1328.557, 1346.381, 1379.411, 1324.937, 1325.198, 1310.335, 1307.313, 1305.189, 1303.666]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [1.0, 603.0, 1139.0, 1676.0, 2213.0, 2750.0, 3286.0, 3824.0, 4361.0, 4898.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 419.848038324
varray: [0.0, 1328.557, 1346.381, 1379.411, 1324.937, 1325.198, 1310.335, 1307.313, 1305.189, 1303.666]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [1.0, 603.0, 1139.0, 1676.0, 2213.0, 2750.0, 3286.0, 3824.0, 4361.0, 4898.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.9034
varray: [0.0, 0.608, -0.261, -0.716, 1.217, 0.995, 1.986, 1.63, 1.818, 1.757]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [1.0, 603.0, 1139.0, 1676.0, 2213.0, 2750.0, 3286.0, 3824.0, 4361.0, 4898.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.957630664134
varray: [0.0, 0.608, -0.261, -0.716, 1.217, 0.995, 1.986, 1.63, 1.818, 1.757]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [1.0, 603.0, 1139.0, 1676.0, 2213.0, 2750.0, 3286.0, 3824.0, 4361.0, 4898.0]
source_col: slope_shift
SAMP_d18O_STD | 2.34948415379
varray: [0.0, -7.973, -7.588, -7.563, -7.361, -7.255, -7.238, -7.171, -7.121, -7.183]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 603.0, 1139.0, 1676.0, 2213.0, 2750.0, 3286.0, 3824.0, 4361.0, 4898.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 19.416032214
varray: [0.0, -66.916, -62.272, -62.339, -60.824, -59.622, -59.615, -58.914, -58.741, -58.693]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 603.0, 1139.0, 1676.0, 2213.0, 2750.0, 3286.0, 3824.0, 4361.0, 4898.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.15
varray: [0.0, 121.892, 124.447, 131.309, 117.086, 121.274, 115.269, 116.445, 116.74, 117.038]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [1.0, 603.0, 1139.0, 1676.0, 2213.0, 2750.0, 3286.0, 3824.0, 4361.0, 4898.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 38.3138499124
varray: [0.0, 121.892, 124.447, 131.309, 117.086, 121.274, 115.269, 116.445, 116.74, 117.038]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [1.0, 603.0, 1139.0, 1676.0, 2213.0, 2750.0, 3286.0, 3824.0, 4361.0, 4898.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.8232
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4019
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1193.0987
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.9034
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.22325
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.22325
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -7.32820167736
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -8.8561
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 0
varray: [0.0, -7.973, -7.588, -7.563, -7.361, -7.255, -7.238, -7.171, -7.121, -7.183]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [1.0, 603.0, 1139.0, 1676.0, 2213.0, 2750.0, 3286.0, 3824.0, 4361.0, 4898.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 0
varray: [0.0, -66.916, -62.272, -62.339, -60.824, -59.622, -59.615, -58.914, -58.741, -58.693]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [1.0, 603.0, 1139.0, 1676.0, 2213.0, 2750.0, 3286.0, 3824.0, 4361.0, 4898.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.6453
varray: [0.0, -7.973, -7.588, -7.563, -7.361, -7.255, -7.238, -7.171, -7.121, -7.183]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 603.0, 1139.0, 1676.0, 2213.0, 2750.0, 3286.0, 3824.0, 4361.0, 4898.0]
source_col: d(18_16)Mean
dD_mean | -54.7936
varray: [0.0, -66.916, -62.272, -62.339, -60.824, -59.622, -59.615, -58.914, -58.741, -58.693]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 603.0, 1139.0, 1676.0, 2213.0, 2750.0, 3286.0, 3824.0, 4361.0, 4898.0]
source_col: d(D_H)Mean
d18O_stdev | 2.34948415379
varray: [0.0, -7.973, -7.588, -7.563, -7.361, -7.255, -7.238, -7.171, -7.121, -7.183]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 603.0, 1139.0, 1676.0, 2213.0, 2750.0, 3286.0, 3824.0, 4361.0, 4898.0]
source_col: d(18_16)Mean
dD_stdev | 19.416032214
varray: [0.0, -66.916, -62.272, -62.339, -60.824, -59.622, -59.615, -58.914, -58.741, -58.693]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 603.0, 1139.0, 1676.0, 2213.0, 2750.0, 3286.0, 3824.0, 4361.0, 4898.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.6453
varray: [0.0, -7.973, -7.588, -7.563, -7.361, -7.255, -7.238, -7.171, -7.121, -7.183]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 603.0, 1139.0, 1676.0, 2213.0, 2750.0, 3286.0, 3824.0, 4361.0, 4898.0]
source_col: d(18_16)Mean
measured_dD_mean | -54.7936
varray: [0.0, -66.916, -62.272, -62.339, -60.824, -59.622, -59.615, -58.914, -58.741, -58.693]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 603.0, 1139.0, 1676.0, 2213.0, 2750.0, 3286.0, 3824.0, 4361.0, 4898.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 2.34948415379
varray: [0.0, -7.973, -7.588, -7.563, -7.361, -7.255, -7.238, -7.171, -7.121, -7.183]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 603.0, 1139.0, 1676.0, 2213.0, 2750.0, 3286.0, 3824.0, 4361.0, 4898.0]
source_col: d(18_16)Mean
measured_dD_stdev | 19.416032214
varray: [0.0, -66.916, -62.272, -62.339, -60.824, -59.622, -59.615, -58.914, -58.741, -58.693]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 603.0, 1139.0, 1676.0, 2213.0, 2750.0, 3286.0, 3824.0, 4361.0, 4898.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.45637973597
actual_array: [ 1.93 -15.84]
varray: [0.0, -7.973, -7.588, -7.563, -7.361, -7.255, -7.238, -7.171, -7.121, -7.183]
fit_rsq: 1.0
fit_intercept: 2.77505289814
except: True
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -6.6453
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [1.0, 603.0, 1139.0, 1676.0, 2213.0, 2750.0, 3286.0, 3824.0, 4361.0, 4898.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -69.7837355296
actual_array: [ 18.1 -119. ]
varray: [0.0, -66.916, -62.272, -62.339, -60.824, -59.622, -59.615, -58.914, -58.741, -58.693]
fit_rsq: 1.0
fit_intercept: 19.4350184537
except: True
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -54.7936
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [1.0, 603.0, 1139.0, 1676.0, 2213.0, 2750.0, 3286.0, 3824.0, 4361.0, 4898.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.0137
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -108.7382
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 2 - Sample 2 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 2
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.0891
varray: [3.072, 3.036, 3.097, 3.079, 3.15, 3.067, 3.079, 3.064, 3.116, 3.131]
inst: mean_all
operation: mean_all(residuals)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0]
source_col: residuals
SAMP_residuals_STD | 0.0344559299847
varray: [3.072, 3.036, 3.097, 3.079, 3.15, 3.067, 3.079, 3.064, 3.116, 3.131]
inst: std_all
operation: std_all(residuals)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4292
varray: [0.36, 0.182, 0.258, 0.352, 0.808, 0.314, 0.628, 0.379, 0.495, 0.516]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.186139613075
varray: [0.36, 0.182, 0.258, 0.352, 0.808, 0.314, 0.628, 0.379, 0.495, 0.516]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1302.6749
varray: [1301.698, 1302.669, 1304.042, 1303.083, 1302.412, 1303.232, 1302.083, 1302.591, 1302.382, 1302.557]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.652041997625
varray: [1301.698, 1302.669, 1304.042, 1303.083, 1302.412, 1303.232, 1302.083, 1302.591, 1302.382, 1302.557]
inst: std_all
operation: std_all(baseline_shift)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.0347
varray: [1.05, 0.88, 1.132, 1.026, 1.362, 1.102, 1.135, 0.739, 1.003, 0.918]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.169143627331
varray: [1.05, 0.88, 1.132, 1.026, 1.362, 1.102, 1.135, 0.739, 1.003, 0.918]
inst: std_all
operation: std_all(slope_shift)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0]
source_col: slope_shift
SAMP_d18O_STD | 0.286954138806
varray: [3.937, 4.541, 4.7, 4.777, 4.769, 4.784, 4.832, 4.852, 4.918, 4.867]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 4.19060293859
varray: [27.51, 36.325, 38.866, 39.653, 40.114, 40.396, 40.695, 41.027, 41.052, 41.241]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 117.8026
varray: [117.406, 116.372, 117.625, 117.268, 118.966, 117.206, 117.939, 117.682, 118.904, 118.658]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.830569978321
varray: [117.406, 116.372, 117.625, 117.268, 118.966, 117.206, 117.939, 117.682, 118.904, 118.658]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.0891
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4292
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1302.6749
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.0347
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.04265
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.04265
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.39999015247
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.7965
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 4.6977
varray: [3.937, 4.541, 4.7, 4.777, 4.769, 4.784, 4.832, 4.852, 4.918, 4.867]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 38.6879
varray: [27.51, 36.325, 38.866, 39.653, 40.114, 40.396, 40.695, 41.027, 41.052, 41.241]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0]
source_col: d(D_H)Mean
known_dD | 18.1
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 1.93
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 4.6977
varray: [3.937, 4.541, 4.7, 4.777, 4.769, 4.784, 4.832, 4.852, 4.918, 4.867]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0]
source_col: d(18_16)Mean
dD_mean | 38.6879
varray: [27.51, 36.325, 38.866, 39.653, 40.114, 40.396, 40.695, 41.027, 41.052, 41.241]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0]
source_col: d(D_H)Mean
d18O_stdev | 0.286954138806
varray: [3.937, 4.541, 4.7, 4.777, 4.769, 4.784, 4.832, 4.852, 4.918, 4.867]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0]
source_col: d(18_16)Mean
dD_stdev | 4.19060293859
varray: [27.51, 36.325, 38.866, 39.653, 40.114, 40.396, 40.695, 41.027, 41.052, 41.241]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0]
source_col: d(D_H)Mean
measured_d18O_mean | 4.6977
varray: [3.937, 4.541, 4.7, 4.777, 4.769, 4.784, 4.832, 4.852, 4.918, 4.867]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0]
source_col: d(18_16)Mean
measured_dD_mean | 38.6879
varray: [27.51, 36.325, 38.866, 39.653, 40.114, 40.396, 40.695, 41.027, 41.052, 41.241]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.286954138806
varray: [3.937, 4.541, 4.7, 4.777, 4.769, 4.784, 4.832, 4.852, 4.918, 4.867]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0]
source_col: d(18_16)Mean
measured_dD_stdev | 4.19060293859
varray: [27.51, 36.325, 38.866, 39.653, 40.114, 40.396, 40.695, 41.027, 41.052, 41.241]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.93
actual_array: [ 1.93 -15.84]
varray: [3.937, 4.541, 4.7, 4.777, 4.769, 4.784, 4.832, 4.852, 4.918, 4.867]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 4.6977
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 18.1
actual_array: [ 18.1 -119. ]
varray: [27.51, 36.325, 38.866, 39.653, 40.114, 40.396, 40.695, 41.027, 41.052, 41.241]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: 38.6879
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [5486.0, 6023.0, 6560.0, 7097.0, 7634.0, 8172.0, 8709.0, 9246.0, 9783.0, 10321.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1298.7478
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.838
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 3 - Sample 3 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 3
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.0038
varray: [3.042, 2.983, 3.005, 2.982, 3.045, 3.003, 2.993, 3.001, 2.976, 3.008]
inst: mean_all
operation: mean_all(residuals)
tarray: [10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: residuals
SAMP_residuals_STD | 0.0235268168504
varray: [3.042, 2.983, 3.005, 2.982, 3.045, 3.003, 2.993, 3.001, 2.976, 3.008]
inst: std_all
operation: std_all(residuals)
tarray: [10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.4507
varray: [0.675, 0.414, 0.522, 0.345, 0.508, 0.474, 0.407, 0.425, 0.333, 0.404]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.100190984292
varray: [0.675, 0.414, 0.522, 0.345, 0.508, 0.474, 0.407, 0.425, 0.333, 0.404]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1300.9989
varray: [1299.562, 1301.321, 1301.135, 1301.284, 1301.216, 1301.536, 1301.078, 1301.359, 1301.224, 1300.274]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.606989190825
varray: [1299.562, 1301.321, 1301.135, 1301.284, 1301.216, 1301.536, 1301.078, 1301.359, 1301.224, 1300.274]
inst: std_all
operation: std_all(baseline_shift)
tarray: [10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 2.0238
varray: [2.287, 2.107, 2.897, 1.963, 1.934, 1.556, 1.569, 2.156, 1.803, 1.966]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.386678965322
varray: [2.287, 2.107, 2.897, 1.963, 1.934, 1.556, 1.569, 2.156, 1.803, 1.966]
inst: std_all
operation: std_all(slope_shift)
tarray: [10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: slope_shift
SAMP_d18O_STD | 0.391069530618
varray: [-11.957, -12.806, -13.051, -13.053, -13.144, -13.151, -13.224, -13.199, -13.175, -13.286]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 6.27677580096
varray: [-90.283, -104.086, -107.145, -108.517, -109.271, -109.877, -110.253, -110.5, -110.672, -110.843]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 116.2096
varray: [115.334, 116.053, 116.106, 115.447, 116.843, 117.281, 116.468, 116.446, 116.35, 115.768]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.60276791185
varray: [115.334, 116.053, 116.106, 115.447, 116.843, 117.281, 116.468, 116.446, 116.35, 115.768]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.0038
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4507
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1300.9989
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 2.0238
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.04265
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.04265
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.39999015247
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.7965
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [-11.957, -12.806, -13.051, -13.053, -13.144, -13.151, -13.224, -13.199, -13.175, -13.286]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-90.283, -104.086, -107.145, -108.517, -109.271, -109.877, -110.253, -110.5, -110.672, -110.843]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: d(D_H)Mean
known_dD | -119.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | -15.84
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -13.0046
varray: [-11.957, -12.806, -13.051, -13.053, -13.144, -13.151, -13.224, -13.199, -13.175, -13.286]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: d(18_16)Mean
dD_mean | -107.1447
varray: [-90.283, -104.086, -107.145, -108.517, -109.271, -109.877, -110.253, -110.5, -110.672, -110.843]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: d(D_H)Mean
d18O_stdev | 0.391069530618
varray: [-11.957, -12.806, -13.051, -13.053, -13.144, -13.151, -13.224, -13.199, -13.175, -13.286]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: d(18_16)Mean
dD_stdev | 6.27677580096
varray: [-90.283, -104.086, -107.145, -108.517, -109.271, -109.877, -110.253, -110.5, -110.672, -110.843]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: d(D_H)Mean
measured_d18O_mean | -13.0046
varray: [-11.957, -12.806, -13.051, -13.053, -13.144, -13.151, -13.224, -13.199, -13.175, -13.286]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: d(18_16)Mean
measured_dD_mean | -107.1447
varray: [-90.283, -104.086, -107.145, -108.517, -109.271, -109.877, -110.253, -110.5, -110.672, -110.843]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.391069530618
varray: [-11.957, -12.806, -13.051, -13.053, -13.144, -13.151, -13.224, -13.199, -13.175, -13.286]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: d(18_16)Mean
measured_dD_stdev | 6.27677580096
varray: [-90.283, -104.086, -107.145, -108.517, -109.271, -109.877, -110.253, -110.5, -110.672, -110.843]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -15.84
actual_array: [ 1.93 -15.84]
varray: [-11.957, -12.806, -13.051, -13.053, -13.144, -13.151, -13.224, -13.199, -13.175, -13.286]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -13.0046
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -119.0
actual_array: [ 18.1 -119. ]
varray: [-90.283, -104.086, -107.145, -108.517, -109.271, -109.877, -110.253, -110.5, -110.672, -110.843]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -107.1447
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [10910.0, 11447.0, 11984.0, 12522.0, 13058.0, 13596.0, 14134.0, 14671.0, 15208.0, 15745.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1298.8331
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.838
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 4 - Sample 4 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 4
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9421
varray: [2.995, 3.046, 2.932, 2.922, 2.903, 2.979, 2.928, 2.867, 2.876, 2.973]
inst: mean_all
operation: mean_all(residuals)
tarray: [16334.0, 16871.0, 17409.0, 17946.0, 18482.0, 19019.0, 19556.0, 20093.0, 20630.0, 21168.0]
source_col: residuals
SAMP_residuals_STD | 0.0559294198075
varray: [2.995, 3.046, 2.932, 2.922, 2.903, 2.979, 2.928, 2.867, 2.876, 2.973]
inst: std_all
operation: std_all(residuals)
tarray: [16334.0, 16871.0, 17409.0, 17946.0, 18482.0, 19019.0, 19556.0, 20093.0, 20630.0, 21168.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3064
varray: [0.23, 0.603, 0.274, 0.198, 0.038, 0.362, 0.298, 0.319, 0.258, 0.484]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [16334.0, 16871.0, 17409.0, 17946.0, 18482.0, 19019.0, 19556.0, 20093.0, 20630.0, 21168.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.15486639403
varray: [0.23, 0.603, 0.274, 0.198, 0.038, 0.362, 0.298, 0.319, 0.258, 0.484]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [16334.0, 16871.0, 17409.0, 17946.0, 18482.0, 19019.0, 19556.0, 20093.0, 20630.0, 21168.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1298.1313
varray: [1300.303, 1300.315, 1298.504, 1298.801, 1297.719, 1297.981, 1296.921, 1296.548, 1297.248, 1296.973]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [16334.0, 16871.0, 17409.0, 17946.0, 18482.0, 19019.0, 19556.0, 20093.0, 20630.0, 21168.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.34818207227
varray: [1300.303, 1300.315, 1298.504, 1298.801, 1297.719, 1297.981, 1296.921, 1296.548, 1297.248, 1296.973]
inst: std_all
operation: std_all(baseline_shift)
tarray: [16334.0, 16871.0, 17409.0, 17946.0, 18482.0, 19019.0, 19556.0, 20093.0, 20630.0, 21168.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.7587
varray: [2.123, 1.512, 1.557, 2.132, 1.82, 1.864, 1.634, 1.401, 1.927, 1.617]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [16334.0, 16871.0, 17409.0, 17946.0, 18482.0, 19019.0, 19556.0, 20093.0, 20630.0, 21168.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.253937897045
varray: [2.123, 1.512, 1.557, 2.132, 1.82, 1.864, 1.634, 1.401, 1.927, 1.617]
inst: std_all
operation: std_all(slope_shift)
tarray: [16334.0, 16871.0, 17409.0, 17946.0, 18482.0, 19019.0, 19556.0, 20093.0, 20630.0, 21168.0]
source_col: slope_shift
SAMP_d18O_STD | 0.120392921534
varray: [-7.945, -7.578, -7.585, -7.595, -7.565, -7.536, -7.573, -7.532, -7.577, -7.591]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16334.0, 16871.0, 17409.0, 17946.0, 18482.0, 19019.0, 19556.0, 20093.0, 20630.0, 21168.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.94230384681
varray: [-66.123, -61.595, -60.611, -60.407, -60.11, -59.737, -59.848, -59.885, -60.05, -59.857]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16334.0, 16871.0, 17409.0, 17946.0, 18482.0, 19019.0, 19556.0, 20093.0, 20630.0, 21168.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 114.1365
varray: [115.997, 117.612, 114.346, 113.473, 112.269, 114.202, 113.131, 112.919, 112.67, 114.746]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [16334.0, 16871.0, 17409.0, 17946.0, 18482.0, 19019.0, 19556.0, 20093.0, 20630.0, 21168.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.65099224374
varray: [115.997, 117.612, 114.346, 113.473, 112.269, 114.202, 113.131, 112.919, 112.67, 114.746]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [16334.0, 16871.0, 17409.0, 17946.0, 18482.0, 19019.0, 19556.0, 20093.0, 20630.0, 21168.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9421
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3064
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1298.1313
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.7587
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.10435
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.10435
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.42529829802
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.8696
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [-7.945, -7.578, -7.585, -7.595, -7.565, -7.536, -7.573, -7.532, -7.577, -7.591]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [16334.0, 16871.0, 17409.0, 17946.0, 18482.0, 19019.0, 19556.0, 20093.0, 20630.0, 21168.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-66.123, -61.595, -60.611, -60.407, -60.11, -59.737, -59.848, -59.885, -60.05, -59.857]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [16334.0, 16871.0, 17409.0, 17946.0, 18482.0, 19019.0, 19556.0, 20093.0, 20630.0, 21168.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.6077
varray: [-7.945, -7.578, -7.585, -7.595, -7.565, -7.536, -7.573, -7.532, -7.577, -7.591]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16334.0, 16871.0, 17409.0, 17946.0, 18482.0, 19019.0, 19556.0, 20093.0, 20630.0, 21168.0]
source_col: d(18_16)Mean
dD_mean | -60.8223
varray: [-66.123, -61.595, -60.611, -60.407, -60.11, -59.737, -59.848, -59.885, -60.05, -59.857]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16334.0, 16871.0, 17409.0, 17946.0, 18482.0, 19019.0, 19556.0, 20093.0, 20630.0, 21168.0]
source_col: d(D_H)Mean
d18O_stdev | 0.120392921534
varray: [-7.945, -7.578, -7.585, -7.595, -7.565, -7.536, -7.573, -7.532, -7.577, -7.591]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16334.0, 16871.0, 17409.0, 17946.0, 18482.0, 19019.0, 19556.0, 20093.0, 20630.0, 21168.0]
source_col: d(18_16)Mean
dD_stdev | 1.94230384681
varray: [-66.123, -61.595, -60.611, -60.407, -60.11, -59.737, -59.848, -59.885, -60.05, -59.857]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16334.0, 16871.0, 17409.0, 17946.0, 18482.0, 19019.0, 19556.0, 20093.0, 20630.0, 21168.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.6077
varray: [-7.945, -7.578, -7.585, -7.595, -7.565, -7.536, -7.573, -7.532, -7.577, -7.591]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16334.0, 16871.0, 17409.0, 17946.0, 18482.0, 19019.0, 19556.0, 20093.0, 20630.0, 21168.0]
source_col: d(18_16)Mean
measured_dD_mean | -60.8223
varray: [-66.123, -61.595, -60.611, -60.407, -60.11, -59.737, -59.848, -59.885, -60.05, -59.857]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16334.0, 16871.0, 17409.0, 17946.0, 18482.0, 19019.0, 19556.0, 20093.0, 20630.0, 21168.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.120392921534
varray: [-7.945, -7.578, -7.585, -7.595, -7.565, -7.536, -7.573, -7.532, -7.577, -7.591]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16334.0, 16871.0, 17409.0, 17946.0, 18482.0, 19019.0, 19556.0, 20093.0, 20630.0, 21168.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.94230384681
varray: [-66.123, -61.595, -60.611, -60.407, -60.11, -59.737, -59.848, -59.885, -60.05, -59.857]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16334.0, 16871.0, 17409.0, 17946.0, 18482.0, 19019.0, 19556.0, 20093.0, 20630.0, 21168.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.4224603018
actual_array: [ 1.93 -15.84]
varray: [-7.945, -7.578, -7.585, -7.595, -7.565, -7.536, -7.573, -7.532, -7.577, -7.591]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.6077
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [16334.0, 16871.0, 17409.0, 17946.0, 18482.0, 19019.0, 19556.0, 20093.0, 20630.0, 21168.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -75.4514310243
actual_array: [ 18.1 -119. ]
varray: [-66.123, -61.595, -60.611, -60.407, -60.11, -59.737, -59.848, -59.885, -60.05, -59.857]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -60.8223
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [16334.0, 16871.0, 17409.0, 17946.0, 18482.0, 19019.0, 19556.0, 20093.0, 20630.0, 21168.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1298.8948
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -3.7056
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 5 - Sample 5 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 5
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.94375
varray: [2.957, 2.951, 2.937, 2.93]
inst: mean_all
operation: mean_all(residuals)
tarray: [21758.0, 22295.0, 22832.0, 23370.0]
source_col: residuals
SAMP_residuals_STD | 0.0124197423484
varray: [2.957, 2.951, 2.937, 2.93]
inst: std_all
operation: std_all(residuals)
tarray: [21758.0, 22295.0, 22832.0, 23370.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.41225
varray: [0.382, 0.277, 0.589, 0.401]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [21758.0, 22295.0, 22832.0, 23370.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.129839323781
varray: [0.382, 0.277, 0.589, 0.401]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [21758.0, 22295.0, 22832.0, 23370.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1298.002
varray: [1297.568, 1299.544, 1297.768, 1297.128]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [21758.0, 22295.0, 22832.0, 23370.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.06219144539
varray: [1297.568, 1299.544, 1297.768, 1297.128]
inst: std_all
operation: std_all(baseline_shift)
tarray: [21758.0, 22295.0, 22832.0, 23370.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.968
varray: [1.864, 1.928, 2.43, 1.65]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [21758.0, 22295.0, 22832.0, 23370.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.330143403185
varray: [1.864, 1.928, 2.43, 1.65]
inst: std_all
operation: std_all(slope_shift)
tarray: [21758.0, 22295.0, 22832.0, 23370.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0597013400185
varray: [-8.926, -8.969, -9.025, -9.061]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21758.0, 22295.0, 22832.0, 23370.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.41354271248
varray: [-81.38, -83.399, -84.178, -84.545]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21758.0, 22295.0, 22832.0, 23370.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.99225
varray: [113.931, 114.198, 114.056, 113.784]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [21758.0, 22295.0, 22832.0, 23370.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.176556648133
varray: [113.931, 114.198, 114.056, 113.784]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [21758.0, 22295.0, 22832.0, 23370.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.94375
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.41225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1298.002
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.968
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1027
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1027
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.37113689704
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.01385
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [-8.926, -8.969, -9.025, -9.061]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [21758.0, 22295.0, 22832.0, 23370.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-81.38, -83.399, -84.178, -84.545]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [21758.0, 22295.0, 22832.0, 23370.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.99525
varray: [-8.926, -8.969, -9.025, -9.061]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21758.0, 22295.0, 22832.0, 23370.0]
source_col: d(18_16)Mean
dD_mean | -83.3755
varray: [-81.38, -83.399, -84.178, -84.545]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21758.0, 22295.0, 22832.0, 23370.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0597013400185
varray: [-8.926, -8.969, -9.025, -9.061]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21758.0, 22295.0, 22832.0, 23370.0]
source_col: d(18_16)Mean
dD_stdev | 1.41354271248
varray: [-81.38, -83.399, -84.178, -84.545]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21758.0, 22295.0, 22832.0, 23370.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.99525
varray: [-8.926, -8.969, -9.025, -9.061]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21758.0, 22295.0, 22832.0, 23370.0]
source_col: d(18_16)Mean
measured_dD_mean | -83.3755
varray: [-81.38, -83.399, -84.178, -84.545]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21758.0, 22295.0, 22832.0, 23370.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0597013400185
varray: [-8.926, -8.969, -9.025, -9.061]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21758.0, 22295.0, 22832.0, 23370.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.41354271248
varray: [-81.38, -83.399, -84.178, -84.545]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21758.0, 22295.0, 22832.0, 23370.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -11.815316795
actual_array: [ 1.93 -15.84]
varray: [-8.926, -8.969, -9.025, -9.061]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -8.99525
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [21758.0, 22295.0, 22832.0, 23370.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -96.6541231522
actual_array: [ 18.1 -119. ]
varray: [-81.38, -83.399, -84.178, -84.545]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -83.3755
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [21758.0, 22295.0, 22832.0, 23370.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1298.89315
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -3.8349
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 6 - Sample 6 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 6
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.94575
varray: [3.021, 2.883, 2.954, 2.925]
inst: mean_all
operation: mean_all(residuals)
tarray: [23960.0, 24496.0, 25034.0, 25571.0]
source_col: residuals
SAMP_residuals_STD | 0.0580193933095
varray: [3.021, 2.883, 2.954, 2.925]
inst: std_all
operation: std_all(residuals)
tarray: [23960.0, 24496.0, 25034.0, 25571.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.195
varray: [0.392, -0.092, 0.143, 0.337]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [23960.0, 24496.0, 25034.0, 25571.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.219124013594
varray: [0.392, -0.092, 0.143, 0.337]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [23960.0, 24496.0, 25034.0, 25571.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1299.1925
varray: [1300.071, 1298.901, 1299.358, 1298.44]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [23960.0, 24496.0, 25034.0, 25571.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.695313118914
varray: [1300.071, 1298.901, 1299.358, 1298.44]
inst: std_all
operation: std_all(baseline_shift)
tarray: [23960.0, 24496.0, 25034.0, 25571.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.5015
varray: [1.302, 1.538, 1.501, 1.665]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [23960.0, 24496.0, 25034.0, 25571.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.150405008782
varray: [1.302, 1.538, 1.501, 1.665]
inst: std_all
operation: std_all(slope_shift)
tarray: [23960.0, 24496.0, 25034.0, 25571.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0594495584508
varray: [-7.238, -7.184, -7.119, -7.112]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23960.0, 24496.0, 25034.0, 25571.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.245767505311
varray: [-78.453, -78.214, -77.927, -77.96]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23960.0, 24496.0, 25034.0, 25571.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 114.4575
varray: [116.261, 113.007, 114.118, 114.444]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [23960.0, 24496.0, 25034.0, 25571.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.35055605832
varray: [116.261, 113.007, 114.118, 114.444]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [23960.0, 24496.0, 25034.0, 25571.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.94575
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.195
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1299.1925
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.5015
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1007
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1007
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.30548671404
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.5486
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [-7.238, -7.184, -7.119, -7.112]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [23960.0, 24496.0, 25034.0, 25571.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-78.453, -78.214, -77.927, -77.96]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [23960.0, 24496.0, 25034.0, 25571.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.16325
varray: [-7.238, -7.184, -7.119, -7.112]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23960.0, 24496.0, 25034.0, 25571.0]
source_col: d(18_16)Mean
dD_mean | -78.1385
varray: [-78.453, -78.214, -77.927, -77.96]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23960.0, 24496.0, 25034.0, 25571.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0594495584508
varray: [-7.238, -7.184, -7.119, -7.112]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23960.0, 24496.0, 25034.0, 25571.0]
source_col: d(18_16)Mean
dD_stdev | 0.245767505311
varray: [-78.453, -78.214, -77.927, -77.96]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23960.0, 24496.0, 25034.0, 25571.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.16325
varray: [-7.238, -7.184, -7.119, -7.112]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23960.0, 24496.0, 25034.0, 25571.0]
source_col: d(18_16)Mean
measured_dD_mean | -78.1385
varray: [-78.453, -78.214, -77.927, -77.96]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23960.0, 24496.0, 25034.0, 25571.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0594495584508
varray: [-7.238, -7.184, -7.119, -7.112]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23960.0, 24496.0, 25034.0, 25571.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.245767505311
varray: [-78.453, -78.214, -77.927, -77.96]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23960.0, 24496.0, 25034.0, 25571.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.97631056416
actual_array: [ 1.93 -15.84]
varray: [-7.238, -7.184, -7.119, -7.112]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.16325
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [23960.0, 24496.0, 25034.0, 25571.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -91.7307198802
actual_array: [ 18.1 -119. ]
varray: [-78.453, -78.214, -77.927, -77.96]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -78.1385
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [23960.0, 24496.0, 25034.0, 25571.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1298.89115
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -2.6444
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 7 - Sample 7 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 7
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 6.50225
varray: [6.673, 6.435, 6.421, 6.48]
inst: mean_all
operation: mean_all(residuals)
tarray: [26162.0, 26699.0, 27236.0, 27774.0]
source_col: residuals
SAMP_residuals_STD | 0.116582946151
varray: [6.673, 6.435, 6.421, 6.48]
inst: std_all
operation: std_all(residuals)
tarray: [26162.0, 26699.0, 27236.0, 27774.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -13.6635
varray: [-13.878, -13.437, -13.825, -13.514]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [26162.0, 26699.0, 27236.0, 27774.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.220412492084
varray: [-13.878, -13.437, -13.825, -13.514]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [26162.0, 26699.0, 27236.0, 27774.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1670.54825
varray: [1680.041, 1667.745, 1668.101, 1666.306]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [26162.0, 26699.0, 27236.0, 27774.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 6.37589812628
varray: [1680.041, 1667.745, 1668.101, 1666.306]
inst: std_all
operation: std_all(baseline_shift)
tarray: [26162.0, 26699.0, 27236.0, 27774.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -51.82225
varray: [-52.729, -51.8, -51.45, -51.31]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [26162.0, 26699.0, 27236.0, 27774.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.63866025136
varray: [-52.729, -51.8, -51.45, -51.31]
inst: std_all
operation: std_all(slope_shift)
tarray: [26162.0, 26699.0, 27236.0, 27774.0]
source_col: slope_shift
SAMP_d18O_STD | 0.154706981097
varray: [15.521, 15.746, 15.854, 15.844]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26162.0, 26699.0, 27236.0, 27774.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.04607855027
varray: [-28.048, -26.456, -25.865, -25.801]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26162.0, 26699.0, 27236.0, 27774.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 139.80425
varray: [143.299, 139.445, 138.14, 138.333]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [26162.0, 26699.0, 27236.0, 27774.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.39976684631
varray: [143.299, 139.445, 138.14, 138.333]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [26162.0, 26699.0, 27236.0, 27774.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 6.50225
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 13.6635
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1670.54825
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 51.82225
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 3.4558
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 3.4558
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 113.436951206
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 22.79815
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [15.521, 15.746, 15.854, 15.844]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [26162.0, 26699.0, 27236.0, 27774.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-28.048, -26.456, -25.865, -25.801]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [26162.0, 26699.0, 27236.0, 27774.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 15.74125
varray: [15.521, 15.746, 15.854, 15.844]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26162.0, 26699.0, 27236.0, 27774.0]
source_col: d(18_16)Mean
dD_mean | -26.5425
varray: [-28.048, -26.456, -25.865, -25.801]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26162.0, 26699.0, 27236.0, 27774.0]
source_col: d(D_H)Mean
d18O_stdev | 0.154706981097
varray: [15.521, 15.746, 15.854, 15.844]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26162.0, 26699.0, 27236.0, 27774.0]
source_col: d(18_16)Mean
dD_stdev | 1.04607855027
varray: [-28.048, -26.456, -25.865, -25.801]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26162.0, 26699.0, 27236.0, 27774.0]
source_col: d(D_H)Mean
measured_d18O_mean | 15.74125
varray: [15.521, 15.746, 15.854, 15.844]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26162.0, 26699.0, 27236.0, 27774.0]
source_col: d(18_16)Mean
measured_dD_mean | -26.5425
varray: [-28.048, -26.456, -25.865, -25.801]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26162.0, 26699.0, 27236.0, 27774.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.154706981097
varray: [15.521, 15.746, 15.854, 15.844]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26162.0, 26699.0, 27236.0, 27774.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.04607855027
varray: [-28.048, -26.456, -25.865, -25.801]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26162.0, 26699.0, 27236.0, 27774.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 13.0157845308
actual_array: [ 1.93 -15.84]
varray: [15.521, 15.746, 15.854, 15.844]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 15.74125
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [26162.0, 26699.0, 27236.0, 27774.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -43.2243392767
actual_array: [ 18.1 -119. ]
varray: [-28.048, -26.456, -25.865, -25.801]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -26.5425
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [26162.0, 26699.0, 27236.0, 27774.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1295.33465
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 368.71135
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 8 - Sample 8 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 8
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 4.60525
varray: [4.752, 4.513, 4.57, 4.586]
inst: mean_all
operation: mean_all(residuals)
tarray: [28360.0, 28897.0, 29434.0, 29971.0]
source_col: residuals
SAMP_residuals_STD | 0.102727390051
varray: [4.752, 4.513, 4.57, 4.586]
inst: std_all
operation: std_all(residuals)
tarray: [28360.0, 28897.0, 29434.0, 29971.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -9.46225
varray: [-9.664, -9.259, -9.279, -9.647]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [28360.0, 28897.0, 29434.0, 29971.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.223403036386
varray: [-9.664, -9.259, -9.279, -9.647]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [28360.0, 28897.0, 29434.0, 29971.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 2231.6015
varray: [2255.738, 2218.743, 2225.417, 2226.508]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [28360.0, 28897.0, 29434.0, 29971.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 16.4529980753
varray: [2255.738, 2218.743, 2225.417, 2226.508]
inst: std_all
operation: std_all(baseline_shift)
tarray: [28360.0, 28897.0, 29434.0, 29971.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -57.3105
varray: [-58.201, -56.811, -56.968, -57.262]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [28360.0, 28897.0, 29434.0, 29971.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.622400996143
varray: [-58.201, -56.811, -56.968, -57.262]
inst: std_all
operation: std_all(slope_shift)
tarray: [28360.0, 28897.0, 29434.0, 29971.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0919941121304
varray: [8.079, 7.93, 7.883, 7.887]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28360.0, 28897.0, 29434.0, 29971.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.416050778151
varray: [-41.103, -41.734, -42.103, -41.749]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28360.0, 28897.0, 29434.0, 29971.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 121.41875
varray: [123.722, 119.968, 121.432, 120.553]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [28360.0, 28897.0, 29434.0, 29971.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.64917502103
varray: [123.722, 119.968, 121.432, 120.553]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [28360.0, 28897.0, 29434.0, 29971.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 4.60525
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 9.46225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 2231.6015
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 57.3105
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 1.5588
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.5588
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 51.1677526301
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 4.41265
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [8.079, 7.93, 7.883, 7.887]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [28360.0, 28897.0, 29434.0, 29971.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-41.103, -41.734, -42.103, -41.749]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [28360.0, 28897.0, 29434.0, 29971.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 7.94475
varray: [8.079, 7.93, 7.883, 7.887]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28360.0, 28897.0, 29434.0, 29971.0]
source_col: d(18_16)Mean
dD_mean | -41.67225
varray: [-41.103, -41.734, -42.103, -41.749]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28360.0, 28897.0, 29434.0, 29971.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0919941121304
varray: [8.079, 7.93, 7.883, 7.887]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28360.0, 28897.0, 29434.0, 29971.0]
source_col: d(18_16)Mean
dD_stdev | 0.416050778151
varray: [-41.103, -41.734, -42.103, -41.749]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28360.0, 28897.0, 29434.0, 29971.0]
source_col: d(D_H)Mean
measured_d18O_mean | 7.94475
varray: [8.079, 7.93, 7.883, 7.887]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28360.0, 28897.0, 29434.0, 29971.0]
source_col: d(18_16)Mean
measured_dD_mean | -41.67225
varray: [-41.103, -41.734, -42.103, -41.749]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28360.0, 28897.0, 29434.0, 29971.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0919941121304
varray: [8.079, 7.93, 7.883, 7.887]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28360.0, 28897.0, 29434.0, 29971.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.416050778151
varray: [-41.103, -41.734, -42.103, -41.749]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28360.0, 28897.0, 29434.0, 29971.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 5.189467894
actual_array: [ 1.93 -15.84]
varray: [8.079, 7.93, 7.883, 7.887]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 7.94475
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [28360.0, 28897.0, 29434.0, 29971.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -57.4481049162
actual_array: [ 18.1 -119. ]
varray: [-41.103, -41.734, -42.103, -41.749]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -41.67225
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [28360.0, 28897.0, 29434.0, 29971.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1297.23165
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 929.7646
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 9 - Sample 9 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 9
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.89975
varray: [2.928, 2.893, 2.855, 2.923]
inst: mean_all
operation: mean_all(residuals)
tarray: [30560.0, 31097.0, 31634.0, 32172.0]
source_col: residuals
SAMP_residuals_STD | 0.0335993551525
varray: [2.928, 2.893, 2.855, 2.923]
inst: std_all
operation: std_all(residuals)
tarray: [30560.0, 31097.0, 31634.0, 32172.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.29075
varray: [0.148, 0.313, 0.218, 0.484]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [30560.0, 31097.0, 31634.0, 32172.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.1455
varray: [0.148, 0.313, 0.218, 0.484]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [30560.0, 31097.0, 31634.0, 32172.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1303.12475
varray: [1311.727, 1301.922, 1300.054, 1298.796]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [30560.0, 31097.0, 31634.0, 32172.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 5.87687260227
varray: [1311.727, 1301.922, 1300.054, 1298.796]
inst: std_all
operation: std_all(baseline_shift)
tarray: [30560.0, 31097.0, 31634.0, 32172.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.283
varray: [0.489, 1.2, 1.782, 1.661]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [30560.0, 31097.0, 31634.0, 32172.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.585721776956
varray: [0.489, 1.2, 1.782, 1.661]
inst: std_all
operation: std_all(slope_shift)
tarray: [30560.0, 31097.0, 31634.0, 32172.0]
source_col: slope_shift
SAMP_d18O_STD | 0.094164749243
varray: [-5.765, -5.945, -5.966, -5.946]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30560.0, 31097.0, 31634.0, 32172.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.699565758167
varray: [-71.707, -72.919, -73.162, -73.177]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30560.0, 31097.0, 31634.0, 32172.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 113.316
varray: [115.519, 113.444, 111.557, 112.744]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [30560.0, 31097.0, 31634.0, 32172.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.66241370703
varray: [115.519, 113.444, 111.557, 112.744]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [30560.0, 31097.0, 31634.0, 32172.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.89975
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.29075
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1303.12475
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.283
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1467
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1467
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.81544092304
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.6901
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [-5.765, -5.945, -5.966, -5.946]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [30560.0, 31097.0, 31634.0, 32172.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-71.707, -72.919, -73.162, -73.177]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [30560.0, 31097.0, 31634.0, 32172.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.9055
varray: [-5.765, -5.945, -5.966, -5.946]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30560.0, 31097.0, 31634.0, 32172.0]
source_col: d(18_16)Mean
dD_mean | -72.74125
varray: [-71.707, -72.919, -73.162, -73.177]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30560.0, 31097.0, 31634.0, 32172.0]
source_col: d(D_H)Mean
d18O_stdev | 0.094164749243
varray: [-5.765, -5.945, -5.966, -5.946]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30560.0, 31097.0, 31634.0, 32172.0]
source_col: d(18_16)Mean
dD_stdev | 0.699565758167
varray: [-71.707, -72.919, -73.162, -73.177]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30560.0, 31097.0, 31634.0, 32172.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.9055
varray: [-5.765, -5.945, -5.966, -5.946]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30560.0, 31097.0, 31634.0, 32172.0]
source_col: d(18_16)Mean
measured_dD_mean | -72.74125
varray: [-71.707, -72.919, -73.162, -73.177]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30560.0, 31097.0, 31634.0, 32172.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.094164749243
varray: [-5.765, -5.945, -5.966, -5.946]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30560.0, 31097.0, 31634.0, 32172.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.699565758167
varray: [-71.707, -72.919, -73.162, -73.177]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30560.0, 31097.0, 31634.0, 32172.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.7137504731
actual_array: [ 1.93 -15.84]
varray: [-5.765, -5.945, -5.966, -5.946]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -5.9055
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [30560.0, 31097.0, 31634.0, 32172.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -86.6566625364
actual_array: [ 18.1 -119. ]
varray: [-71.707, -72.919, -73.162, -73.177]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -72.74125
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [30560.0, 31097.0, 31634.0, 32172.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1298.93715
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1.28785
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 10 - Sample 10 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 10
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 6.0455
varray: [5.996, 6.083, 6.049, 6.054]
inst: mean_all
operation: mean_all(residuals)
tarray: [32760.0, 33297.0, 33835.0, 34372.0]
source_col: residuals
SAMP_residuals_STD | 0.0362445398187
varray: [5.996, 6.083, 6.049, 6.054]
inst: std_all
operation: std_all(residuals)
tarray: [32760.0, 33297.0, 33835.0, 34372.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -12.729
varray: [-12.55, -12.938, -12.942, -12.486]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [32760.0, 33297.0, 33835.0, 34372.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.245044213698
varray: [-12.55, -12.938, -12.942, -12.486]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [32760.0, 33297.0, 33835.0, 34372.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 2028.29025
varray: [2028.928, 2032.711, 2025.968, 2025.554]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [32760.0, 33297.0, 33835.0, 34372.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.30805404783
varray: [2028.928, 2032.711, 2025.968, 2025.554]
inst: std_all
operation: std_all(baseline_shift)
tarray: [32760.0, 33297.0, 33835.0, 34372.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -61.55375
varray: [-60.902, -61.693, -62.185, -61.435]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [32760.0, 33297.0, 33835.0, 34372.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.534398961451
varray: [-60.902, -61.693, -62.185, -61.435]
inst: std_all
operation: std_all(slope_shift)
tarray: [32760.0, 33297.0, 33835.0, 34372.0]
source_col: slope_shift
SAMP_d18O_STD | 0.121574394234
varray: [12.915, 13.11, 13.157, 13.184]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32760.0, 33297.0, 33835.0, 34372.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.512596982694
varray: [-36.736, -35.73, -35.714, -35.69]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32760.0, 33297.0, 33835.0, 34372.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 133.01225
varray: [132.745, 133.165, 134.06, 132.079]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [32760.0, 33297.0, 33835.0, 34372.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.829355321922
varray: [132.745, 133.165, 134.06, 132.079]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [32760.0, 33297.0, 33835.0, 34372.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 6.0455
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 12.729
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 2028.29025
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 61.55375
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 2.99905
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.99905
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 98.4440906629
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 16.00615
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [12.915, 13.11, 13.157, 13.184]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [32760.0, 33297.0, 33835.0, 34372.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-36.736, -35.73, -35.714, -35.69]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [32760.0, 33297.0, 33835.0, 34372.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 13.0915
varray: [12.915, 13.11, 13.157, 13.184]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32760.0, 33297.0, 33835.0, 34372.0]
source_col: d(18_16)Mean
dD_mean | -35.9675
varray: [-36.736, -35.73, -35.714, -35.69]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32760.0, 33297.0, 33835.0, 34372.0]
source_col: d(D_H)Mean
d18O_stdev | 0.121574394234
varray: [12.915, 13.11, 13.157, 13.184]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32760.0, 33297.0, 33835.0, 34372.0]
source_col: d(18_16)Mean
dD_stdev | 0.512596982694
varray: [-36.736, -35.73, -35.714, -35.69]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32760.0, 33297.0, 33835.0, 34372.0]
source_col: d(D_H)Mean
measured_d18O_mean | 13.0915
varray: [12.915, 13.11, 13.157, 13.184]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32760.0, 33297.0, 33835.0, 34372.0]
source_col: d(18_16)Mean
measured_dD_mean | -35.9675
varray: [-36.736, -35.73, -35.714, -35.69]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32760.0, 33297.0, 33835.0, 34372.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.121574394234
varray: [12.915, 13.11, 13.157, 13.184]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32760.0, 33297.0, 33835.0, 34372.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.512596982694
varray: [-36.736, -35.73, -35.714, -35.69]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32760.0, 33297.0, 33835.0, 34372.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 10.3559009281
actual_array: [ 1.93 -15.84]
varray: [12.915, 13.11, 13.157, 13.184]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 13.0915
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [32760.0, 33297.0, 33835.0, 34372.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -52.0849609758
actual_array: [ 18.1 -119. ]
varray: [-36.736, -35.73, -35.714, -35.69]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -35.9675
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [32760.0, 33297.0, 33835.0, 34372.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1295.7914
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 726.45335
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 11 - Sample 11 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 11
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.63875
varray: [2.628, 2.64, 2.64, 2.647]
inst: mean_all
operation: mean_all(residuals)
tarray: [34961.0, 35498.0, 36034.0, 36572.0]
source_col: residuals
SAMP_residuals_STD | 0.00788986691903
varray: [2.628, 2.64, 2.64, 2.647]
inst: std_all
operation: std_all(residuals)
tarray: [34961.0, 35498.0, 36034.0, 36572.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.28525
varray: [0.114, 0.228, 0.404, 0.395]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [34961.0, 35498.0, 36034.0, 36572.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.139941356765
varray: [0.114, 0.228, 0.404, 0.395]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [34961.0, 35498.0, 36034.0, 36572.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1291.08325
varray: [1296.462, 1290.26, 1289.341, 1288.27]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [34961.0, 35498.0, 36034.0, 36572.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.67688739516
varray: [1296.462, 1290.26, 1289.341, 1288.27]
inst: std_all
operation: std_all(baseline_shift)
tarray: [34961.0, 35498.0, 36034.0, 36572.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.20625
varray: [0.785, 1.305, 1.354, 1.381]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [34961.0, 35498.0, 36034.0, 36572.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.282589661288
varray: [0.785, 1.305, 1.354, 1.381]
inst: std_all
operation: std_all(slope_shift)
tarray: [34961.0, 35498.0, 36034.0, 36572.0]
source_col: slope_shift
SAMP_d18O_STD | 0.200735937656
varray: [-8.186, -8.545, -8.548, -8.64]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34961.0, 35498.0, 36034.0, 36572.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.69507096705
varray: [-88.77, -91.492, -92.185, -92.49]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34961.0, 35498.0, 36034.0, 36572.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 104.0855
varray: [103.475, 102.943, 105.32, 104.604]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [34961.0, 35498.0, 36034.0, 36572.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.07561594757
varray: [103.475, 102.943, 105.32, 104.604]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [34961.0, 35498.0, 36034.0, 36572.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.63875
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.28525
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1291.08325
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.20625
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.4077
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.4077
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -13.3827898045
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -12.9206
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [-8.186, -8.545, -8.548, -8.64]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [34961.0, 35498.0, 36034.0, 36572.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-88.77, -91.492, -92.185, -92.49]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [34961.0, 35498.0, 36034.0, 36572.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.47975
varray: [-8.186, -8.545, -8.548, -8.64]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [34961.0, 35498.0, 36034.0, 36572.0]
source_col: d(18_16)Mean
dD_mean | -91.23425
varray: [-88.77, -91.492, -92.185, -92.49]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [34961.0, 35498.0, 36034.0, 36572.0]
source_col: d(D_H)Mean
d18O_stdev | 0.200735937656
varray: [-8.186, -8.545, -8.548, -8.64]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34961.0, 35498.0, 36034.0, 36572.0]
source_col: d(18_16)Mean
dD_stdev | 1.69507096705
varray: [-88.77, -91.492, -92.185, -92.49]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34961.0, 35498.0, 36034.0, 36572.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.47975
varray: [-8.186, -8.545, -8.548, -8.64]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [34961.0, 35498.0, 36034.0, 36572.0]
source_col: d(18_16)Mean
measured_dD_mean | -91.23425
varray: [-88.77, -91.492, -92.185, -92.49]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [34961.0, 35498.0, 36034.0, 36572.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.200735937656
varray: [-8.186, -8.545, -8.548, -8.64]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34961.0, 35498.0, 36034.0, 36572.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.69507096705
varray: [-88.77, -91.492, -92.185, -92.49]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34961.0, 35498.0, 36034.0, 36572.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -11.2978453365
actual_array: [ 1.93 -15.84]
varray: [-8.186, -8.545, -8.548, -8.64]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -8.47975
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [34961.0, 35498.0, 36034.0, 36572.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -104.042283447
actual_array: [ 18.1 -119. ]
varray: [-88.77, -91.492, -92.185, -92.49]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -91.23425
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [34961.0, 35498.0, 36034.0, 36572.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.19815
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.75365
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 12 - Sample 12 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 12
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.75525
varray: [2.851, 2.741, 2.738, 2.691]
inst: mean_all
operation: mean_all(residuals)
tarray: [37162.0, 37699.0, 38239.0, 38776.0]
source_col: residuals
SAMP_residuals_STD | 0.0678153129217
varray: [2.851, 2.741, 2.738, 2.691]
inst: std_all
operation: std_all(residuals)
tarray: [37162.0, 37699.0, 38239.0, 38776.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.19225
varray: [0.585, 0.266, -0.118, 0.036]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [37162.0, 37699.0, 38239.0, 38776.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.30570178824
varray: [0.585, 0.266, -0.118, 0.036]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [37162.0, 37699.0, 38239.0, 38776.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1292.8095
varray: [1291.536, 1292.706, 1295.176, 1291.82]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [37162.0, 37699.0, 38239.0, 38776.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.6544835045
varray: [1291.536, 1292.706, 1295.176, 1291.82]
inst: std_all
operation: std_all(baseline_shift)
tarray: [37162.0, 37699.0, 38239.0, 38776.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.83975
varray: [1.516, 2.06, 1.905, 1.878]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [37162.0, 37699.0, 38239.0, 38776.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.230249683315
varray: [1.516, 2.06, 1.905, 1.878]
inst: std_all
operation: std_all(slope_shift)
tarray: [37162.0, 37699.0, 38239.0, 38776.0]
source_col: slope_shift
SAMP_d18O_STD | 0.111649749962
varray: [-12.128, -12.334, -12.363, -12.353]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37162.0, 37699.0, 38239.0, 38776.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.990656474263
varray: [-106.41, -108.026, -108.321, -108.636]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37162.0, 37699.0, 38239.0, 38776.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 107.767
varray: [110.109, 107.535, 107.825, 105.599]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [37162.0, 37699.0, 38239.0, 38776.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.84773519026
varray: [110.109, 107.535, 107.825, 105.599]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [37162.0, 37699.0, 38239.0, 38776.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.75525
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.19225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1292.8095
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.83975
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.2912
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.2912
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -9.55866664478
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -9.2391
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [-12.128, -12.334, -12.363, -12.353]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [37162.0, 37699.0, 38239.0, 38776.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-106.41, -108.026, -108.321, -108.636]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [37162.0, 37699.0, 38239.0, 38776.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -12.2945
varray: [-12.128, -12.334, -12.363, -12.353]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37162.0, 37699.0, 38239.0, 38776.0]
source_col: d(18_16)Mean
dD_mean | -107.84825
varray: [-106.41, -108.026, -108.321, -108.636]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37162.0, 37699.0, 38239.0, 38776.0]
source_col: d(D_H)Mean
d18O_stdev | 0.111649749962
varray: [-12.128, -12.334, -12.363, -12.353]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37162.0, 37699.0, 38239.0, 38776.0]
source_col: d(18_16)Mean
dD_stdev | 0.990656474263
varray: [-106.41, -108.026, -108.321, -108.636]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37162.0, 37699.0, 38239.0, 38776.0]
source_col: d(D_H)Mean
measured_d18O_mean | -12.2945
varray: [-12.128, -12.334, -12.363, -12.353]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37162.0, 37699.0, 38239.0, 38776.0]
source_col: d(18_16)Mean
measured_dD_mean | -107.84825
varray: [-106.41, -108.026, -108.321, -108.636]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37162.0, 37699.0, 38239.0, 38776.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.111649749962
varray: [-12.128, -12.334, -12.363, -12.353]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37162.0, 37699.0, 38239.0, 38776.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.990656474263
varray: [-106.41, -108.026, -108.321, -108.636]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37162.0, 37699.0, 38239.0, 38776.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -15.1271843207
actual_array: [ 1.93 -15.84]
varray: [-12.128, -12.334, -12.363, -12.353]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -12.2945
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [37162.0, 37699.0, 38239.0, 38776.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -119.661420732
actual_array: [ 18.1 -119. ]
varray: [-106.41, -108.026, -108.321, -108.636]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -107.84825
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [37162.0, 37699.0, 38239.0, 38776.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.08165
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -9.0274
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 13 - Sample 13 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 13
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.086
varray: [0.0, 2.79, 2.75, 2.804]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [39366.0, 39903.0, 40440.0, 40977.0]
source_col: residuals
SAMP_residuals_STD | 1.39085489298
varray: [0.0, 2.79, 2.75, 2.804]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [39366.0, 39903.0, 40440.0, 40977.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.254
varray: [0.0, 0.197, 0.354, 0.465]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [39366.0, 39903.0, 40440.0, 40977.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.201896012838
varray: [0.0, 0.197, 0.354, 0.465]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [39366.0, 39903.0, 40440.0, 40977.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 968.84775
varray: [0.0, 1292.921, 1291.144, 1291.326]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [39366.0, 39903.0, 40440.0, 40977.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 645.898993273
varray: [0.0, 1292.921, 1291.144, 1291.326]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [39366.0, 39903.0, 40440.0, 40977.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.24125
varray: [0.0, 1.765, 1.647, 1.553]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [39366.0, 39903.0, 40440.0, 40977.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.83203300215
varray: [0.0, 1.765, 1.647, 1.553]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [39366.0, 39903.0, 40440.0, 40977.0]
source_col: slope_shift
SAMP_d18O_STD | 3.82822738571
varray: [0.0, -7.685, -7.651, -7.633]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39366.0, 39903.0, 40440.0, 40977.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 30.8638063042
varray: [0.0, -62.481, -61.489, -61.183]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39366.0, 39903.0, 40440.0, 40977.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 81.28775
varray: [0.0, 108.188, 108.261, 108.702]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [39366.0, 39903.0, 40440.0, 40977.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 54.1923090138
varray: [0.0, 108.188, 108.261, 108.702]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [39366.0, 39903.0, 40440.0, 40977.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.086
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.254
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 968.84775
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.24125
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.96045
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.96045
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -31.5268591311
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -35.71835
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [0.0, -7.685, -7.651, -7.633]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [39366.0, 39903.0, 40440.0, 40977.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [0.0, -62.481, -61.489, -61.183]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [39366.0, 39903.0, 40440.0, 40977.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.74225
varray: [0.0, -7.685, -7.651, -7.633]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39366.0, 39903.0, 40440.0, 40977.0]
source_col: d(18_16)Mean
dD_mean | -46.28825
varray: [0.0, -62.481, -61.489, -61.183]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39366.0, 39903.0, 40440.0, 40977.0]
source_col: d(D_H)Mean
d18O_stdev | 3.82822738571
varray: [0.0, -7.685, -7.651, -7.633]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39366.0, 39903.0, 40440.0, 40977.0]
source_col: d(18_16)Mean
dD_stdev | 30.8638063042
varray: [0.0, -62.481, -61.489, -61.183]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39366.0, 39903.0, 40440.0, 40977.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.74225
varray: [0.0, -7.685, -7.651, -7.633]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39366.0, 39903.0, 40440.0, 40977.0]
source_col: d(18_16)Mean
measured_dD_mean | -46.28825
varray: [0.0, -62.481, -61.489, -61.183]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39366.0, 39903.0, 40440.0, 40977.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 3.82822738571
varray: [0.0, -7.685, -7.651, -7.633]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39366.0, 39903.0, 40440.0, 40977.0]
source_col: d(18_16)Mean
measured_dD_stdev | 30.8638063042
varray: [0.0, -62.481, -61.489, -61.183]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39366.0, 39903.0, 40440.0, 40977.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.54987614604
actual_array: [ 1.93 -15.84]
varray: [0.0, -7.685, -7.651, -7.633]
fit_rsq: 1.0
fit_intercept: 2.77505289814
except: True
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -5.74225
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [39366.0, 39903.0, 40440.0, 40977.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -61.787694281
actual_array: [ 18.1 -119. ]
varray: [0.0, -62.481, -61.489, -61.183]
fit_rsq: 1.0
fit_intercept: 19.4350184537
except: True
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -46.28825
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [39366.0, 39903.0, 40440.0, 40977.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.7509
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -332.98915
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 14 - Sample 14 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 14
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0875
varray: [0.0, 2.816, 2.769, 2.765]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [41567.0, 42104.0, 42641.0, 43178.0]
source_col: residuals
SAMP_residuals_STD | 1.39185930802
varray: [0.0, 2.816, 2.769, 2.765]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [41567.0, 42104.0, 42641.0, 43178.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.213
varray: [0.0, 0.419, 0.232, 0.201]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [41567.0, 42104.0, 42641.0, 43178.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.171571170849
varray: [0.0, 0.419, 0.232, 0.201]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [41567.0, 42104.0, 42641.0, 43178.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 970.0305
varray: [0.0, 1293.668, 1293.349, 1293.105]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [41567.0, 42104.0, 42641.0, 43178.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 646.687041087
varray: [0.0, 1293.668, 1293.349, 1293.105]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [41567.0, 42104.0, 42641.0, 43178.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.53325
varray: [0.0, 1.775, 2.183, 2.175]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [41567.0, 42104.0, 42641.0, 43178.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.03976227251
varray: [0.0, 1.775, 2.183, 2.175]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [41567.0, 42104.0, 42641.0, 43178.0]
source_col: slope_shift
SAMP_d18O_STD | 5.80333990044
varray: [0.0, -11.595, -11.609, -11.616]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41567.0, 42104.0, 42641.0, 43178.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 50.8743057045
varray: [0.0, -101.044, -101.844, -102.341]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41567.0, 42104.0, 42641.0, 43178.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 81.64025
varray: [0.0, 110.024, 108.808, 107.729]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [41567.0, 42104.0, 42641.0, 43178.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 54.4349066952
varray: [0.0, 110.024, 108.808, 107.729]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [41567.0, 42104.0, 42641.0, 43178.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0875
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.213
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 970.0305
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.53325
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.95895
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.95895
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -31.4776214939
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -35.36585
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [0.0, -11.595, -11.609, -11.616]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [41567.0, 42104.0, 42641.0, 43178.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [0.0, -101.044, -101.844, -102.341]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [41567.0, 42104.0, 42641.0, 43178.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.705
varray: [0.0, -11.595, -11.609, -11.616]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [41567.0, 42104.0, 42641.0, 43178.0]
source_col: d(18_16)Mean
dD_mean | -76.30725
varray: [0.0, -101.044, -101.844, -102.341]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [41567.0, 42104.0, 42641.0, 43178.0]
source_col: d(D_H)Mean
d18O_stdev | 5.80333990044
varray: [0.0, -11.595, -11.609, -11.616]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41567.0, 42104.0, 42641.0, 43178.0]
source_col: d(18_16)Mean
dD_stdev | 50.8743057045
varray: [0.0, -101.044, -101.844, -102.341]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41567.0, 42104.0, 42641.0, 43178.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.705
varray: [0.0, -11.595, -11.609, -11.616]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [41567.0, 42104.0, 42641.0, 43178.0]
source_col: d(18_16)Mean
measured_dD_mean | -76.30725
varray: [0.0, -101.044, -101.844, -102.341]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [41567.0, 42104.0, 42641.0, 43178.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 5.80333990044
varray: [0.0, -11.595, -11.609, -11.616]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41567.0, 42104.0, 42641.0, 43178.0]
source_col: d(18_16)Mean
measured_dD_stdev | 50.8743057045
varray: [0.0, -101.044, -101.844, -102.341]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41567.0, 42104.0, 42641.0, 43178.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -11.523956774
actual_array: [ 1.93 -15.84]
varray: [0.0, -11.595, -11.609, -11.616]
fit_rsq: 1.0
fit_intercept: 2.77505289814
except: True
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -8.705
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [41567.0, 42104.0, 42641.0, 43178.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -90.0091269373
actual_array: [ 18.1 -119. ]
varray: [0.0, -101.044, -101.844, -102.341]
fit_rsq: 1.0
fit_intercept: 19.4350184537
except: True
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -76.30725
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [41567.0, 42104.0, 42641.0, 43178.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.7494
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -331.8064
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 15 - Sample 15 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 15
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.76
varray: [2.801, 2.755, 2.727, 2.757]
inst: mean_all
operation: mean_all(residuals)
tarray: [43766.0, 44302.0, 44839.0, 45377.0]
source_col: residuals
SAMP_residuals_STD | 0.0305723186341
varray: [2.801, 2.755, 2.727, 2.757]
inst: std_all
operation: std_all(residuals)
tarray: [43766.0, 44302.0, 44839.0, 45377.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.29875
varray: [0.552, 0.37, 0.168, 0.105]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [43766.0, 44302.0, 44839.0, 45377.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.203180338616
varray: [0.552, 0.37, 0.168, 0.105]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [43766.0, 44302.0, 44839.0, 45377.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1292.2895
varray: [1291.803, 1292.287, 1292.406, 1292.662]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [43766.0, 44302.0, 44839.0, 45377.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.360100448949
varray: [1291.803, 1292.287, 1292.406, 1292.662]
inst: std_all
operation: std_all(baseline_shift)
tarray: [43766.0, 44302.0, 44839.0, 45377.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.83025
varray: [1.17, 2.292, 2.356, 1.503]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [43766.0, 44302.0, 44839.0, 45377.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.586699454576
varray: [1.17, 2.292, 2.356, 1.503]
inst: std_all
operation: std_all(slope_shift)
tarray: [43766.0, 44302.0, 44839.0, 45377.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0191898410624
varray: [-11.48, -11.438, -11.447, -11.468]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43766.0, 44302.0, 44839.0, 45377.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.064812550225
varray: [-102.117, -102.127, -102.088, -102.236]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43766.0, 44302.0, 44839.0, 45377.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.33725
varray: [110.05, 108.008, 107.021, 108.27]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [43766.0, 44302.0, 44839.0, 45377.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.26213506277
varray: [110.05, 108.008, 107.021, 108.27]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [43766.0, 44302.0, 44839.0, 45377.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.76
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.29875
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1292.2895
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.83025
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.28645
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.28645
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -9.40274746016
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -8.66885
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [-11.48, -11.438, -11.447, -11.468]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [43766.0, 44302.0, 44839.0, 45377.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-102.117, -102.127, -102.088, -102.236]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [43766.0, 44302.0, 44839.0, 45377.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -11.45825
varray: [-11.48, -11.438, -11.447, -11.468]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [43766.0, 44302.0, 44839.0, 45377.0]
source_col: d(18_16)Mean
dD_mean | -102.142
varray: [-102.117, -102.127, -102.088, -102.236]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [43766.0, 44302.0, 44839.0, 45377.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0191898410624
varray: [-11.48, -11.438, -11.447, -11.468]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43766.0, 44302.0, 44839.0, 45377.0]
source_col: d(18_16)Mean
dD_stdev | 0.064812550225
varray: [-102.117, -102.127, -102.088, -102.236]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43766.0, 44302.0, 44839.0, 45377.0]
source_col: d(D_H)Mean
measured_d18O_mean | -11.45825
varray: [-11.48, -11.438, -11.447, -11.468]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [43766.0, 44302.0, 44839.0, 45377.0]
source_col: d(18_16)Mean
measured_dD_mean | -102.142
varray: [-102.117, -102.127, -102.088, -102.236]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [43766.0, 44302.0, 44839.0, 45377.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0191898410624
varray: [-11.48, -11.438, -11.447, -11.468]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43766.0, 44302.0, 44839.0, 45377.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.064812550225
varray: [-102.117, -102.127, -102.088, -102.236]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43766.0, 44302.0, 44839.0, 45377.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -14.2877361981
actual_array: [ 1.93 -15.84]
varray: [-11.48, -11.438, -11.447, -11.468]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -11.45825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [43766.0, 44302.0, 44839.0, 45377.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -114.296866613
actual_array: [ 18.1 -119. ]
varray: [-102.117, -102.127, -102.088, -102.236]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -102.142
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [43766.0, 44302.0, 44839.0, 45377.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.0769
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -9.5474
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 16 - Sample 16 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 16
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 5.4565
varray: [5.421, 5.492, 5.436, 5.477]
inst: mean_all
operation: mean_all(residuals)
tarray: [45965.0, 46502.0, 47039.0, 47575.0]
source_col: residuals
SAMP_residuals_STD | 0.0334713808101
varray: [5.421, 5.492, 5.436, 5.477]
inst: std_all
operation: std_all(residuals)
tarray: [45965.0, 46502.0, 47039.0, 47575.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -11.22875
varray: [-11.292, -11.297, -11.315, -11.011]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [45965.0, 46502.0, 47039.0, 47575.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.145502290933
varray: [-11.292, -11.297, -11.315, -11.011]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [45965.0, 46502.0, 47039.0, 47575.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1916.765
varray: [1909.159, 1920.665, 1919.074, 1918.162]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [45965.0, 46502.0, 47039.0, 47575.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 5.1750789366
varray: [1909.159, 1920.665, 1919.074, 1918.162]
inst: std_all
operation: std_all(baseline_shift)
tarray: [45965.0, 46502.0, 47039.0, 47575.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -53.088
varray: [-53.029, -52.852, -53.525, -52.946]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [45965.0, 46502.0, 47039.0, 47575.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.300172172816
varray: [-53.029, -52.852, -53.525, -52.946]
inst: std_all
operation: std_all(slope_shift)
tarray: [45965.0, 46502.0, 47039.0, 47575.0]
source_col: slope_shift
SAMP_d18O_STD | 0.293926294616
varray: [11.604, 12.056, 12.181, 12.263]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45965.0, 46502.0, 47039.0, 47575.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.43409550278
varray: [-41.673, -37.586, -36.643, -36.481]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45965.0, 46502.0, 47039.0, 47575.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 128.93075
varray: [128.561, 128.857, 129.401, 128.904]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [45965.0, 46502.0, 47039.0, 47575.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.348331618816
varray: [128.561, 128.857, 129.401, 128.904]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [45965.0, 46502.0, 47039.0, 47575.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 5.4565
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 11.22875
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1916.765
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 53.088
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 2.41005
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.41005
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 79.1101117694
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 11.92465
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [11.604, 12.056, 12.181, 12.263]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [45965.0, 46502.0, 47039.0, 47575.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-41.673, -37.586, -36.643, -36.481]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [45965.0, 46502.0, 47039.0, 47575.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 12.026
varray: [11.604, 12.056, 12.181, 12.263]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45965.0, 46502.0, 47039.0, 47575.0]
source_col: d(18_16)Mean
dD_mean | -38.09575
varray: [-41.673, -37.586, -36.643, -36.481]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45965.0, 46502.0, 47039.0, 47575.0]
source_col: d(D_H)Mean
d18O_stdev | 0.293926294616
varray: [11.604, 12.056, 12.181, 12.263]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45965.0, 46502.0, 47039.0, 47575.0]
source_col: d(18_16)Mean
dD_stdev | 2.43409550278
varray: [-41.673, -37.586, -36.643, -36.481]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45965.0, 46502.0, 47039.0, 47575.0]
source_col: d(D_H)Mean
measured_d18O_mean | 12.026
varray: [11.604, 12.056, 12.181, 12.263]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45965.0, 46502.0, 47039.0, 47575.0]
source_col: d(18_16)Mean
measured_dD_mean | -38.09575
varray: [-41.673, -37.586, -36.643, -36.481]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45965.0, 46502.0, 47039.0, 47575.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.293926294616
varray: [11.604, 12.056, 12.181, 12.263]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45965.0, 46502.0, 47039.0, 47575.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.43409550278
varray: [-41.673, -37.586, -36.643, -36.481]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45965.0, 46502.0, 47039.0, 47575.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 9.28632607062
actual_array: [ 1.93 -15.84]
varray: [11.604, 12.056, 12.181, 12.263]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 12.026
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [45965.0, 46502.0, 47039.0, 47575.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -54.0857692656
actual_array: [ 18.1 -119. ]
varray: [-41.673, -37.586, -36.643, -36.481]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -38.09575
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [45965.0, 46502.0, 47039.0, 47575.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1296.3804
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 614.9281
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 17 - Sample 17 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 17
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.77475
varray: [2.731, 2.818, 2.767, 2.783]
inst: mean_all
operation: mean_all(residuals)
tarray: [48163.0, 48701.0, 49238.0, 49775.0]
source_col: residuals
SAMP_residuals_STD | 0.0361144015595
varray: [2.731, 2.818, 2.767, 2.783]
inst: std_all
operation: std_all(residuals)
tarray: [48163.0, 48701.0, 49238.0, 49775.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.1595
varray: [0.376, 0.11, 0.056, 0.096]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [48163.0, 48701.0, 49238.0, 49775.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.14613578161
varray: [0.376, 0.11, 0.056, 0.096]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [48163.0, 48701.0, 49238.0, 49775.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1296.94775
varray: [1299.306, 1297.846, 1295.553, 1295.086]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [48163.0, 48701.0, 49238.0, 49775.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.98155130727
varray: [1299.306, 1297.846, 1295.553, 1295.086]
inst: std_all
operation: std_all(baseline_shift)
tarray: [48163.0, 48701.0, 49238.0, 49775.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.275
varray: [1.101, 1.32, 1.431, 1.248]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [48163.0, 48701.0, 49238.0, 49775.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.138282319911
varray: [1.101, 1.32, 1.431, 1.248]
inst: std_all
operation: std_all(slope_shift)
tarray: [48163.0, 48701.0, 49238.0, 49775.0]
source_col: slope_shift
SAMP_d18O_STD | 0.15158138628
varray: [-7.862, -8.052, -8.166, -8.197]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48163.0, 48701.0, 49238.0, 49775.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.798583902918
varray: [-78.137, -79.414, -79.736, -79.898]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48163.0, 48701.0, 49238.0, 49775.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.7745
varray: [107.896, 109.348, 108.739, 109.115]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [48163.0, 48701.0, 49238.0, 49775.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.637145980133
varray: [107.896, 109.348, 108.739, 109.115]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [48163.0, 48701.0, 49238.0, 49775.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.77475
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.1595
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1296.94775
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.275
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.2717
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.2717
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -8.91857736053
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -8.2316
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [-7.862, -8.052, -8.166, -8.197]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [48163.0, 48701.0, 49238.0, 49775.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-78.137, -79.414, -79.736, -79.898]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [48163.0, 48701.0, 49238.0, 49775.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.06925
varray: [-7.862, -8.052, -8.166, -8.197]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48163.0, 48701.0, 49238.0, 49775.0]
source_col: d(18_16)Mean
dD_mean | -79.29625
varray: [-78.137, -79.414, -79.736, -79.898]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48163.0, 48701.0, 49238.0, 49775.0]
source_col: d(D_H)Mean
d18O_stdev | 0.15158138628
varray: [-7.862, -8.052, -8.166, -8.197]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48163.0, 48701.0, 49238.0, 49775.0]
source_col: d(18_16)Mean
dD_stdev | 0.798583902918
varray: [-78.137, -79.414, -79.736, -79.898]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48163.0, 48701.0, 49238.0, 49775.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.06925
varray: [-7.862, -8.052, -8.166, -8.197]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48163.0, 48701.0, 49238.0, 49775.0]
source_col: d(18_16)Mean
measured_dD_mean | -79.29625
varray: [-78.137, -79.414, -79.736, -79.898]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48163.0, 48701.0, 49238.0, 49775.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.15158138628
varray: [-7.862, -8.052, -8.166, -8.197]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48163.0, 48701.0, 49238.0, 49775.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.798583902918
varray: [-78.137, -79.414, -79.736, -79.898]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48163.0, 48701.0, 49238.0, 49775.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.885775436
actual_array: [ 1.93 -15.84]
varray: [-7.862, -8.052, -8.166, -8.197]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -8.06925
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [48163.0, 48701.0, 49238.0, 49775.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -92.8191426677
actual_array: [ 18.1 -119. ]
varray: [-78.137, -79.414, -79.736, -79.898]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -79.29625
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [48163.0, 48701.0, 49238.0, 49775.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.06215
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -4.88915
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 18 - Sample 18 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 18
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.77275
varray: [2.757, 2.755, 2.804, 2.775]
inst: mean_all
operation: mean_all(residuals)
tarray: [50363.0, 50900.0, 51437.0, 51974.0]
source_col: residuals
SAMP_residuals_STD | 0.0226917753088
varray: [2.757, 2.755, 2.804, 2.775]
inst: std_all
operation: std_all(residuals)
tarray: [50363.0, 50900.0, 51437.0, 51974.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3505
varray: [0.381, 0.384, 0.323, 0.314]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [50363.0, 50900.0, 51437.0, 51974.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.037152837487
varray: [0.381, 0.384, 0.323, 0.314]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [50363.0, 50900.0, 51437.0, 51974.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1293.2435
varray: [1292.692, 1292.989, 1293.314, 1293.979]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [50363.0, 50900.0, 51437.0, 51974.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.552223686562
varray: [1292.692, 1292.989, 1293.314, 1293.979]
inst: std_all
operation: std_all(baseline_shift)
tarray: [50363.0, 50900.0, 51437.0, 51974.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.80025
varray: [1.912, 1.91, 1.707, 1.672]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [50363.0, 50900.0, 51437.0, 51974.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.128681454245
varray: [1.912, 1.91, 1.707, 1.672]
inst: std_all
operation: std_all(slope_shift)
tarray: [50363.0, 50900.0, 51437.0, 51974.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0809135340966
varray: [-9.635, -9.823, -9.686, -9.746]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50363.0, 50900.0, 51437.0, 51974.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.992332437577
varray: [-94.839, -96.679, -96.893, -96.871]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50363.0, 50900.0, 51437.0, 51974.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.77725
varray: [108.492, 108.238, 108.798, 109.581]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [50363.0, 50900.0, 51437.0, 51974.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.582695675289
varray: [108.492, 108.238, 108.798, 109.581]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [50363.0, 50900.0, 51437.0, 51974.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.77275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3505
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1293.2435
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.80025
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.2737
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.2737
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -8.98422754353
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -8.22885
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [-9.635, -9.823, -9.686, -9.746]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [50363.0, 50900.0, 51437.0, 51974.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-94.839, -96.679, -96.893, -96.871]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [50363.0, 50900.0, 51437.0, 51974.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.7225
varray: [-9.635, -9.823, -9.686, -9.746]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50363.0, 50900.0, 51437.0, 51974.0]
source_col: d(18_16)Mean
dD_mean | -96.3205
varray: [-94.839, -96.679, -96.893, -96.871]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50363.0, 50900.0, 51437.0, 51974.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0809135340966
varray: [-9.635, -9.823, -9.686, -9.746]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50363.0, 50900.0, 51437.0, 51974.0]
source_col: d(18_16)Mean
dD_stdev | 0.992332437577
varray: [-94.839, -96.679, -96.893, -96.871]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50363.0, 50900.0, 51437.0, 51974.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.7225
varray: [-9.635, -9.823, -9.686, -9.746]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50363.0, 50900.0, 51437.0, 51974.0]
source_col: d(18_16)Mean
measured_dD_mean | -96.3205
varray: [-94.839, -96.679, -96.893, -96.871]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50363.0, 50900.0, 51437.0, 51974.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0809135340966
varray: [-9.635, -9.823, -9.686, -9.746]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50363.0, 50900.0, 51437.0, 51974.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.992332437577
varray: [-94.839, -96.679, -96.893, -96.871]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50363.0, 50900.0, 51437.0, 51974.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -12.5453480621
actual_array: [ 1.93 -15.84]
varray: [-9.635, -9.823, -9.686, -9.746]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -9.7225
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [50363.0, 50900.0, 51437.0, 51974.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -108.823963778
actual_array: [ 18.1 -119. ]
varray: [-94.839, -96.679, -96.893, -96.871]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -96.3205
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [50363.0, 50900.0, 51437.0, 51974.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.06415
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -8.5934
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 19 - Sample 19 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 19
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.63575
varray: [2.671, 2.633, 2.555, 2.684]
inst: mean_all
operation: mean_all(residuals)
tarray: [52562.0, 53099.0, 53635.0, 54172.0]
source_col: residuals
SAMP_residuals_STD | 0.0580193933095
varray: [2.671, 2.633, 2.555, 2.684]
inst: std_all
operation: std_all(residuals)
tarray: [52562.0, 53099.0, 53635.0, 54172.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.1585
varray: [0.358, 0.214, -0.127, 0.189]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [52562.0, 53099.0, 53635.0, 54172.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.204386072585
varray: [0.358, 0.214, -0.127, 0.189]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [52562.0, 53099.0, 53635.0, 54172.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1285.90075
varray: [1287.857, 1284.385, 1284.442, 1286.919]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [52562.0, 53099.0, 53635.0, 54172.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.75965875006
varray: [1287.857, 1284.385, 1284.442, 1286.919]
inst: std_all
operation: std_all(baseline_shift)
tarray: [52562.0, 53099.0, 53635.0, 54172.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.488
varray: [1.277, 1.384, 1.859, 1.432]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [52562.0, 53099.0, 53635.0, 54172.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.255678183139
varray: [1.277, 1.384, 1.859, 1.432]
inst: std_all
operation: std_all(slope_shift)
tarray: [52562.0, 53099.0, 53635.0, 54172.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0399113601205
varray: [-9.309, -9.386, -9.307, -9.365]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52562.0, 53099.0, 53635.0, 54172.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.436655088905
varray: [-88.153, -87.665, -87.237, -87.235]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52562.0, 53099.0, 53635.0, 54172.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 103.157
varray: [105.56, 102.483, 100.437, 104.148]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [52562.0, 53099.0, 53635.0, 54172.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.20674466126
varray: [105.56, 102.483, 100.437, 104.148]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [52562.0, 53099.0, 53635.0, 54172.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.63575
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.1585
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1285.90075
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.488
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.4107
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.4107
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -13.481265079
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -13.8491
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [-9.309, -9.386, -9.307, -9.365]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [52562.0, 53099.0, 53635.0, 54172.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-88.153, -87.665, -87.237, -87.235]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [52562.0, 53099.0, 53635.0, 54172.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.34175
varray: [-9.309, -9.386, -9.307, -9.365]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [52562.0, 53099.0, 53635.0, 54172.0]
source_col: d(18_16)Mean
dD_mean | -87.5725
varray: [-88.153, -87.665, -87.237, -87.235]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [52562.0, 53099.0, 53635.0, 54172.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0399113601205
varray: [-9.309, -9.386, -9.307, -9.365]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52562.0, 53099.0, 53635.0, 54172.0]
source_col: d(18_16)Mean
dD_stdev | 0.436655088905
varray: [-88.153, -87.665, -87.237, -87.235]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52562.0, 53099.0, 53635.0, 54172.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.34175
varray: [-9.309, -9.386, -9.307, -9.365]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [52562.0, 53099.0, 53635.0, 54172.0]
source_col: d(18_16)Mean
measured_dD_mean | -87.5725
varray: [-88.153, -87.665, -87.237, -87.235]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [52562.0, 53099.0, 53635.0, 54172.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0399113601205
varray: [-9.309, -9.386, -9.307, -9.365]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52562.0, 53099.0, 53635.0, 54172.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.436655088905
varray: [-88.153, -87.665, -87.237, -87.235]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52562.0, 53099.0, 53635.0, 54172.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -12.1631419364
actual_array: [ 1.93 -15.84]
varray: [-9.309, -9.386, -9.307, -9.365]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -9.34175
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [52562.0, 53099.0, 53635.0, 54172.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -100.59980265
actual_array: [ 18.1 -119. ]
varray: [-88.153, -87.665, -87.237, -87.235]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -87.5725
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [52562.0, 53099.0, 53635.0, 54172.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.20115
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -15.93615
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 20 - Sample 20 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 20
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.7605
varray: [2.763, 2.706, 2.753, 2.82]
inst: mean_all
operation: mean_all(residuals)
tarray: [54761.0, 55297.0, 55835.0, 56372.0]
source_col: residuals
SAMP_residuals_STD | 0.0468081189539
varray: [2.763, 2.706, 2.753, 2.82]
inst: std_all
operation: std_all(residuals)
tarray: [54761.0, 55297.0, 55835.0, 56372.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.176
varray: [0.049, 0.153, 0.158, 0.344]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [54761.0, 55297.0, 55835.0, 56372.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.122754497542
varray: [0.049, 0.153, 0.158, 0.344]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [54761.0, 55297.0, 55835.0, 56372.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1291.832
varray: [1291.86, 1290.954, 1291.883, 1292.631]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [54761.0, 55297.0, 55835.0, 56372.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.686214252257
varray: [1291.86, 1290.954, 1291.883, 1292.631]
inst: std_all
operation: std_all(baseline_shift)
tarray: [54761.0, 55297.0, 55835.0, 56372.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.80075
varray: [2.112, 1.759, 1.802, 1.53]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [54761.0, 55297.0, 55835.0, 56372.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.239392808302
varray: [2.112, 1.759, 1.802, 1.53]
inst: std_all
operation: std_all(slope_shift)
tarray: [54761.0, 55297.0, 55835.0, 56372.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0241246761636
varray: [-8.392, -8.377, -8.335, -8.368]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54761.0, 55297.0, 55835.0, 56372.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.550087568786
varray: [-78.632, -77.981, -77.563, -77.398]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54761.0, 55297.0, 55835.0, 56372.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 107.56525
varray: [106.595, 106.454, 107.789, 109.423]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [54761.0, 55297.0, 55835.0, 56372.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.37568900919
varray: [106.595, 106.454, 107.789, 109.423]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [54761.0, 55297.0, 55835.0, 56372.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.7605
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.176
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1291.832
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.80075
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.28595
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.28595
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -9.38633491441
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -9.44085
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [-8.392, -8.377, -8.335, -8.368]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [54761.0, 55297.0, 55835.0, 56372.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-78.632, -77.981, -77.563, -77.398]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [54761.0, 55297.0, 55835.0, 56372.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.368
varray: [-8.392, -8.377, -8.335, -8.368]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [54761.0, 55297.0, 55835.0, 56372.0]
source_col: d(18_16)Mean
dD_mean | -77.8935
varray: [-78.632, -77.981, -77.563, -77.398]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [54761.0, 55297.0, 55835.0, 56372.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0241246761636
varray: [-8.392, -8.377, -8.335, -8.368]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54761.0, 55297.0, 55835.0, 56372.0]
source_col: d(18_16)Mean
dD_stdev | 0.550087568786
varray: [-78.632, -77.981, -77.563, -77.398]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54761.0, 55297.0, 55835.0, 56372.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.368
varray: [-8.392, -8.377, -8.335, -8.368]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [54761.0, 55297.0, 55835.0, 56372.0]
source_col: d(18_16)Mean
measured_dD_mean | -77.8935
varray: [-78.632, -77.981, -77.563, -77.398]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [54761.0, 55297.0, 55835.0, 56372.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0241246761636
varray: [-8.392, -8.377, -8.335, -8.368]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54761.0, 55297.0, 55835.0, 56372.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.550087568786
varray: [-78.632, -77.981, -77.563, -77.398]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54761.0, 55297.0, 55835.0, 56372.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -11.185667964
actual_array: [ 1.93 -15.84]
varray: [-8.392, -8.377, -8.335, -8.368]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -8.368
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [54761.0, 55297.0, 55835.0, 56372.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -91.500390722
actual_array: [ 18.1 -119. ]
varray: [-78.632, -77.981, -77.563, -77.398]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -77.8935
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [54761.0, 55297.0, 55835.0, 56372.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.0764
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.0049
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 21 - Sample 21 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 21
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.089
varray: [0.0, 2.79, 2.77, 2.796]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [56962.0, 57499.0, 58036.0, 58573.0]
source_col: residuals
SAMP_residuals_STD | 1.3927110253
varray: [0.0, 2.79, 2.77, 2.796]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [56962.0, 57499.0, 58036.0, 58573.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.2635
varray: [0.0, 0.298, 0.395, 0.361]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [56962.0, 57499.0, 58036.0, 58573.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.180204513447
varray: [0.0, 0.298, 0.395, 0.361]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [56962.0, 57499.0, 58036.0, 58573.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 969.3125
varray: [0.0, 1292.165, 1292.525, 1292.56]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [56962.0, 57499.0, 58036.0, 58573.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 646.208357994
varray: [0.0, 1292.165, 1292.525, 1292.56]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [56962.0, 57499.0, 58036.0, 58573.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.4155
varray: [0.0, 2.078, 1.485, 2.099]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [56962.0, 57499.0, 58036.0, 58573.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.985655281188
varray: [0.0, 2.078, 1.485, 2.099]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [56962.0, 57499.0, 58036.0, 58573.0]
source_col: slope_shift
SAMP_d18O_STD | 5.9307232274
varray: [0.0, -11.857, -11.832, -11.895]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56962.0, 57499.0, 58036.0, 58573.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 52.0258230393
varray: [0.0, -103.508, -104.014, -104.621]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56962.0, 57499.0, 58036.0, 58573.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 81.792
varray: [0.0, 108.34, 109.757, 109.071]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [56962.0, 57499.0, 58036.0, 58573.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 54.5310695353
varray: [0.0, 108.34, 109.757, 109.071]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [56962.0, 57499.0, 58036.0, 58573.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.089
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.2635
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 969.3125
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.4155
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.95745
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.95745
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -31.4283838566
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -35.2141
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [0.0, -11.857, -11.832, -11.895]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [56962.0, 57499.0, 58036.0, 58573.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [0.0, -103.508, -104.014, -104.621]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [56962.0, 57499.0, 58036.0, 58573.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.896
varray: [0.0, -11.857, -11.832, -11.895]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56962.0, 57499.0, 58036.0, 58573.0]
source_col: d(18_16)Mean
dD_mean | -78.03575
varray: [0.0, -103.508, -104.014, -104.621]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56962.0, 57499.0, 58036.0, 58573.0]
source_col: d(D_H)Mean
d18O_stdev | 5.9307232274
varray: [0.0, -11.857, -11.832, -11.895]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56962.0, 57499.0, 58036.0, 58573.0]
source_col: d(18_16)Mean
dD_stdev | 52.0258230393
varray: [0.0, -103.508, -104.014, -104.621]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56962.0, 57499.0, 58036.0, 58573.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.896
varray: [0.0, -11.857, -11.832, -11.895]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56962.0, 57499.0, 58036.0, 58573.0]
source_col: d(18_16)Mean
measured_dD_mean | -78.03575
varray: [0.0, -103.508, -104.014, -104.621]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56962.0, 57499.0, 58036.0, 58573.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 5.9307232274
varray: [0.0, -11.857, -11.832, -11.895]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56962.0, 57499.0, 58036.0, 58573.0]
source_col: d(18_16)Mean
measured_dD_stdev | 52.0258230393
varray: [0.0, -103.508, -104.014, -104.621]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56962.0, 57499.0, 58036.0, 58573.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -11.7156872271
actual_array: [ 1.93 -15.84]
varray: [0.0, -11.857, -11.832, -11.895]
fit_rsq: 1.0
fit_intercept: 2.77505289814
except: True
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -8.896
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [56962.0, 57499.0, 58036.0, 58573.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -91.6341226516
actual_array: [ 18.1 -119. ]
varray: [0.0, -103.508, -104.014, -104.621]
fit_rsq: 1.0
fit_intercept: 19.4350184537
except: True
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -78.03575
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [56962.0, 57499.0, 58036.0, 58573.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.7479
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -332.5244
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 22 - Sample 22 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 22
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.08625
varray: [0.0, 2.78, 2.749, 2.816]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [59160.0, 59696.0, 60233.0, 60770.0]
source_col: residuals
SAMP_residuals_STD | 1.39110276999
varray: [0.0, 2.78, 2.749, 2.816]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [59160.0, 59696.0, 60233.0, 60770.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.30875
varray: [0.0, 0.485, 0.076, 0.674]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [59160.0, 59696.0, 60233.0, 60770.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.323507212697
varray: [0.0, 0.485, 0.076, 0.674]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [59160.0, 59696.0, 60233.0, 60770.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 968.82775
varray: [0.0, 1291.487, 1292.429, 1291.395]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [59160.0, 59696.0, 60233.0, 60770.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 645.885335684
varray: [0.0, 1291.487, 1292.429, 1291.395]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [59160.0, 59696.0, 60233.0, 60770.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.31325
varray: [0.0, 2.023, 1.079, 2.151]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [59160.0, 59696.0, 60233.0, 60770.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.997508354184
varray: [0.0, 2.023, 1.079, 2.151]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [59160.0, 59696.0, 60233.0, 60770.0]
source_col: slope_shift
SAMP_d18O_STD | 3.83406268424
varray: [0.0, -7.699, -7.653, -7.652]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59160.0, 59696.0, 60233.0, 60770.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 30.767912803
varray: [0.0, -62.364, -61.144, -61.065]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59160.0, 59696.0, 60233.0, 60770.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 81.77025
varray: [0.0, 109.013, 108.537, 109.531]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [59160.0, 59696.0, 60233.0, 60770.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 54.5150112622
varray: [0.0, 109.013, 108.537, 109.531]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [59160.0, 59696.0, 60233.0, 60770.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.08625
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.30875
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 968.82775
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.31325
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.9602
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.9602
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -31.5186528582
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -35.23585
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [0.0, -7.699, -7.653, -7.652]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [59160.0, 59696.0, 60233.0, 60770.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [0.0, -62.364, -61.144, -61.065]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [59160.0, 59696.0, 60233.0, 60770.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.751
varray: [0.0, -7.699, -7.653, -7.652]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [59160.0, 59696.0, 60233.0, 60770.0]
source_col: d(18_16)Mean
dD_mean | -46.14325
varray: [0.0, -62.364, -61.144, -61.065]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [59160.0, 59696.0, 60233.0, 60770.0]
source_col: d(D_H)Mean
d18O_stdev | 3.83406268424
varray: [0.0, -7.699, -7.653, -7.652]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59160.0, 59696.0, 60233.0, 60770.0]
source_col: d(18_16)Mean
dD_stdev | 30.767912803
varray: [0.0, -62.364, -61.144, -61.065]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59160.0, 59696.0, 60233.0, 60770.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.751
varray: [0.0, -7.699, -7.653, -7.652]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [59160.0, 59696.0, 60233.0, 60770.0]
source_col: d(18_16)Mean
measured_dD_mean | -46.14325
varray: [0.0, -62.364, -61.144, -61.065]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [59160.0, 59696.0, 60233.0, 60770.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 3.83406268424
varray: [0.0, -7.699, -7.653, -7.652]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59160.0, 59696.0, 60233.0, 60770.0]
source_col: d(18_16)Mean
measured_dD_stdev | 30.767912803
varray: [0.0, -62.364, -61.144, -61.065]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59160.0, 59696.0, 60233.0, 60770.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.5586596092
actual_array: [ 1.93 -15.84]
varray: [0.0, -7.699, -7.653, -7.652]
fit_rsq: 1.0
fit_intercept: 2.77505289814
except: True
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -5.751
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [59160.0, 59696.0, 60233.0, 60770.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -61.6513770241
actual_array: [ 18.1 -119. ]
varray: [0.0, -62.364, -61.144, -61.065]
fit_rsq: 1.0
fit_intercept: 19.4350184537
except: True
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -46.14325
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [59160.0, 59696.0, 60233.0, 60770.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.75065
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -333.00915
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 23 - Sample 23 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 23
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.71825
varray: [3.671, 3.713, 3.793, 3.696]
inst: mean_all
operation: mean_all(residuals)
tarray: [61356.0, 61894.0, 62431.0, 62967.0]
source_col: residuals
SAMP_residuals_STD | 0.0527343973765
varray: [3.671, 3.713, 3.793, 3.696]
inst: std_all
operation: std_all(residuals)
tarray: [61356.0, 61894.0, 62431.0, 62967.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -6.371
varray: [-6.003, -6.532, -6.377, -6.572]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [61356.0, 61894.0, 62431.0, 62967.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.259346614913
varray: [-6.003, -6.532, -6.377, -6.572]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [61356.0, 61894.0, 62431.0, 62967.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1739.53025
varray: [1732.762, 1740.72, 1740.737, 1743.902]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [61356.0, 61894.0, 62431.0, 62967.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 4.75370580881
varray: [1732.762, 1740.72, 1740.737, 1743.902]
inst: std_all
operation: std_all(baseline_shift)
tarray: [61356.0, 61894.0, 62431.0, 62967.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -32.336
varray: [-32.07, -32.226, -32.393, -32.655]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [61356.0, 61894.0, 62431.0, 62967.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.250243881044
varray: [-32.07, -32.226, -32.393, -32.655]
inst: std_all
operation: std_all(slope_shift)
tarray: [61356.0, 61894.0, 62431.0, 62967.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0376519145153
varray: [-1.648, -1.723, -1.713, -1.73]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61356.0, 61894.0, 62431.0, 62967.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.28157277625
varray: [-82.813, -86.743, -87.423, -87.735]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61356.0, 61894.0, 62431.0, 62967.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 117.7625
varray: [117.479, 117.136, 118.809, 117.626]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [61356.0, 61894.0, 62431.0, 62967.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.72724800905
varray: [117.479, 117.136, 118.809, 117.626]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [61356.0, 61894.0, 62431.0, 62967.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.71825
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 6.371
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1739.53025
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 32.336
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.6718
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.6718
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 22.0518964697
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.7564
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [-1.648, -1.723, -1.713, -1.73]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [61356.0, 61894.0, 62431.0, 62967.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-82.813, -86.743, -87.423, -87.735]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [61356.0, 61894.0, 62431.0, 62967.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.7035
varray: [-1.648, -1.723, -1.713, -1.73]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [61356.0, 61894.0, 62431.0, 62967.0]
source_col: d(18_16)Mean
dD_mean | -86.1785
varray: [-82.813, -86.743, -87.423, -87.735]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [61356.0, 61894.0, 62431.0, 62967.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0376519145153
varray: [-1.648, -1.723, -1.713, -1.73]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61356.0, 61894.0, 62431.0, 62967.0]
source_col: d(18_16)Mean
dD_stdev | 2.28157277625
varray: [-82.813, -86.743, -87.423, -87.735]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61356.0, 61894.0, 62431.0, 62967.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.7035
varray: [-1.648, -1.723, -1.713, -1.73]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [61356.0, 61894.0, 62431.0, 62967.0]
source_col: d(18_16)Mean
measured_dD_mean | -86.1785
varray: [-82.813, -86.743, -87.423, -87.735]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [61356.0, 61894.0, 62431.0, 62967.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0376519145153
varray: [-1.648, -1.723, -1.713, -1.73]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61356.0, 61894.0, 62431.0, 62967.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.28157277625
varray: [-82.813, -86.743, -87.423, -87.735]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61356.0, 61894.0, 62431.0, 62967.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.49568050479
actual_array: [ 1.93 -15.84]
varray: [-1.648, -1.723, -1.713, -1.73]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -1.7035
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [61356.0, 61894.0, 62431.0, 62967.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -99.2892767461
actual_array: [ 18.1 -119. ]
varray: [-82.813, -86.743, -87.423, -87.735]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -86.1785
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [61356.0, 61894.0, 62431.0, 62967.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1298.11865
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 437.69335
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 24 - Sample 24 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 24
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.75725
varray: [2.762, 2.78, 2.719, 2.768]
inst: mean_all
operation: mean_all(residuals)
tarray: [63555.0, 64092.0, 64628.0, 65166.0]
source_col: residuals
SAMP_residuals_STD | 0.0265753645318
varray: [2.762, 2.78, 2.719, 2.768]
inst: std_all
operation: std_all(residuals)
tarray: [63555.0, 64092.0, 64628.0, 65166.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3705
varray: [0.238, 0.445, 0.464, 0.335]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [63555.0, 64092.0, 64628.0, 65166.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.105053954392
varray: [0.238, 0.445, 0.464, 0.335]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [63555.0, 64092.0, 64628.0, 65166.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1294.31375
varray: [1298.73, 1293.879, 1292.894, 1291.752]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [63555.0, 64092.0, 64628.0, 65166.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.06977332225
varray: [1298.73, 1293.879, 1292.894, 1291.752]
inst: std_all
operation: std_all(baseline_shift)
tarray: [63555.0, 64092.0, 64628.0, 65166.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.91075
varray: [1.81, 1.981, 2.107, 1.745]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [63555.0, 64092.0, 64628.0, 65166.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.16439054109
varray: [1.81, 1.981, 2.107, 1.745]
inst: std_all
operation: std_all(slope_shift)
tarray: [63555.0, 64092.0, 64628.0, 65166.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0335298374586
varray: [-11.005, -11.052, -11.085, -11.061]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63555.0, 64092.0, 64628.0, 65166.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0681028633759
varray: [-101.333, -101.226, -101.171, -101.222]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63555.0, 64092.0, 64628.0, 65166.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.41825
varray: [108.236, 108.649, 108.011, 108.777]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [63555.0, 64092.0, 64628.0, 65166.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.356377491807
varray: [108.236, 108.649, 108.011, 108.777]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [63555.0, 64092.0, 64628.0, 65166.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.75725
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3705
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1294.31375
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.91075
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.2892
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.2892
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -9.49301646178
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -8.58785
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [-11.005, -11.052, -11.085, -11.061]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [63555.0, 64092.0, 64628.0, 65166.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-101.333, -101.226, -101.171, -101.222]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [63555.0, 64092.0, 64628.0, 65166.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -11.05075
varray: [-11.005, -11.052, -11.085, -11.061]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63555.0, 64092.0, 64628.0, 65166.0]
source_col: d(18_16)Mean
dD_mean | -101.238
varray: [-101.333, -101.226, -101.171, -101.222]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63555.0, 64092.0, 64628.0, 65166.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0335298374586
varray: [-11.005, -11.052, -11.085, -11.061]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63555.0, 64092.0, 64628.0, 65166.0]
source_col: d(18_16)Mean
dD_stdev | 0.0681028633759
varray: [-101.333, -101.226, -101.171, -101.222]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63555.0, 64092.0, 64628.0, 65166.0]
source_col: d(D_H)Mean
measured_d18O_mean | -11.05075
varray: [-11.005, -11.052, -11.085, -11.061]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63555.0, 64092.0, 64628.0, 65166.0]
source_col: d(18_16)Mean
measured_dD_mean | -101.238
varray: [-101.333, -101.226, -101.171, -101.222]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63555.0, 64092.0, 64628.0, 65166.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0335298374586
varray: [-11.005, -11.052, -11.085, -11.061]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63555.0, 64092.0, 64628.0, 65166.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0681028633759
varray: [-101.333, -101.226, -101.171, -101.222]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63555.0, 64092.0, 64628.0, 65166.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -13.8786777707
actual_array: [ 1.93 -15.84]
varray: [-11.005, -11.052, -11.085, -11.061]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -11.05075
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [63555.0, 64092.0, 64628.0, 65166.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -113.446999025
actual_array: [ 18.1 -119. ]
varray: [-101.333, -101.226, -101.171, -101.222]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -101.238
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [63555.0, 64092.0, 64628.0, 65166.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.07965
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -7.52315
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 25 - Sample 25 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 25
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.78325
varray: [2.785, 2.757, 2.761, 2.83]
inst: mean_all
operation: mean_all(residuals)
tarray: [65753.0, 66290.0, 66828.0, 67365.0]
source_col: residuals
SAMP_residuals_STD | 0.0335298374586
varray: [2.785, 2.757, 2.761, 2.83]
inst: std_all
operation: std_all(residuals)
tarray: [65753.0, 66290.0, 66828.0, 67365.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.3335
varray: [0.236, 0.35, 0.391, 0.357]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [65753.0, 66290.0, 66828.0, 67365.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.067421559361
varray: [0.236, 0.35, 0.391, 0.357]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [65753.0, 66290.0, 66828.0, 67365.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1291.50475
varray: [1291.449, 1291.315, 1291.613, 1291.642]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [65753.0, 66290.0, 66828.0, 67365.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.152390671193
varray: [1291.449, 1291.315, 1291.613, 1291.642]
inst: std_all
operation: std_all(baseline_shift)
tarray: [65753.0, 66290.0, 66828.0, 67365.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.57225
varray: [1.242, 1.55, 1.789, 1.708]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [65753.0, 66290.0, 66828.0, 67365.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.241501380258
varray: [1.242, 1.55, 1.789, 1.708]
inst: std_all
operation: std_all(slope_shift)
tarray: [65753.0, 66290.0, 66828.0, 67365.0]
source_col: slope_shift
SAMP_d18O_STD | 0.205498580693
varray: [-5.162, -4.801, -4.758, -4.713]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65753.0, 66290.0, 66828.0, 67365.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.00114973203
varray: [-72.855, -69.605, -68.799, -68.49]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65753.0, 66290.0, 66828.0, 67365.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.194
varray: [108.254, 107.883, 108.373, 108.266]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [65753.0, 66290.0, 66828.0, 67365.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.214123017601
varray: [108.254, 107.883, 108.373, 108.266]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [65753.0, 66290.0, 66828.0, 67365.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.78325
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.3335
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1291.50475
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.57225
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.2632
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.2632
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -8.63956408278
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -8.8121
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [-5.162, -4.801, -4.758, -4.713]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [65753.0, 66290.0, 66828.0, 67365.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-72.855, -69.605, -68.799, -68.49]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [65753.0, 66290.0, 66828.0, 67365.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.8585
varray: [-5.162, -4.801, -4.758, -4.713]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65753.0, 66290.0, 66828.0, 67365.0]
source_col: d(18_16)Mean
dD_mean | -69.93725
varray: [-72.855, -69.605, -68.799, -68.49]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65753.0, 66290.0, 66828.0, 67365.0]
source_col: d(D_H)Mean
d18O_stdev | 0.205498580693
varray: [-5.162, -4.801, -4.758, -4.713]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65753.0, 66290.0, 66828.0, 67365.0]
source_col: d(18_16)Mean
dD_stdev | 2.00114973203
varray: [-72.855, -69.605, -68.799, -68.49]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65753.0, 66290.0, 66828.0, 67365.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.8585
varray: [-5.162, -4.801, -4.758, -4.713]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65753.0, 66290.0, 66828.0, 67365.0]
source_col: d(18_16)Mean
measured_dD_mean | -69.93725
varray: [-72.855, -69.605, -68.799, -68.49]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65753.0, 66290.0, 66828.0, 67365.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.205498580693
varray: [-5.162, -4.801, -4.758, -4.713]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65753.0, 66290.0, 66828.0, 67365.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.00114973203
varray: [-72.855, -69.605, -68.799, -68.49]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65753.0, 66290.0, 66828.0, 67365.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.66274636629
actual_array: [ 1.93 -15.84]
varray: [-5.162, -4.801, -4.758, -4.713]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -4.8585
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [65753.0, 66290.0, 66828.0, 67365.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -84.0205688234
actual_array: [ 18.1 -119. ]
varray: [-72.855, -69.605, -68.799, -68.49]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -69.93725
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [65753.0, 66290.0, 66828.0, 67365.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.05365
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -10.33215
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 26 - Sample 26 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 26
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.73525
varray: [2.76, 2.737, 2.703, 2.741]
inst: mean_all
operation: mean_all(residuals)
tarray: [67953.0, 68491.0, 69028.0, 69564.0]
source_col: residuals
SAMP_residuals_STD | 0.0237258649298
varray: [2.76, 2.737, 2.703, 2.741]
inst: std_all
operation: std_all(residuals)
tarray: [67953.0, 68491.0, 69028.0, 69564.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.20225
varray: [0.475, 0.118, 0.122, 0.094]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [67953.0, 68491.0, 69028.0, 69564.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.182253257858
varray: [0.475, 0.118, 0.122, 0.094]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [67953.0, 68491.0, 69028.0, 69564.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1290.017
varray: [1289.687, 1290.347, 1289.831, 1290.203]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [67953.0, 68491.0, 69028.0, 69564.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.309295974756
varray: [1289.687, 1290.347, 1289.831, 1290.203]
inst: std_all
operation: std_all(baseline_shift)
tarray: [67953.0, 68491.0, 69028.0, 69564.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.5215
varray: [1.402, 1.287, 1.619, 1.778]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [67953.0, 68491.0, 69028.0, 69564.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.219521449218
varray: [1.402, 1.287, 1.619, 1.778]
inst: std_all
operation: std_all(slope_shift)
tarray: [67953.0, 68491.0, 69028.0, 69564.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0424578222082
varray: [-4.102, -4.13, -4.17, -4.07]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67953.0, 68491.0, 69028.0, 69564.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.181330269582
varray: [-69.415, -69.476, -69.729, -69.3]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67953.0, 68491.0, 69028.0, 69564.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 106.94675
varray: [107.637, 106.454, 106.892, 106.804]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [67953.0, 68491.0, 69028.0, 69564.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.497534839651
varray: [107.637, 106.454, 106.892, 106.804]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [67953.0, 68491.0, 69028.0, 69564.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.73525
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.20225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1290.017
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.5215
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.3112
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.3112
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -10.2151684748
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -10.05935
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [-4.102, -4.13, -4.17, -4.07]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [67953.0, 68491.0, 69028.0, 69564.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-69.415, -69.476, -69.729, -69.3]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [67953.0, 68491.0, 69028.0, 69564.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.118
varray: [-4.102, -4.13, -4.17, -4.07]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [67953.0, 68491.0, 69028.0, 69564.0]
source_col: d(18_16)Mean
dD_mean | -69.48
varray: [-69.415, -69.476, -69.729, -69.3]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [67953.0, 68491.0, 69028.0, 69564.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0424578222082
varray: [-4.102, -4.13, -4.17, -4.07]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67953.0, 68491.0, 69028.0, 69564.0]
source_col: d(18_16)Mean
dD_stdev | 0.181330269582
varray: [-69.415, -69.476, -69.729, -69.3]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67953.0, 68491.0, 69028.0, 69564.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.118
varray: [-4.102, -4.13, -4.17, -4.07]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [67953.0, 68491.0, 69028.0, 69564.0]
source_col: d(18_16)Mean
measured_dD_mean | -69.48
varray: [-69.415, -69.476, -69.729, -69.3]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [67953.0, 68491.0, 69028.0, 69564.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0424578222082
varray: [-4.102, -4.13, -4.17, -4.07]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67953.0, 68491.0, 69028.0, 69564.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.181330269582
varray: [-69.415, -69.476, -69.729, -69.3]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67953.0, 68491.0, 69028.0, 69564.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.91941442637
actual_array: [ 1.93 -15.84]
varray: [-4.102, -4.13, -4.17, -4.07]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -4.118
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [67953.0, 68491.0, 69028.0, 69564.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -83.5906994047
actual_array: [ 18.1 -119. ]
varray: [-69.415, -69.476, -69.729, -69.3]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -69.48
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [67953.0, 68491.0, 69028.0, 69564.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.10165
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -11.8199
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 27 - Sample 27 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 27
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.6085
varray: [2.59, 2.596, 2.581, 2.667]
inst: mean_all
operation: mean_all(residuals)
tarray: [70153.0, 70690.0, 71227.0, 71763.0]
source_col: residuals
SAMP_residuals_STD | 0.0394841740448
varray: [2.59, 2.596, 2.581, 2.667]
inst: std_all
operation: std_all(residuals)
tarray: [70153.0, 70690.0, 71227.0, 71763.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.4265
varray: [-0.291, -0.36, -0.408, -0.647]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [70153.0, 70690.0, 71227.0, 71763.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.154644754195
varray: [-0.291, -0.36, -0.408, -0.647]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [70153.0, 70690.0, 71227.0, 71763.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1709.83
varray: [1697.542, 1706.821, 1706.552, 1728.405]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [70153.0, 70690.0, 71227.0, 71763.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 13.1126520836
varray: [1697.542, 1706.821, 1706.552, 1728.405]
inst: std_all
operation: std_all(baseline_shift)
tarray: [70153.0, 70690.0, 71227.0, 71763.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -14.07175
varray: [-13.641, -13.652, -14.035, -14.959]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [70153.0, 70690.0, 71227.0, 71763.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.619219602941
varray: [-13.641, -13.652, -14.035, -14.959]
inst: std_all
operation: std_all(slope_shift)
tarray: [70153.0, 70690.0, 71227.0, 71763.0]
source_col: slope_shift
SAMP_d18O_STD | 0.103225884997
varray: [-6.706, -6.923, -6.923, -6.823]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70153.0, 70690.0, 71227.0, 71763.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.300664126006
varray: [-72.508, -72.783, -72.778, -72.146]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70153.0, 70690.0, 71227.0, 71763.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 100.49725
varray: [100.055, 99.374, 99.466, 103.094]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [70153.0, 70690.0, 71227.0, 71763.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.75725759352
varray: [100.055, 99.374, 99.466, 103.094]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [70153.0, 70690.0, 71227.0, 71763.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.6085
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.4265
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1709.83
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 14.07175
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.43795
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.43795
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -14.3757488224
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -16.50885
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [-6.706, -6.923, -6.923, -6.823]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [70153.0, 70690.0, 71227.0, 71763.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-72.508, -72.783, -72.778, -72.146]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [70153.0, 70690.0, 71227.0, 71763.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.84375
varray: [-6.706, -6.923, -6.923, -6.823]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [70153.0, 70690.0, 71227.0, 71763.0]
source_col: d(18_16)Mean
dD_mean | -72.55375
varray: [-72.508, -72.783, -72.778, -72.146]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [70153.0, 70690.0, 71227.0, 71763.0]
source_col: d(D_H)Mean
d18O_stdev | 0.103225884997
varray: [-6.706, -6.923, -6.923, -6.823]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70153.0, 70690.0, 71227.0, 71763.0]
source_col: d(18_16)Mean
dD_stdev | 0.300664126006
varray: [-72.508, -72.783, -72.778, -72.146]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70153.0, 70690.0, 71227.0, 71763.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.84375
varray: [-6.706, -6.923, -6.923, -6.823]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [70153.0, 70690.0, 71227.0, 71763.0]
source_col: d(18_16)Mean
measured_dD_mean | -72.55375
varray: [-72.508, -72.783, -72.778, -72.146]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [70153.0, 70690.0, 71227.0, 71763.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.103225884997
varray: [-6.706, -6.923, -6.923, -6.823]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70153.0, 70690.0, 71227.0, 71763.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.300664126006
varray: [-72.508, -72.783, -72.778, -72.146]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70153.0, 70690.0, 71227.0, 71763.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.65558868057
actual_array: [ 1.93 -15.84]
varray: [-6.706, -6.923, -6.923, -6.823]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -6.84375
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [70153.0, 70690.0, 71227.0, 71763.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -86.4803902214
actual_array: [ 18.1 -119. ]
varray: [-72.508, -72.783, -72.778, -72.146]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -72.55375
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [70153.0, 70690.0, 71227.0, 71763.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.2284
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 407.9931
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 28 - Sample 28 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 28
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.62825
varray: [2.691, 2.586, 2.583, 2.653]
inst: mean_all
operation: mean_all(residuals)
tarray: [72352.0, 72888.0, 73424.0, 73962.0]
source_col: residuals
SAMP_residuals_STD | 0.0528606659058
varray: [2.691, 2.586, 2.583, 2.653]
inst: std_all
operation: std_all(residuals)
tarray: [72352.0, 72888.0, 73424.0, 73962.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.22075
varray: [-0.889, -1.312, -1.398, -1.284]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [72352.0, 72888.0, 73424.0, 73962.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.22642346021
varray: [-0.889, -1.312, -1.398, -1.284]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [72352.0, 72888.0, 73424.0, 73962.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 2327.09425
varray: [2317.416, 2334.196, 2321.801, 2334.964]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [72352.0, 72888.0, 73424.0, 73962.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 8.83279564558
varray: [2317.416, 2334.196, 2321.801, 2334.964]
inst: std_all
operation: std_all(baseline_shift)
tarray: [72352.0, 72888.0, 73424.0, 73962.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -38.824
varray: [-38.544, -39.573, -38.511, -38.668]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [72352.0, 72888.0, 73424.0, 73962.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.50388689207
varray: [-38.544, -39.573, -38.511, -38.668]
inst: std_all
operation: std_all(slope_shift)
tarray: [72352.0, 72888.0, 73424.0, 73962.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0297993288515
varray: [-5.339, -5.363, -5.345, -5.405]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72352.0, 72888.0, 73424.0, 73962.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.589172512937
varray: [-62.766, -61.666, -61.563, -61.548]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72352.0, 72888.0, 73424.0, 73962.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 96.56075
varray: [98.556, 96.828, 94.857, 96.002]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [72352.0, 72888.0, 73424.0, 73962.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.55642868452
varray: [98.556, 96.828, 94.857, 96.002]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [72352.0, 72888.0, 73424.0, 73962.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.62825
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.22075
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 2327.09425
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 38.824
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.4182
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.4182
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -13.7274532653
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -20.44535
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [-5.339, -5.363, -5.345, -5.405]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [72352.0, 72888.0, 73424.0, 73962.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-62.766, -61.666, -61.563, -61.548]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [72352.0, 72888.0, 73424.0, 73962.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.363
varray: [-5.339, -5.363, -5.345, -5.405]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72352.0, 72888.0, 73424.0, 73962.0]
source_col: d(18_16)Mean
dD_mean | -61.88575
varray: [-62.766, -61.666, -61.563, -61.548]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72352.0, 72888.0, 73424.0, 73962.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0297993288515
varray: [-5.339, -5.363, -5.345, -5.405]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72352.0, 72888.0, 73424.0, 73962.0]
source_col: d(18_16)Mean
dD_stdev | 0.589172512937
varray: [-62.766, -61.666, -61.563, -61.548]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72352.0, 72888.0, 73424.0, 73962.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.363
varray: [-5.339, -5.363, -5.345, -5.405]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72352.0, 72888.0, 73424.0, 73962.0]
source_col: d(18_16)Mean
measured_dD_mean | -61.88575
varray: [-62.766, -61.666, -61.563, -61.548]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72352.0, 72888.0, 73424.0, 73962.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0297993288515
varray: [-5.339, -5.363, -5.345, -5.405]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72352.0, 72888.0, 73424.0, 73962.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.589172512937
varray: [-62.766, -61.666, -61.563, -61.548]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72352.0, 72888.0, 73424.0, 73962.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.16917575682
actual_array: [ 1.93 -15.84]
varray: [-5.339, -5.363, -5.345, -5.405]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -5.363
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [72352.0, 72888.0, 73424.0, 73962.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -76.4512005889
actual_array: [ 18.1 -119. ]
varray: [-62.766, -61.666, -61.563, -61.548]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -61.88575
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [72352.0, 72888.0, 73424.0, 73962.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.20865
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1025.25735
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 29 - Sample 29 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 29
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.99725
varray: [0.0, 2.636, 2.658, 2.695]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [74547.0, 75084.0, 75621.0, 76157.0]
source_col: residuals
SAMP_residuals_STD | 1.33172253742
varray: [0.0, 2.636, 2.658, 2.695]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [74547.0, 75084.0, 75621.0, 76157.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.081
varray: [0.0, 0.032, 0.082, 0.21]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [74547.0, 75084.0, 75621.0, 76157.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0923832596668
varray: [0.0, 0.032, 0.082, 0.21]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [74547.0, 75084.0, 75621.0, 76157.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 970.26275
varray: [0.0, 1296.241, 1293.246, 1291.564]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [74547.0, 75084.0, 75621.0, 76157.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 646.844725454
varray: [0.0, 1296.241, 1293.246, 1291.564]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [74547.0, 75084.0, 75621.0, 76157.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.6845
varray: [0.0, 0.948, 0.722, 1.068]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [74547.0, 75084.0, 75621.0, 76157.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.478348199537
varray: [0.0, 0.948, 0.722, 1.068]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [74547.0, 75084.0, 75621.0, 76157.0]
source_col: slope_shift
SAMP_d18O_STD | 3.83651251138
varray: [0.0, -7.661, -7.673, -7.685]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74547.0, 75084.0, 75621.0, 76157.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 42.2898286225
varray: [0.0, -84.119, -84.7, -84.912]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74547.0, 75084.0, 75621.0, 76157.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 78.7115
varray: [0.0, 103.695, 105.418, 105.733]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [74547.0, 75084.0, 75621.0, 76157.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 52.4819782052
varray: [0.0, 103.695, 105.418, 105.733]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [74547.0, 75084.0, 75621.0, 76157.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.99725
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.081
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 970.26275
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.6845
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.0492
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.0492
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -34.4400860017
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -38.2946
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [0.0, -7.661, -7.673, -7.685]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [74547.0, 75084.0, 75621.0, 76157.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [0.0, -84.119, -84.7, -84.912]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [74547.0, 75084.0, 75621.0, 76157.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.75475
varray: [0.0, -7.661, -7.673, -7.685]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [74547.0, 75084.0, 75621.0, 76157.0]
source_col: d(18_16)Mean
dD_mean | -63.43275
varray: [0.0, -84.119, -84.7, -84.912]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [74547.0, 75084.0, 75621.0, 76157.0]
source_col: d(D_H)Mean
d18O_stdev | 3.83651251138
varray: [0.0, -7.661, -7.673, -7.685]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74547.0, 75084.0, 75621.0, 76157.0]
source_col: d(18_16)Mean
dD_stdev | 42.2898286225
varray: [0.0, -84.119, -84.7, -84.912]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74547.0, 75084.0, 75621.0, 76157.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.75475
varray: [0.0, -7.661, -7.673, -7.685]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [74547.0, 75084.0, 75621.0, 76157.0]
source_col: d(18_16)Mean
measured_dD_mean | -63.43275
varray: [0.0, -84.119, -84.7, -84.912]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [74547.0, 75084.0, 75621.0, 76157.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 3.83651251138
varray: [0.0, -7.661, -7.673, -7.685]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74547.0, 75084.0, 75621.0, 76157.0]
source_col: d(18_16)Mean
measured_dD_stdev | 42.2898286225
varray: [0.0, -84.119, -84.7, -84.912]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74547.0, 75084.0, 75621.0, 76157.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.56242395056
actual_array: [ 1.93 -15.84]
varray: [0.0, -7.661, -7.673, -7.685]
fit_rsq: 1.0
fit_intercept: 2.77505289814
except: True
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -5.75475
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [74547.0, 75084.0, 75621.0, 76157.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -77.9055647023
actual_array: [ 18.1 -119. ]
varray: [0.0, -84.119, -84.7, -84.912]
fit_rsq: 1.0
fit_intercept: 19.4350184537
except: True
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -63.43275
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [74547.0, 75084.0, 75621.0, 76157.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.83965
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -331.57415
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 30 - Sample 30 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 30
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.03575
varray: [0.0, 2.639, 2.706, 2.798]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [76743.0, 77279.0, 77816.0, 78352.0]
source_col: residuals
SAMP_residuals_STD | 1.3587308723
varray: [0.0, 2.639, 2.706, 2.798]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [76743.0, 77279.0, 77816.0, 78352.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.20425
varray: [0.0, 0.232, 0.144, 0.441]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [76743.0, 77279.0, 77816.0, 78352.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.18454335534
varray: [0.0, 0.232, 0.144, 0.441]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [76743.0, 77279.0, 77816.0, 78352.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 970.752
varray: [0.0, 1291.865, 1295.162, 1295.981]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [76743.0, 77279.0, 77816.0, 78352.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 647.170445044
varray: [0.0, 1291.865, 1295.162, 1295.981]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [76743.0, 77279.0, 77816.0, 78352.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.05825
varray: [0.0, 1.076, 1.711, 1.446]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [76743.0, 77279.0, 77816.0, 78352.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.752028534476
varray: [0.0, 1.076, 1.711, 1.446]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [76743.0, 77279.0, 77816.0, 78352.0]
source_col: slope_shift
SAMP_d18O_STD | 1.15840436233
varray: [0.0, -2.381, -2.302, -2.261]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76743.0, 77279.0, 77816.0, 78352.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 28.311213307
varray: [0.0, -57.412, -56.387, -56.032]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76743.0, 77279.0, 77816.0, 78352.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 79.832
varray: [0.0, 104.812, 106.173, 108.343]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [76743.0, 77279.0, 77816.0, 78352.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 53.2411934064
varray: [0.0, 104.812, 106.173, 108.343]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [76743.0, 77279.0, 77816.0, 78352.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.03575
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.20425
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 970.752
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.05825
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -1.0107
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.0107
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -33.176319979
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -37.1741
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [0.0, -2.381, -2.302, -2.261]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [76743.0, 77279.0, 77816.0, 78352.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [0.0, -57.412, -56.387, -56.032]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [76743.0, 77279.0, 77816.0, 78352.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -1.736
varray: [0.0, -2.381, -2.302, -2.261]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [76743.0, 77279.0, 77816.0, 78352.0]
source_col: d(18_16)Mean
dD_mean | -42.45775
varray: [0.0, -57.412, -56.387, -56.032]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [76743.0, 77279.0, 77816.0, 78352.0]
source_col: d(D_H)Mean
d18O_stdev | 1.15840436233
varray: [0.0, -2.381, -2.302, -2.261]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76743.0, 77279.0, 77816.0, 78352.0]
source_col: d(18_16)Mean
dD_stdev | 28.311213307
varray: [0.0, -57.412, -56.387, -56.032]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76743.0, 77279.0, 77816.0, 78352.0]
source_col: d(D_H)Mean
measured_d18O_mean | -1.736
varray: [0.0, -2.381, -2.302, -2.261]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [76743.0, 77279.0, 77816.0, 78352.0]
source_col: d(18_16)Mean
measured_dD_mean | -42.45775
varray: [0.0, -57.412, -56.387, -56.032]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [76743.0, 77279.0, 77816.0, 78352.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 1.15840436233
varray: [0.0, -2.381, -2.302, -2.261]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76743.0, 77279.0, 77816.0, 78352.0]
source_col: d(18_16)Mean
measured_dD_stdev | 28.311213307
varray: [0.0, -57.412, -56.387, -56.032]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76743.0, 77279.0, 77816.0, 78352.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.52830479655
actual_array: [ 1.93 -15.84]
varray: [0.0, -2.381, -2.302, -2.261]
fit_rsq: 1.0
fit_intercept: 2.77505289814
except: True
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -1.736
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [76743.0, 77279.0, 77816.0, 78352.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -58.186568401
actual_array: [ 18.1 -119. ]
varray: [0.0, -57.412, -56.387, -56.032]
fit_rsq: 1.0
fit_intercept: 19.4350184537
except: True
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -42.45775
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [76743.0, 77279.0, 77816.0, 78352.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.80115
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -331.0849
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 31 - Sample 31 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 31
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.761
varray: [2.722, 2.744, 2.806, 2.772]
inst: mean_all
operation: mean_all(residuals)
tarray: [78938.0, 79475.0, 80012.0, 80549.0]
source_col: residuals
SAMP_residuals_STD | 0.0363134502171
varray: [2.722, 2.744, 2.806, 2.772]
inst: std_all
operation: std_all(residuals)
tarray: [78938.0, 79475.0, 80012.0, 80549.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.204
varray: [-0.008, 0.094, 0.425, 0.305]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [78938.0, 79475.0, 80012.0, 80549.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.196711294372
varray: [-0.008, 0.094, 0.425, 0.305]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [78938.0, 79475.0, 80012.0, 80549.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1291.93025
varray: [1291.901, 1292.306, 1291.95, 1291.564]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [78938.0, 79475.0, 80012.0, 80549.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.303629571463
varray: [1291.901, 1292.306, 1291.95, 1291.564]
inst: std_all
operation: std_all(baseline_shift)
tarray: [78938.0, 79475.0, 80012.0, 80549.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.7935
varray: [1.95, 2.02, 1.764, 1.44]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [78938.0, 79475.0, 80012.0, 80549.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.259246986482
varray: [1.95, 2.02, 1.764, 1.44]
inst: std_all
operation: std_all(slope_shift)
tarray: [78938.0, 79475.0, 80012.0, 80549.0]
source_col: slope_shift
SAMP_d18O_STD | 0.179187053104
varray: [-7.224, -7.506, -7.56, -7.634]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78938.0, 79475.0, 80012.0, 80549.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.294001558953
varray: [-60.002, -60.366, -60.396, -60.721]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78938.0, 79475.0, 80012.0, 80549.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 107.55175
varray: [106.501, 106.394, 108.812, 108.5]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [78938.0, 79475.0, 80012.0, 80549.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.28216857316
varray: [106.501, 106.394, 108.812, 108.5]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [78938.0, 79475.0, 80012.0, 80549.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.761
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.204
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1291.93025
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.7935
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.28545
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.28545
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -9.36992236866
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -9.45435
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [-7.224, -7.506, -7.56, -7.634]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [78938.0, 79475.0, 80012.0, 80549.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-60.002, -60.366, -60.396, -60.721]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [78938.0, 79475.0, 80012.0, 80549.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.481
varray: [-7.224, -7.506, -7.56, -7.634]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [78938.0, 79475.0, 80012.0, 80549.0]
source_col: d(18_16)Mean
dD_mean | -60.37125
varray: [-60.002, -60.366, -60.396, -60.721]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [78938.0, 79475.0, 80012.0, 80549.0]
source_col: d(D_H)Mean
d18O_stdev | 0.179187053104
varray: [-7.224, -7.506, -7.56, -7.634]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78938.0, 79475.0, 80012.0, 80549.0]
source_col: d(18_16)Mean
dD_stdev | 0.294001558953
varray: [-60.002, -60.366, -60.396, -60.721]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78938.0, 79475.0, 80012.0, 80549.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.481
varray: [-7.224, -7.506, -7.56, -7.634]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [78938.0, 79475.0, 80012.0, 80549.0]
source_col: d(18_16)Mean
measured_dD_mean | -60.37125
varray: [-60.002, -60.366, -60.396, -60.721]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [78938.0, 79475.0, 80012.0, 80549.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.179187053104
varray: [-7.224, -7.506, -7.56, -7.634]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78938.0, 79475.0, 80012.0, 80549.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.294001558953
varray: [-60.002, -60.366, -60.396, -60.721]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78938.0, 79475.0, 80012.0, 80549.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.2952757551
actual_array: [ 1.93 -15.84]
varray: [-7.224, -7.506, -7.56, -7.634]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.481
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [78938.0, 79475.0, 80012.0, 80549.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -75.0273903434
actual_array: [ 18.1 -119. ]
varray: [-60.002, -60.366, -60.396, -60.721]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -60.37125
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [78938.0, 79475.0, 80012.0, 80549.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.0759
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -9.90665
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 32 - Sample 32 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 32
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.67975
varray: [2.738, 2.669, 2.64, 2.672]
inst: mean_all
operation: mean_all(residuals)
tarray: [81135.0, 81671.0, 82208.0, 82745.0]
source_col: residuals
SAMP_residuals_STD | 0.041427647773
varray: [2.738, 2.669, 2.64, 2.672]
inst: std_all
operation: std_all(residuals)
tarray: [81135.0, 81671.0, 82208.0, 82745.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.023
varray: [0.44, -0.255, -0.179, 0.086]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [81135.0, 81671.0, 82208.0, 82745.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.314083852923
varray: [0.44, -0.255, -0.179, 0.086]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [81135.0, 81671.0, 82208.0, 82745.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1290.31375
varray: [1289.445, 1290.805, 1290.309, 1290.696]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [81135.0, 81671.0, 82208.0, 82745.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.617033426971
varray: [1289.445, 1290.805, 1290.309, 1290.696]
inst: std_all
operation: std_all(baseline_shift)
tarray: [81135.0, 81671.0, 82208.0, 82745.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.933
varray: [0.909, 0.811, 1.07, 0.942]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [81135.0, 81671.0, 82208.0, 82745.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.106942352072
varray: [0.909, 0.811, 1.07, 0.942]
inst: std_all
operation: std_all(slope_shift)
tarray: [81135.0, 81671.0, 82208.0, 82745.0]
source_col: slope_shift
SAMP_d18O_STD | 0.142876111836
varray: [-3.454, -3.214, -3.165, -3.144]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [81135.0, 81671.0, 82208.0, 82745.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.152177527907
varray: [-62.305, -62.332, -62.521, -62.622]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [81135.0, 81671.0, 82208.0, 82745.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 104.7315
varray: [106.23, 104.245, 103.145, 105.306]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [81135.0, 81671.0, 82208.0, 82745.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.33281869235
varray: [106.23, 104.245, 103.145, 105.306]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [81135.0, 81671.0, 82208.0, 82745.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.67975
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.023
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1290.31375
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.933
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.3667
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.3667
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -12.036961053
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -12.2746
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [-3.454, -3.214, -3.165, -3.144]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [81135.0, 81671.0, 82208.0, 82745.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-62.305, -62.332, -62.521, -62.622]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [81135.0, 81671.0, 82208.0, 82745.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.24425
varray: [-3.454, -3.214, -3.165, -3.144]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [81135.0, 81671.0, 82208.0, 82745.0]
source_col: d(18_16)Mean
dD_mean | -62.445
varray: [-62.305, -62.332, -62.521, -62.622]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [81135.0, 81671.0, 82208.0, 82745.0]
source_col: d(D_H)Mean
d18O_stdev | 0.142876111836
varray: [-3.454, -3.214, -3.165, -3.144]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [81135.0, 81671.0, 82208.0, 82745.0]
source_col: d(18_16)Mean
dD_stdev | 0.152177527907
varray: [-62.305, -62.332, -62.521, -62.622]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [81135.0, 81671.0, 82208.0, 82745.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.24425
varray: [-3.454, -3.214, -3.165, -3.144]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [81135.0, 81671.0, 82208.0, 82745.0]
source_col: d(18_16)Mean
measured_dD_mean | -62.445
varray: [-62.305, -62.332, -62.521, -62.622]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [81135.0, 81671.0, 82208.0, 82745.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.142876111836
varray: [-3.454, -3.214, -3.165, -3.144]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [81135.0, 81671.0, 82208.0, 82745.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.152177527907
varray: [-62.305, -62.332, -62.521, -62.622]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [81135.0, 81671.0, 82208.0, 82745.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.04232289025
actual_array: [ 1.93 -15.84]
varray: [-3.454, -3.214, -3.165, -3.144]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -3.24425
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [81135.0, 81671.0, 82208.0, 82745.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -76.976962147
actual_array: [ 18.1 -119. ]
varray: [-62.305, -62.332, -62.521, -62.622]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -62.445
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [81135.0, 81671.0, 82208.0, 82745.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.15715
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -11.52315
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 33 - Sample 33 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 33
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.54875
varray: [2.59, 2.631, 2.714, 2.26]
inst: mean_all
operation: mean_all(residuals)
tarray: [83331.0, 83868.0, 84404.0, 84941.0]
source_col: residuals
SAMP_residuals_STD | 0.199291035089
varray: [2.59, 2.631, 2.714, 2.26]
inst: std_all
operation: std_all(residuals)
tarray: [83331.0, 83868.0, 84404.0, 84941.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.02225
varray: [-0.203, -0.077, 0.284, 0.085]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [83331.0, 83868.0, 84404.0, 84941.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.210585493328
varray: [-0.203, -0.077, 0.284, 0.085]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [83331.0, 83868.0, 84404.0, 84941.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1287.18525
varray: [1289.577, 1291.153, 1292.376, 1275.635]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [83331.0, 83868.0, 84404.0, 84941.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 7.78493563985
varray: [1289.577, 1291.153, 1292.376, 1275.635]
inst: std_all
operation: std_all(baseline_shift)
tarray: [83331.0, 83868.0, 84404.0, 84941.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 0.99525
varray: [0.606, 1.272, 0.858, 1.245]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [83331.0, 83868.0, 84404.0, 84941.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.321101619429
varray: [0.606, 1.272, 0.858, 1.245]
inst: std_all
operation: std_all(slope_shift)
tarray: [83331.0, 83868.0, 84404.0, 84941.0]
source_col: slope_shift
SAMP_d18O_STD | 0.122057363563
varray: [-5.071, -5.146, -5.138, -5.353]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83331.0, 83868.0, 84404.0, 84941.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.887898079737
varray: [-70.368, -71.151, -71.168, -72.507]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83331.0, 83868.0, 84404.0, 84941.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 100.61025
varray: [102.371, 103.318, 105.929, 90.823]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [83331.0, 83868.0, 84404.0, 84941.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 6.69605642026
varray: [102.371, 103.318, 105.929, 90.823]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [83331.0, 83868.0, 84404.0, 84941.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.54875
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.02225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1287.18525
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 0.99525
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.4977
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.4977
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -16.3370480395
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -16.39585
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [-5.071, -5.146, -5.138, -5.353]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [83331.0, 83868.0, 84404.0, 84941.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-70.368, -71.151, -71.168, -72.507]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [83331.0, 83868.0, 84404.0, 84941.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.177
varray: [-5.071, -5.146, -5.138, -5.353]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [83331.0, 83868.0, 84404.0, 84941.0]
source_col: d(18_16)Mean
dD_mean | -71.2985
varray: [-70.368, -71.151, -71.168, -72.507]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [83331.0, 83868.0, 84404.0, 84941.0]
source_col: d(D_H)Mean
d18O_stdev | 0.122057363563
varray: [-5.071, -5.146, -5.138, -5.353]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83331.0, 83868.0, 84404.0, 84941.0]
source_col: d(18_16)Mean
dD_stdev | 0.887898079737
varray: [-70.368, -71.151, -71.168, -72.507]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83331.0, 83868.0, 84404.0, 84941.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.177
varray: [-5.071, -5.146, -5.138, -5.353]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [83331.0, 83868.0, 84404.0, 84941.0]
source_col: d(18_16)Mean
measured_dD_mean | -71.2985
varray: [-70.368, -71.151, -71.168, -72.507]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [83331.0, 83868.0, 84404.0, 84941.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.122057363563
varray: [-5.071, -5.146, -5.138, -5.353]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83331.0, 83868.0, 84404.0, 84941.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.887898079737
varray: [-70.368, -71.151, -71.168, -72.507]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83331.0, 83868.0, 84404.0, 84941.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.98246442553
actual_array: [ 1.93 -15.84]
varray: [-5.071, -5.146, -5.138, -5.353]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -5.177
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [83331.0, 83868.0, 84404.0, 84941.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -85.3003058301
actual_array: [ 18.1 -119. ]
varray: [-70.368, -71.151, -71.168, -72.507]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -71.2985
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [83331.0, 83868.0, 84404.0, 84941.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.28815
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -14.65165
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 34 - Sample 34 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 34
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.6815
varray: [2.731, 2.687, 2.684, 2.624]
inst: mean_all
operation: mean_all(residuals)
tarray: [85528.0, 86065.0, 86601.0, 87138.0]
source_col: residuals
SAMP_residuals_STD | 0.0439431450854
varray: [2.731, 2.687, 2.684, 2.624]
inst: std_all
operation: std_all(residuals)
tarray: [85528.0, 86065.0, 86601.0, 87138.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.129
varray: [0.468, -0.016, 0.188, -0.124]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [85528.0, 86065.0, 86601.0, 87138.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.260407373168
varray: [0.468, -0.016, 0.188, -0.124]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [85528.0, 86065.0, 86601.0, 87138.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1288.7875
varray: [1288.312, 1289.242, 1288.536, 1289.06]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [85528.0, 86065.0, 86601.0, 87138.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.435958331342
varray: [1288.312, 1289.242, 1288.536, 1289.06]
inst: std_all
operation: std_all(baseline_shift)
tarray: [85528.0, 86065.0, 86601.0, 87138.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.28
varray: [1.047, 1.491, 1.672, 0.91]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [85528.0, 86065.0, 86601.0, 87138.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.360265642732
varray: [1.047, 1.491, 1.672, 0.91]
inst: std_all
operation: std_all(slope_shift)
tarray: [85528.0, 86065.0, 86601.0, 87138.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0422009083631
varray: [-4.612, -4.56, -4.625, -4.536]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85528.0, 86065.0, 86601.0, 87138.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.558236732937
varray: [-63.442, -62.422, -62.346, -62.239]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85528.0, 86065.0, 86601.0, 87138.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 104.966
varray: [106.328, 104.263, 104.808, 104.465]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [85528.0, 86065.0, 86601.0, 87138.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.935453187854
varray: [106.328, 104.263, 104.808, 104.465]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [85528.0, 86065.0, 86601.0, 87138.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.6815
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.129
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1288.7875
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.28
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.36495
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.36495
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -11.9795171429
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -12.0401
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [-4.612, -4.56, -4.625, -4.536]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [85528.0, 86065.0, 86601.0, 87138.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-63.442, -62.422, -62.346, -62.239]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [85528.0, 86065.0, 86601.0, 87138.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.58325
varray: [-4.612, -4.56, -4.625, -4.536]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [85528.0, 86065.0, 86601.0, 87138.0]
source_col: d(18_16)Mean
dD_mean | -62.61225
varray: [-63.442, -62.422, -62.346, -62.239]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [85528.0, 86065.0, 86601.0, 87138.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0422009083631
varray: [-4.612, -4.56, -4.625, -4.536]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85528.0, 86065.0, 86601.0, 87138.0]
source_col: d(18_16)Mean
dD_stdev | 0.558236732937
varray: [-63.442, -62.422, -62.346, -62.239]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85528.0, 86065.0, 86601.0, 87138.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.58325
varray: [-4.612, -4.56, -4.625, -4.536]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [85528.0, 86065.0, 86601.0, 87138.0]
source_col: d(18_16)Mean
measured_dD_mean | -62.61225
varray: [-63.442, -62.422, -62.346, -62.239]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [85528.0, 86065.0, 86601.0, 87138.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0422009083631
varray: [-4.612, -4.56, -4.625, -4.536]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85528.0, 86065.0, 86601.0, 87138.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.558236732937
varray: [-63.442, -62.422, -62.346, -62.239]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85528.0, 86065.0, 86601.0, 87138.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.3864437107
actual_array: [ 1.93 -15.84]
varray: [-4.612, -4.56, -4.625, -4.536]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -4.58325
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [85528.0, 86065.0, 86601.0, 87138.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -77.134197052
actual_array: [ 18.1 -119. ]
varray: [-63.442, -62.422, -62.346, -62.239]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -62.61225
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [85528.0, 86065.0, 86601.0, 87138.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.1554
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -13.0494
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 35 - Sample 35 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 35
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.73625
varray: [2.71, 2.743, 2.718, 2.774]
inst: mean_all
operation: mean_all(residuals)
tarray: [87725.0, 88262.0, 88798.0, 89335.0]
source_col: residuals
SAMP_residuals_STD | 0.0288256251739
varray: [2.71, 2.743, 2.718, 2.774]
inst: std_all
operation: std_all(residuals)
tarray: [87725.0, 88262.0, 88798.0, 89335.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.258
varray: [0.197, 0.204, -0.017, 0.648]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [87725.0, 88262.0, 88798.0, 89335.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.279500745378
varray: [0.197, 0.204, -0.017, 0.648]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [87725.0, 88262.0, 88798.0, 89335.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1290.6005
varray: [1290.538, 1290.459, 1291.023, 1290.382]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [87725.0, 88262.0, 88798.0, 89335.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.288777307511
varray: [1290.538, 1290.459, 1291.023, 1290.382]
inst: std_all
operation: std_all(baseline_shift)
tarray: [87725.0, 88262.0, 88798.0, 89335.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.45875
varray: [1.773, 1.463, 1.28, 1.319]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [87725.0, 88262.0, 88798.0, 89335.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.223795107185
varray: [1.773, 1.463, 1.28, 1.319]
inst: std_all
operation: std_all(slope_shift)
tarray: [87725.0, 88262.0, 88798.0, 89335.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0535443741209
varray: [-6.029, -6.141, -6.138, -6.078]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87725.0, 88262.0, 88798.0, 89335.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.727499828179
varray: [-71.772, -73.01, -73.11, -73.428]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87725.0, 88262.0, 88798.0, 89335.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 107.301
varray: [107.108, 106.902, 106.547, 108.647]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [87725.0, 88262.0, 88798.0, 89335.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.926765342468
varray: [107.108, 106.902, 106.547, 108.647]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [87725.0, 88262.0, 88798.0, 89335.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.73625
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.258
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1290.6005
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.45875
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.3102
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.3102
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -10.1823433833
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -9.7051
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [-6.029, -6.141, -6.138, -6.078]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [87725.0, 88262.0, 88798.0, 89335.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-71.772, -73.01, -73.11, -73.428]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [87725.0, 88262.0, 88798.0, 89335.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.0965
varray: [-6.029, -6.141, -6.138, -6.078]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [87725.0, 88262.0, 88798.0, 89335.0]
source_col: d(18_16)Mean
dD_mean | -72.83
varray: [-71.772, -73.01, -73.11, -73.428]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [87725.0, 88262.0, 88798.0, 89335.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0535443741209
varray: [-6.029, -6.141, -6.138, -6.078]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87725.0, 88262.0, 88798.0, 89335.0]
source_col: d(18_16)Mean
dD_stdev | 0.727499828179
varray: [-71.772, -73.01, -73.11, -73.428]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87725.0, 88262.0, 88798.0, 89335.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.0965
varray: [-6.029, -6.141, -6.138, -6.078]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [87725.0, 88262.0, 88798.0, 89335.0]
source_col: d(18_16)Mean
measured_dD_mean | -72.83
varray: [-71.772, -73.01, -73.11, -73.428]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [87725.0, 88262.0, 88798.0, 89335.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0535443741209
varray: [-6.029, -6.141, -6.138, -6.078]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87725.0, 88262.0, 88798.0, 89335.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.727499828179
varray: [-71.772, -73.01, -73.11, -73.428]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87725.0, 88262.0, 88798.0, 89335.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.90548092621
actual_array: [ 1.93 -15.84]
varray: [-6.029, -6.141, -6.138, -6.078]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -6.0965
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [87725.0, 88262.0, 88798.0, 89335.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -86.7400980988
actual_array: [ 18.1 -119. ]
varray: [-71.772, -73.01, -73.11, -73.428]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -72.83
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [87725.0, 88262.0, 88798.0, 89335.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.10065
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -11.2364
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 36 - Sample 36 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 36
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.7495
varray: [2.798, 2.729, 2.734, 2.737]
inst: mean_all
operation: mean_all(residuals)
tarray: [89918.0, 90455.0, 90991.0, 91527.0]
source_col: residuals
SAMP_residuals_STD | 0.032501282026
varray: [2.798, 2.729, 2.734, 2.737]
inst: std_all
operation: std_all(residuals)
tarray: [89918.0, 90455.0, 90991.0, 91527.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.13475
varray: [0.239, -0.059, 0.116, 0.243]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [89918.0, 90455.0, 90991.0, 91527.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.141982099811
varray: [0.239, -0.059, 0.116, 0.243]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [89918.0, 90455.0, 90991.0, 91527.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1293.295
varray: [1293.618, 1293.241, 1293.202, 1293.119]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [89918.0, 90455.0, 90991.0, 91527.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.221261534539
varray: [1293.618, 1293.241, 1293.202, 1293.119]
inst: std_all
operation: std_all(baseline_shift)
tarray: [89918.0, 90455.0, 90991.0, 91527.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.55875
varray: [1.286, 1.817, 1.326, 1.806]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [89918.0, 90455.0, 90991.0, 91527.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.292341552526
varray: [1.286, 1.817, 1.326, 1.806]
inst: std_all
operation: std_all(slope_shift)
tarray: [89918.0, 90455.0, 90991.0, 91527.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0734688142457
varray: [-4.628, -4.488, -4.513, -4.461]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89918.0, 90455.0, 90991.0, 91527.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.615305005668
varray: [-66.473, -65.608, -65.348, -65.042]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89918.0, 90455.0, 90991.0, 91527.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 107.409
varray: [109.521, 106.065, 107.229, 106.821]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [89918.0, 90455.0, 90991.0, 91527.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.48829029426
varray: [109.521, 106.065, 107.229, 106.821]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [89918.0, 90455.0, 90991.0, 91527.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.7495
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.13475
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1293.295
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.55875
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.29695
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.29695
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -9.74741092091
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -9.5971
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [-4.628, -4.488, -4.513, -4.461]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [89918.0, 90455.0, 90991.0, 91527.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-66.473, -65.608, -65.348, -65.042]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [89918.0, 90455.0, 90991.0, 91527.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.5225
varray: [-4.628, -4.488, -4.513, -4.461]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [89918.0, 90455.0, 90991.0, 91527.0]
source_col: d(18_16)Mean
dD_mean | -65.61775
varray: [-66.473, -65.608, -65.348, -65.042]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [89918.0, 90455.0, 90991.0, 91527.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0734688142457
varray: [-4.628, -4.488, -4.513, -4.461]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89918.0, 90455.0, 90991.0, 91527.0]
source_col: d(18_16)Mean
dD_stdev | 0.615305005668
varray: [-66.473, -65.608, -65.348, -65.042]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89918.0, 90455.0, 90991.0, 91527.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.5225
varray: [-4.628, -4.488, -4.513, -4.461]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [89918.0, 90455.0, 90991.0, 91527.0]
source_col: d(18_16)Mean
measured_dD_mean | -65.61775
varray: [-66.473, -65.608, -65.348, -65.042]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [89918.0, 90455.0, 90991.0, 91527.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0734688142457
varray: [-4.628, -4.488, -4.513, -4.461]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89918.0, 90455.0, 90991.0, 91527.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.615305005668
varray: [-66.473, -65.608, -65.348, -65.042]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89918.0, 90455.0, 90991.0, 91527.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.32546138072
actual_array: [ 1.93 -15.84]
varray: [-4.628, -4.488, -4.513, -4.461]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -4.5225
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [89918.0, 90455.0, 90991.0, 91527.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -79.959724746
actual_array: [ 18.1 -119. ]
varray: [-66.473, -65.608, -65.348, -65.042]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -65.61775
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [89918.0, 90455.0, 90991.0, 91527.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.0874
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -8.5419
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 37 - Sample 37 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 37
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 4.56825
varray: [6.273, 0.0, 5.721, 6.279]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [92111.0, 92648.0, 93184.0, 93720.0]
source_col: residuals
SAMP_residuals_STD | 3.05671821567
varray: [6.273, 0.0, 5.721, 6.279]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [92111.0, 92648.0, 93184.0, 93720.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -9.56125
varray: [-13.055, 0.0, -11.911, -13.279]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [92111.0, 92648.0, 93184.0, 93720.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 6.4022595165
varray: [-13.055, 0.0, -11.911, -13.279]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [92111.0, 92648.0, 93184.0, 93720.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1204.32575
varray: [1618.085, 0.0, 1579.545, 1619.673]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [92111.0, 92648.0, 93184.0, 93720.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 803.098178458
varray: [1618.085, 0.0, 1579.545, 1619.673]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [92111.0, 92648.0, 93184.0, 93720.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -35.93075
varray: [-49.476, 0.0, -44.571, -49.676]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [92111.0, 92648.0, 93184.0, 93720.0]
source_col: slope_shift
SAMP_slope_shift_STD | 24.0698872089
varray: [-49.476, 0.0, -44.571, -49.676]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [92111.0, 92648.0, 93184.0, 93720.0]
source_col: slope_shift
SAMP_d18O_STD | 8.78304210776
varray: [17.411, 0.0, 17.541, 17.74]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [92111.0, 92648.0, 93184.0, 93720.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 10.2793597523
varray: [-21.328, 0.0, -20.739, -19.424]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [92111.0, 92648.0, 93184.0, 93720.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 92.8395
varray: [128.907, 0.0, 116.261, 126.19]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [92111.0, 92648.0, 93184.0, 93720.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 62.1312047981
varray: [128.907, 0.0, 116.261, 126.19]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [92111.0, 92648.0, 93184.0, 93720.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 4.56825
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 9.56125
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1204.32575
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 35.93075
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 1.5218
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.5218
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 49.9532242446
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -24.1666
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [17.411, 0.0, 17.541, 17.74]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [92111.0, 92648.0, 93184.0, 93720.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-21.328, 0.0, -20.739, -19.424]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [92111.0, 92648.0, 93184.0, 93720.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 13.173
varray: [17.411, 0.0, 17.541, 17.74]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [92111.0, 92648.0, 93184.0, 93720.0]
source_col: d(18_16)Mean
dD_mean | -15.37275
varray: [-21.328, 0.0, -20.739, -19.424]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [92111.0, 92648.0, 93184.0, 93720.0]
source_col: d(D_H)Mean
d18O_stdev | 8.78304210776
varray: [17.411, 0.0, 17.541, 17.74]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [92111.0, 92648.0, 93184.0, 93720.0]
source_col: d(18_16)Mean
dD_stdev | 10.2793597523
varray: [-21.328, 0.0, -20.739, -19.424]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [92111.0, 92648.0, 93184.0, 93720.0]
source_col: d(D_H)Mean
measured_d18O_mean | 13.173
varray: [17.411, 0.0, 17.541, 17.74]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [92111.0, 92648.0, 93184.0, 93720.0]
source_col: d(18_16)Mean
measured_dD_mean | -15.37275
varray: [-21.328, 0.0, -20.739, -19.424]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [92111.0, 92648.0, 93184.0, 93720.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 8.78304210776
varray: [17.411, 0.0, 17.541, 17.74]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [92111.0, 92648.0, 93184.0, 93720.0]
source_col: d(18_16)Mean
measured_dD_stdev | 10.2793597523
varray: [-21.328, 0.0, -20.739, -19.424]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [92111.0, 92648.0, 93184.0, 93720.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 10.4377126136
actual_array: [ 1.93 -15.84]
varray: [17.411, 0.0, 17.541, 17.74]
fit_rsq: 1.0
fit_intercept: 2.77505289814
except: True
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 13.173
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [92111.0, 92648.0, 93184.0, 93720.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -32.7234449293
actual_array: [ 18.1 -119. ]
varray: [-21.328, 0.0, -20.739, -19.424]
fit_rsq: 1.0
fit_intercept: 19.4350184537
except: True
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -15.37275
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [92111.0, 92648.0, 93184.0, 93720.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1297.26865
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -97.51115
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 38 - Sample 38 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 38
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.72825
varray: [2.634, 2.774, 2.704, 2.801]
inst: mean_all
operation: mean_all(residuals)
tarray: [94305.0, 94841.0, 95378.0, 95914.0]
source_col: residuals
SAMP_residuals_STD | 0.0749594334735
varray: [2.634, 2.774, 2.704, 2.801]
inst: std_all
operation: std_all(residuals)
tarray: [94305.0, 94841.0, 95378.0, 95914.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.126
varray: [-0.273, 0.172, -0.485, 0.082]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [94305.0, 94841.0, 95378.0, 95914.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.306897376985
varray: [-0.273, 0.172, -0.485, 0.082]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [94305.0, 94841.0, 95378.0, 95914.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1415.852
varray: [1413.403, 1416.409, 1416.382, 1417.214]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [94305.0, 94841.0, 95378.0, 95914.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.67767636927
varray: [1413.403, 1416.409, 1416.382, 1417.214]
inst: std_all
operation: std_all(baseline_shift)
tarray: [94305.0, 94841.0, 95378.0, 95914.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -4.185
varray: [-4.341, -4.203, -4.216, -3.98]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [94305.0, 94841.0, 95378.0, 95914.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.150162134597
varray: [-4.341, -4.203, -4.216, -3.98]
inst: std_all
operation: std_all(slope_shift)
tarray: [94305.0, 94841.0, 95378.0, 95914.0]
source_col: slope_shift
SAMP_d18O_STD | 0.200327024304
varray: [-6.406, -6.714, -6.801, -6.854]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94305.0, 94841.0, 95378.0, 95914.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.02197305085
varray: [-68.159, -69.882, -70.132, -70.44]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94305.0, 94841.0, 95378.0, 95914.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 106.701
varray: [105.007, 108.246, 105.493, 108.058]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [94305.0, 94841.0, 95378.0, 95914.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.68892214149
varray: [105.007, 108.246, 105.493, 108.058]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [94305.0, 94841.0, 95378.0, 95914.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.72825
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.126
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1415.852
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 4.185
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.3182
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.3182
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -10.4449441153
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -10.3051
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [-6.406, -6.714, -6.801, -6.854]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [94305.0, 94841.0, 95378.0, 95914.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-68.159, -69.882, -70.132, -70.44]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [94305.0, 94841.0, 95378.0, 95914.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.69375
varray: [-6.406, -6.714, -6.801, -6.854]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [94305.0, 94841.0, 95378.0, 95914.0]
source_col: d(18_16)Mean
dD_mean | -69.65325
varray: [-68.159, -69.882, -70.132, -70.44]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [94305.0, 94841.0, 95378.0, 95914.0]
source_col: d(D_H)Mean
d18O_stdev | 0.200327024304
varray: [-6.406, -6.714, -6.801, -6.854]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94305.0, 94841.0, 95378.0, 95914.0]
source_col: d(18_16)Mean
dD_stdev | 1.02197305085
varray: [-68.159, -69.882, -70.132, -70.44]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94305.0, 94841.0, 95378.0, 95914.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.69375
varray: [-6.406, -6.714, -6.801, -6.854]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [94305.0, 94841.0, 95378.0, 95914.0]
source_col: d(18_16)Mean
measured_dD_mean | -69.65325
varray: [-68.159, -69.882, -70.132, -70.44]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [94305.0, 94841.0, 95378.0, 95914.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.200327024304
varray: [-6.406, -6.714, -6.801, -6.854]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94305.0, 94841.0, 95378.0, 95914.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.02197305085
varray: [-68.159, -69.882, -70.132, -70.44]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94305.0, 94841.0, 95378.0, 95914.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.5050150263
actual_array: [ 1.93 -15.84]
varray: [-6.406, -6.714, -6.801, -6.854]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -6.69375
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [94305.0, 94841.0, 95378.0, 95914.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -83.7535750237
actual_array: [ 18.1 -119. ]
varray: [-68.159, -69.882, -70.132, -70.44]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -69.65325
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [94305.0, 94841.0, 95378.0, 95914.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.10865
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 114.0151
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 39 - Sample 39 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 39
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.731
varray: [2.716, 2.73, 2.715, 2.763]
inst: mean_all
operation: mean_all(residuals)
tarray: [96498.0, 97035.0, 97571.0, 98108.0]
source_col: residuals
SAMP_residuals_STD | 0.0224053565024
varray: [2.716, 2.73, 2.715, 2.763]
inst: std_all
operation: std_all(residuals)
tarray: [96498.0, 97035.0, 97571.0, 98108.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.1565
varray: [0.22, 0.138, 0.189, 0.079]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [96498.0, 97035.0, 97571.0, 98108.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0617440955342
varray: [0.22, 0.138, 0.189, 0.079]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [96498.0, 97035.0, 97571.0, 98108.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1291.97025
varray: [1291.466, 1292.496, 1291.945, 1291.974]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [96498.0, 97035.0, 97571.0, 98108.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.420845478373
varray: [1291.466, 1292.496, 1291.945, 1291.974]
inst: std_all
operation: std_all(baseline_shift)
tarray: [96498.0, 97035.0, 97571.0, 98108.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.15425
varray: [1.524, 1.117, 1.385, 0.591]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [96498.0, 97035.0, 97571.0, 98108.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.411743751218
varray: [1.524, 1.117, 1.385, 0.591]
inst: std_all
operation: std_all(slope_shift)
tarray: [96498.0, 97035.0, 97571.0, 98108.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0969273783132
varray: [-5.978, -5.787, -5.795, -5.773]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96498.0, 97035.0, 97571.0, 98108.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0577032639169
varray: [-72.968, -73.018, -73.041, -73.107]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96498.0, 97035.0, 97571.0, 98108.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 106.561
varray: [105.197, 106.518, 106.698, 107.831]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [96498.0, 97035.0, 97571.0, 98108.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.07919939461
varray: [105.197, 106.518, 106.698, 107.831]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [96498.0, 97035.0, 97571.0, 98108.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.731
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.1565
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1291.97025
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.15425
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.31545
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.31545
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -10.3546751137
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -10.4451
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [-5.978, -5.787, -5.795, -5.773]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [96498.0, 97035.0, 97571.0, 98108.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-72.968, -73.018, -73.041, -73.107]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [96498.0, 97035.0, 97571.0, 98108.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.83325
varray: [-5.978, -5.787, -5.795, -5.773]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [96498.0, 97035.0, 97571.0, 98108.0]
source_col: d(18_16)Mean
dD_mean | -73.0335
varray: [-72.968, -73.018, -73.041, -73.107]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [96498.0, 97035.0, 97571.0, 98108.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0969273783132
varray: [-5.978, -5.787, -5.795, -5.773]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96498.0, 97035.0, 97571.0, 98108.0]
source_col: d(18_16)Mean
dD_stdev | 0.0577032639169
varray: [-72.968, -73.018, -73.041, -73.107]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96498.0, 97035.0, 97571.0, 98108.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.83325
varray: [-5.978, -5.787, -5.795, -5.773]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [96498.0, 97035.0, 97571.0, 98108.0]
source_col: d(18_16)Mean
measured_dD_mean | -73.0335
varray: [-72.968, -73.018, -73.041, -73.107]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [96498.0, 97035.0, 97571.0, 98108.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0969273783132
varray: [-5.978, -5.787, -5.795, -5.773]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96498.0, 97035.0, 97571.0, 98108.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0577032639169
varray: [-72.968, -73.018, -73.041, -73.107]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96498.0, 97035.0, 97571.0, 98108.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.64122416296
actual_array: [ 1.93 -15.84]
varray: [-5.978, -5.787, -5.795, -5.773]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -5.83325
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [96498.0, 97035.0, 97571.0, 98108.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -86.931412318
actual_array: [ 18.1 -119. ]
varray: [-72.968, -73.018, -73.041, -73.107]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -73.0335
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [96498.0, 97035.0, 97571.0, 98108.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.1059
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -9.86665
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 40 - Sample 40 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 40
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0885
varray: [0.0, 2.808, 2.829, 2.717]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [98693.0, 99230.0, 99766.0, 100302.0]
source_col: residuals
SAMP_residuals_STD | 1.39318161056
varray: [0.0, 2.808, 2.829, 2.717]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [98693.0, 99230.0, 99766.0, 100302.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.352
varray: [0.0, 0.374, 0.788, 0.246]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [98693.0, 99230.0, 99766.0, 100302.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.329504678773
varray: [0.0, 0.374, 0.788, 0.246]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [98693.0, 99230.0, 99766.0, 100302.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 968.1575
varray: [0.0, 1293.237, 1290.023, 1289.37]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [98693.0, 99230.0, 99766.0, 100302.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 645.440546287
varray: [0.0, 1293.237, 1290.023, 1289.37]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [98693.0, 99230.0, 99766.0, 100302.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.08975
varray: [0.0, 1.213, 0.826, 2.32]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [98693.0, 99230.0, 99766.0, 100302.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.963643217171
varray: [0.0, 1.213, 0.826, 2.32]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [98693.0, 99230.0, 99766.0, 100302.0]
source_col: slope_shift
SAMP_d18O_STD | 3.75916147565
varray: [0.0, -7.469, -7.547, -7.538]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [98693.0, 99230.0, 99766.0, 100302.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 30.2483041618
varray: [0.0, -60.451, -60.68, -60.357]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [98693.0, 99230.0, 99766.0, 100302.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 81.42525
varray: [0.0, 109.959, 110.201, 105.541]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [98693.0, 99230.0, 99766.0, 100302.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 54.3257440899
varray: [0.0, 109.959, 110.201, 105.541]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [98693.0, 99230.0, 99766.0, 100302.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0885
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.352
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 968.1575
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.08975
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.95795
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.95795
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -31.4447964024
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -35.58085
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [0.0, -7.469, -7.547, -7.538]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [98693.0, 99230.0, 99766.0, 100302.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [0.0, -60.451, -60.68, -60.357]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [98693.0, 99230.0, 99766.0, 100302.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.6385
varray: [0.0, -7.469, -7.547, -7.538]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [98693.0, 99230.0, 99766.0, 100302.0]
source_col: d(18_16)Mean
dD_mean | -45.372
varray: [0.0, -60.451, -60.68, -60.357]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [98693.0, 99230.0, 99766.0, 100302.0]
source_col: d(D_H)Mean
d18O_stdev | 3.75916147565
varray: [0.0, -7.469, -7.547, -7.538]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [98693.0, 99230.0, 99766.0, 100302.0]
source_col: d(18_16)Mean
dD_stdev | 30.2483041618
varray: [0.0, -60.451, -60.68, -60.357]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [98693.0, 99230.0, 99766.0, 100302.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.6385
varray: [0.0, -7.469, -7.547, -7.538]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [98693.0, 99230.0, 99766.0, 100302.0]
source_col: d(18_16)Mean
measured_dD_mean | -45.372
varray: [0.0, -60.451, -60.68, -60.357]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [98693.0, 99230.0, 99766.0, 100302.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 3.75916147565
varray: [0.0, -7.469, -7.547, -7.538]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [98693.0, 99230.0, 99766.0, 100302.0]
source_col: d(18_16)Mean
measured_dD_stdev | 30.2483041618
varray: [0.0, -60.451, -60.68, -60.357]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [98693.0, 99230.0, 99766.0, 100302.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.4457293685
actual_array: [ 1.93 -15.84]
varray: [0.0, -7.469, -7.547, -7.538]
fit_rsq: 1.0
fit_intercept: 2.77505289814
except: True
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -5.6385
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [98693.0, 99230.0, 99766.0, 100302.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -60.9263102352
actual_array: [ 18.1 -119. ]
varray: [0.0, -60.451, -60.68, -60.357]
fit_rsq: 1.0
fit_intercept: 19.4350184537
except: True
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -45.372
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [98693.0, 99230.0, 99766.0, 100302.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.7484
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -333.6794
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 41 - Sample 41 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 41
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.7375
varray: [2.705, 2.751, 2.718, 2.776]
inst: mean_all
operation: mean_all(residuals)
tarray: [100887.0, 101423.0, 101960.0, 102497.0]
source_col: residuals
SAMP_residuals_STD | 0.0321506868771
varray: [2.705, 2.751, 2.718, 2.776]
inst: std_all
operation: std_all(residuals)
tarray: [100887.0, 101423.0, 101960.0, 102497.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.26375
varray: [0.06, 0.165, 0.51, 0.32]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [100887.0, 101423.0, 101960.0, 102497.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.195847517217
varray: [0.06, 0.165, 0.51, 0.32]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [100887.0, 101423.0, 101960.0, 102497.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1290.61475
varray: [1292.016, 1290.937, 1288.852, 1290.654]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [100887.0, 101423.0, 101960.0, 102497.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.31354161842
varray: [1292.016, 1290.937, 1288.852, 1290.654]
inst: std_all
operation: std_all(baseline_shift)
tarray: [100887.0, 101423.0, 101960.0, 102497.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.3835
varray: [1.711, 1.285, 1.327, 1.211]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [100887.0, 101423.0, 101960.0, 102497.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.223537468895
varray: [1.711, 1.285, 1.327, 1.211]
inst: std_all
operation: std_all(slope_shift)
tarray: [100887.0, 101423.0, 101960.0, 102497.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0945599457135
varray: [-4.887, -4.71, -4.727, -4.673]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100887.0, 101423.0, 101960.0, 102497.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.488304123131
varray: [-67.081, -67.706, -68.188, -68.022]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100887.0, 101423.0, 101960.0, 102497.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 106.98525
varray: [106.96, 107.519, 105.941, 107.521]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [100887.0, 101423.0, 101960.0, 102497.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.744538503594
varray: [106.96, 107.519, 105.941, 107.521]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [100887.0, 101423.0, 101960.0, 102497.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.7375
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.26375
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1290.61475
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.3835
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.30895
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.30895
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -10.1413120189
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -10.02085
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [-4.887, -4.71, -4.727, -4.673]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [100887.0, 101423.0, 101960.0, 102497.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-67.081, -67.706, -68.188, -68.022]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [100887.0, 101423.0, 101960.0, 102497.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.74925
varray: [-4.887, -4.71, -4.727, -4.673]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [100887.0, 101423.0, 101960.0, 102497.0]
source_col: d(18_16)Mean
dD_mean | -67.74925
varray: [-67.081, -67.706, -68.188, -68.022]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [100887.0, 101423.0, 101960.0, 102497.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0945599457135
varray: [-4.887, -4.71, -4.727, -4.673]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100887.0, 101423.0, 101960.0, 102497.0]
source_col: d(18_16)Mean
dD_stdev | 0.488304123131
varray: [-67.081, -67.706, -68.188, -68.022]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100887.0, 101423.0, 101960.0, 102497.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.74925
varray: [-4.887, -4.71, -4.727, -4.673]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [100887.0, 101423.0, 101960.0, 102497.0]
source_col: d(18_16)Mean
measured_dD_mean | -67.74925
varray: [-67.081, -67.706, -68.188, -68.022]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [100887.0, 101423.0, 101960.0, 102497.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0945599457135
varray: [-4.887, -4.71, -4.727, -4.673]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100887.0, 101423.0, 101960.0, 102497.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.488304123131
varray: [-67.081, -67.706, -68.188, -68.022]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100887.0, 101423.0, 101960.0, 102497.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.55307855476
actual_array: [ 1.93 -15.84]
varray: [-4.887, -4.71, -4.727, -4.673]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -4.74925
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [100887.0, 101423.0, 101960.0, 102497.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -81.9635884226
actual_array: [ 18.1 -119. ]
varray: [-67.081, -67.706, -68.188, -68.022]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -67.74925
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [100887.0, 101423.0, 101960.0, 102497.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.0994
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -11.22215
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 42 - Sample 42 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 42
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.80575
varray: [3.48, 3.7, 4.082, 3.961]
inst: mean_all
operation: mean_all(residuals)
tarray: [103082.0, 103619.0, 104155.0, 104692.0]
source_col: residuals
SAMP_residuals_STD | 0.269389897113
varray: [3.48, 3.7, 4.082, 3.961]
inst: std_all
operation: std_all(residuals)
tarray: [103082.0, 103619.0, 104155.0, 104692.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -7.1145
varray: [-6.383, -6.808, -7.518, -7.749]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [103082.0, 103619.0, 104155.0, 104692.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.630987849857
varray: [-6.383, -6.808, -7.518, -7.749]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [103082.0, 103619.0, 104155.0, 104692.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1529.85
varray: [1497.647, 1518.962, 1553.769, 1549.022]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [103082.0, 103619.0, 104155.0, 104692.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 26.427674245
varray: [1497.647, 1518.962, 1553.769, 1549.022]
inst: std_all
operation: std_all(baseline_shift)
tarray: [103082.0, 103619.0, 104155.0, 104692.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -27.75475
varray: [-25.127, -26.371, -29.839, -29.682]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [103082.0, 103619.0, 104155.0, 104692.0]
source_col: slope_shift
SAMP_slope_shift_STD | 2.37193484663
varray: [-25.127, -26.371, -29.839, -29.682]
inst: std_all
operation: std_all(slope_shift)
tarray: [103082.0, 103619.0, 104155.0, 104692.0]
source_col: slope_shift
SAMP_d18O_STD | 0.102895983077
varray: [5.169, 5.308, 5.382, 5.392]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [103082.0, 103619.0, 104155.0, 104692.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.459007897536
varray: [-54.857, -54.754, -54.002, -54.026]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [103082.0, 103619.0, 104155.0, 104692.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 111.00475
varray: [102.747, 107.032, 118.888, 115.352]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [103082.0, 103619.0, 104155.0, 104692.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 7.41658930484
varray: [102.747, 107.032, 118.888, 115.352]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [103082.0, 103619.0, 104155.0, 104692.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.80575
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 7.1145
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1529.85
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.75475
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.7593
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.7593
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 24.9240919759
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -6.00135
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [5.169, 5.308, 5.382, 5.392]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [103082.0, 103619.0, 104155.0, 104692.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-54.857, -54.754, -54.002, -54.026]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [103082.0, 103619.0, 104155.0, 104692.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 5.31275
varray: [5.169, 5.308, 5.382, 5.392]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [103082.0, 103619.0, 104155.0, 104692.0]
source_col: d(18_16)Mean
dD_mean | -54.40975
varray: [-54.857, -54.754, -54.002, -54.026]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [103082.0, 103619.0, 104155.0, 104692.0]
source_col: d(D_H)Mean
d18O_stdev | 0.102895983077
varray: [5.169, 5.308, 5.382, 5.392]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [103082.0, 103619.0, 104155.0, 104692.0]
source_col: d(18_16)Mean
dD_stdev | 0.459007897536
varray: [-54.857, -54.754, -54.002, -54.026]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [103082.0, 103619.0, 104155.0, 104692.0]
source_col: d(D_H)Mean
measured_d18O_mean | 5.31275
varray: [5.169, 5.308, 5.382, 5.392]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [103082.0, 103619.0, 104155.0, 104692.0]
source_col: d(18_16)Mean
measured_dD_mean | -54.40975
varray: [-54.857, -54.754, -54.002, -54.026]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [103082.0, 103619.0, 104155.0, 104692.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.102895983077
varray: [5.169, 5.308, 5.382, 5.392]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [103082.0, 103619.0, 104155.0, 104692.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.459007897536
varray: [-54.857, -54.754, -54.002, -54.026]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [103082.0, 103619.0, 104155.0, 104692.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 2.54740217373
actual_array: [ 1.93 -15.84]
varray: [5.169, 5.308, 5.382, 5.392]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 5.31275
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [103082.0, 103619.0, 104155.0, 104692.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -69.4228708464
actual_array: [ 18.1 -119. ]
varray: [-54.857, -54.754, -54.002, -54.026]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -54.40975
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [103082.0, 103619.0, 104155.0, 104692.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1298.03115
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 228.0131
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 43 - Sample 43 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 43
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.65175
varray: [2.657, 2.655, 2.596, 2.699]
inst: mean_all
operation: mean_all(residuals)
tarray: [105277.0, 105813.0, 106350.0, 106886.0]
source_col: residuals
SAMP_residuals_STD | 0.0423428467001
varray: [2.657, 2.655, 2.596, 2.699]
inst: std_all
operation: std_all(residuals)
tarray: [105277.0, 105813.0, 106350.0, 106886.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.14475
varray: [0.11, 0.127, -0.124, 0.466]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [105277.0, 105813.0, 106350.0, 106886.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.242865360505
varray: [0.11, 0.127, -0.124, 0.466]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [105277.0, 105813.0, 106350.0, 106886.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1290.08125
varray: [1292.784, 1290.059, 1289.33, 1288.152]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [105277.0, 105813.0, 106350.0, 106886.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.96568349012
varray: [1292.784, 1290.059, 1289.33, 1288.152]
inst: std_all
operation: std_all(baseline_shift)
tarray: [105277.0, 105813.0, 106350.0, 106886.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.405
varray: [1.107, 1.305, 1.635, 1.573]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [105277.0, 105813.0, 106350.0, 106886.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.244899979584
varray: [1.107, 1.305, 1.635, 1.573]
inst: std_all
operation: std_all(slope_shift)
tarray: [105277.0, 105813.0, 106350.0, 106886.0]
source_col: slope_shift
SAMP_d18O_STD | 0.176071954988
varray: [-9.22, -9.492, -9.52, -9.636]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [105277.0, 105813.0, 106350.0, 106886.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.69456651291
varray: [-95.403, -98.079, -98.749, -99.179]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [105277.0, 105813.0, 106350.0, 106886.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 104.823
varray: [105.381, 104.719, 103.174, 106.018]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [105277.0, 105813.0, 106350.0, 106886.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.22057445492
varray: [105.381, 104.719, 103.174, 106.018]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [105277.0, 105813.0, 106350.0, 106886.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.65175
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.14475
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1290.08125
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.405
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.3947
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.3947
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -12.956063615
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -12.1831
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [-9.22, -9.492, -9.52, -9.636]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [105277.0, 105813.0, 106350.0, 106886.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-95.403, -98.079, -98.749, -99.179]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [105277.0, 105813.0, 106350.0, 106886.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.467
varray: [-9.22, -9.492, -9.52, -9.636]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [105277.0, 105813.0, 106350.0, 106886.0]
source_col: d(18_16)Mean
dD_mean | -97.8525
varray: [-95.403, -98.079, -98.749, -99.179]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [105277.0, 105813.0, 106350.0, 106886.0]
source_col: d(D_H)Mean
d18O_stdev | 0.176071954988
varray: [-9.22, -9.492, -9.52, -9.636]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [105277.0, 105813.0, 106350.0, 106886.0]
source_col: d(18_16)Mean
dD_stdev | 1.69456651291
varray: [-95.403, -98.079, -98.749, -99.179]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [105277.0, 105813.0, 106350.0, 106886.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.467
varray: [-9.22, -9.492, -9.52, -9.636]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [105277.0, 105813.0, 106350.0, 106886.0]
source_col: d(18_16)Mean
measured_dD_mean | -97.8525
varray: [-95.403, -98.079, -98.749, -99.179]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [105277.0, 105813.0, 106350.0, 106886.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.176071954988
varray: [-9.22, -9.492, -9.52, -9.636]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [105277.0, 105813.0, 106350.0, 106886.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.69456651291
varray: [-95.403, -98.079, -98.749, -99.179]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [105277.0, 105813.0, 106350.0, 106886.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -12.2888709377
actual_array: [ 1.93 -15.84]
varray: [-9.22, -9.492, -9.52, -9.636]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -9.467
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [105277.0, 105813.0, 106350.0, 106886.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -110.264226106
actual_array: [ 18.1 -119. ]
varray: [-95.403, -98.079, -98.749, -99.179]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -97.8525
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [105277.0, 105813.0, 106350.0, 106886.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.18515
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -11.75565
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 44 - Sample 44 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 44
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 5.07875
varray: [5.008, 4.762, 5.321, 5.224]
inst: mean_all
operation: mean_all(residuals)
tarray: [107471.0, 108007.0, 108543.0, 109080.0]
source_col: residuals
SAMP_residuals_STD | 0.248407427425
varray: [5.008, 4.762, 5.321, 5.224]
inst: std_all
operation: std_all(residuals)
tarray: [107471.0, 108007.0, 108543.0, 109080.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -10.9565
varray: [-10.827, -10.659, -11.395, -10.945]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [107471.0, 108007.0, 108543.0, 109080.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.315008465495
varray: [-10.827, -10.659, -11.395, -10.945]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [107471.0, 108007.0, 108543.0, 109080.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1972.48725
varray: [1962.888, 1924.072, 2009.786, 1993.203]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [107471.0, 108007.0, 108543.0, 109080.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 37.6674974071
varray: [1962.888, 1924.072, 2009.786, 1993.203]
inst: std_all
operation: std_all(baseline_shift)
tarray: [107471.0, 108007.0, 108543.0, 109080.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -54.77475
varray: [-54.186, -51.517, -56.485, -56.911]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [107471.0, 108007.0, 108543.0, 109080.0]
source_col: slope_shift
SAMP_slope_shift_STD | 2.47979036412
varray: [-54.186, -51.517, -56.485, -56.911]
inst: std_all
operation: std_all(slope_shift)
tarray: [107471.0, 108007.0, 108543.0, 109080.0]
source_col: slope_shift
SAMP_d18O_STD | 0.255438674963
varray: [13.562, 14.035, 13.977, 14.147]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [107471.0, 108007.0, 108543.0, 109080.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.498982643
varray: [-40.006, -36.46, -35.076, -34.399]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [107471.0, 108007.0, 108543.0, 109080.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 104.551
varray: [103.402, 97.67, 108.736, 108.396]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [107471.0, 108007.0, 108543.0, 109080.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 5.19508139173
varray: [103.402, 97.67, 108.736, 108.396]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [107471.0, 108007.0, 108543.0, 109080.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 5.07875
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 10.9565
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1972.48725
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 54.77475
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 2.0323
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.0323
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 66.7104334553
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -12.4551
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [13.562, 14.035, 13.977, 14.147]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [107471.0, 108007.0, 108543.0, 109080.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-40.006, -36.46, -35.076, -34.399]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [107471.0, 108007.0, 108543.0, 109080.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 13.93025
varray: [13.562, 14.035, 13.977, 14.147]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [107471.0, 108007.0, 108543.0, 109080.0]
source_col: d(18_16)Mean
dD_mean | -36.48525
varray: [-40.006, -36.46, -35.076, -34.399]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [107471.0, 108007.0, 108543.0, 109080.0]
source_col: d(D_H)Mean
d18O_stdev | 0.255438674963
varray: [13.562, 14.035, 13.977, 14.147]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [107471.0, 108007.0, 108543.0, 109080.0]
source_col: d(18_16)Mean
dD_stdev | 2.498982643
varray: [-40.006, -36.46, -35.076, -34.399]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [107471.0, 108007.0, 108543.0, 109080.0]
source_col: d(D_H)Mean
measured_d18O_mean | 13.93025
varray: [13.562, 14.035, 13.977, 14.147]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [107471.0, 108007.0, 108543.0, 109080.0]
source_col: d(18_16)Mean
measured_dD_mean | -36.48525
varray: [-40.006, -36.46, -35.076, -34.399]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [107471.0, 108007.0, 108543.0, 109080.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.255438674963
varray: [13.562, 14.035, 13.977, 14.147]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [107471.0, 108007.0, 108543.0, 109080.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.498982643
varray: [-40.006, -36.46, -35.076, -34.399]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [107471.0, 108007.0, 108543.0, 109080.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 11.1978586116
actual_array: [ 1.93 -15.84]
varray: [13.562, 14.035, 13.977, 14.147]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 13.93025
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [107471.0, 108007.0, 108543.0, 109080.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -52.5717075949
actual_array: [ 18.1 -119. ]
varray: [-40.006, -36.46, -35.076, -34.399]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -36.48525
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [107471.0, 108007.0, 108543.0, 109080.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1296.75815
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | True
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | True
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 670.65035
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 45 - Sample 45 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 45
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.54075
varray: [2.392, 2.661, 2.525, 2.585]
inst: mean_all
operation: mean_all(residuals)
tarray: [109665.0, 110201.0, 110737.0, 111274.0]
source_col: residuals
SAMP_residuals_STD | 0.113714188502
varray: [2.392, 2.661, 2.525, 2.585]
inst: std_all
operation: std_all(residuals)
tarray: [109665.0, 110201.0, 110737.0, 111274.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.00275
varray: [-0.263, 0.103, 0.024, 0.147]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [109665.0, 110201.0, 110737.0, 111274.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.184330093401
varray: [-0.263, 0.103, 0.024, 0.147]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [109665.0, 110201.0, 110737.0, 111274.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1289.00325
varray: [1290.305, 1291.347, 1286.487, 1287.874]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [109665.0, 110201.0, 110737.0, 111274.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.22069334143
varray: [1290.305, 1291.347, 1286.487, 1287.874]
inst: std_all
operation: std_all(baseline_shift)
tarray: [109665.0, 110201.0, 110737.0, 111274.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.12225
varray: [0.387, 1.295, 1.434, 1.373]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [109665.0, 110201.0, 110737.0, 111274.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.493456769468
varray: [0.387, 1.295, 1.434, 1.373]
inst: std_all
operation: std_all(slope_shift)
tarray: [109665.0, 110201.0, 110737.0, 111274.0]
source_col: slope_shift
SAMP_d18O_STD | 0.209656067565
varray: [-9.084, -9.37, -9.498, -9.554]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [109665.0, 110201.0, 110737.0, 111274.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.92466626977
varray: [-94.109, -97.015, -98.05, -98.311]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [109665.0, 110201.0, 110737.0, 111274.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 101.02925
varray: [96.24, 104.41, 100.346, 103.121]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [109665.0, 110201.0, 110737.0, 111274.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.61518255648
varray: [96.24, 104.41, 100.346, 103.121]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [109665.0, 110201.0, 110737.0, 111274.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.54075
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.00275
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1289.00325
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.12225
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.5057
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.5057
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -16.5996487715
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -15.97685
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [-9.084, -9.37, -9.498, -9.554]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [109665.0, 110201.0, 110737.0, 111274.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-94.109, -97.015, -98.05, -98.311]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [109665.0, 110201.0, 110737.0, 111274.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.3765
varray: [-9.084, -9.37, -9.498, -9.554]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [109665.0, 110201.0, 110737.0, 111274.0]
source_col: d(18_16)Mean
dD_mean | -96.87125
varray: [-94.109, -97.015, -98.05, -98.311]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [109665.0, 110201.0, 110737.0, 111274.0]
source_col: d(D_H)Mean
d18O_stdev | 0.209656067565
varray: [-9.084, -9.37, -9.498, -9.554]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [109665.0, 110201.0, 110737.0, 111274.0]
source_col: d(18_16)Mean
dD_stdev | 1.92466626977
varray: [-94.109, -97.015, -98.05, -98.311]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [109665.0, 110201.0, 110737.0, 111274.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.3765
varray: [-9.084, -9.37, -9.498, -9.554]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [109665.0, 110201.0, 110737.0, 111274.0]
source_col: d(18_16)Mean
measured_dD_mean | -96.87125
varray: [-94.109, -97.015, -98.05, -98.311]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [109665.0, 110201.0, 110737.0, 111274.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.209656067565
varray: [-9.084, -9.37, -9.498, -9.554]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [109665.0, 110201.0, 110737.0, 111274.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.92466626977
varray: [-94.109, -97.015, -98.05, -98.311]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [109665.0, 110201.0, 110737.0, 111274.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -12.1980248329
actual_array: [ 1.93 -15.84]
varray: [-9.084, -9.37, -9.498, -9.554]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -9.3765
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [109665.0, 110201.0, 110737.0, 111274.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -109.341734324
actual_array: [ 18.1 -119. ]
varray: [-94.109, -97.015, -98.05, -98.311]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -96.87125
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [109665.0, 110201.0, 110737.0, 111274.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.29615
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -12.83365
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 46 - Sample 46 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 46
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.08725
varray: [0.0, 2.79, 2.813, 2.746]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [111860.0, 112396.0, 112932.0, 113469.0]
source_col: residuals
SAMP_residuals_STD | 1.39177761035
varray: [0.0, 2.79, 2.813, 2.746]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [111860.0, 112396.0, 112932.0, 113469.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.12
varray: [0.0, 0.225, 0.3, -0.045]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [111860.0, 112396.0, 112932.0, 113469.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.168374582405
varray: [0.0, 0.225, 0.3, -0.045]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [111860.0, 112396.0, 112932.0, 113469.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 971.3735
varray: [0.0, 1295.948, 1294.829, 1294.717]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [111860.0, 112396.0, 112932.0, 113469.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 647.582571833
varray: [0.0, 1295.948, 1294.829, 1294.717]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [111860.0, 112396.0, 112932.0, 113469.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.04175
varray: [0.0, 1.435, 1.658, 1.074]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [111860.0, 112396.0, 112932.0, 113469.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.735004024932
varray: [0.0, 1.435, 1.658, 1.074]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [111860.0, 112396.0, 112932.0, 113469.0]
source_col: slope_shift
SAMP_d18O_STD | 2.16151958508
varray: [0.0, -4.316, -4.317, -4.336]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [111860.0, 112396.0, 112932.0, 113469.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 31.117389747
varray: [0.0, -62.77, -62.267, -61.647]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [111860.0, 112396.0, 112932.0, 113469.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 81.28175
varray: [0.0, 109.127, 108.707, 107.293]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [111860.0, 112396.0, 112932.0, 113469.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 54.1935122032
varray: [0.0, 109.127, 108.707, 107.293]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [111860.0, 112396.0, 112932.0, 113469.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.08725
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.12
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 971.3735
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.04175
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.9592
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.9592
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -31.4858277667
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -35.72435
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [0.0, -4.316, -4.317, -4.336]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [111860.0, 112396.0, 112932.0, 113469.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [0.0, -62.77, -62.267, -61.647]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [111860.0, 112396.0, 112932.0, 113469.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.24225
varray: [0.0, -4.316, -4.317, -4.336]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [111860.0, 112396.0, 112932.0, 113469.0]
source_col: d(18_16)Mean
dD_mean | -46.671
varray: [0.0, -62.77, -62.267, -61.647]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [111860.0, 112396.0, 112932.0, 113469.0]
source_col: d(D_H)Mean
d18O_stdev | 2.16151958508
varray: [0.0, -4.316, -4.317, -4.336]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [111860.0, 112396.0, 112932.0, 113469.0]
source_col: d(18_16)Mean
dD_stdev | 31.117389747
varray: [0.0, -62.77, -62.267, -61.647]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [111860.0, 112396.0, 112932.0, 113469.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.24225
varray: [0.0, -4.316, -4.317, -4.336]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [111860.0, 112396.0, 112932.0, 113469.0]
source_col: d(18_16)Mean
measured_dD_mean | -46.671
varray: [0.0, -62.77, -62.267, -61.647]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [111860.0, 112396.0, 112932.0, 113469.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 2.16151958508
varray: [0.0, -4.316, -4.317, -4.336]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [111860.0, 112396.0, 112932.0, 113469.0]
source_col: d(18_16)Mean
measured_dD_stdev | 31.117389747
varray: [0.0, -62.77, -62.267, -61.647]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [111860.0, 112396.0, 112932.0, 113469.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.04031524152
actual_array: [ 1.93 -15.84]
varray: [0.0, -4.316, -4.317, -4.336]
fit_rsq: 1.0
fit_intercept: 2.77505289814
except: True
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -3.24225
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [111860.0, 112396.0, 112932.0, 113469.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -62.1475248333
actual_array: [ 18.1 -119. ]
varray: [0.0, -62.77, -62.267, -61.647]
fit_rsq: 1.0
fit_intercept: 19.4350184537
except: True
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -46.671
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [111860.0, 112396.0, 112932.0, 113469.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.74965
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -330.4634
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 47 - Sample 47 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 47
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.78475
varray: [2.767, 2.881, 2.749, 2.742]
inst: mean_all
operation: mean_all(residuals)
tarray: [114055.0, 114592.0, 115128.0, 115665.0]
source_col: residuals
SAMP_residuals_STD | 0.0650249951942
varray: [2.767, 2.881, 2.749, 2.742]
inst: std_all
operation: std_all(residuals)
tarray: [114055.0, 114592.0, 115128.0, 115665.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.091
varray: [0.173, 0.419, -0.015, -0.213]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [114055.0, 114592.0, 115128.0, 115665.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.269542822819
varray: [0.173, 0.419, -0.015, -0.213]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [114055.0, 114592.0, 115128.0, 115665.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 1292.36225
varray: [1292.128, 1292.662, 1291.404, 1293.255]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [114055.0, 114592.0, 115128.0, 115665.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.787394172784
varray: [1292.128, 1292.662, 1291.404, 1293.255]
inst: std_all
operation: std_all(baseline_shift)
tarray: [114055.0, 114592.0, 115128.0, 115665.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.12875
varray: [1.325, 0.78, 1.056, 1.354]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [114055.0, 114592.0, 115128.0, 115665.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.268434194295
varray: [1.325, 0.78, 1.056, 1.354]
inst: std_all
operation: std_all(slope_shift)
tarray: [114055.0, 114592.0, 115128.0, 115665.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0622220486109
varray: [-2.223, -2.109, -2.096, -2.093]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [114055.0, 114592.0, 115128.0, 115665.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.00755160665
varray: [-17.985, -13.103, -11.951, -11.377]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [114055.0, 114592.0, 115128.0, 115665.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 108.133
varray: [108.195, 110.408, 107.042, 106.887]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [114055.0, 114592.0, 115128.0, 115665.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.62504010207
varray: [108.195, 110.408, 107.042, 106.887]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [114055.0, 114592.0, 115128.0, 115665.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.78475
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.091
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 1292.36225
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.12875
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.2617
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.2617
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -8.59032644553
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -8.8731
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [-2.223, -2.109, -2.096, -2.093]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [114055.0, 114592.0, 115128.0, 115665.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [-17.985, -13.103, -11.951, -11.377]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [114055.0, 114592.0, 115128.0, 115665.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.13025
varray: [-2.223, -2.109, -2.096, -2.093]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [114055.0, 114592.0, 115128.0, 115665.0]
source_col: d(18_16)Mean
dD_mean | -13.604
varray: [-17.985, -13.103, -11.951, -11.377]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [114055.0, 114592.0, 115128.0, 115665.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0622220486109
varray: [-2.223, -2.109, -2.096, -2.093]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [114055.0, 114592.0, 115128.0, 115665.0]
source_col: d(18_16)Mean
dD_stdev | 3.00755160665
varray: [-17.985, -13.103, -11.951, -11.377]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [114055.0, 114592.0, 115128.0, 115665.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.13025
varray: [-2.223, -2.109, -2.096, -2.093]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [114055.0, 114592.0, 115128.0, 115665.0]
source_col: d(18_16)Mean
measured_dD_mean | -13.604
varray: [-17.985, -13.103, -11.951, -11.377]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [114055.0, 114592.0, 115128.0, 115665.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0622220486109
varray: [-2.223, -2.109, -2.096, -2.093]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [114055.0, 114592.0, 115128.0, 115665.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.00755160665
varray: [-17.985, -13.103, -11.951, -11.377]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [114055.0, 114592.0, 115128.0, 115665.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.92406255119
actual_array: [ 1.93 -15.84]
varray: [-2.223, -2.109, -2.096, -2.093]
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -2.13025
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [114055.0, 114592.0, 115128.0, 115665.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -31.0606094248
actual_array: [ 18.1 -119. ]
varray: [-17.985, -13.103, -11.951, -11.377]
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -13.604
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [114055.0, 114592.0, 115128.0, 115665.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.05215
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -9.47465
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 48 - Sample 48 Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Sample Variables
Sample 48
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0645
varray: [0.0, 2.812, 2.73, 2.716]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [116251.0, 116788.0, 117324.0, 117860.0]
source_col: residuals
SAMP_residuals_STD | 1.3769845073
varray: [0.0, 2.812, 2.73, 2.716]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [116251.0, 116788.0, 117324.0, 117860.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | 0.152
varray: [0.0, 0.411, 0.087, 0.11]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [116251.0, 116788.0, 117324.0, 117860.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.179047479737
varray: [0.0, 0.411, 0.087, 0.11]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [116251.0, 116788.0, 117324.0, 117860.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 968.55625
varray: [0.0, 1291.48, 1291.777, 1290.968]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [116251.0, 116788.0, 117324.0, 117860.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 645.704253121
varray: [0.0, 1291.48, 1291.777, 1290.968]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [116251.0, 116788.0, 117324.0, 117860.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 1.296
varray: [0.0, 1.541, 2.031, 1.612]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [116251.0, 116788.0, 117324.0, 117860.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.890640593431
varray: [0.0, 1.541, 2.031, 1.612]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [116251.0, 116788.0, 117324.0, 117860.0]
source_col: slope_shift
SAMP_d18O_STD | 3.70059701265
varray: [0.0, -7.329, -7.405, -7.467]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [116251.0, 116788.0, 117324.0, 117860.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 29.3640318232
varray: [0.0, -57.801, -58.97, -59.368]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [116251.0, 116788.0, 117324.0, 117860.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 80.30975
varray: [0.0, 108.506, 106.003, 106.73]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [116251.0, 116788.0, 117324.0, 117860.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 53.5501545742
varray: [0.0, 108.506, 106.003, 106.73]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [116251.0, 116788.0, 117324.0, 117860.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0645
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.152
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 968.55625
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.296
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.98195
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.98195
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -32.2325985984
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -36.69635
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.0046
varray: [0.0, -7.329, -7.405, -7.467]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [116251.0, 116788.0, 117324.0, 117860.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.1447
varray: [0.0, -57.801, -58.97, -59.368]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [116251.0, 116788.0, 117324.0, 117860.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.55025
varray: [0.0, -7.329, -7.405, -7.467]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [116251.0, 116788.0, 117324.0, 117860.0]
source_col: d(18_16)Mean
dD_mean | -44.03475
varray: [0.0, -57.801, -58.97, -59.368]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [116251.0, 116788.0, 117324.0, 117860.0]
source_col: d(D_H)Mean
d18O_stdev | 3.70059701265
varray: [0.0, -7.329, -7.405, -7.467]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [116251.0, 116788.0, 117324.0, 117860.0]
source_col: d(18_16)Mean
dD_stdev | 29.3640318232
varray: [0.0, -57.801, -58.97, -59.368]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [116251.0, 116788.0, 117324.0, 117860.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.55025
varray: [0.0, -7.329, -7.405, -7.467]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [116251.0, 116788.0, 117324.0, 117860.0]
source_col: d(18_16)Mean
measured_dD_mean | -44.03475
varray: [0.0, -57.801, -58.97, -59.368]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [116251.0, 116788.0, 117324.0, 117860.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 3.70059701265
varray: [0.0, -7.329, -7.405, -7.467]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [116251.0, 116788.0, 117324.0, 117860.0]
source_col: d(18_16)Mean
measured_dD_stdev | 29.3640318232
varray: [0.0, -57.801, -58.97, -59.368]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [116251.0, 116788.0, 117324.0, 117860.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.35714186857
actual_array: [ 1.93 -15.84]
varray: [0.0, -7.329, -7.405, -7.467]
fit_rsq: 1.0
fit_intercept: 2.77505289814
except: True
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -5.55025
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: [116251.0, 116788.0, 117324.0, 117860.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -59.6691360848
actual_array: [ 18.1 -119. ]
varray: [0.0, -57.801, -58.97, -59.368]
fit_rsq: 1.0
fit_intercept: 19.4350184537
except: True
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -44.03475
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: [116251.0, 116788.0, 117324.0, 117860.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 1299.7724
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -333.28065
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 1 - Sample 1 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 1 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.78566570446
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.2712303696
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 1 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.973
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.916
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.7891573411
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.973
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.1802342549
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -66.916
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 1 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.588
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.272
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4026849618
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.588
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.8143215577
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -62.272
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 1 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.563
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.339
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3775893528
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.563
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.8773095316
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -62.339
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 1 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.361
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.824
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1748168317
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.361
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.4530292267
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -60.824
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 1 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.255
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.622
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0684114494
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.255
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3230061728
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -59.622
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 1 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.238
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.615
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0513464352
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.238
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3164253397
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -59.615
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 1 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.171
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.914
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.98409020297
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.171
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.6574019115
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -58.914
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 1 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.121
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.741
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.93389898488
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.121
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.4947613222
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -58.741
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 1 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.183
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.693
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.99613609531
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.183
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.4496356096
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -58.693
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 2 - Sample 2 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 2 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.867
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 41.241
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6977
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 38.6879
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.937
operation: col(d(18_16)Mean)
inst: col
dD_actual | 27.51
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.16639080797
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 3.937
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 7.59144368269
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: 27.51
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 2 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.867
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 41.241
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6977
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 38.6879
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.541
operation: col(d(18_16)Mean)
inst: col
dD_actual | 36.325
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.7727007225
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 4.541
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 15.8785927838
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: 36.325
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 2 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.867
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 41.241
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6977
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 38.6879
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.7
operation: col(d(18_16)Mean)
inst: col
dD_actual | 38.866
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.93230879603
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 4.7
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 18.2674351962
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: 38.866
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 2 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.867
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 41.241
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6977
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 38.6879
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.777
operation: col(d(18_16)Mean)
inst: col
dD_actual | 39.653
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.00960327189
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 4.777
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.0073088596
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: 39.653
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 2 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.867
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 41.241
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6977
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 38.6879
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.769
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.114
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.001572677
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 4.769
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.4407037247
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: 40.114
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 2 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.867
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 41.241
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6977
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 38.6879
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.784
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.396
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.01663004242
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 4.784
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.7058172864
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: 40.396
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 2 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.867
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 41.241
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6977
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 38.6879
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.832
operation: col(d(18_16)Mean)
inst: col
dD_actual | 40.695
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.06481361179
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 4.832
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.9869128713
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: 40.695
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 2 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.867
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 41.241
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6977
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 38.6879
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.852
operation: col(d(18_16)Mean)
inst: col
dD_actual | 41.027
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.08489009903
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 4.852
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.2990323837
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: 41.027
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 2 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.867
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 41.241
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6977
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 38.6879
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.918
operation: col(d(18_16)Mean)
inst: col
dD_actual | 41.052
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.15114250691
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 4.918
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.322535359
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: 41.052
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 2 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 4.867
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 41.241
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 4.6977
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 38.6879
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.867
operation: col(d(18_16)Mean)
inst: col
dD_actual | 41.241
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.09994746445
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 4.867
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.5002178525
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: 41.241
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 3 - Sample 3 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 3 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.957
operation: col(d(18_16)Mean)
inst: col
dD_actual | -90.283
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.7883935986
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -11.957
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -103.147995236
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -90.283
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 3 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.806
operation: col(d(18_16)Mean)
inst: col
dD_actual | -104.086
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.6406404817
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -12.806
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -116.124457974
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -104.086
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 3 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.051
operation: col(d(18_16)Mean)
inst: col
dD_actual | -107.145
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.8865774504
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -13.051
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -119.000282036
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -107.145
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 3 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.053
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.517
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.8885850991
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -13.053
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.290125322
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -108.517
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 3 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.144
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.271
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.979933116
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -13.144
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.998975058
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -109.271
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 3 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.151
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.877
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9869598866
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -13.151
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.56868718
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -109.877
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 3 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.224
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.253
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.060239065
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -13.224
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.922171929
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -110.253
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 3 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.199
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.5
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0351434559
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -13.199
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.154381325
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -110.5
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 3 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.175
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.672
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0110516713
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -13.175
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.316081795
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -110.672
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 3 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.286
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.843
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.1224761754
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -13.286
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.476842146
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -110.843
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 4 - Sample 4 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 4 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.945
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.123
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.761050259
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.945
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.4347198774
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -66.123
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 4 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.578
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.595
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3926467182
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.578
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.1778609858
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -61.595
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 4 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.585
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.611
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3996734888
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.585
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.2527838769
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -60.611
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 4 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.595
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.407
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4097117324
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.595
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.0609995982
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -60.407
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 4 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.565
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.11
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3795970015
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.565
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.7817842513
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -60.11
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 4 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.536
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.737
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.350486095
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.536
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.4311198593
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -59.737
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 4 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.573
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.848
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3876275964
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.573
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.5354730698
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -59.848
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 4 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.532
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.885
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3464707976
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.532
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.5702574733
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -59.885
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 4 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.577
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.05
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3916428939
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.577
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.7253771105
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -60.05
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 4 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.591
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.857
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4056964349
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.591
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.5439341409
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -59.857
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 5 - Sample 5 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 5 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.926
operation: col(d(18_16)Mean)
inst: col
dD_actual | -81.38
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.7458019579
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -8.926
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -94.7781156614
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -81.38
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 5 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.969
operation: col(d(18_16)Mean)
inst: col
dD_actual | -83.399
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.7889664055
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -8.969
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -96.676215949
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -83.399
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 5 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.025
operation: col(d(18_16)Mean)
inst: col
dD_actual | -84.178
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.8451805698
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -9.025
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.4085686602
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -84.178
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 5 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.061
operation: col(d(18_16)Mean)
inst: col
dD_actual | -84.545
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.8813182468
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -9.061
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.7535923381
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -84.545
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 6 - Sample 6 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 6 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.238
operation: col(d(18_16)Mean)
inst: col
dD_actual | -78.453
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0513464352
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.238
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -92.0263873098
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -78.453
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 6 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.184
operation: col(d(18_16)Mean)
inst: col
dD_actual | -78.214
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.99713991967
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.184
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -91.8016988657
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -78.214
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 6 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.119
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.927
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.93189133615
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.119
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -91.5318847089
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -77.927
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 6 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.112
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.96
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.92486456562
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.112
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -91.5629086363
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -77.96
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 7 - Sample 7 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 7 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 15.521
operation: col(d(18_16)Mean)
inst: col
dD_actual | -28.048
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 12.7946922151
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 15.521
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -44.639688451
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -28.048
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 7 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 15.746
operation: col(d(18_16)Mean)
inst: col
dD_actual | -26.456
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 13.0205526965
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 15.746
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -43.143018982
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -26.456
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 7 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 15.854
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.865
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 13.1289657276
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 15.854
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -42.5874086453
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -25.865
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 7 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 15.844
operation: col(d(18_16)Mean)
inst: col
dD_actual | -25.801
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 13.118927484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 15.844
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -42.5272410284
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -25.801
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 8 - Sample 8 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 8 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 8.079
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.103
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 5.32423131457
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 8.079
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.9129421679
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -41.103
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 8 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 7.93
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.734
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 5.17466148467
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 7.93
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.5061572652
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -41.734
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 8 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 7.883
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.103
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 5.12748173966
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 7.883
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.8530611811
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -42.103
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 8 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 7.887
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.749
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 5.13149703711
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 7.887
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.5202590504
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -41.749
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 9 - Sample 9 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 9 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.765
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.707
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.57271315027
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -5.765
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -85.684344447
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -71.707
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 9 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.945
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.919
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.75340153539
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -5.945
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.823768691
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -72.919
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 9 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.966
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.162
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.77448184699
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -5.966
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -87.0522176112
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -73.162
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 9 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.946
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.177
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.75440535976
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -5.946
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -87.0663193963
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -73.177
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 10 - Sample 10 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 10 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 12.915
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.736
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 10.1787259283
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 12.915
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.8074424374
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -36.736
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 10 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 13.11
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.73
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 10.3744716788
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 13.11
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.8616827102
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -35.73
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 10 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 13.157
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.714
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 10.4216514238
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 13.157
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.846640806
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -35.714
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 10 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 13.184
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.69
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 10.4487546816
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 13.184
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.8240779496
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -35.69
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 11 - Sample 11 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 11 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.186
operation: col(d(18_16)Mean)
inst: col
dD_actual | -88.77
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.0029719302
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -8.186
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -101.725595169
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -88.77
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 11 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.545
operation: col(d(18_16)Mean)
inst: col
dD_actual | -91.492
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.3633448761
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -8.545
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -104.284599123
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -91.492
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 11 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.548
operation: col(d(18_16)Mean)
inst: col
dD_actual | -92.185
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.3663563492
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -8.548
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -104.936101599
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -92.185
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 11 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.64
operation: col(d(18_16)Mean)
inst: col
dD_actual | -92.49
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.4587081905
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -8.64
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -105.222837898
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -92.49
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 12 - Sample 12 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 12 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.128
operation: col(d(18_16)Mean)
inst: col
dD_actual | -106.41
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.9600475644
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -12.128
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -118.309294561
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -106.41
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 12 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.334
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.026
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.166835383
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -12.334
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -119.828526886
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -108.026
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 12 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.363
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.321
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.1959462895
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -12.363
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.105861995
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -108.321
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 12 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -12.353
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.636
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.1859080458
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -12.353
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.401999484
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -108.636
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 13 - Sample 13 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 13 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.78566570446
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.2712303696
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 13 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.685
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.481
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5000559249
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.685
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.0108064315
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -62.481
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 13 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.651
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.489
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4659258966
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.651
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.0782083704
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -61.489
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 13 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.633
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.183
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4478570581
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.633
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.7905319524
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -61.183
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 14 - Sample 14 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 14 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.78566570446
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.2712303696
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 14 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.595
operation: col(d(18_16)Mean)
inst: col
dD_actual | -101.044
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.4250091796
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -11.595
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -113.264615936
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -101.044
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 14 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.609
operation: col(d(18_16)Mean)
inst: col
dD_actual | -101.844
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.4390627207
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -11.609
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -114.016711147
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -101.844
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 14 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.616
operation: col(d(18_16)Mean)
inst: col
dD_actual | -102.341
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.4460894912
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -11.616
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -114.483950296
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -102.341
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 15 - Sample 15 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 15 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.48
operation: col(d(18_16)Mean)
inst: col
dD_actual | -102.117
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.309569378
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -11.48
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -114.273363637
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -102.117
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 15 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.438
operation: col(d(18_16)Mean)
inst: col
dD_actual | -102.127
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.2674087548
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -11.438
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -114.282764828
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -102.127
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 15 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.447
operation: col(d(18_16)Mean)
inst: col
dD_actual | -102.088
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.2764431741
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -11.447
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -114.246100186
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -102.088
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 15 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.468
operation: col(d(18_16)Mean)
inst: col
dD_actual | -102.236
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.2975234856
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -11.468
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -114.3852378
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -102.236
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 16 - Sample 16 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 16 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 11.604
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.673
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 8.86271218994
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 11.604
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.4488100054
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -41.673
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 16 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 12.056
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.586
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 9.31644080148
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 12.056
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.6065435986
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -37.586
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 16 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 12.181
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.643
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 9.4419188467
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 12.181
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.7200113692
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -36.643
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 16 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 12.263
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.481
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 9.52423244437
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 12.263
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.5677120891
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -36.481
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 17 - Sample 17 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 17 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.862
operation: col(d(18_16)Mean)
inst: col
dD_actual | -78.137
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.677732837
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.862
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -91.7293097017
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -78.137
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 17 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.052
operation: col(d(18_16)Mean)
inst: col
dD_actual | -79.414
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.8684594657
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -8.052
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -92.9298416815
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -79.414
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 17 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.166
operation: col(d(18_16)Mean)
inst: col
dD_actual | -79.736
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.982895443
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -8.166
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -93.2325600037
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -79.736
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 17 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.197
operation: col(d(18_16)Mean)
inst: col
dD_actual | -79.898
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.0140139982
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -8.197
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -93.3848592839
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -79.898
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 18 - Sample 18 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 18 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.635
operation: col(d(18_16)Mean)
inst: col
dD_actual | -94.839
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.4575134305
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -9.635
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -107.43117746
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -94.839
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 18 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.823
operation: col(d(18_16)Mean)
inst: col
dD_actual | -96.679
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.6462324105
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -9.823
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -109.160996444
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -96.679
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 18 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.686
operation: col(d(18_16)Mean)
inst: col
dD_actual | -96.893
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.5087084729
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -9.686
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -109.362181913
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -96.893
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 18 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.746
operation: col(d(18_16)Mean)
inst: col
dD_actual | -96.871
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.5689379346
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -9.746
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -109.341499294
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -96.871
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 19 - Sample 19 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 19 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.309
operation: col(d(18_16)Mean)
inst: col
dD_actual | -88.153
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.1302666885
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -9.309
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -101.145541738
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -88.153
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 19 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.386
operation: col(d(18_16)Mean)
inst: col
dD_actual | -87.665
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.2075611644
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -9.386
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -100.686763659
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -87.665
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 19 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.307
operation: col(d(18_16)Mean)
inst: col
dD_actual | -87.237
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.1282590398
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -9.307
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -100.284392722
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -87.237
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 19 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.365
operation: col(d(18_16)Mean)
inst: col
dD_actual | -87.235
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.1864808528
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -9.365
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -100.282512483
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -87.235
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 20 - Sample 20 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 20 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.392
operation: col(d(18_16)Mean)
inst: col
dD_actual | -78.632
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.2097597487
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -8.392
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -92.1946686132
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -78.632
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 20 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.377
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.981
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.1947023833
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -8.377
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -91.5826511356
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -77.981
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 20 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.335
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.563
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.1525417601
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -8.335
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -91.1896813881
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -77.563
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 20 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.368
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.398
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.185667964
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -8.368
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -91.0345617509
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -77.398
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 21 - Sample 21 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 21 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.78566570446
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.2712303696
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 21 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.857
operation: col(d(18_16)Mean)
inst: col
dD_actual | -103.508
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.6880111624
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -11.857
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -115.581069185
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -103.508
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 21 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.832
operation: col(d(18_16)Mean)
inst: col
dD_actual | -104.014
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.6629155533
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -11.832
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -116.056769405
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -104.014
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 21 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.895
operation: col(d(18_16)Mean)
inst: col
dD_actual | -104.621
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.7261564881
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -11.895
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -116.627421646
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -104.621
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 22 - Sample 22 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 22 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.78566570446
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.2712303696
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 22 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.699
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.364
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.514109466
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.699
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.9008125069
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -62.364
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 22 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.653
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.144
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4679335454
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.653
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.7538673109
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -61.144
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 22 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.652
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.065
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.466929721
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.652
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.6795979088
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -61.065
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 23 - Sample 23 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 23 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.648
operation: col(d(18_16)Mean)
inst: col
dD_actual | -82.813
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.43996825271
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -1.648
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -96.1253062073
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -82.813
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 23 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.723
operation: col(d(18_16)Mean)
inst: col
dD_actual | -86.743
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.51525507985
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -1.723
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -99.819973929
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -86.743
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 23 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.713
operation: col(d(18_16)Mean)
inst: col
dD_actual | -87.423
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.50521683623
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -1.713
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -100.459254858
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -87.423
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 23 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -1.73
operation: col(d(18_16)Mean)
inst: col
dD_actual | -87.735
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.52228185038
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -1.73
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -100.75257199
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -87.735
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 24 - Sample 24 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 24 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.005
operation: col(d(18_16)Mean)
inst: col
dD_actual | -101.333
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -13.8327528061
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -11.005
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -113.536310331
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -101.333
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 24 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.052
operation: col(d(18_16)Mean)
inst: col
dD_actual | -101.226
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -13.8799325511
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -11.052
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -113.435717597
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -101.226
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 24 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.085
operation: col(d(18_16)Mean)
inst: col
dD_actual | -101.171
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -13.9130587551
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -11.085
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -113.384011051
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -101.171
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 24 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.061
operation: col(d(18_16)Mean)
inst: col
dD_actual | -101.222
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -13.8889669704
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -11.061
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -113.431957121
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -101.222
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 25 - Sample 25 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 25 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.162
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.855
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.9674070601
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -5.162
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.7636010741
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -72.855
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 25 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.801
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.605
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.60502646549
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -4.801
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.7082142813
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -69.605
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 25 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.758
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.799
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.56186201793
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -4.758
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.9504783567
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -68.799
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 25 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.713
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.49
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.51668992165
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -4.713
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.6599815816
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -68.49
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 26 - Sample 26 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 26 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.102
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.415
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.90335323659
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -4.102
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.5295916688
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -69.415
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 26 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.13
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.476
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.93146031872
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -4.13
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.5869389286
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -69.476
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 26 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.17
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.729
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.97161329319
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -4.17
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.8247890389
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -69.729
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 26 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.07
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.3
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.87123085701
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -4.07
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.4214779823
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -69.3
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 27 - Sample 27 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 27 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.706
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.508
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.51731187473
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -6.706
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.4373797765
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -72.508
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 27 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.923
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.783
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.73514176124
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -6.923
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.6959125052
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -72.783
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 27 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.923
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.778
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.73514176124
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -6.923
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.6912119101
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -72.778
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 27 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.823
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.146
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.63475932506
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -6.823
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.0970566938
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -72.146
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 28 - Sample 28 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 28 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.339
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.766
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.14508397214
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -5.339
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.2787403502
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -62.766
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 28 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.363
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.666
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.16917575682
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -5.363
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.2446094358
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -61.666
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 28 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.345
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.563
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.15110691831
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -5.345
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.1477771774
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -61.563
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 28 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.405
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.548
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.21133638002
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -5.405
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.1336753922
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -61.548
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 29 - Sample 29 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 29 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.78566570446
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.2712303696
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 29 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.661
operation: col(d(18_16)Mean)
inst: col
dD_actual | -84.119
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4759641403
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.661
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.3531016385
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -84.119
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 29 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.673
operation: col(d(18_16)Mean)
inst: col
dD_actual | -84.7
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4880100326
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.673
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.8993107851
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -84.7
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 29 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.685
operation: col(d(18_16)Mean)
inst: col
dD_actual | -84.912
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5000559249
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.685
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -98.0986160159
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -84.912
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 30 - Sample 30 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 30 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.78566570446
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.2712303696
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 30 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.381
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.412
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.17577150992
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -2.381
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.2453431537
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -57.412
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 30 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.302
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.387
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.09646938533
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -2.302
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.2817211652
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -56.387
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 30 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.261
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.032
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.0553125865
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -2.261
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.9479789156
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -56.032
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 31 - Sample 31 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 31 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.224
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.002
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0372928941
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.224
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.6802513978
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -60.002
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 31 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.506
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.366
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3203713642
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.506
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.0224547186
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -60.366
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 31 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.56
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.396
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3745778797
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.56
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.050658289
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -60.396
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 31 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.634
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.721
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4488608825
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.634
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.3561969683
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -60.721
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 32 - Sample 32 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 32 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.454
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.305
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.25287505013
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -3.454
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.8453454852
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -62.305
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 32 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.214
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.332
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.0119572033
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -3.214
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.8707286985
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -62.332
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 32 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.165
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.521
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.96276980957
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -3.165
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.048411192
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -62.521
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 32 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.144
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.622
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.94168949797
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -3.144
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.1433632123
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -62.622
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 33 - Sample 33 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 33 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.071
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.368
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.87605904318
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -5.071
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -84.4255250884
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -70.368
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 33 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.146
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.151
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.95134587031
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -5.146
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -85.1616382757
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -71.151
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 33 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.138
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.168
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.94331527542
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -5.138
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -85.1776202989
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -71.168
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 33 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.353
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.507
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.1591375132
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -5.353
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.4364396575
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -72.507
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 34 - Sample 34 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 34 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.612
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.442
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.41530366111
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -4.612
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.9142608031
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -63.442
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 34 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.56
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.422
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.36310479429
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -4.56
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.9553394097
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -62.422
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 34 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.625
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.346
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.42835337781
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -4.625
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.8838903647
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -62.346
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 34 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.536
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.239
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.33901300961
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -4.536
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.7832976303
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -62.239
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 35 - Sample 35 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 35 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.029
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.772
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.83772278179
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -6.029
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -85.7454521828
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -71.772
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 35 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.141
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.01
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.95015111031
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -6.141
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.9093195211
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -73.01
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 35 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.138
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.11
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.94713963722
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -6.138
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -87.0033314225
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -73.11
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 35 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.078
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.428
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.88691017551
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -6.078
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -87.3022892687
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -73.428
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 36 - Sample 36 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 36 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.628
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.473
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.4313648509
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -4.628
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.7637615321
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -66.473
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 36 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.488
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.608
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.29082944024
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -4.488
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.9505585857
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -65.608
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 36 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.513
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.348
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.31592504929
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -4.513
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.7061276422
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -65.348
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 36 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.461
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.042
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.26372618247
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -4.461
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.4184512242
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -65.042
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 37 - Sample 37 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 37 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 17.411
operation: col(d(18_16)Mean)
inst: col
dD_actual | -21.328
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 14.6919202589
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 17.411
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -38.3220886825
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -21.328
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 37 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.78566570446
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.2712303696
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 37 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 17.541
operation: col(d(18_16)Mean)
inst: col
dD_actual | -20.739
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 14.822417426
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 17.541
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -37.7683585837
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -20.739
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 37 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 17.74
operation: col(d(18_16)Mean)
inst: col
dD_actual | -19.424
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 15.022178474
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 17.74
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -36.5321020814
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -19.424
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 38 - Sample 38 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 38 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.406
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.159
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.21616456619
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -6.406
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.3488021883
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -68.159
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 38 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.714
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.882
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.52534246962
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -6.714
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.968627248
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -69.882
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 38 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.801
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.132
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.6126751891
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -6.801
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -84.2036570012
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -70.132
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 38 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.854
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.44
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.66587788028
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -6.854
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -84.4932136573
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -70.44
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 39 - Sample 39 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 39 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.978
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.968
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.78652773933
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -5.978
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.8698345226
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -72.968
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 39 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.787
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.018
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.59479728623
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -5.787
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.9168404733
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -73.018
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 39 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.795
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.041
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.60282788112
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -5.795
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.9384632106
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -73.041
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 39 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.773
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.107
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.58074374516
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -5.773
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -87.0005110654
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -73.107
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 40 - Sample 40 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 40 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.78566570446
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.2712303696
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 40 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.469
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.451
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2832298628
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.469
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.1023648347
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -60.451
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 40 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.547
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.68
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.361528163
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.547
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.3176520888
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -60.68
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 40 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.538
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.357
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3524937438
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.538
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.0139936475
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -60.357
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 41 - Sample 41 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 41 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.887
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.081
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.6913553606
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -4.887
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.3353538921
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -67.081
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 41 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.71
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.706
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.51367844856
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -4.71
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.9229282753
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -67.706
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 41 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.727
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.188
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.53074346271
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -4.727
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.3760656396
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -68.188
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 41 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.673
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.022
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.47653694718
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -4.673
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.2200058835
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -68.022
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 42 - Sample 42 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 42 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 5.169
operation: col(d(18_16)Mean)
inst: col
dD_actual | -54.857
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.40310242172
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 5.169
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.8433390751
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -54.857
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 42 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 5.308
operation: col(d(18_16)Mean)
inst: col
dD_actual | -54.754
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.54263400801
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 5.308
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.7465068167
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -54.754
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 42 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 5.382
operation: col(d(18_16)Mean)
inst: col
dD_actual | -54.002
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.61691701078
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 5.382
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.0395373188
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -54.002
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 42 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 5.392
operation: col(d(18_16)Mean)
inst: col
dD_actual | -54.026
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.6269552544
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 5.392
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.0621001751
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -54.026
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 43 - Sample 43 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 43 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.22
operation: col(d(18_16)Mean)
inst: col
dD_actual | -95.403
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.0409263203
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -9.22
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -107.961404583
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -95.403
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 43 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.492
operation: col(d(18_16)Mean)
inst: col
dD_actual | -98.079
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.3139665467
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -9.492
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -110.477163062
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -98.079
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 43 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.52
operation: col(d(18_16)Mean)
inst: col
dD_actual | -98.749
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.3420736289
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -9.52
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -111.107042801
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -98.749
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 43 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.636
operation: col(d(18_16)Mean)
inst: col
dD_actual | -99.179
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.4585172548
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -9.636
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -111.511293977
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -99.179
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 44 - Sample 44 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 44 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 13.562
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.006
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 10.8282002904
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 13.562
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.8816316105
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -40.006
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 44 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 14.035
operation: col(d(18_16)Mean)
inst: col
dD_actual | -36.46
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 11.3030092135
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 14.035
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.5479695898
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -36.46
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 44 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 13.977
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.076
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 11.2447874005
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 13.977
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.2468448756
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -35.076
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 44 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 14.147
operation: col(d(18_16)Mean)
inst: col
dD_actual | -34.399
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 11.415437542
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 14.147
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -50.6103843036
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -34.399
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 45 - Sample 45 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 45 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.084
operation: col(d(18_16)Mean)
inst: col
dD_actual | -94.109
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.9044062071
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -9.084
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -106.74489058
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -94.109
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 45 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.37
operation: col(d(18_16)Mean)
inst: col
dD_actual | -97.015
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.1914999746
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -9.37
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -109.476876432
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -97.015
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 45 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.498
operation: col(d(18_16)Mean)
inst: col
dD_actual | -98.05
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.3199894929
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -9.498
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -110.449899611
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -98.05
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 45 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.554
operation: col(d(18_16)Mean)
inst: col
dD_actual | -98.311
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.3762036572
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -9.554
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -110.695270673
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -98.311
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 46 - Sample 46 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 46 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.78566570446
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.2712303696
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 46 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.316
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.77
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.11817165001
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -4.316
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.2825008263
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -62.77
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 46 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.317
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.267
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.11917547437
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -4.317
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.8096209627
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -62.267
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 46 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.336
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.647
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.13824813725
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -4.336
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.2267471745
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -61.647
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 47 - Sample 47 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 47 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.223
operation: col(d(18_16)Mean)
inst: col
dD_actual | -17.985
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.01716726075
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -2.223
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -35.1792708215
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -17.985
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 47 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.109
operation: col(d(18_16)Mean)
inst: col
dD_actual | -13.103
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.90273128351
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -2.109
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -30.5896097992
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -13.103
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 47 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.096
operation: col(d(18_16)Mean)
inst: col
dD_actual | -11.951
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.8896815668
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -2.096
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -29.506592696
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -11.951
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 47 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.093
operation: col(d(18_16)Mean)
inst: col
dD_actual | -11.377
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.88667009372
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -2.093
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -28.9669643824
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -11.377
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean

## 48 - Sample 48 Detail Vars
ChemCorrect(tm) - Mon, May 19, 2025, 14:21:29
Detail Variables
Detail 48 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -2.78566570446
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -18.2712303696
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 48 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.329
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.801
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1426944521
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.329
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.6110494498
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -57.801
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 48 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.405
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.97
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2189851036
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.405
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.7100485762
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -58.97
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean
Detail 48 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 4.6977
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 38.6879
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.286
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.843
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.0046
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.1447
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.467
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.368
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2812222141
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 2.77505289814
measure_array: [ 4.6977 -13.0046]
inst: fit_type1
actual_mean: -7.467
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.996190208216
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.0842159435
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 19.4350184537
measure_array: [ 38.6879 -107.1447]
inst: fit_type1
actual_mean: -59.368
operation: fit_type1(d(D_H)Mean)
fit_slope: 1.06369511306
tarray: []
source_col: d(D_H)Mean